## Events
Event ID | Report Number | MDR Report Key | MAUDE Report Link | Event Date | Report Date | Report Date | Date Received | Date Report To FDA | Date Report To Manufacturer | Date Manufacturer Received | Device Date Received 1 | Device Date Received 2 | Device Date Received 3 | Device Expiration Date 1 | Device Expiration Date 2 | Patient Date Received 1 | Patient Date Received 2 | Product Class 1 | Product Class 2 | Product Class 3 | Brand Name 1 | Brand Name 2 | Brand Name 3 | Product Code 1 | Product Code 2 | Product Code 3 | Model Number 1 | Model Number 2 | Model Number 3 | Manufacturer 1 | Manufacturer 2 | Manufacturer 3 | Manufacturer Country 1 | Manufacturer Country 2 | Lot Number 1 | Lot Number 2 | Catalog Number 1 | Catalog Number 2 | Device Availability 1 | Device Availability 2 | Device Availability 3 | Device Evaluated By Manufacturer 1 | Device Evaluated By Manufacturer 2 | Single Use Flag | Reprocessed And Reused Flag | Device Operator 1 | Device Operator 2 | Device Operator 3 | Report Source Code | Health Professional | Reporter Occupation Code | Source Type 1 | Source Type 2 | Source Type 3 | Source Type 4 | Source Type 5 | Source Type 6 | Patient Age 1 | Patient Age 2 | Patient Sex 1 | Patient Weight 1 | Patient Ethnicity 1 | Patient Race 1 | Event Type | Adverse Event Flag | Product Problem Flag | Device Problem 1 | Device Problem 2 | Device Problem 3 | Device Problem 4 | Device Problem 5 | Device Problem 6 | Device Problem 7 | Patient Problem 1 | Patient Problem 2 | Patient Problem 3 | Patient Problem 4 | Patient Problem 5 | Patient Problem 6 | Patient Problem 7 | Patient Problem 8 | Patient Problem 9 | Patient Problem 10 | Patient Problem 11 | Patient Problem 12 | Patient Problem 13 | Patient Outcome 1 | Patient Outcome 2 | Patient Outcome 3 | Patient Outcome 4 | Patient Outcome 5 | Patient Outcome 6 | Patient Treatment 1 | Patient Treatment 2 | Patient Treatment 3 | Patient Treatment 4 | Patient Treatment 5 | Patient Treatment 6 | Patient Treatment 7 | Patient Treatment 8 | Patient Treatment 9 | Patient Treatment 10 | Patient Treatment 11 | Patient Treatment 12 | Patient Treatment 13 | Patient Treatment 14 | Patient Treatment 15 | Patient Treatment 16 | Patient Treatment 17 | Patient Treatment 18 | Patient Sequence Number Treatment 2 1 | Previous Use Code | Remedial Action 1 | Removal Correction Number | Initial Report To Fda | Type Of Report 1 | Type Of Report 2 | Type Of Report 3 | Type Of Report 4 | Type Of Report 5 | Date Facility Aware | Report To Fda | Event Location | Device Other Id Number 1 | Device Other Id Number 2 | Description of Event or Problem | Description of Event or Problem 2 | Description of Event or Problem 3 | Description of Event or Problem 4 | Description of Event or Problem 5 | Description of Event or Problem 6 | Additional Manufacturer Narrative | Additional Manufacturer Narrative 2 | Additional Manufacturer Narrative 3 | Additional Manufacturer Narrative 4 | Additional Manufacturer Narrative 5 | Additional Manufacturer Narrative 6
1 | 1220648-2026-04181 | 24466139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24466139&pc=OZD | 11/02/2025 |  |  | 02/27/2026 |  |  | 11/02/2025 | 02/27/2026 |  |  |  |  | 02/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745774 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A (b)(6) male presented with acute myocardial infarction complicated by cardiogenic shock (scai stage d). the patient was evaluated for surgical revascularization and the device was implanted as a bridge to coronary artery bypass grafting (cabg).the device was successfully implanted without reported procedural complications. no additional detailed clinical course or hospital documentation was provided to the manufacturer. during follow-up communication with the treating team, it was mentioned that the patient experienced a stroke during the hospitalization. the team also reported that the patient was successfully weaned from the device on day 8 post-implant.no further information regarding diagnostics, treatments, patient condition after weaning, or ultimate clinical outcome was made available. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
2 | 1220648-2026-04179 | 24466165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24466165&pc=OZD | 11/01/2025 |  |  | 02/27/2026 |  |  | 11/01/2025 | 02/27/2026 |  |  |  |  | 02/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026738971 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 99 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The patient is a (b)(6) male with a history heart failure with reduced ejection fraction secondary to non-ischemic cardiomyopathy, chronic kidney disease stage 3b, and polysubstance use (alcohol and cocaine). the patient presented in cardiogenic shock after running out of prescribed medications. prior to device implantation, the patient experienced worsening cardiogenic shock and renal failure requiring initiation of continuous renal replacement therapy (crrt) and extracorporeal life support (ecls). the initial planned device implantation was postponed for approximately 11 days due to cadiac clot formation. no documentation was provided regarding specific antithrombotic management during this interval. the patient subsequently underwent successful implantation of an impella 5.5 device as a bridge to left ventricular assist device (lvad) therapy. the patient was successfully bridged to lvad implantation. during the support period, the aortic placement signal was reported as unreliable at times. the field representative relayed information that he ao pressure on the impella was reset. after it was reset, it correlated w aline. over the 15-day duration of support (noted as off-label use), one repositioning of the device was performed due to arrhythmia. on day 28 of support, minor bleeding at the access site occurred after patient mobilization. according to the clinical team, the issue was managed following the instructions for use, including application of manual pressure and dressing changes. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3 | 1220648-2026-04185 | 24468666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24468666&pc=OZD | 02/17/2026 | 02/27/2026 |  | 02/27/2026 |  |  | 02/17/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026774679 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
4 | 1220648-2026-04190 | 24468769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24468769&pc=OZD | 02/18/2026 | 02/27/2026 |  | 02/27/2026 |  |  | 02/18/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026813005 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
5 | 1220648-2026-04189 | 24468771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24468771&pc=OZD | 02/17/2026 |  |  | 02/27/2026 |  |  | 02/17/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801120 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
6 | 1220648-2026-04191 | 24468777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24468777&pc=OZD | 02/17/2026 |  |  | 02/27/2026 |  |  | 02/17/2026 | 02/27/2026 |  |  |  |  | 02/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1815943 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was inserted via right axillary artery in a 70 year old male who was admitted for acute myocardial infarction with cardiogenic shock. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage c. the patient received support from the 5.5 while in the intensive care unit. on day 28 of support, the patient became septic. a tear in the sterile sheath cover was observed, cleaned with chlorhexidine gluconate, and an occlusive tegaderm applied. the 5.5 was not considered to be a contributing factor to the sepsis as the patient had already had an increase in white blood cell count. that same day, there was an error with the automated impella controller during purge change, which resolved without intervention. this event is conservatively reported as the sleeve tear prompted intervention, but there was no lasting harm or injury reported and no consequence to support. the controller error will be conservatively reported for hemodynamic instability during purge change, but there was no intervention needed and no lasting harm or injury reported. | [5 empty] | A4 weight is unknown this report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate. related medical device reports: this automated impella controller device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5.
7 | 1220648-2026-04196 | 24469063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24469063&pc=OZD | 02/20/2026 |  |  | 02/27/2026 |  |  | 02/20/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1678954 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
8 | 1220648-2026-04197 | 24469107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24469107&pc=OZD | 02/20/2026 |  |  | 02/27/2026 |  |  | 02/20/2026 | 02/27/2026 |  |  |  |  | 02/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1253378 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the autotest could not be completed on an automated impella controller. a new controller was used to complete the procedure. no patient harm was reported. | [5 empty] | Patient information was not provided. this report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc or its employees that the report constitutes an admission that the product, abiomed inc, or its employees, caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
9 | 1220648-2026-04201 | 24469279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24469279&pc=OZD | 02/19/2026 |  |  | 02/27/2026 |  |  | 02/20/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779058 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Increase in Pressure | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
10 | 1220648-2026-04202 | 24469490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24469490&pc=OZD | 02/18/2026 |  |  | 02/27/2026 |  |  | 02/18/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026656861 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
11 | 1220648-2026-04205 | 24469601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24469601&pc=OZD | 02/18/2026 |  |  | 02/27/2026 |  |  | 02/18/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2027818913 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
12 | 1220648-2026-04208 | 24469684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24469684&pc=OZD | 02/20/2026 |  |  | 02/27/2026 |  |  | 02/20/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026815946 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Increase in Pressure | [6 empty] | Perforation of Vessels | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
13 | 1220648-2026-04207 | 24469903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24469903&pc=OZD | 02/20/2026 | 02/27/2026 |  | 02/27/2026 |  |  | 02/20/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026778434 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Restricted Flow rate | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
14 | 1220648-2026-04217 | 24469943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24469943&pc=OZD | 02/18/2026 |  |  | 02/27/2026 |  |  | 02/18/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1456617 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
15 | 1220648-2026-04221 | 24470078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24470078&pc=OZD | 02/20/2026 |  |  | 02/27/2026 |  |  | 02/20/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026727040 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Ischemia | Vascular Dissection | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
16 | 1220648-2026-04096 | 24470155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24470155&pc=OZD | 02/18/2026 |  |  | 02/27/2026 |  |  | 02/18/2026 | 02/27/2026 |  |  |  |  | 02/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026751888 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male | 80 KG |  |  | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella cp was attempted via right femoral artery in a 71-year-old male admitted for a cardiac ablation. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was identified as scai shock stage a. during implant, the device could not be implanted due to patient anatomy. due to tortuosity of the vessels, the introducer sheath kinked and attempts to pass through vasculature were unsuccessful. the cp implant was aborted and the ablation proceeded without impella support. the delivery issue will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the position attempts; however, the attempts and removal were performed with no consequence to the patient. | [5 empty] | This is one of two related reports. this report represents the impella cp and another report was submitted to represent the introducer. this report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
17 | 1220648-2026-04223 | 24470157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24470157&pc=OZD | 02/17/2026 |  |  | 02/27/2026 |  |  | 02/20/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026780326 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
18 | 1220648-2026-04225 | 24470441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24470441&pc=OZD | 02/17/2026 |  |  | 02/27/2026 |  |  | 02/17/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1925275 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
19 | 1220648-2026-04227 | 24470679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24470679&pc=OZD | 02/18/2026 | 02/27/2026 |  | 02/27/2026 |  |  | 02/18/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1177958 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
20 | 1220648-2026-04230 | 24471062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24471062&pc=OZD | 02/20/2026 |  |  | 02/27/2026 |  |  | 02/20/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2027821527 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
21 | 1220648-2026-04232 | 24471081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24471081&pc=OZD | 02/19/2026 |  |  | 02/27/2026 |  |  | 02/19/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026618478 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
22 | 1220648-2026-04238 | 24471586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24471586&pc=OZD | 02/21/2026 |  |  | 02/27/2026 |  |  | 02/21/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797267 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
23 | 1220648-2026-04240 | 24471659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24471659&pc=OZD | 02/21/2026 |  |  | 02/27/2026 |  |  | 02/21/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026812430 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Unspecified Infection | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
24 | 1220648-2026-04243 | 24471812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24471812&pc=OZD | 02/21/2026 |  |  | 02/27/2026 |  |  | 02/21/2026 | 02/27/2026 |  |  |  |  | 02/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801721 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in a 57-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs), scai shock stage e, with a history of prior coronary artery bypass grafting (cabg) and known coronary artery disease (cad). during the procedure, cannula kinking occurred within the body, preventing the team from achieving adequate device position. after multiple unsuccessful attempts to correct the positioning, the physician made the decision to abort the placement and explant the pump. the patient was subsequently cannulated for ecmo for ongoing circulatory support. the reported events include product damage, medical device removal, additional device required and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during device removal; however, the removal was performed with no consequence to the patient. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
25 | 1220648-2026-04248 | 24472146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24472146&pc=OZD | 02/20/2026 |  |  | 02/27/2026 |  |  | 02/21/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026757921 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
26 | 1220648-2026-04251 | 24472303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24472303&pc=OZD | 02/18/2026 |  |  | 02/27/2026 |  |  | 02/18/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1955422 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
27 | 1220648-2026-04252 | 24472395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24472395&pc=OZD | 10/13/2025 |  |  | 02/27/2026 |  |  | 10/14/2025 | 02/27/2026 |  |  |  |  | 02/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026658602 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 110 KG |  |  | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 54-year-old male with a history of coronary artery disease and multiple percutaneous coronary interventions treated for advanced heart failure presented with chest pain and shortness of breath. an impella 5.5 was inserted and the patient prepared for heart transplantation. after an off-label prolonged support duration of 5 months, the patient suffered a thrombotic stroke treated by interventional thrombectomy. shortly after, a "purge pressure high" alarm occurred in conjunction with low purge flows. the off-label administration of a lytic agent via the purge was performed but did not resolve the issue so that the device was weaned and removed. on day 151 of impella 5.5 support, the patient experienced high purge pressure with an active purge system blocked alarm. medication (tpa) was started in response and the clinical support center adviced the nurse to continue monitoring motor current while waiting for purge flow and purge pressure to restore to normal. approximately 30 minutes later, the nurse called back to report a purge system failure alarm. the alarm was resolved by replacing the purge cassette. the hospital staff continued running tpa through the purge line and monitoring motor current. on day 119 of impella 5.5 support, the patient experienced a placement signal issue. the abiomed representative notes that placement signal was lost. there was no futher mention of this issue so, with the available information, there was no patient consequence due to this issue. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
28 | 1220648-2026-04254 | 24472487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24472487&pc=OZD | 02/17/2026 |  |  | 02/27/2026 |  |  | 02/17/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1726714 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
29 | 1220648-2026-04253 | 24472842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24472842&pc=OZD | 12/20/2025 |  |  | 02/27/2026 |  |  | 12/20/2025 | 02/27/2026 |  |  |  |  | 02/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026740932 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 101 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Unspecified Infection | Hypovolemia | Respiratory Failure | Multiple Organ Dysfunction Syndrome | Pericardial Effusion | Unspecified Tissue Injury |  |  |  |  | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A fifty-two-year-old male patient with a medical history of coronary artery disease received an impella cp device for acute myocardial infarction¿related cardiogenic shock, classified as society for cardiovascular angiography and interventions cardiogenic shock stage e. prior to device placement, the patient was receiving three intravenous medications to support heart muscle contraction, an intra-aortic balloon pump, and mechanical ventilation. due to profound shock, the patient was escalated from impella cp to an impella 5.5 device on the same day. on the first day of support, the patient showed mild improvement with impella cp but subsequently deteriorated, leading to initiation of venoarterial extracorporeal membrane oxygenation. on the fourth day, the patient was extubated and escalated to full impella 5.5 support. on the fifteenth day, the patient developed a fever, blood cultures were drawn, and minimal output persisted from the pericardial drain; the patient also received blood products, likely related to the combined use of mechanical circulatory support devices. on the sixteenth day, renal function worsened and dialysis was initiated. on the seventeenth day, the patient developed gastrointestinal bleeding, followed by hypovolemia and hypotension that required transfusion; no surgical intervention was needed. on the twenty-first day, the patient underwent an exploratory laparotomy with partial resection of the gastrointestinal tract to control bleeding. in the following days, gastrointestinal bleeding continued and additional blood products were administered. on the thirtieth day, the patient developed respiratory acidosis. on the thirty-sixth day, veno-venous extracorporeal membrane oxygenation was initiated for respiratory failure, and the circuit was removed on the fiftieth day. on the sixty-fifth day, a high purge-pressure alarm occurred on the impella system; engineering review will provide further evaluation, and no direct adverse patient outcome was reported. over the subsequent weeks, the patient demonstrated gradual clinical improvement. as of the current date, after approximately four months of support, the patient continues on impella 5.5 therapy. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
30 | 1220648-2026-04257 | 24473857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24473857&pc=OZD | 12/20/2025 |  |  | 02/27/2026 |  |  | 12/20/2025 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794861 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
31 | 1220648-2026-04258 | 24473914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24473914&pc=OZD | 02/19/2026 |  |  | 02/27/2026 |  |  | 02/19/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798624 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
32 | 1220648-2026-04255 | 24473967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24473967&pc=OZD | 02/18/2026 |  |  | 02/27/2026 |  |  | 02/18/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026814057 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Fluid/Blood Leak | Decrease in Pressure | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
33 | 1220648-2026-04259 | 24474447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24474447&pc=OZD | 12/20/2025 |  |  | 02/27/2026 |  |  | 12/20/2025 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026690954 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Thrombocytopenia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
34 | 1220648-2026-04262 | 24474700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24474700&pc=OZD | 02/20/2026 |  |  | 02/27/2026 |  |  | 02/20/2026 | 02/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1571192 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
35 | 1220648-2026-04072 | 24453961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24453961&pc=OZD | 02/12/2026 |  |  | 02/26/2026 |  |  | 02/12/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1116075 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
36 | 1220648-2026-04027 | 24453964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24453964&pc=OZD | 02/17/2026 |  |  | 02/26/2026 |  |  | 02/18/2026 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806666 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female | 70 KG |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp via percutaneous right femoral artery for mechanical circulatory support. the placement signal was disturbed at the begin of the case for about 25 minutes. the placement signal waveform was visible on the automated impella controller screen after the issue occurred. pump position was confirmed via echocardiogram. the pump flow appeared ok. the procedure was not affected and was successful. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
37 | 1220648-2026-04080 | 24454274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24454274&pc=OZD | 02/17/2026 |  |  | 02/26/2026 |  |  | 02/17/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1210879 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
38 | 1220648-2026-04082 | 24454302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24454302&pc=OZD | 02/12/2026 |  |  | 02/26/2026 |  |  | 02/12/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779310 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Device Sensing Problem | [6 empty] | Cardiac Perforation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
39 | 1220648-2026-04084 | 24454380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24454380&pc=OZD | 02/19/2026 |  |  | 02/26/2026 |  |  | 02/19/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1109485 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | False Alarm | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
40 | 1220648-2026-04085 | 24454406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24454406&pc=OZD | 02/19/2026 |  |  | 02/26/2026 |  |  | 02/19/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1087349 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Application Program Problem | Connection Problem | Device Sensing Problem |  |  |  |  | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
41 | 1220648-2026-04087 | 24454425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24454425&pc=OZD | 02/19/2026 |  |  | 02/26/2026 |  |  | 02/19/2026 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802626 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 64 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted surgically into the right axillary artery in a 66-year-old female patient with a past medical history of renal insufficiency presenting with acute decompensated chronic non-ischemic cardiomyopathy in scai stage c shock on inotropes and vasopressors prior to initiation of support. while the patient was in the intensive care unit (icu), the patient complained of weakness and numbness in the right arm and hand. there was no apparent swelling, hematoma, or bleeding. the patient later reported the symptoms seemed to be improving. the post-procedure outcome was unknown. the device functioned at p-7 at 3.7 l/min with no issues. limb ischemia is listed as a potential adverse event and is consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc or its employees that the report constitutes an admission that the product, abiomed inc, or its employees, caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
42 | 1220648-2026-04094 | 24454689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24454689&pc=OZD | 02/18/2026 |  |  | 02/26/2026 |  |  | 02/18/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1154073 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
43 | 1220648-2026-04099 | 24454807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24454807&pc=OZD | 02/18/2026 |  |  | 02/26/2026 |  |  | 02/18/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026812424 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | [10 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
44 | 1220648-2026-04102 | 24454877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24454877&pc=OZD | 02/16/2026 |  |  | 02/26/2026 |  |  | 02/17/2026 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026813024 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 135 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Vascular Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was inserted via right axillary artery in a 54 year old male who was admitted for cardiomyopathy. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage c. on day 7 of support, while in the intensive care unit, puffiness and a hematoma was observed at the right axillary insertion site. the site remained stable. vascular injury is a known potential adverse event of the impella 5.5 per the instructions for use. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
45 | 1220648-2026-04107 | 24455142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24455142&pc=OZD | 02/18/2026 |  |  | 02/26/2026 |  |  | 02/18/2026 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026805577 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male | 83 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pain | Numbness | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 was inserted via right axillary artery in a 40 year old male who was admitted for cardiomyopathy. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage d. the patient received support from the 5.5 while in the intensive care unit. on day 6 of support, the patient complained of numbness to his right upper extremity, which occurred off and on for a few days. there was no paresthesia, and the right arm and hand were pink, warm and with full range of motion. the axillary site was stable and within normal limits. tylenol and neurontin were prescribed for pain management if needed. this event is conservatively reported as numbness prompted intervention, but there was no lasting harm or injury reported. the patient currently /remains on support. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
46 | 1220648-2026-04108 | 24455191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24455191&pc=OZD | 02/16/2026 |  |  | 02/26/2026 |  |  | 02/17/2026 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026743371 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female | 77 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted into the right axillary/subclavian artery in a 44-year-old female patient presenting in cardiomyopathy, and in scai stage d shock, and on multiple inotropes and vasopressors, prior to initiation of support. while the patient was in the intensive care unit (icu), the patient experienced an ischemic stroke. it is unknown if the patient received any treatment. the post-procedure outcome is unknown. clots can also occur due to inadequate anticoagulation. cerebral vascular accident/stroke is listed as a potential adverse event, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
47 | 1220648-2026-04112 | 24455390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24455390&pc=OZD | 02/17/2026 |  |  | 02/26/2026 |  |  | 02/17/2026 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026813091 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted into the right axillary/subclavian artery in a 64-year-old male patient with a past medical history of a prior coronary artery disease (cad), presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), and in scai stage a shock, prior to initiation of support. the impella was inserted for ventricular support for a high-risk surgery: coronary artery bypass graft (cabg). while the patient was in the intensive care unit (icu), the patient had an episode of ventricular tachycardia (vt) and was debrillated once. impella position was verified under echocardiogram, and was confirmed to be in good position. the physician attributed the vt as possibly related to the cabg from the previous day. the post-procedure outcome is unknown. the impella functioned at p-4 at 2.7 without issues. coronary artery bypass graft surgery can cause ventricular tachycardia, particularly in the early postoperative recovery period. as blood flow returns to previously ischemic heart muscle, it can cause transient ventricular tachycardia. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
48 | 1220648-2026-04113 | 24455588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24455588&pc=OZD | 08/29/2025 |  |  | 02/26/2026 |  |  | 08/29/2025 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1953067 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | A 69 year old male with a past medical history significant for ischemic cardiomyopathy, systolic heart failure, cad, prior stemi, icd, and hypertension presented to the emergency department with chest pain, arrhythmias, and shortness of breath. the impella 5.5 (indication for use: cardiomyopathy; model ic9731; serial number (b)(6) was placed following hemodynamic decline after an iabp failed to augment overnight. the patient was taken to the or where the device was implanted via right axillary graft with tee and fluoroscopic guidance, positioned appropriately with inlet mid ventricle at 4 cm, and secured using standard three point fixation. based on the information provided, the reported aic battery behavior occurred only when the controller was unplugged and charging transitioned to internal battery power. the device resumed normal charging function when reconnected, and an exchange was recommended out of caution. heparin was stopped due to a spontaneous nosebleed and hemostasis was achieved; this was not attributed to the impella device. a suction alarm occurred which resolved after discontinuation of a lasix infusion. no further alarms occurred. no impella 5.5 pump malfunction, performance issue, or patient harm was identified. the suction alarm on 7/30 resolved with clinical management and was not attributed to device failure. the impella 5.5 continued to function as intended while the patient awaited transplant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
49 | 1220648-2026-04116 | 24455787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24455787&pc=OZD | 07/28/2025 |  |  | 02/26/2026 |  |  | 07/29/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1885195 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Noise, Audible | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
50 | 1220648-2026-04117 | 24455888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24455888&pc=OZD | 08/23/2025 |  |  | 02/26/2026 |  |  | 08/23/2025 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026661774 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | 60m history of pulmonary fibrosis, copd, cocaine abuse. presented with cp/ nstemi. patient had an impella assisted opcab x3 (lima to lad/ svg to d1, svg to plb) there were no noted device related alarms or issues with impella during support. there was a note made of transfusion, this does not appear to be related to any bleeding event/ likely transfused for perioperative anemia. the device was successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
51 | 1220648-2026-04118 | 24455940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24455940&pc=OZD | 09/20/2025 |  |  | 02/26/2026 |  |  | 09/20/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026710513 |  | 1000100 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
52 | 1220648-2026-04120 | 24456424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24456424&pc=OZD | 02/19/2026 |  |  | 02/26/2026 |  |  | 02/19/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026807949 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Multiple Organ Dysfunction Syndrome | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
53 | 1220648-2026-04121 | 24456427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24456427&pc=OZD | 02/18/2026 |  |  | 02/26/2026 |  |  | 02/18/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026699501 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
54 | 1220648-2026-04122 | 24456445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24456445&pc=OZD | 09/24/2025 |  |  | 02/26/2026 |  |  | 09/24/2025 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026717743 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 140 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Fever | Hemolysis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A (b)(6) m hx of familial cardiomyopathy, htn, aki in the setting of ckd, aicd, non-compliance. presented with acutely decompensated chronic hf. on 9/23 underwent impella 5.5 implant (r axillary artery) as a btlvad. on 9/23 the patient returned to or 6 hours post implant for worsening metabolic acidosis. impella p level was increased with no impella related concerns, appropriate position confirmed on echo and crrt was initiated. on 9/25 the patient experienced ventricular tachycardia and the impella 5.5 repositioned at bedside with subsequent resolution of ventricular tachycardia. on 9/28 impella 5.5 repositioned for rotation towards mitral valve. on 9/29 the patient was noted to be febrile and diagnosed with pneumonia. on 9/30 ldh elevated/ pfhgb30 which resolved with repositioning based on repeat labs. on 10/2 the patient returned to the or for lvad placement and successful explant of impella 5.5 | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
55 | 1220648-2026-04123 | 24456694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24456694&pc=OZD | 09/25/2025 |  |  | 02/26/2026 |  |  | 09/25/2025 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026717757 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 73 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 63 year old female with unknown history presented with ami-cs on (b)(6) and underwent placement of femoral impella cp and protected pci of lad. on (b)(6) they were taken to or and escalated to impella 5.5 with explant of impella cp. on (b)(6) there were noted changes in neuro status and the head ct showed small bleed/ heparin held/ stable device function. on 9/25 ther was a sudden drop in purge flow/ increased purge pressure/ tpa started in purge with resolution, which should be noted as off-label use. on 9/26 taken to or for lvad, successful explant of impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
56 | 1220648-2026-04124 | 24456887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24456887&pc=OZD | 09/26/2025 |  |  | 02/26/2026 |  |  | 09/26/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026709935 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
57 | 1220648-2026-04125 | 24456933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24456933&pc=OZD | 02/16/2026 |  |  | 02/26/2026 |  |  | 02/19/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026785642 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | Material Integrity Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
58 | 1220648-2026-04077 | 24457025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24457025&pc=OZD | 02/14/2026 |  |  | 02/26/2026 |  |  | 02/17/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026804712 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
59 | 1220648-2026-04127 | 24457042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24457042&pc=OZD | 03/14/2025 |  |  | 02/26/2026 |  |  | 09/26/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025486895 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Obstruction/Occlusion | Thrombocytopenia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
60 | 1220648-2026-04128 | 24457328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24457328&pc=OZD | 09/26/2025 |  |  | 02/26/2026 |  |  | 09/26/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026700168 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
61 | 1220648-2026-04129 | 24457479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24457479&pc=OZD | 09/26/2025 |  |  | 02/26/2026 |  |  | 09/26/2025 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026678493 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 60 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 34-year-old male with history of nonischemic cardiomyopathy presented with acute/chronic with decompensated heart failure had an impella 5.5 inserted via right axillary artery as bridge to transplant. on 9/26 a defibrillator was implanted. on 9/29 there was device repositioning and rv dysfunction noted. on 10/6 worsening right ventricle function noted and on 10/9 underwent cardiac transplant with explant of impella 5.5 without incident. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
62 | 1220648-2026-04131 | 24457596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24457596&pc=OZD | 02/21/2026 |  |  | 02/26/2026 |  |  | 02/21/2026 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2027816762 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp was initially inserted via percutaneous femoral arterial access in an 81-year-old male who presented with acute myocardial infarction complicated by cardiogenic shock classified as society for cardiovascular angiography and interventions stage d. the patient required intubation, mechanical ventilation for respiratory support, and vasopressor therapy due to hemodynamic instability. following transfer to the intensive care unit, the patient developed nausea and emesis. a hematoma was noted at the groin access site. the patient experienced an access site¿related complication associated with the impella cp, and device revision or replacement was performed in the setting of hematoma formation. access site hematoma is a known potential complication of large-bore femoral arterial access, particularly in critically ill patients receiving anticoagulation and vasopressor support. due to persistent hemodynamic instability and clinical deterioration, escalation of mechanical circulatory support was pursued. the patient was taken to the operating room for placement of an impella 5.5 via surgical cutdown and graft to the subclavian artery. although graft anastomosis was successfully completed, significant subclavian tortuosity prevented full advancement of the impella 5.5 into the ascending aorta despite multiple attempts and use of adjunctive techniques. after unsuccessful advancement, the impella 5.5 was removed. the patient experienced hemodynamic instability in the setting of advanced cardiogenic shock and failure to advance the 5.5 device, necessitating device revision and replacement. failure to advance in the presence of vascular tortuosity is a known procedural risk and is described in the instructions for use. no information was provided to suggest a device-related issue. a subsequent 5.5 device placement was reported without complaint. the patient survived. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
63 | 1220648-2026-04132 | 24457611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24457611&pc=OZD | 09/26/2025 |  |  | 02/26/2026 |  |  | 09/26/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026703574 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
64 | 1220648-2026-04137 | 24458170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24458170&pc=OZD | 02/19/2026 |  |  | 02/26/2026 |  |  | 02/19/2026 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766839 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted via the femoral artery in a patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). upon arrival to the icu, the aic displayed the notification ¿impella position aorta.¿ the cause was unknown. repositioning was discussed, and best practices were reviewed with the care team. after consultation with the icu nurse and further discussion among the medical team, the pump was determined to be positioned in the aorta, and the decision was made to explant the impella cp. the patient remained on ecmo support, and because the impella cp could not be re-advanced into the ventricle, the device was explanted and escalated to an impella 5.5. the patient survived to explant. the reported events include device in wrong position, device revision or replacement, and hemodynamic instability. the impella cp will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
65 | 1220648-2026-04079 | 24459092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24459092&pc=OZD | 02/18/2026 |  |  | 02/26/2026 |  |  | 02/18/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1112882 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Operating System Becomes Nonfunctional | Key or Button Unresponsive/not Working | [5 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
66 | 1220648-2026-04138 | 24459257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24459257&pc=OZD | 09/26/2025 |  |  | 02/26/2026 |  |  | 09/26/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026734190 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Fever | Hematoma | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
67 | 1220648-2026-04139 | 24459328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24459328&pc=OZD | 02/18/2026 |  |  | 02/26/2026 |  |  | 02/18/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026733491 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
68 | 1220648-2026-04140 | 24459726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24459726&pc=OZD | 02/17/2026 |  |  | 02/26/2026 |  |  | 02/17/2026 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1075485 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Break | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: biomed contacted the team to report that the aic purge door was broken, and the unit was being shipped back to headquarters for evaluation. the device was replaced, and support was continued. the reported events include product damage, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
69 | 1220648-2026-04141 | 24459941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24459941&pc=OZD | 09/26/2025 |  |  | 02/26/2026 |  |  | 09/26/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026653454 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
70 | 1220648-2026-04143 | 24460643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24460643&pc=OZD | 09/26/2025 |  |  | 02/26/2026 |  |  | 09/26/2025 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026727007 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 52-year-old male presented with acute on chronic decompensated heart failure on (b)(6) and was placed on vaecmo and impella cp for lv unloading. on (b)(6) there was noted pump malrotation towards mitral valve (impella cp) and attempts at impella cp repositioning were made. implantation of impella 5.5, second pump, on (b)(6) and removal of impella cp occurred and blood transfusion from explant of impella cp (left groin site). on (b)(6) ecmo decannulation and the impella 5.5 was repositioned secondary to ventricular tachycardia. patient experienced vt with hypotension, requiring defibrillation once. added lidocaine, levophed and vaso overnight. patient also had rp bleed. positioning issues were encountered and impella 5.5 was exchanged for a new on (b)(6). on (b)(6) impella 5.5 explanted without incident. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
71 | 1220648-2026-04144 | 24460697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24460697&pc=OZD | 02/20/2026 | 02/26/2026 |  | 02/26/2026 |  |  | 02/20/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1484655 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Connection Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
72 | 1220648-2026-04147 | 24460876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24460876&pc=OZD | 09/30/2025 |  |  | 02/26/2026 |  |  | 09/30/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026733800 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
73 | 1220648-2026-04148 | 24460879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24460879&pc=OZD | 02/16/2026 | 02/26/2026 |  | 02/26/2026 |  |  | 02/17/2026 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1386344 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 58 YR |  | Male | 165 KG | Hispanic | White | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | An impella cp was placed via the femoral artery to support the 58 year old male patient who had been admitted in with acute myocardial infarction and cardiogenic shock, presenting in scai stage d shock. the underlying medical history was not shared. the pump was being used in combination with an automated impella controller (aic) when a yellow alarm took place which prompted the team to replace the aic. the alarm persisted on the backup aic, but the pump was not explanted nor replaced. support continued til wean and explant after 2 days of support. the malfunction of the pump-aic will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange, however the exchange was performed with no consequence to the patient and support. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
74 | 1220648-2026-04149 | 24460973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24460973&pc=OZD | 02/17/2026 | 02/26/2026 |  | 02/26/2026 |  |  | 02/17/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795577 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
75 | 1220648-2026-04152 | 24461076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24461076&pc=OZD | 09/30/2025 |  |  | 02/26/2026 |  |  | 09/30/2025 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026712269 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female | 134 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 38-year-old female with history of nonischemic cardiomyopathy, type 2 diabetes, acute kidney injury on crrt had an impella 5.5 placement as a bridge to transplant or lvad. on (b)(6) placement of central rvad for severe rv dysfunction on (b)(6) the pa catheter was removed. on (b)(6) there was a placement signal alarm which was believed to have caused erroneous suction alarms. on (b)(6) there were positive blood cultures requiring antibiotics were noted which were attributed to line infection. on (b)(6) the patient was transplanted with explant of imp5.5 and rvad without incident. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
76 | 1220648-2026-04153 | 24461097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24461097&pc=OZD | 10/02/2025 |  |  | 02/26/2026 |  |  | 10/02/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025563086 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Heart Block | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
77 | 1220648-2026-04155 | 24461209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24461209&pc=OZD | 02/13/2026 |  |  | 02/26/2026 |  |  | 02/06/2026 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026812362 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Arrhythmia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 was surgically implanted via the right axillary/subclavian artery in a 64-year-old male patient with cardiomyopathy presenting in scai stage d shock. during support, the patient initially experienced ventricular tachycardia (vt), which was attributed to temporary pump malposition; bedside echocardiography confirmed that the impella was lodged against the papillary muscle and oriented toward the mitral valve. following repositioning, the arrhythmia resolved, pump flows normalized, and the device continued to function appropriately with no further recurrence of vt. in addition, a purge leak was noted coming from the red plug. the local team was alerted and dispatched to support onsite. a data download and full product return were requested, and a replacement device was requested. at the time of reporting, it was unknown whether the device was removed due to this event. no additional contributing factors beyond the observed leakage were identified. based on the available information, the earlier arrhythmia was consistent with known risks of impella pump migration and resolved fully with repositioning, indicating no underlying device malfunction. the later purge system leak represents an isolated mechanical interface issue of unclear origin pending device evaluation. there is currently no evidence to suggest a systemic device performance failure. final assessment will remain pending full product return and analysis. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees, that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
78 | 1220648-2026-04156 | 24461236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24461236&pc=OZD | 10/01/2025 |  |  | 02/26/2026 |  |  | 10/01/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026731184 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
79 | 1220648-2026-04157 | 24461262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24461262&pc=OZD | 02/21/2026 | 02/26/2026 |  | 02/26/2026 |  |  | 02/21/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026681846 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
80 | 1220648-2026-04158 | 24461363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24461363&pc=OZD | 02/22/2026 | 02/26/2026 |  | 02/26/2026 |  |  | 02/22/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026814936 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
81 | 1220648-2026-04161 | 24461616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24461616&pc=OZD | 10/03/2025 |  |  | 02/26/2026 |  |  | 10/03/2025 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026733474 |  | 000000000001000100 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Priming Problem | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was inserted via surgical cutdown of the right axillary artery in a 72-year-old female who presented with cardiomyopathy, heart failure, and cardiogenic shock classified as society for cardiovascular angiography and interventions stage e. the patient was receiving inotropic support, vasopressors, and mechanical ventilation for respiratory support at the time of implantation. the purge solution consisted of dextrose 5 percent in water with 25 milliequivalents of sodium bicarbonate. during support, the system generated purge pressure high alarms and priming-related issues associated with the purge cassette. hemodynamic instability was noted in the setting of advanced cardiogenic shock; however, no specific signs, symptoms, or direct patient harm were reported in association with the purge-related alerts. troubleshooting was performed per standard protocol. due to persistent purge pressure high alarms and inability to resolve the priming issue, the device was revised and replaced. purge pressure alarms and priming-related issues are known potential occurrences in mechanical circulatory support systems and may be associated with purge cassette function, driveline factors, purge fluid dynamics, or system setup variables. these conditions are described in the instructions for use and are managed through established troubleshooting pathways, including cassette evaluation and device exchange when indicated. based on the information provided, there were no reported clinical sequelae directly attributed to the purge-related condition. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
82 | 1220648-2026-04162 | 24461632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24461632&pc=OZD | 10/02/2025 |  |  | 02/26/2026 |  |  | 10/02/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026704460 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Inaccurate Flow Rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
83 | 1220648-2026-04163 | 24461890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24461890&pc=OZD | 10/06/2025 |  |  | 02/26/2026 |  |  | 10/06/2025 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1420348 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 70 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | An impella cp device was inserted into the right femoral artery in a (b)(6) male patient with a past medical history of coronary artery disease (cad), cerebrovascular accident (cva), hypertension, prostate cancer, dialysis, chronic obstructive pulmonary disease (copd)/emphysema, and deaf. the patient was presenting in scai stage b shock, prior to initiation of support. the impella was inserted for ventricular support for a protected percutaneous coronary intervention (pci). during the procedure, the impella was inserted with resistance, so shockwave treatment was performed to improve the passage of the device. during the pci, the patient developed frothy sputum and coffee ground emesis. the patient was intubated for respiratory support. heparin and integrillin were stopped, and protamine 50mg (anti-coagulation reversal) was given. the patient received 2 units of packed red blood cells (prbcs). the activated clotting time (act) was 262. the impella was left in place and the automated impella controller (aic) had an impella in aorta alarm. the impella was advanced, and when in the left ventricle (lv), the placement signal (ps) was lost, and a controller failure alarm displayed on the aic. the ps signal reappeared, but the alarm was off. the decision was made to swap aic consoles. the following day, the impella was successfully weaned. the post-procedure outcome is survived at explant. the automated impella controller will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange however the exchange was performed with no consequence to the patient and support. high-risk interventions require intense blood thinners, increasing the risk of active bleeding. bleeding is also a known risk due to the impella's anticoagulation and purge requirements. bleeding is listed as a potential adverse event, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
84 | 1220648-2026-04164 | 24462337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24462337&pc=OZD | 12/23/2025 |  |  | 02/26/2026 |  |  | 12/23/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1109485 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
85 | 1220648-2026-04167 | 24463004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24463004&pc=PYX | 12/03/2025 |  |  | 02/26/2026 |  |  | 12/03/2025 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PUMP, RP-IJ, STERILE |  |  | PYX |  |  | IMPELLA PUMP, RP-IJ, STERILE |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1901670 |  | 1000186 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Anemia | Hemolysis | Renal Failure | Sepsis | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission
86 | 1220648-2026-04170 | 24463601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24463601&pc=OZD | 02/17/2026 | 02/26/2026 |  | 02/26/2026 |  |  | 02/17/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026812479 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
87 | 1220648-2026-04175 | 24464693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24464693&pc=OZD | 02/17/2026 | 02/26/2026 |  | 02/26/2026 |  |  | 02/17/2026 | 02/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798737 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
88 | 1220648-2026-04178 | 24465010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24465010&pc=OZD | 02/17/2026 |  |  | 02/26/2026 |  |  | 02/17/2026 | 02/26/2026 |  |  |  |  | 02/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775835 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Material Split, Cut or Torn | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was inserted via right axillary artery in a (b)(6) male who was admitted for acute myocardial infarction with cardiogenic shock. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage c. the patient received support from the 5.5 while in the intensive care unit. on day 28 of support, the patient became septic. a tear in the sterile sheath cover was observed, cleaned with chlorhexidine gluconate, and an occlusive tegaderm applied. the 5.5 was not considered to be a contributing factor to the sepsis as the patient had already had an increase in white blood cell count. that same day, there was an error with the automated impella controller during purge change, which resolved without intervention. this event is conservatively reported as the sleeve tear prompted intervention, but there was no lasting harm or injury reported and no consequence to support. the controller error will be conservatively reported for hemodynamic instability during purge change, but there was no intervention needed and no lasting harm or injury reported. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate. a4 is unknown.
89 | 1220648-2026-03968 | 24441899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24441899&pc=OZD | 01/29/2026 |  |  | 02/25/2026 |  |  | 01/30/2026 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026808770 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a purge system blockage. no kinks were identified. five percent dextrose with bicarb was added to the purge. no patient harm was reported. | [5 empty] | A4 is unknown. this report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc or its employees that the report constitutes an admission that the product, abiomed inc, or its employees, caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
90 | 1220648-2026-03969 | 24441906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24441906&pc=OZD | 02/14/2026 |  |  | 02/25/2026 |  |  | 02/14/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026807611 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
91 | 1220648-2026-03975 | 24442058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24442058&pc=OZD | 02/03/2026 |  |  | 02/25/2026 |  |  | 02/03/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026776003 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
92 | 1220648-2026-03977 | 24442131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24442131&pc=OZD | 01/30/2026 |  |  | 02/25/2026 |  |  | 01/30/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1249361 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Unintended Application Program Shut Down | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
93 | 1220648-2026-03978 | 24442158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24442158&pc=OZD | 02/16/2026 |  |  | 02/25/2026 |  |  | 02/16/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1814730 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
94 | 1220648-2026-03982 | 24442256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24442256&pc=OZD | 02/17/2026 |  |  | 02/25/2026 |  |  | 02/17/2026 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1126217 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: after switching on, the aic displayed a ¿controller error¿ alarm. the aic was replaced, and support continued. the reported events include controller failure, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
95 | 1220648-2026-03990 | 24442453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24442453&pc=OZD | 02/14/2026 |  |  | 02/25/2026 |  |  | 02/15/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026803437 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
96 | 1220648-2026-03991 | 24442507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24442507&pc=OZD | 02/16/2026 |  |  | 02/25/2026 |  |  | 02/16/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1739661 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
97 | 1220648-2026-03999 | 24443513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24443513&pc=OZD | 02/17/2026 |  |  | 02/25/2026 |  |  | 02/17/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781108 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
98 | 1220648-2026-04000 | 24443712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24443712&pc=OZD | 10/02/2025 |  |  | 02/25/2026 |  |  | 10/02/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713964 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
99 | 1220648-2026-04002 | 24444069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24444069&pc=OZD | 12/02/2025 |  |  | 02/25/2026 |  |  | 12/02/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026710590 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
100 | 1220648-2026-04005 | 24444366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24444366&pc=OZD | 10/22/2025 |  |  | 02/25/2026 |  |  | 10/22/2025 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026684341 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 111 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Renal Failure | Pericardial Effusion | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 64 year old male presented with chest discomfort and was diagnosed with coronary artery disease and aortic valve stenosis with an ejection fraction of 10% and alcohol abuse with liver cirrhosis. during cardiac surgery, an impella 5.5 was electively placed in addition to an extracorporeal membrane oxygenation. the patient also required continuous renal replacement therapy due to renal failure. on the first day of support, a pericardiocentesis was required to drain a pericardial effusion. a controller error occurred and resolved without any documented action taken. systemic anticoagulation was withheld following thrombocytopenia and a suspected heparin-induced thrombocytopenia that was later ruled out. after 11 days of support, the patient developed atrial fibrillation that intermittently re-occurred. after an off-label prolonged support duration of 27 days, the patient was successfully weaned and the impella 5.5 was removed. aic - controller failure/error during impella 5.5 support (day 2), there was an issue with the automated impella console (aic) that triggered a controller error (yellow alarm). there was no documented action taken in response to this issue. there were no further mentions of this issue or any additional aic issues. 5.5 (source: 11/3 pt. update note) on day 13 of impella 5.5 support, it was determined that the device was in the wrong position. the abiomed representative documented that the impella was too deep. the pump was subsequently pulled back to reposition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
101 | 1220648-2026-04003 | 24444367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24444367&pc=OZD | 10/22/2025 |  |  | 02/25/2026 |  |  | 10/22/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1664272 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
102 | 1220648-2026-04006 | 24444389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24444389&pc=OZD | 02/17/2026 |  |  | 02/25/2026 |  |  | 02/17/2026 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1175897 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 88 KG | Non Hispanic | White | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 79-year-old female patient presenting for pre-operative placement, with a history of known coronary artery disease (cad) and diabetes mellitus (dm). during support, the customer reported a ¿controller error. switch to back-up controller¿ yellow alarm. the aic was exchanged, after which no further alarms were noted. the removed aic was isolated for further investigation. the reported events include controller failure, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
103 | 1220648-2026-04008 | 24444626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24444626&pc=OZD | 11/20/2025 |  |  | 02/25/2026 |  |  | 11/20/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
104 | 1220648-2026-04010 | 24444643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24444643&pc=OZD | 02/17/2026 |  |  | 02/25/2026 |  |  | 02/17/2026 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026814962 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 64 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via the left axillary artery in a 74-year-old male patient presenting with post-cardiotomy cardiogenic shock (pccs/lcos), with a history of diabetes mellitus (dm). during the procedure, the impella 5.5 could not be advanced due to patient anatomy. an 8 mm × 30 mm graft was utilized per physician preference. a v-18 wire was attempted without success. significant calcification was noted, and the impella 5.5 would not pass the anastomosis. the case was therefore aborted, and the patient remained supported on va-ecmo. the reported events include failure to advance, medical device removal, additional device required, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the device removal; however, the removal was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
105 | 1220648-2026-04012 | 24444764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24444764&pc=OZD | 01/24/2026 |  |  | 02/25/2026 |  |  | 01/27/2026 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1135506 |  | 0042-0000-EU |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted via the right femoral artery in a 64-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). during support, the automated impella controller (aic) required replacement. subsequently, the aic was exchanged for another console, and the exchange was performed with no harm to the patient. the reported events are controller failure, medical device removal, device revision or replacement, and hemodynamic instability. the aic will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was completed without any patient consequence, and hemodynamic support was promptly resumed. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
106 | 1220648-2026-04013 | 24444775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24444775&pc=OZD | 02/17/2026 |  |  | 02/25/2026 |  |  | 02/17/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2027816743 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
107 | 1220648-2026-04014 | 24444822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24444822&pc=OZD | 11/23/2025 |  |  | 02/25/2026 |  |  | 11/23/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026741761 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Fever | Hemolysis | Pulmonary Edema | Respiratory Failure | Thrombocytopenia | [32 empty] | Invalid/Incomplete |  |  |  | Initial submission
108 | 1220648-2026-04011 | 24444827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24444827&pc=OZD | 11/23/2025 |  |  | 02/25/2026 |  |  | 11/23/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1083048 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
109 | 1220648-2026-04015 | 24444893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24444893&pc=OZD | 10/03/2025 |  |  | 02/25/2026 |  |  | 10/03/2025 | 02/25/2026 |  |  | 02/28/2026 |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025463870 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 14 YR |  | Female | 95 KG | Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | Device Damaged by Another Device |  |  |  |  | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | 14 year old female iwth history of nonischemic cardiomyopaty, myocarditis, acute kidney injury and itp presented with acute decompensated heart failure. an impella cp was implanted and later escalated to impella 5.5. on (b)(6) pump stop/ restarted to p8/ patient stable. on (b)(6) it was noted through the sleeve that there was kinked tubing. on (b)(6) suction alarms (position/ volume) and on (b)(6) impella 5.5 was explanted without incident. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
110 | 1220648-2026-04019 | 24445264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445264&pc=OZD | 02/14/2026 |  |  | 02/25/2026 |  |  | 02/14/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1736295 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
111 | 1220648-2026-04020 | 24445272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445272&pc=OZD | 10/04/2025 |  |  | 02/25/2026 |  |  | 10/04/2025 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026653609 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | 55-year-old male with history of cad, diabetes and hyperlipedimia presented with ami-cs. on (b)(6) they underwent cp placement with plan for staged hrpci on (b)(6) escalated to imp5.5. on (b)(6) low flow 1l/min unresponsive to p-level increases. it is unclear from the notes it the device was replaced. on (b)(6) they underwent hrpci with impella 5.5 support and on (b)(6) they were successfully weaned and imell 5.5 explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
112 | 1220648-2026-04021 | 24445408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445408&pc=OZD | 02/18/2026 |  |  | 02/25/2026 |  |  | 02/18/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737071 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
113 | 1220648-2026-04022 | 24445443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445443&pc=OZD | 02/17/2026 |  |  | 02/25/2026 |  |  | 02/17/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026815772 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Fluid/Blood Leak | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
114 | 1220648-2026-04024 | 24445528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445528&pc=OZD | 10/07/2025 |  |  | 02/25/2026 |  |  | 10/07/2025 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026722345 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 136 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 47 year old male patient and a medical history of human immunodeficiency virus, a colonostomy, diabetes mellitus and heart failure with reduced ejection fraction of 15% was admitted after failing home therapy with milrinon. an impella 5.5 was placed. after three weeks, the colonostomy was reversed in the operation room. after one month of support, the patient developed an intestinal bleeding, systemic anticoagulation was temporarily halted and a colonoscopy revealed an ulcer that was not actively bleeding at the time. patient also received one unit of packed red blood cells. after 37 days, an "air in purge" alarm occurred that was resolved by an exchange of the purge cassette. after an off-label prolonged support duration of 41 days, the patient was successfully bridged to a permanent left ventricular assist device and the impella 5.5 was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
115 | 1220648-2026-04026 | 24445561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445561&pc=OZD | 02/14/2026 |  |  | 02/25/2026 |  |  | 02/14/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026711512 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | Device Sensing Problem | Material Split, Cut or Torn |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
116 | 1220648-2026-04029 | 24445638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445638&pc=OZD | 02/13/2026 |  |  | 02/25/2026 |  |  | 02/13/2026 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2027833728 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male | 71 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: a 63 year old male with an acute myocardial infarction and cardiogenic shock, pre support with scai shock stage d, was supported with an impella pump inserted via the right femoral artery using a percutaneous approach. pump 1 was implanted on (b)(6) 2026 at 6:05 am. the device initially functioned as intended; however, later in the day, the patient experienced deteriorating support characterized by low flow and suction alarms at p2. echocardiography confirmed adequate lv filling, and fluid resuscitation did not resolve the alarms. due to persistent low flow and suction events, a pump exchange was recommended. the exchange was performed quickly and successfully. the reported low-flow and suction alarms on pump 1 were associated with a small thrombus identified at the device inflow upon explant. the exchange to a new pump resolved the issue, and the patient experienced no adverse clinical outcome. the event did not result in patient harm, and the patient remained supported post exchange. further device analysis is pending product return. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
117 | 1220648-2026-04030 | 24445651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445651&pc=OZD | 10/03/2025 |  |  | 02/25/2026 |  |  | 10/03/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026734927 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
118 | 1220648-2026-04032 | 24445730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445730&pc=OZD | 10/02/2025 |  |  | 02/25/2026 |  |  | 10/02/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737493 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
119 | 1220648-2026-04033 | 24445746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445746&pc=OZD | 05/22/2024 |  |  | 02/25/2026 |  |  | 12/12/2025 | 02/25/2026 |  |  | 02/28/2026 |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025471111 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | Nerve Damage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 64-year-old male patient with advanced stage d heart failure and multiple comorbidities was admitted with worsening symptoms and significantly elevated filling pressures. despite aggressive medical therapy, he remained hemodynamically unstable, and the impella 5.5 device was implanted as a bridge to transplant. on the second day of support, the device migrated into the aorta and required surgical repositioning. the patient subsequently improved and was successfully bridged to transplantation on the sixteenth day. more than twenty months later, additional adverse event information was reported. the patient experienced a right brachial plexus injury following axillary impella insertion and removal, associated with an axillary hematoma. symptoms persisted over the following year despite a carpal tunnel release procedure. the hematoma, noted the day after the injury, improved with conservative management and resolved. additionally, it was noted that the impella had a few suction events during patient movement, which self-resolved. separate loqi complaint with three ae's around access site, limb ischemia. chronic issues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
120 | 1220648-2026-04034 | 24445828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445828&pc=OZD | 12/20/2025 |  |  | 02/25/2026 |  |  | 12/20/2025 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026749029 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A seventy-nine-year-old male patient presented with acute heart failure secondary to non-ischemic cardiomyopathy. placement of an impella 5.5 device via the axillary artery was planned; however, the surgeon was unable to advance the device into the ascending aorta due to the patient¿s anatomy. as a result, the procedure was aborted and no device was implanted. unable to implant due to anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
121 | 1220648-2026-04035 | 24445841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445841&pc=OZD | 02/16/2026 |  |  | 02/25/2026 |  |  | 02/17/2026 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026804483 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was placed via the axillary graft site to support the 63-year-old male patient who had been supported in with acute myocardial infarction and cardiogenic shock, presenting in scai stage e shock. the underlying medical history was not shared. the pump was supporting when after over 2 weeks of support the patient was observed to have a retroperitoneal bleed, at the t12/l1 intercostal and lumbar arteries, that prompted the team to treat the bleed with embolization and turning off the heparin. these measures were taken to achieve hemostasis. the pump remains on for support despite the harm. bleeding is a known risk associated with impella support as described in the instructions for use. the connection and definitive relationship of the axillary placed 5.5 and the retroperitoneal bleed was not shared. the bleed is conservatively being reported despite no allegation. | [5 empty] | A4. weight, a5. ethnicity, a6 race are unknown. this report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
122 | 1220648-2026-04037 | 24445981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24445981&pc=OZD | 02/18/2026 |  |  | 02/25/2026 |  |  | 02/18/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025561732 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
123 | 1220648-2026-04038 | 24446019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24446019&pc=OZD | 02/16/2026 |  |  | 02/25/2026 |  |  | 02/17/2026 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026812381 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female | 93 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during support, the impella 5.5 generated a ¿placement signal unreliable¿ alarm. motor current and pump flow were reported to be consistent with expectations for the performance level, and purge flow and purge pressure were also within the expected ranges. due to the unreliable placement signal, the placement signal alarm monitoring was disabled, and the medical team proceeded to manage the device by monitoring motor current, pump flow, and the purge system. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the patient was successfully weaned from the impella 5.5 during the transplant procedure and underwent a successful heart transplant. the patient survived the procedure. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
124 | 1220648-2026-04041 | 24446345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24446345&pc=OZD | 10/03/2025 |  |  | 02/25/2026 |  |  | 10/03/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026727006 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
125 | 1220648-2026-04042 | 24446427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24446427&pc=OZD | 10/29/2025 |  |  | 02/25/2026 |  |  | 10/29/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026739650 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
126 | 1220648-2026-04043 | 24446433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24446433&pc=OZD | 12/21/2025 |  |  | 02/25/2026 |  |  | 12/21/2025 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744007 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 119 KG | Non Hispanic |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 54-year-old male patient with a history of diabetes mellitus presented with acute heart failure and cardiogenic shock and was supported with the impella 5.5 device. prior to implantation, the patient was receiving one inotropic medication and was on mechanical ventilation. his medical history included a ventricular septal defect, which represents a relative contraindication for use of the impella device. following implantation, the patient experienced intermittent ventricular tachycardia over several days. on the sixth day, he developed torsades de pointes requiring direct-current cardioversion. in the subsequent days, the patient demonstrated gradual hemodynamic improvement. after a total of 27 days of support, the impella device was successfully weaned. on the 25th day of support, the patient developed high purge pressure alarms and a purge-system blockage. purge flow had been steadily decreasing, and no kinks were identified in the purge line. tissue plasminogen activator was administered in the purge system to resolve the issue. high purge pressure / tpa use. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
127 | 1220648-2026-04045 | 24446588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24446588&pc=OZD | 12/20/2025 |  |  | 02/25/2026 |  |  | 12/21/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025515137 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
128 | 1220648-2026-04046 | 24446622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24446622&pc=OZD | 09/25/2025 |  |  | 02/25/2026 |  |  | 09/25/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026672080 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Syncope/Fainting | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
129 | 1220648-2026-04047 | 24446687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24446687&pc=OZD | 09/26/2025 |  |  | 02/25/2026 |  |  | 09/26/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026666888 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | False Alarm | [6 empty] | Atrial Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
130 | 1220648-2026-04018 | 24447100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24447100&pc=OZD | 11/11/2025 | 02/25/2026 |  | 02/25/2026 |  |  | 11/11/2025 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026756831 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | Cardiac Arrest | Low Blood Pressure/ Hypotension | Renal Failure | Sepsis | Hypovolemia | Cardiogenic Shock | Respiratory Failure | Multiple Organ Dysfunction Syndrome | Heart Failure/Congestive Heart Failure |  |  | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | 68-year-old male presented with nstemi, received iabp in cath lab, then underwent impella 5.5 implantation in or. patient intubated pre-op: multiple blood products administered intraoperatively. post-op course included arrhythmias requiring epinephrine weaning, vf arrest on day 5 (resolved with defibrillation), renal failure requiring crrt, hypotension treated with fluids and methylene blue, and sepsis confirmed by positive blood cultures. impella repositioned for shallow placement; patient later reintubated for respiratory failure. clinical course complicated by cardiogenic and septic shock, liver failure, and persistent hypotension. impella removed without incident on (b)(6) 2025. most complications (shock, organ failure, sepsis) were likely related to underlying critical illness; arrhythmia and anemia possibly device-related but unconfirmed. no definitive adverse events attributed to impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. most complications (shock, organ failure, sepsis) were likely related to underlying critical illness; arrhythmia and anemia possibly device-related but unconfirmed. no definitive adverse events attributed to impella 5.5.
131 | 1220648-2026-04048 | 24447101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24447101&pc=OZD | 12/21/2025 |  |  | 02/25/2026 |  |  | 12/22/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026731520 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
132 | 1220648-2026-04049 | 24448540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24448540&pc=OZD | 02/17/2026 | 02/25/2026 |  | 02/25/2026 |  |  | 02/17/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026736338 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
133 | 1220648-2026-04051 | 24449064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24449064&pc=OZD | 02/16/2026 |  |  | 02/25/2026 |  |  | 02/16/2026 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744117 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
134 | 1220648-2026-04052 | 24449193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24449193&pc=OZD | 02/16/2026 |  |  | 02/25/2026 |  |  | 02/16/2026 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1136985 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 69 KG | Non Hispanic | White | Injury | Yes | No | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted via the right femoral artery in an 82-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). the patient had a history of coronary artery disease. during device setup, the aic displayed an alarm indicating that the optical sensor was not detected. the catheter was inspected, and light was confirmed through the plug-in area. after re-plugging the device, setup was completed successfully. the impella initially ran normally without issues. approximately 20 minutes later, a controller failure alarm occurred, during which the ao and lv waveforms disappeared, although motor current, flows, and purge parameters remained stable with no loss of support. the controller was exchanged from aic (b)(6) to aic (b)(6), after which the device operated normally with no alarms. there was no harm or loss of support during the alarm or during the controller exchange. the reported events are controller failure, medical device removal, device revision or replacement, and hemodynamic instability. the impella cp will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
135 | 1220648-2026-04055 | 24450015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24450015&pc=OZD | 02/16/2026 |  |  | 02/25/2026 |  |  | 02/16/2026 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1250417 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 72-year-old male patient presenting for elective placement. the patient had a history of coronary artery disease and diabetes mellitus. during cabg, a controller error alarm occurred on the console. the team switched to a new console per recommendations without complication. console ic 4489 was quarantined and staff were notified. the replacement console, ic 4490, functioned normally, and support was maintained. the reported events are controller failure, device revision or replacement, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
136 | 1220648-2026-04056 | 24450019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24450019&pc=OZD | 12/05/2025 |  |  | 02/25/2026 |  |  | 12/05/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1972400 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
137 | 1220648-2026-04057 | 24450043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24450043&pc=OZD | 11/22/2025 |  |  | 02/25/2026 |  |  | 11/24/2025 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765525 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical assessment by (b)(6). 60-year-old male for elective placement of impella 5.5 for high-risk surgery (coronary artery bypass surgery and valve replacement). it was noted that the suture line had dehiscence at the anastomosis on aorta. re-exploration in or, suture line was reinforced and 3 units of packed red blood cells given. ecmo required in addition to impella 5.5. impella 5.5 was successfully removed after support without incident. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
138 | 1220648-2026-01527 | 24450057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24450057&pc=OZD | 09/24/2025 |  |  | 02/25/2026 |  |  | 09/24/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1751872 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
139 | 1220648-2026-04058 | 24450077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24450077&pc=OZD | 02/15/2026 | 02/25/2026 |  | 02/25/2026 |  |  | 02/17/2026 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT, PKGD, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1680221 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Computer Operating System Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella cp was inserted via left femoral artery in a 85 year old male who was admitted for acute myocardial infarction with cardiogenic shock. the patient¿s comorbidities and clinical state prior to initiation of support were unknown. while starting up the automated impella controller (aic), it would not proceed beyond the self-check screen. a replacement controller was used. the patient received support from the cp for the percutaneous coronary intervention. on day 3 of support, the device was explanted. the startup issue will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange, however the exchange was performed with no consequence to the patient and support. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
140 | 1220648-2026-03318 | 24450172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24450172&pc=OZD | 02/07/2026 | 02/16/2026 |  | 02/25/2026 |  |  | 02/07/2026 | 02/25/2026 |  |  |  |  | 02/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026763813 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 95 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 73-year-old male with a rp flex and a 5.5 impella in place post surgical. rp flex began alarming that placement signal not reliable multiple times without the ability to clear the alarm. an active flow on monitor and pa waveform were still present. the patient remained on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
141 | 1220648-2026-04067 | 24450871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24450871&pc=OZD | 10/30/2025 |  |  | 02/25/2026 |  |  | 10/30/2025 | 02/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713401 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
142 | 1220648-2026-03875 | 24427825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24427825&pc=OZD | 02/14/2026 |  |  | 02/24/2026 |  |  | 02/14/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026711512 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
143 | 1220648-2026-03878 | 24427923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24427923&pc=OZD | 02/16/2026 |  |  | 02/24/2026 |  |  | 02/16/2026 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798100 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 55 KG |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was placed via the axillary graft access to support the 63 year old male who had been admitted in with cardiomyopathy, and presenting in scai stage d shock. the previously had been supported by a femorally placed impella cp captured in prior complaint (b)(4). the patient was known to be on inotropes and vasopressors prior to pump placement. on the 3rd day of the support the 5.5 access site was seen to have increased drainage from the drain. the surgical team was made aware and they have 25mg of protamine and transfused 2 units of blood cells. the heparin was additionally stopped. these measures have resolved the bleed and the pump remains on for support. anticoagulation necessary for the pump placement and support can contribute to access site bleeding. bleeding is a known risk associated with impella support as described in the instructions for use. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp.
144 | 1220648-2026-03882 | 24428103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24428103&pc=OZD | 02/12/2026 |  |  | 02/24/2026 |  |  | 02/12/2026 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026777399 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp was being placed to support the patient of unknown age and gender who had been admitted in with acute myocardial infarction and cardiogenic shock. the patient's underlying history and comorbidities were not shared. the patient had a va-ecmo, inotropes, and vasopressors ongoing prior to the cp pump being placed. the team was placing the hemostasis device perclose prior to the cp placement, for the "pre-closure", and dissected the artery. the delivery of the cp was aborted and the patient remained on va-ecmo alone. the intervention performed the remedy the harm was not shared. vascular injury is a known risk associated with impella support as described in the instructions for use (ifu), and the "preclosure" technique with a perclose device is noted in the ifu and can be used by experienced operators and may reduce bleeding and vascular complications. | [5 empty] | A2, a3, and a4 are unknown. this report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc or its employees that the report constitutes an admission that the product, abiomed inc, or its employees, caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
145 | 1220648-2026-03890 | 24428333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24428333&pc=OZD | 02/06/2026 |  |  | 02/24/2026 |  |  | 02/06/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801065 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
146 | 1220648-2026-03893 | 24428365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24428365&pc=OZD | 02/16/2026 |  |  | 02/24/2026 |  |  | 02/16/2026 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806226 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was being placed via the access at the left axillary graft site to support the 75 year old male patient with a known cardiac history. the patient was in for the electrophysiology procedure of an ablation. prior to this admission the patient had previously been on another impella, had an aorto-fem bypass, and was known to have peripheral arterial disease with calcification in the aorta. the team attempted the 5.5 placement using all best practices and guidewire exchange, but the pump could not be delivered as the team met resistance at the aortic arch. the team explanted the 5.5 and instead utilized the impella cp pump and new access site. the team did note the 5.5 was damaged/kinked in the attempted insertions. the instructions for use of the 5.5 pump does note the contraindication of severe arterial disease precluding placement of the impella system. | [5 empty] | This is one of three related reports. this report represents the impella 5.5 and other reports were submitted for the impella cp and automated impella controller. this report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
147 | 1220648-2026-03892 | 24428366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24428366&pc=OZD | 02/17/2026 |  |  | 02/24/2026 |  |  | 02/17/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025568322 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
148 | 1220648-2026-03898 | 24428474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24428474&pc=OZD | 02/15/2026 |  |  | 02/24/2026 |  |  | 02/15/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797919 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Restricted Flow rate | Malposition of Device | [5 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
149 | 1220648-2026-03907 | 24430164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24430164&pc=OZD | 02/13/2026 |  |  | 02/24/2026 |  |  | 02/15/2026 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794126 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male | 60 KG | Non Hispanic | Asian | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinician narrative an impella cp was inserted via the left femoral artery in a 57-year-old male who presented for elective placement in the setting of severe mitral regurgitation and high-risk candidacy for mitral valve replacement or intervention. the patient was classified as scai stage d cardiogenic shock and required inotropes, vasopressors, and ventilatory support for respiratory failure at the time of mechanical circulatory support initiation. while supported with the impella cp, the patient developed hemolysis with associated laboratory abnormalities and progressive renal failure, including decreased urine output. the provider reported that device position was confirmed at insertion and remained appropriate by imaging during management. support levels were described as high, and suction findings were noted in the setting of pump support. due to ongoing hemolysis and worsening renal function, the decision was made to remove the impella cp and escalate support. the patient was bridged to an impella 5.5 for continued mechanical circulatory support. following transition to the 5.5, no further complaints were reported. hemolysis parameters improved, renal status stabilized, and the patient remained supported on the impella 5.5 at the time of reporting. hemolysis and renal dysfunction are recognized complications in patients requiring high-level mechanical circulatory support, particularly in the setting of cardiogenic shock, vasopressor use, and critical illness. large-bore femoral access and elevated pump support levels may contribute to shear-related hemolysis and end-organ hypoperfusion in severely compromised patients. based on the information provided, the clinical course appears consistent with the patient¿s underlying hemodynamic instability and severity of illness. the patient survived and continues on 5.5 support. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
150 | 1220648-2026-03908 | 24430187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24430187&pc=OZD | 09/26/2025 |  |  | 02/24/2026 |  |  | 09/26/2025 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026693762 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Nerve Damage | Renal Failure | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 72-year-old female with - among others - a previous medical history of hypertonus, diabetes ii and factor v leiden mutation presented with chest pain and shortness of breath over multiple days. electrocardiography confirmed a myocardial infarction of the anterior wall in conjunction with a ventricular septal defect. an intra-aortic balloon pump was inserted and the patient was transferred to a higher care center for further treatment. patient was brought to the opeation room for placement of an impella 5.5. following removal of the intra-aortic balloon pump, the patient started to decline and a large hematoma to the previous access site was noted and controlled with manual pressure. following transfer to the intensive care unit, the patient developed repetitive episodes of atrial flutter and other arrhytmia including ventricular tachycardia requiring cardioversion and medication. the patient also suffered a gastro-intestinal bleeding requiring a surgical intervention and a blood transfusion as well as renal failure requiring continuous renal replacement therapy. a weakness to the right arm was noted and continuously treated with physiotherapy. after 14 days of support, a "controller error" alarm occurred on the impella controller, requiring a replacement controller. after 20 days of support, surigcal repair of the ventricular septal defect could be carried out. after an off-label prolonged duration of 32 days of support, the patient was successfully weaned and the impella 5.5 was removed. "aic - controller failure/error: during impella 5.5 support, there was an issue with the automated impella console (aic) that triggered a controller error (yellow alarm) on day 14 of support. the abiomed representative immediately instructed the hospital staff to perform an aic-to-aic transfer. there were no further aic issues during the patient's case with the new aic. 5.5 during impella 5.5 support, the patient experiences an access site adverse event. on day 13 of support ((b)(6) 2025), abiomed representative noted a very slow ooze at the access site which prompted a dressing change overnight with no active bleeding since then. as there was no device malfunction, the patient outcome code for this event will be dictated by the physician. " | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
151 | 1220648-2026-03909 | 24430205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24430205&pc=OZD | 02/18/2026 |  |  | 02/24/2026 |  |  | 02/18/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026814960 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
152 | 1220648-2026-03913 | 24430524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24430524&pc=OZD | 09/26/2025 |  |  | 02/24/2026 |  |  | 09/26/2025 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1172893 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | No | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 72-year-old female with - among others - a previous medical history of hypertonus, diabetes ii and factor v leiden mutation presented with chest pain and shortness of breath over multiple days. electrocardiography confirmed a myocardial infarction of the anterior wall in conjunction with a ventricular septal defect. an intra-aortic balloon pump was inserted and the patient was transferred to a higher care center for further treatment. patient was brought to the opeation room for placement of an impella 5.5. following removal of the intra-aortic balloon pump, the patient started to decline and a large hematoma to the previous access site was noted and controlled with manual pressure. following transfer to the intensive care unit, the patient developed repetitive episodes of atrial flutter and other arrhytmia including ventricular tachycardia requiring cardioversion and medication. the patient also suffered a gastro-intestinal bleeding requiring a surgical intervention and a blood transfusion as well as renal failure requiring continuous renal replacement therapy. a weakness to the right arm was noted and continuously treated with physiotherapy. after 14 days of support, a "controller error" alarm occurred on the impella controller, requiring a replacement controller. after 20 days of support, surigcal repair of the ventricular septal defect could be carried out. after an off-label prolonged duration of 32 days of support, the patient was successfully weaned and the impella 5.5 was removed. "aic - controller failure/error: during impella 5.5 support, there was an issue with the automated impella console (aic) that triggered a controller error (yellow alarm) on day 14 of support. the abiomed representative immediately instructed the hospital staff to perform an aic-to-aic transfer. there were no further aic issues during the patient's case with the new aic. 5.5: during impella 5.5 support, the patient experiences an access site adverse event. on day 13 of support ((b)(6) 2025), abiomed representative noted a very slow ooze at the access site which prompted a dressing change overnight with no active bleeding since then. as there was no device malfunction, the patient outcome code for this event will be dictated by the physician. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
153 | 1220648-2026-03914 | 24430736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24430736&pc=OZD | 10/02/2025 |  |  | 02/24/2026 |  |  | 10/02/2025 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1173720 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A 79-year-old female with a medical history of diabetes, a transient ischemic attack, chronic kidney disease, pulmonary embolism and coronary arter disease including a previous myocardial infarction leading to ischemic cardiomyopathy. she presented with concerns of a rapid heart rate. before intervention could be performed, the patient became hypotensive and hypoxic requiring intubation. in the catheterization laboratory, an impella cp was placed and a percutaneous coronary intervention carried out. as the patient continued to deteriorate, the procedure had to be aborted and the patient wass transferred to a higher care level. while rounding on the patient, the aortic placement signal disappeared, shortly after, a ¿controller error¿ alarm popped up and the controller had to be exchanged. soem slight oozing at the groin access was noted managed by a dressing change. the patient developed signs of sepsis, blood cultures later confirmed to be positive for yeast with a previously known yeast infection of the groins. the patient was started on antibiotic treatment as well as continuous renal replacement therapy due to renal failure. on the next day, the patient showed signs of gastro-intestinal bleeding that did not require intervention and decompensated during a device weaning trial, leading to supraventricular then ventricular tachycardia requiring defibrillation. postition of the impella pump was reconfirmed with both echocardiography and fluoroscopy without need for repositioning. over the next days, the patient developed a thrombocytopenia and systemic heparin was halted while heparin induced thrombocytopenia could be ruled out. after 5 days of support, the patient was successfully weaned and the impella cp removed. complications will be coded to the impella cp, as - while presumably heavily related to pre-existing disease (yeast infection of the groin, renal failure based on chronic kidney disease, procedural application of contrast agent and shock-induced organ hypoperfusion) - an association to the device cannot be fully ruled out. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
154 | 1220648-2026-03918 | 24430897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24430897&pc=OZD | 02/13/2026 |  |  | 02/24/2026 |  |  | 02/16/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026771014 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | Hemodynamic instability | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
155 | 1220648-2026-03919 | 24430937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24430937&pc=OZD | 10/03/2025 |  |  | 02/24/2026 |  |  | 10/03/2025 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026721277 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 110 KG |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 54 year old male with a history of hypertension, diabetes, sleep apnea, obesity, bipolar, anxiety, asthma, and heart failure with a residual ejection fraction of 35% presented with worsening shortness of breath and lower extremity edema. while primarily placed on an intra-aortic balloon pump, the decision was made to escalate to an impella 5.5. on the intensive care unit, the patient suffered multiple episodes of ventricular tachycardia requiring medication and device repositioning. after 16 days of support, a controller error alarm occurred and the console was exchanged without documented consequences to the patient. after an off-label prolonged 46 days of support, the patient was successfully bridged to a permanent left ventricular assist device. aic - controller failure/error and false alarm. during impella 5.5 support, there was an issue with the automated impella console (aic) that triggered a controller error (yellow alarm) on day 38 of support. the abiomed representative identified this alarm occurred while reviewing console history after the patients' cases had ended. the hospital staff did not reach out to impella team about the alarm when it occured and no mentions of further issues with the aic were noted in patient update notes. this indicates that it was a false alarm, as it resolved itself with no action taken and did not persist beyond its single occurence. as a result, there was no patient consequence caused by this event. 5.5 - device in wrong position: additionally, the impella 5.5 was determined to be in the wrong position via an echo approximately 7 hours post-implant. the device was too deep and subsequently repositioned by pulling the pump back. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
156 | 1220648-2026-03920 | 24431012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24431012&pc=OZD | 02/16/2026 |  |  | 02/24/2026 |  |  | 02/16/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798750 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
157 | 1220648-2026-03922 | 24431181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24431181&pc=OZD | 10/24/2025 |  |  | 02/24/2026 |  |  | 10/24/2025 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1778242 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
158 | 1220648-2026-03923 | 24431207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24431207&pc=OZD | 10/24/2025 |  |  | 02/24/2026 |  |  | 10/24/2025 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026741783 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 91 KG |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a brief ¿pump defective¿ alarm occurred in the operating room and the issue self-resolved. the automated impella controller subsequently displayed a ¿pump defective¿ alarm and a controller failure alert. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
159 | 1220648-2026-03926 | 24431406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24431406&pc=OZD | 02/15/2026 |  |  | 02/24/2026 |  |  | 02/18/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1420740 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
160 | 1220648-2026-03927 | 24431455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24431455&pc=OZD | 11/16/2025 |  |  | 02/24/2026 |  |  | 11/16/2025 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748987 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
161 | 1220648-2026-03928 | 24431471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24431471&pc=OZD | 11/16/2025 |  |  | 02/24/2026 |  |  | 11/16/2025 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1431034 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | No Audible Alarm | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
162 | 1220648-2026-03931 | 24431709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24431709&pc=OZD | 02/19/2026 |  |  | 02/24/2026 |  |  | 02/19/2026 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796215 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella 5.5 via right surgical axillary/subclavian artery for mechanical circulatory support. a high purge pressure alarm was noted on the impella connect. de-air was performed. reducing the dextrose concentration resolved the issue. the sidearm was cleaned and the sidearm retainer was used. motor currents were monitored for rises. the purge cassette was removed and reinserted into the console and resolved the issue. the procedure was successful. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
163 | 1220648-2026-03932 | 24431815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24431815&pc=OZD | 01/28/2026 |  |  | 02/24/2026 |  |  | 01/28/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026807241 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
164 | 1220648-2026-03840 | 24432287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24432287&pc=OZD | 02/16/2026 |  |  | 02/24/2026 |  |  | 02/16/2026 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1126568 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinician narrative an impella 5.5 was inserted via the right axillary/subclavian artery in a 75-year-old male who presented for electrophysiology ablation. the patient had a prior history of coronary artery bypass grafting, known coronary artery disease, and diabetes mellitus. the patient developed cardiogenic shock (scai stage e) requiring inotropic support, vasopressor therapy, and mechanical ventilation for respiratory support. during support, repeated controller error alarms were reported. due to the frequency of the alarms and concern from the medical team regarding hemodynamic instability and console reliability, the decision was made to exchange the controller console. the controller was removed from service and replaced. the instructions for use provide guidance for standard troubleshooting in the setting of controller alarms, including assessment of connections, power sources, and system checks, and recommend console exchange if alarms persist or reliability is in question. the patient survived. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate. device status. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. related medical device reports: this automated impella controller device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device reports on the impella cp and impella 5.5.
165 | 1220648-2026-03933 | 24432671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24432671&pc=OZD | 02/16/2026 |  |  | 02/24/2026 |  |  | 02/16/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737008 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
166 | 1220648-2026-03934 | 24433063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24433063&pc=OZD | 02/16/2026 |  |  | 02/24/2026 |  |  | 02/16/2026 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806121 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 58-year-old male patient presenting with cardiomyopathy, scai shock stage d. the patient had a history of diabetes mellitus. bleeding was noted from the axillary graft with a hematoma forming. the patient was taken to the operating room for exploration, where the bleeding source was identified and repaired without additional issues. no harm was reported, and the patient remained on support. the reported hematoma and major bleed requiring surgical intervention is consistent with access site complications and the anticoagulation and purge requirements associated with impella support, which are known risks described in the instructions for use. it is unknown whether recommended best practices for access management were followed to mitigate the risk of bleeding. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
167 | 1220648-2026-03941 | 24433400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24433400&pc=OZD | 02/16/2026 |  |  | 02/24/2026 |  |  | 02/16/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2019074 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
168 | 1220648-2026-03943 | 24433549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24433549&pc=OZD | 02/14/2026 |  |  | 02/24/2026 |  |  | 02/14/2026 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026812452 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 76 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella cp was inserted via right femoral artery in a 49 year old male who was admitted for cardiomyopathy. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage d. on day 9 of support, the patient was bridged to impella 5.5 via right axillary artery. on day 10 of 5.5 support, a hematoma was observed at the 5.5. insertion site and heparin suspended. on day 14 of 5.5 support, the device was explanted. vascular injury is a known potential adverse event of the impella 5.5 per the instructions for use. | [5 empty] | This is one of two related reports. this report represents the impella 5.5 and another report was submitted to represent the impella cp. this report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
169 | 1220648-2026-03944 | 24433596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24433596&pc=OZD | 11/17/2025 |  |  | 02/24/2026 |  |  | 11/16/2025 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1691229 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller displayed a sensor failure alert, while the pump continued to operate without interruption. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
170 | 1220648-2026-03946 | 24433696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24433696&pc=OZD | 02/18/2026 |  |  | 02/24/2026 |  |  | 02/18/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802451 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
171 | 1220648-2026-03947 | 24433769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24433769&pc=OZD | 12/02/2025 |  |  | 02/24/2026 |  |  | 12/02/2025 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026733792 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
172 | 1220648-2026-03948 | 24433809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24433809&pc=OZD | 12/22/2025 |  |  | 02/24/2026 |  |  | 12/22/2025 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773628 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Inflammation | Renal Failure | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
173 | 1220648-2026-03951 | 24434185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24434185&pc=OZD | 11/17/2025 |  |  | 02/24/2026 |  |  | 11/17/2025 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1723047 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Priming Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | A 53-year-old female patient treated with impella 5.5 due to heart failure with reduced ejection fraction. patient has factor v deficiency and non-ischemic cardiomyopathy due to breast cancer and chemotherapy and radiation. reduced ejection fraction of 10 % discovered while patients was waiting at home on guideline directed medical therapy for orthotopic heart transplantation. patient having more heart failure symptoms and being hospitalized multiple times - patient admitted to intensive care unit and placed on intra-aortic balloon pump. plan was to bridge to heart transplant with impella 5.5. intra-aortic balloon pump weaned during impella insertion procedure. patient having a high rate of panel-reactive antibodies despite plasmapheresis - therefore transplantation option deemed low for this patient. more then four weeks after support initiation, placement signal issues happened during support and patient also experienced self-resolving suction alarms - support time was longer than recommended in instructions for use. no further information available on how this was addressed or of the consequences for the patient. following purge cassette change there was an alarm noted on purge failure and switch was made to a new automated impella controller - new controller with the previously used purge cassette resolved the issue. hcp noted, that previously purge pressure had been bouncing in values and shared this following the controller switch - no patient consequences were reported and limited information available. patient continued to have alarms for placement and followed to have placement signal unreliable alarms. during this patient stayed asymptomatic as well as the pump flows and motor current were unchanged. hcp gave more volume as low status assumed to be the reasoning for alarms. during support noted that 3 way fixation of cable was not in place and reduced to 2 way fixation - informed hcp about importance of fixating the cable to ensure optimal device performance. after more than six weeks patient successfully bridged to transplant - this is beyond the recommended time of use for the device. impella support was continued beyond the recommended 14 day duration; prolonged use can increase the risk of complications and therefore may have played a contributory role in this adverse event. aic - priming problem: on day 31 of impella 5.5 support, there was a priming problem with the automated impella controller (aic). after completing priming of a new purge cassette for a purge cassette change, the aic alarmed for purge system failure. staff performed an aic exchange and used this same new purge cassette to provide support without issue. the nurse noted that earlier that the purge pressure was oscillating between the 400s and 800s mmhg realm. the purge pressure and purge flow were stable after the aic replacement. this signals that the issue stemmed from the aic's purge driver/controls. 5.5 low or blocked pump flow: during impella 5.5 support, there was a recurrent issue with low or blocked pump flow due to suction. on day 24 of support, there were occasional suction events that were treated with albumin to increase fluid volume in the left ventricle. on day 41 of support, there were also occasional suction alarms that were improved by volume given to the patient due to auto diuresis. there are no further mentions of suction. 5.5 placement signal issue: once the patient reached 29 days of impella 5.5 support, there were multiple placement signal issues. on day 29, the placement signal metric (56/46) was significantly lower than the patient's blood pressure (88/67). on day 39, the aortic placement signal dropped significantly which was accompanied by sustained placement signal low alarms. the patient was asymptomatic with stable blood pressure and echo confirmed correct pump position. the ao signal was adjusted on the console which corrected all waveforms and resolved the issue. the next day, day 40 of support, the placement signal drifted again. the patient was once again asymptomatic with motor current and flows within normal limits. the alarm progressed to a placement signal not reliable (psnr) alarm, which was silenced due to stable patient metrics and no symptoms. the abiomed representative educated the staff on monitoring via the motor current metrics, as the aortic placement signal metrics were suspended due to active psnr. he with the currently suspended aortic placement signal metrics. there was no consequence to the patient, as continued monitoring remained feasible via staff education on motor current console metrics. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
174 | 1220648-2026-03950 | 24434201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24434201&pc=OZD | 11/17/2025 |  |  | 02/24/2026 |  |  | 11/17/2025 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026736476 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
175 | 1220648-2026-03952 | 24434266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24434266&pc=OZD | 11/22/2025 |  |  | 02/24/2026 |  |  | 11/22/2025 | 02/24/2026 |  |  |  |  | 02/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026735820 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female | 96 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Tachycardia | Hypovolemia | Thrombosis/Thrombus | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 50 year old female patient treated with impella 5.5 due postcardiotomy cargiogenic shock in the context of coronary artery bypass graft surgery. patient had a history of known coronary artery disease. on first day of support four units of packed red blood cells were given. patient noted a febrile. patient did experience a couple runs of ventricular tachycardia following trialysis catheter placement. impella position deemed very good, but left ventricle appears to be under filled therefore coded for hypovolemia. cyanokit given last and urine is purplish color - not coded as result expected. patient started on continuous renal replacement therapy. on the day impella was weaned fibrin clot noted at impella outlet and clotting around 9 fr impella catheter closely to the motor. hcp noted it was not surprising to see this as anticoagulation was held for the first week due to bleeding and the patient was coagulopathic. patient successfully weaned after 11 days of support. case conservatively coded for serious injury due to bleeding and complications of this very complex patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
176 | 1220648-2026-03960 | 24438421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24438421&pc=OZD | 02/09/2026 | 02/24/2026 |  | 02/24/2026 |  |  | 02/10/2026 | 02/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | ABIOMED, INC. |  |  | US |  | 2026777167 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
177 | 1220648-2026-03741 | 24412927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24412927&pc=OZD | 02/10/2026 |  |  | 02/23/2026 |  |  | 02/10/2026 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746216 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Pericardial Effusion | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
178 | 1220648-2026-03745 | 24413029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24413029&pc=OZD | 02/15/2026 |  |  | 02/23/2026 |  |  | 02/15/2026 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026677295 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
179 | 1220648-2026-03750 | 24413193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24413193&pc=OZD | 02/13/2026 |  |  | 02/23/2026 |  |  | 02/13/2026 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806170 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
180 | 1220648-2026-03751 | 24413291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24413291&pc=OZD | 02/10/2026 |  |  | 02/23/2026 |  |  | 02/13/2026 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713244 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted into the femoral artery in a patient presenting in cardiomyopathy. while the patient was in the intensive care unit (icu), limb ischemia was noted. it is unknown how the limb ischemia was treated. the post-procedure outcome is unknown. limb ischemia is listed as a potential adverse event, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
181 | 1220648-2026-03752 | 24413313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24413313&pc=OZD | 02/12/2026 |  |  | 02/23/2026 |  |  | 02/12/2026 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026809827 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
182 | 1220648-2026-03754 | 24413423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24413423&pc=OZD | 02/13/2026 |  |  | 02/23/2026 |  |  | 02/13/2026 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1684985 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 57 KG | Non Hispanic | White | Injury | Yes | Yes | Device Damaged by Another Device | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Clinical rationale: an impella 5.5 device was inserted via the right direct surgical approach in a 63-year-old female patient presenting with pccs. the patient¿s medical history included diabetes mellitus (dm) and renal insufficiency. while performing a purge change, the nurse was unable to advance the purge disc into position. a different purge cassette and disc were attempted without success. the aic was then swapped out, and a successful purge changeout was completed. the reported events are mechanical issue, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was successfully maintained. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
183 | 1220648-2026-03760 | 24413639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24413639&pc=OZD | 02/08/2026 |  |  | 02/23/2026 |  |  | 02/13/2026 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801529 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella cp insertion was attempted via right femoral artery in n a 70 year old male for acute myocardial infarction with cardiogenic shock. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was identified as scai shock stage c. during cp implant, prior to percutaneous coronary intervention, it was attempted to detach the cp in the left ventricle using 0.018 inch guidewire, which was unsuccessful. a motor was confirmed in the aortic arch and the device did not advance further. insertion of the cp was aborted, and extracorporeal membrane oxygenation and intra-aortic balloon pump were inserted instead. the failure to advance will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the implant attempt, however the attempt and removal were performed with no consequence to the patient and support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
184 | 1220648-2026-03762 | 24413655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24413655&pc=OZD | 09/27/2025 |  |  | 02/23/2026 |  |  | 09/27/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1281602 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 device was inserted into the right axillary/subclavian artery in a 54-year-old male patient with a past medical history of coronary artery disease (cad) and diabetes, presenting in an out-of-hospital cardiac arrest requiring cardiopulmonary resuscitation (cpr), in acute myocardial infarction (ami) and cardiogenic shock (cgs), and in scai stage d shock, on multiple inotropes and vasopressors, and on a ventilator for respiratory support, prior to initiation of support. the patient was transferred to the intensive care unit (icu) from an outside hospital (osh), where an automated impella controller (aic) to aic exchange occurred. the nurse noted that the aic screen had an ongoing yellow alarm; however, the impella pump was able to function without issues. the abiomed representative reviewed the previous aic and found one placement signal lost (psl) alarm, and multiple suction alarms, but no other critical alarms. since the pump was functioning without issues, the rep advised the nurse not to perform another aic to aic exchange due to the chance that the pump may stop working. the impella functioned at p-8 at 4.1l/min as intended. one month later, the impella was removed, and the patient was bridged to a left ventricular assist device (lvad). the post-procedure outcome is survived at explant. the automated impella controller will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
185 | 1220648-2026-03764 | 24413675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24413675&pc=OZD | 02/13/2026 |  |  | 02/23/2026 |  |  | 02/13/2026 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1245092 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Application Program Problem | Device Sensing Problem | [5 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
186 | 1220648-2026-03766 | 24413709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24413709&pc=OZD | 02/11/2026 |  |  | 02/23/2026 |  |  | 02/12/2026 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026700716 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted surgically into the right axillary artery in a 44-year-old male patient presenting with cardiomyopathy in scai stage d shock on inotropes, vasopressors, and ventilator support prior to initiation of support. the impella 5.5 was inserted for escalation of therapy. while the patient was in the coronary care unit (ccu), the patient began to experience incisional pain. a computed tomography (ct) scan showed hematoma at the impella insertion/cutdown site. the hematoma was treated with surgical evacuation and subsequently resolved. the impella 5.5 was explanted, and the patient was doing well. the post-procedure outcome was survived at explant. bleeding is a known risk due to the impella's anticoagulation and purge requirements. bleeding is listed as a potential adverse event and is consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
187 | 1220648-2026-03768 | 24413724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24413724&pc=OZD | 02/16/2026 |  |  | 02/23/2026 |  |  | 02/16/2026 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802595 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
188 | 1220648-2026-03777 | 24414131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24414131&pc=OZD | 02/12/2026 |  |  | 02/23/2026 |  |  | 02/13/2026 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2027816562 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that there was a alarm: purge system blocked. checked the line if kinked, did de-airing of the system, changed purge bottle and the purge cassette. motor current in normal ranges. after initiating heparin bolus, the problem resolved. the patient was noted in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
189 | 1220648-2026-03785 | 24414492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24414492&pc=OZD | 02/13/2026 |  |  | 02/23/2026 |  |  | 02/13/2026 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026756714 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
190 | 1220648-2026-03786 | 24414519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24414519&pc=OZD | 12/04/2025 |  |  | 02/23/2026 |  |  | 02/16/2026 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1099847 |  | 0042-0000-EU |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
191 | 1220648-2026-03798 | 24414788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24414788&pc=OZD | 02/13/2026 |  |  | 02/23/2026 |  |  | 02/13/2026 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806228 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was inserted via right axillary artery in a 64-year-old male who was admitted for acute myocardial infarction with cardiogenic shock. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage d. the patient received support from the 5.5 for coronary angioplasty with stent placement. on day 3 of support, the patient had acute kidney injury and a dialysis catheter was placed for continuous renal replacement therapy (crrt). renal failure is a known potential adverse event of the impella 5.5 per the instructions for use. | [5 empty] | A4 is unknown. this report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc or its employees that the report constitutes an admission that the product, abiomed inc, or its employees, caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
192 | 1220648-2026-03802 | 24415192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24415192&pc=OZD | 02/13/2026 |  |  | 02/23/2026 |  |  | 02/13/2026 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806227 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: a 71 year old male with postcardiotomy cardiogenic shock (pccs) and a pre support clinical state consistent with scai shock stage e underwent robotic mitral valve repair and asd closure on (b)(6) 2026, which was complicated by rescue sternotomy and lung injury, requiring initiation of va ecmo. an axillary graft was placed; however, concern arose for clot formation in the left atrium (la) and left ventricle (lv). after clot clearance, the axillary graft subsequently thrombosed, prompting the decision to implant an impella 5.5 (sn: (b)(6) via direct aortic graft. significant intraoperative transfusion support was required, including 21 units prbcs, 24 units ffp, 6 units cryoprecipitate, and 6 units platelets. during support, the local team reported suspected thrombus injection involving the impella 5.5, with concern for obstruction of the purge gaps and potential thrombus involvement of the outflow cage. device data review via impella connect identified a rapid purge flow decline from ~9 ml/hr to <4 ml/hr between approximately 05:56¿05:30 am. treating physicians were informed, and per clinical judgment, a tpa purge protocol was initiated following pharmacy guidelines. two tpa purge bags were administered, after which purge flow improved. at the time of reporting, the device remained in use, and the medical team elected not to exchange the pump. the patient remains critically ill on support. failure modes associated with this event include major bleeding, thrombosis, and high purge pressure. the pump remains implanted, and product return is not yet available for evaluation. a definitive root cause cannot be determined at this time; however, patient specific factors¿including coagulopathy, extensive surgical intervention, ecmo use, and high transfusion burden¿likely contributed to the suspected thrombotic event. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
193 | 1220648-2026-03803 | 24415264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24415264&pc=OZD | 11/22/2025 |  |  | 02/23/2026 |  |  | 11/22/2025 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026730016 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
194 | 1220648-2026-03805 | 24415405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24415405&pc=OZD | 10/29/2025 |  |  | 02/23/2026 |  |  | 10/29/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026637869 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female | 88 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Tachycardia | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | A 43-year-old female patient presented with de novo cardiomyopathy and cardiogenic shock, scai stage e (initial lactate 19.6 mmol/l, ef 10¿15%, bnp 22,693 pg/ml). the clinical team elected to implant an impella 5.5 for mechanical circulatory support. backup escalation to ecls was prepared but ultimately not required. the impella 5.5 provided effective support for a total of 24 days. during support, a brief episode of ventricular tachycardia occurred and was treated medically. on day 3, one of three blood culture samples returned positive, and antimicrobial therapy was initiated. on the same day, a gastrointestinal bleed required intervention and transfusion; the patient had pre-existing anemia (hemoglobin 9.2 g/dl). after approximately 10 days of support, the patient developed a platelet decline, and heparin-induced thrombocytopenia was suspected. anticoagulation therapy was adjusted, and transfusions were administered as needed. minor oozing at the access site and at the swan-ganz catheter site was managed with manual compression. suction events occurred and were resolved per ifu, using echocardiography-guided volume administration and at least one device repositioning. however, suction occurred multiple times. additionally, urine output was initially low but gradually improved; continuous renal replacement therapy was not required. the patient was successfully weaned from impella 5.5 support after 24 days. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
195 | 1220648-2026-03806 | 24415425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24415425&pc=OZD | 11/24/2025 |  |  | 02/23/2026 |  |  | 11/24/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026771203 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Patient already on ecmo and impella 5.5. patient to or for chest washout and tracheostomy. the patient was having suction alarms on impella with echo showing catheter too deep and against the left ventricular wall. unable to reposition catheter as flows dropped to zero, so patient taken to or for impella removal. patient remained on ecmo support. during impella 5.5 support, the patient experienced the device in wrong position and low or blocked pump flow. the patient experienced suction alarms and the pump was found to be deep and up against the wall with echo imaging. the pump flows were described as 1.1-2, but when repositioning the displayed flows dropped to zero. the p-level was lowered to p2 and the impella 5.5 was removed. the device in wrong position and low or blocked pump flow are considered reportable malfunctions. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
196 | 1220648-2026-03808 | 24415495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24415495&pc=OZD | 02/16/2026 |  |  | 02/23/2026 |  |  | 02/16/2026 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2027816102 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
197 | 1220648-2026-03807 | 24415504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24415504&pc=OZD | 11/21/2025 |  |  | 02/23/2026 |  |  | 11/24/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025612301 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | 68-year-old male patient with extensive past cardiac medical and surgical history. he was admitted to the hospital following a right heart cath and elevated filling pressures with low cardiac output. his left ventricle was grossly dilated on echo and he was taken to the or for placement of impella 5.5. he was supported until his heart transplant on (b)(6) 2025. it was reported that on explant of the patient¿s native heart, the patient had developed an ulcer on the aorta "more distal than usual due to the impella catheter¿. the ulcer was found after cross clamp & native heart removal. the surgeon had to move the cross clamp which caused damage of the intimal layer of the aorta, and the patient had to undergo circulatory arrest due to aortic anastomosis to the donor heart. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
198 | 1220648-2026-03810 | 24415592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24415592&pc=OZD | 02/09/2026 |  |  | 02/23/2026 |  |  | 02/10/2026 | 02/23/2026 |  |  | 02/28/2026 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025474858 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
199 | 1220648-2026-03811 | 24415613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24415613&pc=OZD | 10/29/2025 |  |  | 02/23/2026 |  |  | 10/29/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026747637 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 65-year-old female patient with acute myocardial infarction complicated by cardiogenic shock was supported with an impella cp. due to hemolysis (ci-47162) the clinical team planned to escalate support to an impella 5.5. during the procedure, angiography demonstrated an occlusion of the right axillary artery. the left axillary artery showed multiple stenoses that could not be crossed or dilated. because vascular access for impella 5.5 placement could not be established, escalation was not achieved, and the patient remained on impella cp for bridge to cabg. no additional device-related adverse events were reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
200 | 1220648-2026-03815 | 24415728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24415728&pc=OZD | 12/20/2025 |  |  | 02/23/2026 |  |  | 12/21/2025 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744948 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
201 | 1220648-2026-03816 | 24415745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24415745&pc=OZD | 02/16/2026 |  |  | 02/23/2026 |  |  | 02/16/2026 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1147694 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male | 122 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella pump was implanted in a patient via percutaneous left femoral artery for mechanical circulatory support. the staff had issues with purge disc detection. the console kept saying purge disc at detected. the purge cassette was removed and reinserted into the console and resolved the issue. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
202 | 1220648-2026-03819 | 24416018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24416018&pc=OZD | 02/16/2026 |  |  | 02/23/2026 |  |  | 02/16/2026 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025589304 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 83 KG | Non Hispanic |  | Injury | Yes | Yes | Failure to Advance | [6 empty] | Intraoperative Cardiac Injury | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was being placed via the direct surgical access for the 75 year old male patient. the patient was in the operating suite for a bypass/cabg and was found to be unable to come off the bypass despite trying to wean two times. the 5.5 was to be placed despite knowledge and discussion of native anatomic issues that would make the access difficult. wire exchanges and all measures/techniques and suggestions were taken, however the pump failed to deliver to the left ventricle. the team observed increasing bleeding at the graft site and the placement of the 5.5 was aborted. anticoagulation necessary for the pump placement and support can contribute to bleeding. if there was any treatment for the bleeding was not shared, nor was the patient's post surgical outcome. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
203 | 1220648-2026-03821 | 24416178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24416178&pc=OZD | 11/08/2025 |  |  | 02/23/2026 |  |  | 11/08/2025 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026631248 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Failure to Advance | [6 empty] | Cardiac Arrest | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
204 | 1220648-2026-03823 | 24416425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24416425&pc=OZD | 02/15/2026 |  |  | 02/23/2026 |  |  | 02/15/2026 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026626875 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 79 KG | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during use, the impella 5.5 was observed to be leaking a small amount of clear fluid from the purge unit near the purge filter. the leak appeared to self-resolve. the medical team was instructed to monitor for any further fluid accumulation or dripping, as well as to closely monitor motor current, purge flow, and purge pressures. no signs or symptoms of patient injury were reported, and no harm was associated with the event. the device was successfully explanted after weaning. the explant was not considered premature and was unrelated to the reported complication. at the time of explant, the patient survived. the device is reportedly available for return; however, as of the time of this report, the device has not yet been returned for evaluation. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
205 | 1220648-2026-03824 | 24416457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24416457&pc=OZD | 11/07/2025 |  |  | 02/23/2026 |  |  | 11/07/2025 | 02/23/2026 |  |  | 10/31/2025 |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2024415387 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hematoma | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | 64 year old femal patient treated with impella 5.5 due to heart failure cardiogenic shock and acute decompansation. patient has a history of valve surgeries prior to this and was treated as bridge to heart transplantation. left ventricular ejection fraction was assesed as 23 % prior to pump insertion. hcp noted the sterile sleeve of the impella pump in the proximal section looked as if it were cut off - this did not have an effect on the function of the device. patient developed a hematoma at the insertion site - hcp concerned about infection, but plan is to compress to keep under control. hematoma developed on support day 12 and the day before heart transplant became available. patient was successfully bridged to transplant. this case was conservatively coded to serious injury due to hematoma and infection. the occurrence of infection during impella support is most plausibly related to the patient¿s critical illness and the presence of multiple invasive devices rather than the impella itself. critically ill patients frequently require central lines, arterial access, mechanical ventilation, urinary catheters, and prolonged icu care, all of which increase the risk of nosocomial infection. impaired immune function due to shock and systemic inflammation further predisposes to infectious complications. therefore, infection in this context is most likely multifactorial and primarily driven by the patient¿s underlying condition and the invasive intensive care unit (icu) environment rather than an intrinsic device-related issue. the rn reported that the proximal section of the anticontamination sleeve "looked as if it were cut off". this damage had no impact to impella function or treatment and therefore had no patient consequence. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
206 | 1220648-2026-03831 | 24417561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24417561&pc=OZD | 12/20/2025 |  |  | 02/23/2026 |  |  | 12/20/2025 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1683279 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
207 | 1220648-2026-03833 | 24417785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24417785&pc=OZD | 12/21/2025 |  |  | 02/23/2026 |  |  | 12/21/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775487 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 93 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Inaccurate Flow Rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 48-year-old male with a history of coronary artery disease, diabetes mellitus, renal insufficiency, and dialysis received an impella 5.5 for acute heart failure due to ischemic cardiomyopathy while supported with two inotropic medications, an intra-aortic balloon pump, and mechanical ventilation. the patient was positive for heparin-induced thrombocytopenia and was treated with argatroban, with an activated clotting time of 193 seconds at the time of insertion. the first impella 5.5 encountered difficulty crossing the aortic valve and later generated purge system blockage alarms, leading to removal and replacement. the second device was more difficult to insert due to repeated loss of wire position but was ultimately placed successfully with normalized activated clotting time. initial issues with left ventricular and aortic pressure waveforms and pump performance resolved by the end of the procedure. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
208 | 1220648-2026-03834 | 24417986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24417986&pc=OZD | 02/15/2026 |  |  | 02/23/2026 |  |  | 02/16/2026 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | OZD | [5 empty] | ABIOMED EUROPE GMBH - 3007904765 |  |  | GM |  | 1398838 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
209 | 1220648-2026-03836 | 24418719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24418719&pc=OZD | 12/23/2025 |  |  | 02/23/2026 |  |  | 12/23/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026743300 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A sixty-three year old male patient was treated for de-novo cardiomyopathy. an impella 5.5 was placed. on the intensive care unit, the patient developed signs of hemolysis. position of the impella was deemed deep in the ventricle but kept. after 7 days of support, the patient was successfully bridged to a left ventricular assist device. the impella 5.5 will be coded for hemoylsis and serious injury to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
210 | 1220648-2026-03839 | 24419887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24419887&pc=OZD | 11/29/2025 |  |  | 02/23/2026 |  |  | 11/29/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1123583 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 103 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A 50-year-old male patient treated with impella 5.5 due to acute decompensated cardiomyopathy. patient being previously supported by impella cp and then switched to impella 5.5 support. after insertion patient not on systemic anticoagulation due to oozing around right axillary site. one unit of packed red blood cells given overnight and additional unit over night. day one of impella 5.5 support purge system blocked alarm and tissue plasminogen activator being used as purge fluid. renal replacement therapy initiated and patient also showing signs of pancreatitis therefore coded for inflammation. after 2.5 weeks controller error alarm occurred, displaying a motor current and flow, but no placement signal or lv waveform. the controller error alarm lasted 3 minutes and then resolved. hcps did not want to make an automated impella controller switched even though advised to do so. later that same day information was shared that alarm happened two more times and controller switch was initiated. when the pump was plugged into switched controller, the pump placement signal was lost and alarm on controller showed placement signal not reliable yellow alarm. the patient's hemodynamics were unchanged. patient experiencing suction alarms and impella measuring 6cm from aortic valve, which is lower than recommended and hcp planning to reposition. after 1.5 months of support pump alarmed that the purge disc was not detected. change of purge cassette was not possible and automated impella controller hung up on insert purge disc screen. with onsite support purge cassette could not be changed as it was not detected by controller. controller change not recommended as previous placement signal issue not reliable of pump is still present. after more than 1.5 months patient care was withdrawn and patient expired, therefore very conservatively coded on death. thrombosis during impella support is most likely due to the combined effects of critical-illness¿related hypercoagulability and device-related blood flow disturbances. death was conservatively coded as most likely due to underlying disease and not device-related. impella support was continued beyond the recommended 14 day duration; prolonged use can increase the risk of complications and therefore may have played a contributory role in this adverse event. 5.5 - purge pressure high complaint noted a purge pressure high alarm trigger on the console, indicate a potential blockage at the rotor gap. purge fluid was changed to tpa as a mitigation. aic 1 (sn: (b)(6) controller failure/error & aic 2 (sn: (b)(6) placement signal issue first controller showed a yellow controller error alarm. console was replaced with a back up. when on the back up console, a placement signal not relaible alarm triggered. purge disk not detected alarm was triggered on the console, 3 hours after a purge cassette change occurred. to mitigate this, a new cassette was used, but the controller was unable to detect the purge disc resulting in not being able to complete the purge cassette replacement procedure. site made decision to not replace the controller because of an active placement signal alarm. complaint notes patient hemodynamics were unchanged, and as a result, there was no patient consequence. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
211 | 1220648-2026-03837 | 24419916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24419916&pc=OZD | 11/29/2025 |  |  | 02/23/2026 |  |  | 11/29/2025 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026735831 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Inflammation | Renal Failure | Thrombosis/Thrombus | Hemodynamic instability | [33 empty] | Invalid/Incomplete |  |  |  | Initial submission
212 | 1220648-2026-03838 | 24419947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24419947&pc=OZD | 11/29/2025 |  |  | 02/23/2026 |  |  | 11/29/2025 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1183120 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
213 | 1220648-2026-03849 | 24421133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24421133&pc=OZD | 11/26/2025 |  |  | 02/23/2026 |  |  | 11/26/2025 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026691733 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
214 | 1220648-2026-03852 | 24421480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24421480&pc=OZD | 12/04/2025 |  |  | 02/23/2026 |  |  | 12/04/2025 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1542953 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
215 | 1220648-2026-03853 | 24421659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24421659&pc=OZD | 12/16/2025 |  |  | 02/23/2026 |  |  | 12/16/2025 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1548998 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
216 | 1220648-2026-03855 | 24423758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24423758&pc=OZD | 02/13/2026 |  |  | 02/23/2026 |  |  | 02/14/2026 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026721352 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [41 empty] | Initial submission
217 | 1220648-2026-03857 | 24424168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24424168&pc=OZD | 02/16/2026 |  |  | 02/23/2026 |  |  | 02/16/2026 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026689023 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
218 | 1220648-2026-03858 | 24424562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24424562&pc=OZD | 12/04/2025 |  |  | 02/23/2026 |  |  | 12/04/2025 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713057 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hypovolemia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
219 | 1220648-2026-03861 | 24424573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24424573&pc=OZD | 11/23/2025 |  |  | 02/23/2026 |  |  | 11/23/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737453 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 80 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Anemia | Unspecified Infection | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 60-year-old female who had a history of complex percutaneous coronary intervention had a follow up visit in clinic. she had not taken her post procedure medications and ultimately presented to the ed with st segment myocardial infarction and cardiac arrest. she was taken to the cath lab for occluded stents and coronary arteries. she underwent a high-risk percutaneous coronary intervention and was placed on an impella cp and transferred to a higher level of care hospital. upon arrival she was started on va ecmo. the icu team was then able to start weaning medications. the impella required repositioning after arrival to the 2nd hospital. the patient required platelet administration. she required continuous renal replacement therapy and was able to come off ecmo on (b)(6) 2025 and placed on the impella 5.5. she received blood and platelets due to anemia and thrombocytopenia. the patient had high purge pressures in which the team treated with tpa through the purge. the patient continues on impella 5.5 support as of on (b)(6) 2026 without complication. medication required for high purge pressures (tpa). it is difficult to determine whether the patient anemia and thrombocytopenia were secondary to ecmo or to impella cp or impella 5.5. during impella 5.5 support, the patient experienced high purge pressure. description stated that purge flow was decreasing gradually with increasing purge pressure. a csc call was made inquiring about the use of tissue plasminogen activator and it was later added to the purge to address the high purge pressure. the purge pressure high is considered a reportable malfunction. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
220 | 1220648-2026-03862 | 24424577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24424577&pc=OZD | 11/18/2025 |  |  | 02/23/2026 |  |  | 11/18/2025 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745728 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
221 | 1220648-2026-03863 | 24424580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24424580&pc=OZD | 12/04/2025 |  |  | 02/23/2026 |  |  | 12/04/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026771234 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Prior to impella 5.5 support, the pump was prepped for use by opening in the sterile field. the user clamped the patient cable to the sterile field prior to connection to the console. this clamp damaged the purge pressure line, leading to a high purge pressure (likely due to clamp squeezing purge line). upon release of clamp, the purge pressure dropped too low, indicating potential damage to the purge line in the patient cable. user attributed this to a device failure out of box and replaced the device with a new impella 5.5 pump prior to insertion. the patient outcome is a delay in receiving treatment, since additional time was required to troubleshoot the problem and open a new pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
222 | 1220648-2026-03864 | 24424582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24424582&pc=OZD | 12/04/2025 |  |  | 02/23/2026 |  |  | 12/04/2025 | 02/23/2026 |  |  | 02/28/2026 |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025479896 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 58 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A sixty-nine-year-old male patient on multiple inotropic medications and mechanical ventilation presented in stage e of the society for cardiovascular angiography and interventions shock classification and received an impella 5.5 device for acute myocardial infarction cardiogenic shock. one day after implantation, the patient developed an axillary hematoma requiring operative intervention, after which his condition initially improved. on the sixteenth day of impella support, a purge system blocked alarm on the automated impella controller occurred while the patient¿s hemodynamics remained stable. systemic heparin therapy had been restarted the prior day after being held earlier in the week following percutaneous coronary intervention. after consultation with the clinical support center, the care team began preparations to exchange the purge system but subsequently received alarms for purge pressure high and impella position in aorta. the patient was taken to the operating room for device exchange and during echocardiographic evaluation after device removal, the physician noted thrombus in the outlet cage of the impella 5.5 pump as well as thrombus in the left ventricle and aorta. after further monitoring, anesthesia identified a malposition and clotted dialysis catheter following catheter removal, transesophageal echocardiography revealed no thrombus in the left ventricle or aorta. the physician elected to proceed with implantation of a new impella 5.5 device. a vascular surgeon performed a new graft anastomosis and the device was inserted per instructions for use with stable waveforms and hemodynamics. after insertion, no flow was detected to the right arm, prompting revascularization of the subclavian artery during which multiple blood products were administered. the patient subsequently suffered cardiac arrest, and cardiopulmonary resuscitation was performed for greater than twenty minutes. the interventional cardiologist then discussed the situation with the family, who elected to withdraw care. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
223 | 1220648-2026-03865 | 24424584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24424584&pc=OZD | 12/05/2025 |  |  | 02/23/2026 |  |  | 12/05/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779269 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female |  | Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp supported patient was scheduled for escalation to an impella 5.5 device. during impella 5.5 insertion, user was unable to advance 5.5 pump across aortic valve. a clinical decision was made to abort the impella 5.5 insertion and remain on impella cp-sa support. because the patient was already supported with an impella cp device and the support was deemed adequate from subsequent actions, the only patient consequence was a delay in therapy, since the user intended to escalate support to a 5.5 (provide more support) but was unable to do so with the 5.5 device. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
224 | 1220648-2026-03868 | 24424590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24424590&pc=OZD | 12/18/2025 |  |  | 02/23/2026 |  |  | 12/18/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026678240 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a sixty five year old male patient with a medical history of coronary artery disease and type two diabetes mellitus received an impella 5.5 device for post cardiotomy cardiogenic shock following high risk coronary artery bypass surgery. prior to device placement, the patient was receiving one intravenous medication to support heart muscle contraction, consistent with society for cardiovascular angiography and interventions cardiogenic shock stage c. after insertion, the impella 5.5 device was providing approximately five liters per minute of flow at performance level eight, with the inlet oriented toward the mitral valve. approximately four hours after initiation of support, the patient developed dark brown urine, and the bedside nurse contacted the on call team. the device was weaned to performance level six, which the patient tolerated well, and the urine color improved significantly. the care team deferred plasma free hemoglobin testing and bedside echocardiography until the following morning. renal function remained stable with a creatinine of 1.1. echocardiography later confirmed the inlet positioned near the mitral valve leaflet, and attempts to reposition the device were unsuccessful; therefore, the device was maintained at performance level four. after four days of impella 5.5 support, the patient demonstrated recovery of native cardiac function, and the device was successfully removed. hemolysis in urine, combined with repositioning and lowering support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
225 | 1220648-2026-03870 | 24424617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24424617&pc=OZD | 12/20/2025 |  |  | 02/23/2026 |  |  | 12/20/2025 | 02/23/2026 |  |  |  |  | 02/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026632528 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 85 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a fifty seven year old male patient with a medical history of coronary artery disease received an impella 5.5 device electively for high risk cardiac surgery involving mitral and tricuspid valve procedures. the surgery was successfully completed under impella 5.5 support, and the patient was transferred to the surgical intensive care unit on impella support for postoperative management. following surgery, the patient received intravenous medications to support heart muscle contraction and mechanical ventilation as part of routine postoperative care. on the fifteenth day of impella support, the bedside nurse reported that an ¿unknown position¿ alarm occurred whenever the patient moved his arms, although impella flow remained stable at approximately 2.1 liters per minute at performance level three, with a mean arterial pressure of sixty-four millimeters of mercury. a stat echocardiogram was recommended to verify device position. the nurse later reported an impella depth measurement of seven centimeters. the medical team was advised to notify the physician for possible repositioning. according to the bedside nurse, the physician reviewed the images and clinical information and elected not to reposition the device as long as the patient remained clinically stable, with plans to reassess during rounds the following day. the patient remained stable on impella support for up to twenty-five days following insertion. no additional clinical updates were available in the record thereafter. position unknown alarm, no patient consequence | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
226 | 1220648-2026-03871 | 24424621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24424621&pc=OZD | 12/21/2025 |  |  | 02/23/2026 |  |  | 12/21/2025 | 02/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765457 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
227 | 1220648-2026-03736 | 24410609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24410609&pc=OZD | 02/14/2026 |  |  | 02/22/2026 |  |  | 02/14/2026 | 02/22/2026 |  |  |  |  | 02/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026809749 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  | White | Injury | Yes | No | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemodynamic instability | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in an 81-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs) with a history of coronary artery disease. the clinical team was unable to advance the impella 5.5 across the aortic valve due to possible stenosis or tortuosity. a second device was opened after the first catheter became kinked while attempting to advance it, resulting in inadequate pushability; however, the second device was also unable to be advanced. the reported events are product damage, failure to advance, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was completed without patient consequence. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees, that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
228 | 1220648-2026-03739 | 24410624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24410624&pc=OZD | 02/15/2026 |  |  | 02/22/2026 |  |  | 02/15/2026 | 02/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026790250 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | Company representation | Distributor | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
229 | 1220648-2026-03740 | 24410626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24410626&pc=OZD | 02/15/2026 |  |  | 02/22/2026 |  |  | 02/16/2026 | 02/22/2026 |  |  |  |  | 02/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026812881 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | Company representation | Distributor |  |  |  | 65 YR |  | Male | 99 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The 65-year-old male with acute myocardial infarction and cardiogenic shock requiring scai stage e mechanical circulatory support underwent sequential impella therapy beginning with an impella cp, followed by transition to an impella 5.5. the initial impella cp support was complicated by hematuria, prompting device explant and replacement. subsequent support with the impella 5.5 via axillary access was notable for surgical site bleed requiring surgical intervention; however, the event was determined not to be device-related. no device return, console data download, or replacement request is expected. the patient survived to explant on (b)(6) 2026. bleeding is a known risk per our ifu (instructions for use) because of our anticoagulation and purge requirements. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees, that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate
230 | 1220648-2026-03720 | 24410278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24410278&pc=OZD | 10/29/2025 |  |  | 02/21/2026 |  |  | 10/29/2025 | 02/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026729448 |  | 000000000001000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Respiratory Failure | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
231 | 1220648-2026-03721 | 24410290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24410290&pc=OZD | 10/29/2025 |  |  | 02/21/2026 |  |  | 10/29/2025 | 02/21/2026 |  |  |  |  | 02/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737038 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 68-year-old male presented with acute-on-chronic decompensation secondary to recurrent atrial fibrillation (scai d) and was initiated on ecls. subsequently, the clinical team decided to unload with an impella 5.5, wean ecmo, and bridge the patient to a durable lvad, which was implanted on day 21. following ecmo removal on day 3 of impella support, the patient required transfusions. during the course of impella 5.5 support, a left atrial appendage (laa) thrombus was identified. given the patient¿s history of atrial fibrillation prior to initiation of any mechanical circulatory support, a causal relationship to the impella device is considered highly unlikely, although the event was reported. the patient was successfully bridged to lvad implantation. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
232 | 1220648-2026-03723 | 24410308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24410308&pc=OZD | 10/29/2025 |  |  | 02/21/2026 |  |  | 10/29/2025 | 02/21/2026 |  |  |  |  | 02/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026735830 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 59-year-old male developed severe cardiogenic shock (scai e) following cabg and experienced cardiac arrest. veno-arterial ecls was initiated and maintained for 10 days. during this period, the patient developed acute kidney injury requiring crrt, as well as shock liver and ards. after 10 days of va-ecmo support, the clinical team decided to unload the left ventricle and facilitate ecmo weaning with impella 5.5 support. the impella 5.5 was implanted successfully. on day 3 of impella support, gastrointestinal bleeding occurred, and anticoagulation was reduced accordingly. after 8 days of support, the impella 5.5 was successfully weaned. no further information was available. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
233 | 1220648-2026-03724 | 24410326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24410326&pc=OZD | 10/30/2025 |  |  | 02/21/2026 |  |  | 10/30/2025 | 02/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026722324 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Device Sensing Problem | [6 empty] | Renal Failure | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
234 | 1220648-2026-03726 | 24410352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24410352&pc=OZD | 10/30/2025 |  |  | 02/21/2026 |  |  | 10/30/2025 | 02/21/2026 |  |  |  |  | 02/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713399 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Increase in Pressure | Pumping Stopped | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 55-year-old male with a two-week history of symptoms presented with delayed acute myocardial infarction complicated by cardiogenic shock (scai stage e). prior to the clinical decision to implant an impella 5.5 device, the patient experienced a cardiac arrest. extracorporeal life support (ecls) was initiated and the patient was stabilized. two days later, an impella 5.5 was implanted successfully, and the patient remained on support for approximately 13 days. during the initial implantation procedure, the first pump was unable to start properly. the device was exchanged for another impella 5.5 pump, which functioned as intended. the procedure was completed without further device-related complications. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
235 | 1220648-2026-03727 | 24410353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24410353&pc=OZD | 10/30/2025 |  |  | 02/21/2026 |  |  | 10/30/2025 | 02/21/2026 |  |  |  |  | 02/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026721349 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Tamponade | Hypovolemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 43-year-old male with morbid obesity was admitted with progressive shortness of breath. he was diagnosed with multivessel coronary artery disease and a left ventricular ejection fraction of 20%. the patient was scheduled for high-risk cabg ×3 with elective impella 5.5 implantation for perioperative hemodynamic support. postoperatively, the patient required two re-operations for cardiac tamponade and bleeding. although transfusions are likely in the context of repeated surgical interventions, no documentation regarding blood product administration was provided. two days after the final re-operation, the patient experienced fluid imbalance, which led to suction events and required volume resuscitation. the subsequent course of impella support was uneventful. after eight days of support, the patient was successfully weaned, with improvement in left ventricular ejection fraction. | [5 empty] | This report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate.
236 | 1220648-2026-03729 | 24410359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24410359&pc=OZD | 10/30/2025 |  |  | 02/21/2026 |  |  | 10/30/2025 | 02/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026663242 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
237 | 1220648-2026-03638 | 24399731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399731&pc=OZD | 11/20/2025 |  |  | 02/20/2026 |  |  | 11/20/2025 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026669818 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hypovolemia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
238 | 1220648-2026-03642 | 24399806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399806&pc=OZD | 11/21/2025 |  |  | 02/20/2026 |  |  | 11/21/2025 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766878 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Ischemia | Renal Failure | Cardiac Tamponade | Thrombocytopenia | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 76-year-old male patient developed post cardiotomy cardiogenic shock following a coronary artery bypass surgery procedure. va ecmo was initiated and an impella 5.5 was placed. the patient was transferred to a higher level of care. upon arrival, the impella position was unknown on aic with very flattened aortic waveform flowing 1.4l on p4. the patient returned to the cvor for mediastinal washout and mediastinal chest tube change out for tamponade. no active bleeding was noted. received multiple blood products. post operatively, the patient was started on continuous renal replacement therapy (crrt). ecmo was decannulated on (b)(6) 2025 with much improved pulsatility with the impella. a head ct showed a small subacute infarction with eeg showing encephalopathy immediately after ecmo decannulation. during transesophageal echo, the patient became hypotensive requiring additional vasopressor support. there was oozing around all invasive line sites so bivalirudin was discontinued. the patient was noted to have a right ij dvt which was remedied in interventional radiology. the patient went for endoscopy to rule out gi bleed with dropping hematocrit requiring blood transfusion and black, tarry stools. the impella was explanted successfully on (b)(6) 2025 without incident. it is difficult to attribute the adverse events to the impella pump as they could have been secondary to the post cardiotomy shock clinical scenario the patient was in after surgery. this would pertain to the renal failure, anemia, and thrombocytopenia. the small head infarct seems to be secondary to the ecmo decannulation as noted. most likely, the cardiac tamponade can be attributed to the surgery itself as it was noted that the chest tubes were clotted off. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
239 | 1220648-2026-03643 | 24399813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399813&pc=OZD | 11/22/2025 |  |  | 02/20/2026 |  |  | 11/22/2025 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765800 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Cardiogenic Shock | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 69 year old female was supported with an impella 5.5 device for postcardiotomy cardiogenic shock (pccs). her pre support clinical condition was consistent with scai shock stage d. the impella 5.5 was implanted on (b)(6) 2025 at 9:30 am while the patient was on veno arterial ecmo for refractory cardiogenic shock requiring inotropic and vasopressor support. during the support period, the patient developed thrombocytopenia with morning platelet count of 77, prompting initiation of bivalirudin per mcs protocol. the patient also required continuous renal replacement therapy (crrt).the impella 5.5 functioned as intended for ventricular unloading during severe cardiogenic shock and ecmo support. there is no evidence of device related malfunction or injury, and the patient was successfully weaned from mechanical circulatory support and survived. tr 13feb2026. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
240 | 1220648-2026-03644 | 24399830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399830&pc=OZD | 11/22/2025 |  |  | 02/20/2026 |  |  | 11/22/2025 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748336 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
241 | 1220648-2026-03645 | 24399833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399833&pc=OZD | 11/22/2025 |  |  | 02/20/2026 |  |  | 11/22/2025 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026716169 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 52-year-old patient was admitted for exacerbation of heart failure. an intra-aortic balloon pump (iabp) was initially placed and the patient was transferred to a higher level of care. due to the poor ejection fraction and ongoing cardiogenic shock, the patient was switched to an impella 5.5 with the iabp removed. fasciotomy performed at the iabp site at the right lower extremity. purge fluid changed from heparin due to bleeding at jp impella insertion site. the patient had gi bleed requiring endoscopy and transfusion of blood and platelets. he was hypotensive during this time and required levophed initiation. suction issues were noted and the patient received additional blood products. the patient was ultimately implanted with a durable left ventricular assist device on (b)(6) 2026 with impella removal. it is difficult to link the gi bleed to the impella device, however, with heparin infusion to support the impella device, it is likely that this may have contributed to the gi bleed. the hypotension most likely was due to the patients cardiac condition. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
242 | 1220648-2026-03646 | 24399845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399845&pc=OZD | 11/23/2025 |  |  | 02/20/2026 |  |  | 11/23/2025 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026757906 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 59-year-old patient underwent a left heart catheterization showing severe multi vessel disease with 25-30% ejection fraction. she underwent coronary artery bypass surgery with successful implant of impella 5.5 for ongoing support post operatively. the patient had some initial episodes of low pusatility due to hypovolemia and poor native heart function and was started on vasopressors. the patient was given platelets for low platelet count post operatively. the patient was successfully weaned and explanted on (b)(6) 2025 without incident. unable to determine whether the thrombocytopenia was due to impella pump or just post operative sequalae. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
243 | 1220648-2026-03648 | 24399852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399852&pc=OZD | 11/23/2025 |  |  | 02/20/2026 |  |  | 11/23/2025 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748270 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 106 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 65-year-old patient was admitted with chest pain and was treated in the cath lab. post procedure, the patient developed severe chest pain, underwent right and left heart cath which showed cardiogenic shock (cs) and an impella cp was placed. the patient was transferred to another hospital for advanced care. upon arrival, the patient was found to have lower leg and foot mottling and was found to have bleeding from the groin insertion site. the catheter and sheath were repositioned and no longer showed any bleeding or hematoma. the labs showed liver failure. inotropes and vasopressors were weaned however, due to the ongoing cs, an impella 5.5 was placed and the cp was removed. the patient was pan cultured due to fever. the patient went to the cath lab on (b)(6) 2025 for multiple cardiac procedures. ultimately, the patient was successfully weaned and the impella was removed on (b)(6) 2025. the lower extremity vascular issues were likely caused by the impella cp. the bleeding from the site was due to the way the cp catheter and sheath was implanted. for the 5.5, the organ failure most likely due to the cs. it is difficult to attribute the fever to the impella. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
244 | 1220648-2026-03650 | 24399861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399861&pc=OZD | 02/13/2026 |  |  | 02/20/2026 |  |  | 02/13/2026 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802420 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 122 KG | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced high purge pressure. no kinks or obstructions were noted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
245 | 1220648-2026-03651 | 24399877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399877&pc=OZD | 11/23/2025 |  |  | 02/20/2026 |  |  | 11/24/2025 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026743983 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 83 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Material Split, Cut or Torn | [5 empty] | Arrhythmia | Cardiogenic Shock | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 57-year-old male patient with dilated cardiomyopathy was admitted with heart failure cardiogenic shock (cs). on arrival to hospital, he was found to be acidotic and hypotensive. the icu team attempted to wean inotropes however; the patient did not tolerate resulting in acute kidney injury and new onset atrial fibrillation. the team decided to place an impella 5.5 due to his ongoing cardiogenic shock. due to the impella being too deep on echo, the team repositioned at the bedside. during the patient¿s heart transplant/ventricular assist device workup, he had a gout flare with knee pain. during the patient¿s support, it was noted that the distal end of anticontamination sleeve was ripped off and the white fastener with external catheter was exposed to air at the distal portion. the staff managed this with dressing change and ongoing monitoring. there was not event noted from this. the patient was without clinical complications until he was ultimately transplanted and the impella 5.5 was removed without incident. as this patient was in cardiogenic shock on admission, the symptoms appear to be prior to the impella placement so most likely unrelated. his gout flare was a comorbid condition and not impella related. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
246 | 1220648-2026-03652 | 24399900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399900&pc=OZD | 11/24/2025 |  |  | 02/20/2026 |  |  | 11/24/2025 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026686439 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Low Blood Pressure/ Hypotension | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
247 | 1220648-2026-03655 | 24399915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399915&pc=OZD | 02/11/2026 |  |  | 02/20/2026 |  |  | 02/11/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745773 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
248 | 1220648-2026-03657 | 24399951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399951&pc=OZD | 02/12/2026 |  |  | 02/20/2026 |  |  | 02/12/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779298 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Increase in Pressure | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
249 | 1220648-2026-03658 | 24399960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24399960&pc=OZD | 02/11/2026 |  |  | 02/20/2026 |  |  | 02/12/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1429644 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
250 | 24400018 | 24400018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400018&pc=OZD | 01/22/2026 | 02/12/2026 | 02/12/2026 | 02/20/2026 | 02/12/2026 | 02/20/2026 |  | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 1000100 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 1000100 |  | Device was returned to manufacturer |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 44 YR |  | Female |  |  |  | Malfunction | No | Yes | Insufficient Flow or Under Infusion | Defective Device | [5 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella device malfunctioned, which then led to a series of unfortunate events. following placement of the initial impella device which in itself was completely straightforward, a few hours later, the icu team noticed reduction of impella purge flows. when the impella rep was contacted, he thought that there might be thrombus formation on the device, which then led him to suggest thrombolysis of this potential clot with tpa [tissue plasminogen activator]. initially, with 1 or 2 hours of tpa administration, the impella purge flows improved, and the tpa was stopped. however, the problem recurred, and tpa was re-administered. this then led to sweating of blood from the dacron graft itself, which then led to formation of a hematoma at the impella insertion site, which then required re-exploration by my colleague over the weekend. this is unarguably a potential problem with administration of thrombolysis with tpa. the device continued to malfunction with intermittent cessation of impella purge flows. further anticoagulation and tpa would not improve the situation any further. given that this patient was a blood type o and was expected to potentially have the impella device for at least 2 months while waiting for suitable heart due to the current unos [united network for organ sharing] waiting times, the icu team and i thought that the best option would be to replace the malfunctioning impella device with a new impella device. the impella device which was explanted in the operating room has subsequently been retrieved by johnson & johnson for close inspection and examination.
251 | 1220648-2026-03662 | 24400061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400061&pc=OZD | 01/31/2026 |  |  | 02/20/2026 |  |  | 01/31/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1288953 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
252 | 1220648-2026-03666 | 24400226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400226&pc=OZD | 02/13/2026 |  |  | 02/20/2026 |  |  | 02/13/2026 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1805843 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 97 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in a 65-year-old male patient presenting with cardiomyopathy and a history of cad and prior cabg. during support, a controller error occurred, and the catheter would not function on any aic despite attempts to troubleshoot. the reported events are failed to detect purge cassette, controller failure, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
253 | 1220648-2026-03667 | 24400273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400273&pc=OZD | 02/13/2026 |  |  | 02/20/2026 |  |  | 02/13/2026 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1776251 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Priming Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | During the initial case, the purge tubing did not prime fluid forward. a different aic was then used, and the same purge tubing primed without issue, resolving the problem. the reported events are priming problem, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
254 | 1220648-2026-03668 | 24400351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400351&pc=OZD | 02/11/2026 |  |  | 02/20/2026 |  |  | 02/11/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1486033 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
255 | 1220648-2026-03669 | 24400374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400374&pc=OZD | 02/14/2026 |  |  | 02/20/2026 |  |  | 02/14/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026809746 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
256 | 1220648-2026-03670 | 24400398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400398&pc=OZD | 02/13/2026 |  |  | 02/20/2026 |  |  | 02/13/2026 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1087351 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Purge disc not detected occurred during setup. troubleshooting was attempted; however, the console continued to indicate that the purge disc was not recognized. the reported events are purge disc not detected, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
257 | 1220648-2026-03671 | 24400416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400416&pc=OZD | 02/13/2026 |  |  | 02/20/2026 |  |  | 02/13/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1420045 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Application Program Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
258 | 1220648-2026-03675 | 24400757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400757&pc=OZD | 02/13/2026 |  |  | 02/20/2026 |  |  | 02/13/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765247 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Vascular Dissection | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
259 | 1220648-2026-03676 | 24400761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400761&pc=OZD | 02/09/2026 |  |  | 02/20/2026 |  |  | 02/12/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802300 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
260 | 1220648-2026-03677 | 24400778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400778&pc=OZD | 02/13/2026 |  |  | 02/20/2026 |  |  | 02/13/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026814416 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
261 | 1220648-2026-03679 | 24400792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400792&pc=OZD | 02/10/2026 |  |  | 02/20/2026 |  |  | 02/14/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806141 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Backflow | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
262 | 1220648-2026-03681 | 24400877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400877&pc=OZD | 02/13/2026 |  |  | 02/20/2026 |  |  | 02/13/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797225 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
263 | 1220648-2026-03682 | 24400897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24400897&pc=OZD | 02/10/2026 |  |  | 02/20/2026 |  |  | 02/10/2026 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026683684 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Inaccurate Flow Rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: a 72-year-old male undergoing high risk percutaneous coronary intervention (hrpci) with a pre support clinical state consistent with scai cardiogenic shock stage c was supported with an impella 5.5 device via a right axillary/subclavian surgical arterial graft. the device exhibited saturated pump flow and waveform abnormalities during use, prompting clinical concern for potential pump failure. biomaterial was identified at the outlet area after explant, indicating a possible obstruction contributing to the observed flow saturation. the device was removed in response to the failure mode and will be returned for investigation to determine the root cause. no patient injury was reported, and the patient survived the procedure. root cause cannot be confirmed until completion of the manufacturer¿s evaluation. | [5 empty] | This is one of two related reports. this report represents the second impella 5.5 and another report was submitted to represent the first impella 5.5. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
264 | 1220648-2026-03606 | 24401057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24401057&pc=OZD | 12/22/2025 |  |  | 02/20/2026 |  |  | 12/22/2025 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1130885 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
265 | 1220648-2026-03688 | 24402094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24402094&pc=OZD | 12/21/2025 |  |  | 02/20/2026 |  |  | 12/21/2025 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026753181 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 93 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Inaccurate Flow Rate | Increase in Pressure | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 48-year-old male with a history of coronary artery disease, diabetes mellitus, renal insufficiency, and dialysis received an impella 5.5 for acute heart failure due to ischemic cardiomyopathy while supported with two inotropic medications, an intra-aortic balloon pump, and mechanical ventilation. the patient was positive for heparin-induced thrombocytopenia and was treated with argatroban, with an activated clotting time of 193 seconds at the time of insertion. the first impella 5.5 encountered difficulty crossing the aortic valve and later generated purge system blockage alarms, leading to removal and replacement. the second device was more difficult to insert due to repeated loss of wire position but was ultimately placed successfully with normalized activated clotting time. initial issues with left ventricular and aortic pressure waveforms and pump performance resolved by the end of the procedure. this will be coded in detail by engineering team. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
266 | 1220648-2026-03701 | 24405173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24405173&pc=OZD | 02/12/2026 |  |  | 02/20/2026 |  |  | 02/12/2026 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1692374 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: a warning flag appeared on the console indicating that the system recommended changing out the console. the warning flag subsequently disappeared and did not reappear. the customer proceeded to switch out the console. the customer initially believed the alert may have been related to the aic being due for preventative maintenance; however, it was later confirmed that maintenance was not due for several months. a return box was requested so the customer can send back the aic to ensure it is not malfunctioning. the reported events are controller failure, medical device removal, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was successfully maintained. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
267 | 1220648-2026-03704 | 24405362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24405362&pc=OZD | 02/13/2026 |  |  | 02/20/2026 |  |  | 02/14/2026 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801104 |  | 1000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was inserted via left axillary artery in a 50 year old female who was admitted for cardiomyopathy. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage d. on day 3 of support, the patient developed a hematoma and heparin was paused. the patient's left should also appeared inflamed, in which a jackson-pratt drain was 1/4 full with sanguinous drainage. on day 8 of support, it was reported the hematoma was resolving. vascular injury is a known potential adverse event of the impella 5.5 per the instructions for use. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
268 | 1220648-2026-03706 | 24405485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24405485&pc=OZD | 12/19/2025 |  |  | 02/20/2026 |  |  | 12/21/2025 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026710575 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
269 | 1220648-2026-03709 | 24405592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24405592&pc=OZD | 12/21/2025 |  |  | 02/20/2026 |  |  | 12/21/2025 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782288 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 74 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 65-year-old male with coronary artery disease received an impella 5.5 for acute myocardial infarction¿related cardiogenic shock while supported with two inotropic medications, an intra-aortic balloon pump, and mechanical ventilation. after right axillary insertion, the patient showed reduced responsiveness to heparin and later developed oozing and a hematoma at the axillary drain site, which improved after additional protamine and local pressure. the device was removed after four days due to recovery of native cardiac function. hematoma day 0. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
270 | 1220648-2026-03708 | 24405619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24405619&pc=OZD | 02/12/2026 |  |  | 02/20/2026 |  |  | 02/12/2026 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1194346 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
271 | 1220648-2026-03711 | 24405761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24405761&pc=OZD | 12/24/2025 |  |  | 02/20/2026 |  |  | 12/24/2025 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1505149 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that after the device was successfully implanted, a placement signal unreliable alarm occurred. upon inspection of the surgical field, it was found that a clamp had been placed across the catheter by the medical staff. after the clamp was removed, the placement signal became intermittently reliable. the physician elected to leave the pump in place, and clinical consulting was contacted regarding the reliability of the placement signal. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
272 | 1220648-2026-03713 | 24405820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24405820&pc=OZD | 02/01/2026 |  |  | 02/20/2026 |  |  | 02/12/2026 | 02/20/2026 |  |  |  |  | 02/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1927390 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 73 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted via the right femoral artery in an 83-year-old female patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). the patient¿s medical history included coronary artery disease (cad) and renal insufficiency. during support, the impella console (b)(6) displayed a purge disc not detecting alert. the reported events are purge disc not detected, device revision or replacement, and hemodynamic instability. the impella cp will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was completed with no consequence to the patient, and support was successfully resumed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
273 | 1220648-2026-03714 | 24405843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24405843&pc=OZD | 12/22/2025 |  |  | 02/20/2026 |  |  | 12/22/2025 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781193 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
274 | 1220648-2026-03719 | 24408409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24408409&pc=OZD | 11/06/2025 |  |  | 02/20/2026 |  |  | 11/06/2025 | 02/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026738763 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Respiratory Failure | [35 empty] | Invalid/Incomplete |  |  | No | Initial submission
275 | 1220648-2026-03531 | 24387417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24387417&pc=OZD | 02/09/2026 |  |  | 02/19/2026 |  |  | 02/09/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026769373 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was placed via the direct aorta/surgical access while in the operating suite for bypass surgery/cabg and tricuspid valve replacement. the patient is a 67 year old female who was known to be presenting in scai stage e shock. the 5.5 was placed as patient was coming off the bypass pump. the team did note the chest cavity to be oozing blood with a large volume of blood loss noted in the atrium. the team infused concentrated fibrinogen, ddvap desmopressin acetate, and multiple blood products inclusive of blood cells and fresh frozen plasma and cryopreciptate. when taken for a washout the bleed was seen to be from the aortic bypass cannulation site, not from the impella pump site. anticoagulation necessary for the pump placement and support can contribute to bleeding. the pump remains on for support as hemodynamic support is still deemed necessary. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
276 | 1220648-2026-03539 | 24387613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24387613&pc=OZD | 02/09/2026 |  |  | 02/19/2026 |  |  | 02/09/2026 | 02/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2019073 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Failure to Read Input Signal | Device Sensing Problem | [5 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
277 | 1220648-2026-03544 | 24387672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24387672&pc=OZD | 02/11/2026 |  |  | 02/19/2026 |  |  | 02/11/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801616 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 66 YR |  | Female | 81 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Pericardial Effusion | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was placed via the axillary access graft site for the support of a 66 year old female patient who had been admitted in with cardiomyopathy and cardiogenic shock, presenting in scai stage d shock. the patient was on inotropes and vasopressors and a vent for respiratory needs. the 5.5 position was confirmed by imaging and the sheath was removed and the pocket was closed. after 10-15 minutes of support she experienced hypotension and the surgical team re-imaged with tee echo to observe there was an effusion. the team performed a pericardial window to drain and the bleed persisted so, team lifted her heart and this revealed there was a large defect in the left atrium. the attempted surgical repair and transfusion, but this failed and the patient expired. the team did acknowledge that the patient had multiple ablations prior to the effusion and left atrium defect, which could have been a contributing factor. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's presenting native critical condition as presented in scai stage d. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
278 | 1220648-2026-03549 | 24387773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24387773&pc=OZD | 01/01/2026 |  |  | 02/19/2026 |  |  | 02/11/2026 | 02/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748359 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Increase in Pressure | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
279 | 1220648-2026-03550 | 24387802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24387802&pc=OZD | 05/15/2025 |  |  | 02/19/2026 |  |  | 02/10/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026651969 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Cardiac Tamponade | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was inserted for mechanical circulatory support in a patient of unknown age and gender as part of a clinical study. specific past medical history, hemodynamic parameters, access site details, and baseline comorbidities were not provided in the available documentation. per study report, the patient experienced cardiac tamponade, hematoma, and required surgical intervention. the event description indicates that a re-thoracotomy was performed for management. no additional procedural details, timing of onset relative to device placement, anticoagulation parameters, imaging findings, or operative findings were included in the available information. based on the limited information provided, the reported events of cardiac tamponade and hematoma requiring surgical intervention were adjudicated within the study as related to the impella 5.5. however, due to the absence of detailed clinical, procedural, and diagnostic data, a comprehensive independent assessment of causality cannot be performed. cardiac tamponade and postoperative bleeding are known potential complications associated with cardiac surgical procedures, re-intervention, anticoagulation therapy, and large-bore access required for temporary mechanical circulatory support. the patient survived following surgical management of the reported complications. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
280 | 1220648-2026-03551 | 24387838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24387838&pc=OZD | 02/10/2026 |  |  | 02/19/2026 |  |  | 02/11/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026714202 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Increase in Pressure | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in a 56-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). the patient had a history of renal insufficiency. a call was received from the local team requesting that the md be contacted regarding a ¿pp high¿ alarm. the md reported a persistent high purge pressure alarm with no kinking identified on inspection. the purge solution had already been changed from bicarbonate to heparin. at the time of the call, the alarm was not active; however, purge pressure remained around 1000 mmhg with a purge flow of 2.2 ml/h. recommendations and best practices were shared. the possibility of tpa lysis was discussed and requested by the physician, and the tpa lysis protocol was sent. guidance was also provided on how to change the purge bag from 500 ml d5 to 50 ml sterile water with tpa. the purge fluid change was performed, and best practices were reinforced. after lysis, the plan was to transition back to standard bicarbonate purge. the patient remains on support. the reported events are purge pressure high, medication required, and hemodynamic instability. the device will be conservatively reported for serious injury due to medication (tpa) being required to address the event; however, there was no consequence to the patient, and support was continued without any known adverse events.. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
281 | 1220648-2026-03552 | 24387872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24387872&pc=OZD | 02/10/2026 |  |  | 02/19/2026 |  |  | 02/11/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782831 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 147 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella cp (b)(4) was inserted via right femoral artery in a 71-year-old male who was admitted for acute myocardial infarction with cardiogenic shock. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage c. on the first day of support, the patient was bridged to impella 5.5. on day 6 of 5.5 support, a computed tomography scan revealed multiple brain infarcts that were terminal in nature. the family made the decision to withdraw care. the 5.5 was explanted and patient expired. stroke and death are known risks associated with impella 5.5 as described in the instructions for use. | [5 empty] | This is one of two related reports. this report represent the impella 5.5 and another report was submitted to represent the impella cp. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
282 | 1220648-2026-03553 | 24387879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24387879&pc=OZD | 02/10/2026 |  |  | 02/19/2026 |  |  | 02/11/2026 | 02/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026793911 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
283 | 1220648-2026-03556 | 24387993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24387993&pc=OZD | 02/11/2026 |  |  | 02/19/2026 |  |  | 02/11/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719221 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 138 KG |  |  | Injury | Yes | No | Increase in Pressure | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in a 47-year-old male patient presenting with cardiomyopathy. the patient had a history of renal insufficiency. an rn called to report high purge pressures in the 1000s. no kinks were identified on inspection of the purge line. the team was advised to change the purge cassette and to call back if the issue did not resolve. they were further advised that tpa infusion through the purge would be the next recommended troubleshooting step. physician requested tpa protocol. patient remains on support. the reported events are purge pressure high, medication required, and hemodynamic instability. the device will be conservatively reported for serious injury due to medication (tpa) being required to address the event; however, there was no consequence to the patient, and support was continued without any known adverse events. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
284 | 1220648-2026-03559 | 24388070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24388070&pc=OZD | 02/11/2026 |  |  | 02/19/2026 |  |  | 02/11/2026 | 02/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026805517 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
285 | 1220648-2026-03564 | 24388758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24388758&pc=OZD | 11/22/2025 |  |  | 02/19/2026 |  |  | 02/13/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026760817 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant provided clinical details that mention that a chest x-ray was performed during a case and the catheter for the impella pump was found to be twisted inside the patient's body. during investigation an image verified the twisted catheter, but further details of how and when the catheter got twisted are unknown. there was no patient harm was noted and the procedure was completed successfully. | [5 empty] | Investigation summary: the investigation was completed. no product returned. product damage (catheter kink): the root cause of the product damage could not be determined due to lack of conclusive evidence from the provided information and unreturned product for further investigation.
286 | 1220648-2026-03568 | 24389094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24389094&pc=OZD | 02/11/2026 |  |  | 02/19/2026 |  |  | 02/12/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | P140003 |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026793428 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Unknown |  |  |  | Death | Yes | Yes | Increase in Pressure | Malposition of Device | [5 empty] | Arrhythmia | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Clinical rationale: an impella cp device was inserted into the femoral artery in a 62-year-old male patient presenting with an out-of-hospital cardiac arrest requiring cardiopulmonary resuscitation (cpr), acute myocardial infarction (ami), and cardiogenic shock (cgs) in scai stage e shock prior to initiation of support. the patient was critically ill in the cardiac catheterization lab and required defibrillation in the hospital environment. the cath lab staff confirmed optimal impella pump position and function. while the patient was in the intensive care unit (icu), the automated impella controller (aic) console had a ¿high purge pressure/purge system blocked¿ alarm. no kinks were detected and the glucose was never above 5%. the purge cassette was replaced but did not resolve the issue. the icu nurses indicated the purge issues were probably not related to the impella pump flow. the device functioned at p-9 at 3.5 l/min with no issues. care was withdrawn and the post-procedure outcome was expired at explant. the device is unlikely to have contributed to the patient's death, which was most likely due to the clinical condition of a critically ill patient with ami complicated by cgs as they presented in scai stage e shock. arrhythmias are common in patients with ami/cgs and are listed as potential adverse events and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). the automated impella controller will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | Updated clinical rationale: an impella cp device was inserted into the femoral artery in a 62-year-old male patient presenting with an out-of-hospital cardiac arrest requiring cardiopulmonary resuscitation (cpr), acute myocardial infarction (ami), and cardiogenic shock (cgs) in scai stage e shock prior to initiation of support. the patient was critically ill in the cardiac catheterization lab and required defibrillation in the hospital environment. the cath lab staff confirmed optimal impella pump position and function. while the patient was in the intensive care unit (icu), the automated impella controller (aic) console had a ¿high purge pressure/purge system blocked¿ alarm. no kinks were detected and the glucose was never above 5%. the purge cassette was replaced but did not resolve the issue. the icu nurses indicated the purge issues were probably not related to the impella pump flow. the device functioned at p-9 at 3.5 l/min with no issues. care was withdrawn and the post-procedure outcome was expired at explant. the device is unlikely to have contributed to the patient's death, which was most likely due to the clinical condition of a critically ill patient with ami complicated by cgs as they presented in scai stage e shock. arrhythmias are common in patients with ami/cgs. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 was updated. h6 codes e2343 and f1905 erroneously entered on the initial manufacturer device report number.
287 | 1220648-2026-03571 | 24389199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24389199&pc=OZD | 02/12/2026 |  |  | 02/19/2026 |  |  | 02/12/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025569103 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 92 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted surgically into the right aorta in a 68-year-old male patient presenting in scai stage c shock on an intra-aortic balloon pump (iabp) and respiratory support via nasal cannula prior to initiation of support. the impella 5.5 was inserted for ventricular support during an elective high-risk coronary artery bypass graft (cabg) surgery. post-procedure, the patient was transferred to the intensive care unit (icu) in stable condition, following protamine administration prior to separation from the cardiopulmonary bypass (cpb) machine. post-operative laboratory results included partial thromboplastin time (ptt) 28 (starting ptt was 34) and platelets 100,000 (starting platelets were 109,000). a decision was made by the physician to remove the iabp at the bedside. immediately upon removal, the patient¿s blood pressure and pulsatility dropped requiring vasopressor initiation. twenty minutes later, a hematoma developed at the right groin extending into the scrotum and penis. vascular surgery was consulted and the patient returned to the operating room (or) for right groin cutdown and exploration. the patient received two units of packed red blood cells (prbc). the post-procedure outcome was unknown. bleeding is a known risk due to the impella's anticoagulation and purge requirements. bleeding and thrombocytopenia are listed as potential adverse events and are consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
288 | 1220648-2026-03573 | 24389378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24389378&pc=OZD | 02/12/2026 |  |  | 02/19/2026 |  |  | 02/12/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802439 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted surgically into the right axillary artery in a 71-year-old male patient with a past medical history of coronary artery disease (cad) presenting in scai stage b shock prior to initiation of support. the impella 5.5 was inserted for ventricular support during an elective high-risk coronary artery bypass graft (cabg) surgery. while the patient was in the intensive care unit (icu), post-procedure, a hemothorax was discovered. anticoagulation was held and the patient received two units of packed red blood cells (prbc). there were no issues reported with the impella device. the post-procedure outcome was survived at explant. bleeding is a known risk due to the impella's anticoagulation and purge requirements. bleeding is listed as a potential adverse event and is consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | The impella device was discarded. therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
289 | 1220648-2026-03574 | 24389383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24389383&pc=OZD | 02/13/2026 |  |  | 02/19/2026 |  |  | 02/13/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1215712 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 78 KG | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during support with an impella device, the automated impella controller (aic) displayed a battery controller error. no additional information regarding the circumstances of the alarm or troubleshooting steps was provided. there were no signs or symptoms of patient injury reported, and no harm was associated with the event. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
290 | 1220648-2026-03575 | 24389491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24389491&pc=OZD | 01/21/2026 |  |  | 02/19/2026 |  |  | 02/12/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026660561 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male | 75 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Edema | Rupture | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp was inserted via an unknown access site in a 75-year-old male patient with acute myocardial infarction and cardiogenic shock (ami/cgs). during the course of treatment, an explant was performed in the cath lab, during which the device was removed under fluoroscopic guidance without noted resistance. later in the icu, a new murmur was detected on echocardiography, and once the patient was extubated, pulmonary edema became apparent. the patient was transferred to another hospital, where evaluation revealed papillary muscle (chordae) rupture, requiring acute surgical repair. no ischemic territory was identified, and although the physician noted that a mechanical cause was most likely, this could not be definitively confirmed or ruled out. the pump had already been discarded, and no product return was possible. no device damage was noted. surgical intervention was required, and no additional contributing factors were identified. based on the available information, the papillary muscle rupture and subsequent pulmonary edema appear more consistent with a clinical or patient specific complication following device explant rather than a malfunction of the impella cp. the absence of resistance during removal, lack of product damage, and unconfirmed causality support that no device defect has been identified. final assessment remains inconclusive, with no evidence at this time indicating that the impella cp contributed directly to the injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
291 | 1220648-2026-03579 | 24389696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24389696&pc=OZD | 02/11/2026 |  |  | 02/19/2026 |  |  | 02/11/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026807743 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted via the femoral artery in a patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). the clinic reported that the implanting physician attempted to place the impella cp; however, during insertion, a kink was identified in the catheter, making it impossible to advance and implant the pump. the team opened a second impella cp, which was successfully inserted without further difficulty. the patient did not experience any harm related to the defective pump, and support was continued appropriately. the reported events include product damage, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
292 | 1220648-2026-03580 | 24389750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24389750&pc=OZD | 02/11/2026 |  |  | 02/19/2026 |  |  | 02/12/2026 | 02/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026757861 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Pain | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
293 | 1220648-2026-03585 | 24390005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24390005&pc=OZD | 02/14/2026 |  |  | 02/19/2026 |  |  | 02/14/2026 | 02/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026811360 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
294 | 1220648-2026-03586 | 24390008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24390008&pc=OZD | 02/09/2026 |  |  | 02/19/2026 |  |  | 02/12/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746502 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a patient presenting with pre-operative support needs. the clinical team attempted to insert the impella 5.5 through the right subclavian artery (rt sca), but the device could not be advanced due to the vessel diameter being too small to allow passage. the impella cp was then inserted through the same access site, and support was initiated without complications. the reported events are failure to advance, medical device removal, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was successfully maintained. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
295 | 1220648-2026-03598 | 24392565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24392565&pc=OZD | 02/07/2026 |  |  | 02/19/2026 |  |  | 02/07/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795470 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 128 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale updated: a 61-year-old male with an underlying indication of cardiomyopathy underwent implantation of an impella 5.5 (pump 2; case (b)(4), gu case (b)(4)). the clinical team reported that the patient developed a gastrointestinal bleed and ischemic bowel. based on the information provided, the patient experienced significant bleeding complications, including gi bleed and ischemic bowel, while supported with the impella 5.5. a device related contribution cannot be determined at this time; further investigation will proceed upon receipt of additional clinical details and/or product evaluation. bleeding was noted and this is a known risk per our ifu (instructions for use) because of our anticoagulation and purge requirements. the patient survived explant and was bridged to transplant. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
296 | 1220648-2026-03605 | 24393210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24393210&pc=OZD | 12/22/2025 |  |  | 02/19/2026 |  |  | 12/22/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782857 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 55-year-old male was electively placed on an impella 5.5 to facilitate weaning after bypass surgery. following 5 days of support, the impella aic console's soft touch buttons became unresponsive while hemodynamic support to the patient kept ongoing. an aic exchange was carried out. a day later, the purge system appeared to be blocked and couldn't be reactivated. as the patient was already successfully weaned and scheduled for device removal, this was carried out as planned. the impella 5.5 was successfully removed after 5 days of support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
297 | 1220648-2026-03603 | 24393224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24393224&pc=OZD | 12/22/2025 |  |  | 02/19/2026 |  |  | 12/22/2025 | 02/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1283307 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Key or Button Unresponsive/not Working | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
298 | 1220648-2026-03609 | 24393555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24393555&pc=OZD | 02/12/2026 |  |  | 02/19/2026 |  |  | 02/12/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806925 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 84 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 was surgically inserted via the right axillary/subclavian artery in an 84-year-old male patient with acute myocardial infarction and cardiogenic shock (ami/cgs). the patient had a history of coronary artery disease (cad) and diabetes. the patient¿s clinical state prior to initiation of support was scai stage d shock. the patient remained on support when clinical staff reported a thrombus that migrated and occluded the right radial artery. there were no allegations of device malfunction, no request for replacement, and no product was returned. pump position remained appropriate, and the patient continued on support until care was ultimately withdrawn. the patient expired. based on the available information, the thrombus migration and subsequent radial artery occlusion appear clinically related to the patient¿s underlying condition and critical illness. no device malfunction was alleged, and no device related factors were identified as contributing to the thrombus event or the patient¿s death. current findings indicate that the device functioned as intended, with final assessment pending any further product evaluation if conducted. the death is conservatively being reported on the impella 5.5 but is unlikely a contributing factor to the cause of death and is more likely due to the patient¿s declining clinical condition. | [5 empty] | E4 was inadvertently omitted on the initial manufacturer device report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the exact date of death is unknown. if additional information becomes known, a supplemental report will be submitted.
299 | 1220648-2026-03612 | 24394120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24394120&pc=OZD | 12/24/2025 |  |  | 02/19/2026 |  |  | 12/24/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026635230 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 106 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 54 year old male was treated for acute myocardial infarction with cardiogenic shock. an impella cp was placed but soon escalated to an impella 5.5. during placement of the impella 5.5, the abiomed 0.018¿ wire kinked, so the pump would not advance into the left ventricle until the wire was replaced, which delayed placement. during and following the pump exchange, the patient developed bleeding and an expanding hematoma at the impella access site, requiring blood products, increasing pressor support and continued efforts until hemostasis was achieved. in an off-label extended duration, patient support is still ongoing at the time of review. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
300 | 1220648-2026-03614 | 24396725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24396725&pc=OZD | 10/28/2025 |  |  | 02/19/2026 |  |  | 10/28/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713917 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 82 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An 80-year-old male with a past medical history significant for former tobacco use, hyperlipidemia (hld), hypertension (htn), prostate cancer (ca), and bladder cancer (ca) presented to the emergency department (ed) with st-segment elevation myocardial infarction (stemi). coronary angiography demonstrated significant disease of the left anterior descending (lad) artery requiring placement of four stents. an impella cp was implanted for bailout mechanical circulatory support during the procedure. subsequently, the patient developed hemolysis, evidenced by clinical and laboratory findings consistent with red blood cell destruction. the patient was transferred to another facility with worsening cardiogenic shock. due to hemodynamic deterioration, the decision was made to escalate support to an impella 5.5. the impella 5.5 was successfully inserted via the right axillary artery using a double-barrel technique. the impella cp was removed, and hemostasis was achieved. device position of the impella 5.5 was confirmed by transesophageal echocardiography (tee), measured at 4.0 cm, and secured at 42.5 cm with a three-point fixation device. the device was maintained at performance level p-6, and the patient was transferred to the cardiovascular intensive care unit (cvicu) on support. documentation indicated plans to initiate continuous renal replacement therapy (crrt) due to renal dysfunction; however, it is unknown whether crrt was initiated. on the following day, the patient expired after care was withdrawn. hemolysis is considered a serious injury and is a known potential complication associated with impella cp support; therefore, the hemolysis event is attributed to the impella cp. the death is being associated with the impella 5.5 because this device was in place at the time of death. renal failure is also being reported in association with the impella 5.5 due to documented plans for continuous renal replacement therapy, although initiation of therapy could not be confirmed. the likely cause of death is the patient's underlying condition (severe coronary artery disease and cardiogenic shock [scai stage d]). | [5 empty] | The impella device was not received from the customer; therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: ppae (renal failure): the root cause of the injury was not determined due to insufficient clinical details.
301 | 1220648-2026-03615 | 24397082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397082&pc=OZD | 02/12/2026 |  |  | 02/19/2026 |  |  | 02/12/2026 | 02/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026701716 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
302 | 1220648-2026-03616 | 24397093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397093&pc=OZD | 02/12/2026 |  |  | 02/19/2026 |  |  | 02/12/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1103511 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 100 KG | Hispanic | White | Injury | Yes | Yes | Display or Visual Feedback Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Clinical rationale: an impella 5.5 device was inserted via the right direct surgical approach in a 42-year-old male patient presenting for elective placement. the patient¿s medical history included coronary artery disease (cad) and diabetes mellitus (dm). a call was received from the unit reporting that the aic had been changed due to the i-waveform appearing artifact-like. the reported events are display issue, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1 product problem and e4 was inadvertently omitted on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated.
303 | 1220648-2026-03617 | 24397254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397254&pc=OZD | 11/13/2025 |  |  | 02/19/2026 |  |  | 11/13/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026757315 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male | 41 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | Pre-procedure documentation described a multimorbid patient with the last remaining coronary vessel requiring intervention. an impella cp was planned and implanted pre-percutaneous coronary intervention (pci) to provide mechanical circulatory support. during insertion via the common femoral artery, the first impella cp (sn (b)(6)) was noted to kink after passage through the peel-away introducer sheath. vascular calcification had not been identified on fluoroscopy prior to insertion. the initial implant attempt was aborted, and a second impella cp (sn (b)(6)) was subsequently implanted successfully. the patient underwent high-risk pci to the left anterior descending and right coronary arteries, with vessel treatment including percutaneous transluminal coronary angioplasty and stent placement. after the procedure the impella cp pump 2 was successfully weaned, explanted, and the access site was closed without complication. some minutes (exact timeframe unnoted) following explant and closure, the patient unexpectedly experienced cardiopulmonary arrest, required cardiopulmonary resuscitation, and expired in the cardiac catheterization laboratory. no additional clinical or procedural details regarding the event were noted. based on the available information, the patient had underlying coronary artery disease and underwent complex high-risk pci involving the last remaining coronary vessel. although the impella cp pump 2 had been successfully weaned and explanted prior to the cardiopulmonary arrest, the event occurred within [unknown] minutes of device removal during the same procedural setting. no device malfunction or performance issue has been reported for impella cp pump 2 at this time. however, given the temporal proximity to impella cp pump 2 support and explantation, the death is being reported on the impella cp pump 2. | [5 empty] | The impella device was not received from the customer; therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | Additional information was received and noted the following: the exact time between device explant and cardiac arrest is unknown; however, it was estimated to be approximately 7¿10 minutes, as hospital documentation was unavailable. no immediate hemodynamic instability was observed following explant. the patient was reportedly awake, responsive, and normotensive during the initial 7¿10 minutes post-explant. clinical deterioration then occurred suddenly without warning, progressing rapidly to cardiac arrest requiring resuscitation. autopsy status remains unknown/pending, and the cause of death is currently undetermined. the clinical consultant reported no known vascular injury at the femoral access site. regarding potential causality, it remains unclear and speculative whether removal of impella mcs unmasked or precipitated cardiovascular collapse, or whether withdrawal of support may have contributed to acute left ventricular failure. it is also undetermined whether the event is more consistent with underlying disease progression, a percutaneous coronary intervention complication, vascular complication, or a device-related complication, and it is unknown whether an autopsy was performed that could further clarify the cause.
304 | 1220648-2026-03618 | 24397406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397406&pc=OZD | 02/12/2026 |  |  | 02/19/2026 |  |  | 02/12/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026805668 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 120 KG | Non Hispanic | White | Injury | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 70-year-old male patient presenting with hrpci. the patient¿s medical history included coronary artery disease (cad) and diabetes mellitus (dm). the impella 5.5 could not be delivered through the anastomosis, so the clinical team proceeded with impella cp placement. the cp device was delivered; however, the surgeon continued to experience difficulty advancing the device. the reported events are failure to advance, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was successfully maintained. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
305 | 1220648-2026-03619 | 24397539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397539&pc=OZD | 02/12/2026 |  |  | 02/19/2026 |  |  | 02/12/2026 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026706423 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted surgically into the right axillary/subclavian artery in a 72-year-old male patient with a past medical history of renal insufficiency presenting in cardiomyopathy, and in scai stage d shock, on multiple inotropes and vasopressors, and on a ventilator for respiratory support, prior to initiation of support. while the patient was in the intensive care unit (icu), the patient had a gastrointestinal (gi) bleed. it is unknown how the patient was treated. ten days later, the family elected to withdraw care, and the patient expired on support. the device is unlikely to have contributed to the patient's death, which was most likely due to the clinical condition of a critically ill patient, along with the complications of stage d shock. bleeding is a known risk due to the impella's anticoagulation and purge requirements. bleeding is listed as a potential adverse event, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | The pump has been disposed of and will not be returned for evaluation. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the exact date of death is unknown. if additional information becomes known, a supplemental report will be submitted.
306 | 1220648-2026-03621 | 24397887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397887&pc=OZD | 11/20/2025 |  |  | 02/19/2026 |  |  | 11/20/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713971 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was in the operating room for surgery (unclear as this is not noted). an intra-aortic balloon pump (iabp) was inserted. the patient developed epistaxis prior to impella insertion. an impella 5.5 was inserted, and the patients¿ chest remained open. continuous renal replacement therapy started. the patient went to the or the following day for a chest washout (no tamponade). as the patient was coagulopathic, no heparin was started. blood products were administered due to bleeding post operatively. ultimately, the patient expired on (b)(6) 2025 with no information noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Section d catalog number was corrected.
307 | 1220648-2026-03622 | 24397890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397890&pc=OZD | 11/20/2025 |  |  | 02/19/2026 |  |  | 11/20/2025 | 02/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026665967 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
308 | 1220648-2026-03623 | 24397892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397892&pc=OZD | 11/20/2025 |  |  | 02/19/2026 |  |  | 11/20/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026732628 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Heart Failure/Congestive Heart Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient had coronary artery bypass surgery with a tricuspid valve repair and a mitral valve replacement. an intra-aortic balloon pump (iabp) was placed initially after the completion of the case and was removed on post op day 1. the patient had a drop in his ejection fraction and hemodynamic compromise after iabp removed so an impella 5.5 was placed after multiple pressors and inotropes had already been initiated. ultimately, the patient was placed on va ecmo with impella 5.5 left ventricular unloading. the patient was deemed hit (heparin induced thrombocytopenia) positive with thrombocytopenia. since the patient¿s right ventricle was not moving well on echo, an rp flex was placed on (b)(6) 2025 and was decannulated from ecmo. he did well post procedure and on (b)(6) 2025, both impella devices were removed without incident. the patient was in cardiogenic shock which warranted multiple medications to support his failing hemodynamics. ultimately, the decision was made to place an impella with ecmo and then rp flex to support his biventricular failure. he was hit positive, so thus the reason for his low platelet count. there were no apparent impella device complications. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
309 | 1220648-2026-03625 | 24397895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397895&pc=OZD | 11/21/2025 |  |  | 02/19/2026 |  |  | 11/21/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026763791 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 60-year-old male va ecmo patient was transferred for advanced therapies. ecmo configuration changed to left atrial veno arterial ecmo and attempts were made to wean the ecmo unsuccessfully. so, an impella 5.5 was placed. at some point during his care, he was placed on continuous renal replacement therapy (crrt). the patient required tracheostomy due to inability to come off ventilator an the need for longer term ventilator support. post procedure, he had bleeding at the stoma site requiring blood administration. impella site also bleeding and imaging showed bleeding in the pocket. the patient remains on support as a bridge to either durable left ventricular assist device or heart transplant. the complication noted was bleeding from both the tracheostomy site as well as the impella site which required blood transfusion. this most likely was due to the patient being on anticoagulation and started on plavix the day prior to the bleeding so the plavix was stopped and there was no further bleeding noted. as there is no indication of when the crrt was started, it cannot be attributed directly to the impella pump. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
310 | 1220648-2026-03624 | 24397898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397898&pc=OZD | 11/20/2025 |  |  | 02/19/2026 |  |  | 11/20/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026716019 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | [6 empty] | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that this patient had a mitral valve repair with left atrial appendage clipping on (b)(6) and was unable to wean off bypass, so he was placed on ecmo. as he had issues with bleeding on ecmo and had to return to the or twice for chest exploration and washout, an impella 5.5 was ultimately placed with the goal to wean ecmo and close the chest. placement was difficult as impella did not advance appropriately and required repositioning under echo guidance. ecmo was removed on (b)(6) /2025 and the patient had multiple strokes post decannulation. hematoma was noted at the old ecmo cannulation site and bleeding was occurring at his tracheostomy site. heparin was temporarily held due to bleeding and hematoma issues. the impella 5.5 was removed on (b)(6) 2025 without incident. it appears that this patient had complications secondary to ecmo (hematoma and multiple strokes on removal) and cannot attribute these to the impella device. he did have some bleeding however, this appears to be due to systemic anticoagulation and not directly caused by the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
311 | 1220648-2026-03626 | 24397900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397900&pc=OZD | 11/22/2025 |  |  | 02/19/2026 |  |  | 11/22/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026627369 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | 63-year-old male patient with extensive cardiac history was admitted with acute decompensated heart failure. he had been on home milrinone prior to admission. the patient was admitted for an impella 5.5 implant as a bridge to heart transplant. the patient was in the icu and experienced ventricular tachycardia (vt) requiring multiple shocks. after the multiple shocks, the impella began having issues with the flattening of the ao and lv signals and suction alarms. an echo was performed and no clot was found and impella placement deemed to be good. the care team opted to add epinephrine, and both the ao and lv waveform signals returned. continuous renal replacement therapy was initiated due to acute kidney injury. the patient had ongoing issues with vt and required antiarrhythmic medication. the patient was ultimately implanted with a durable left ventricular assist device, and the impella was removed without incident. it is difficult to determine whether the vt and renal failure were directly related to the impella device or secondary to the patient¿s cardiogenic shock status since admission. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
312 | 1220648-2026-03627 | 24397906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397906&pc=OZD | 11/23/2025 |  |  | 02/19/2026 |  |  | 11/23/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026688383 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | Pulmonary Edema | Renal Failure | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The patient was transferred from an outside hospital intubated, sedated, and supported with an impella cp. hematuria was observed on arrival and was attributed by the clinical team to traumatic foley catheter placement. initial echocardiography showed the impella cp positioned too deeply, and the device was repositioned. the left ventricle was noted to be hypertrophic, and multiple suction alarms occurred on admission. the patient subsequently developed complete heart block, progressed to cardiac arrest, and underwent acls with return of circulation. post-resuscitation, the impella became malpositioned and required repositioning in the catheterization lab. due to renal failure, the patient was started on crrt; chart review indicated a history of chronic kidney disease with an existing hemodialysis catheter and ongoing evaluation for heart¿kidney transplantation. given persistent cardiogenic shock, support was escalated from impella cp to impella 5.5. after placement through a surgical graft, the patient experienced significant bleeding into the device pocket. the patient required multiple blood products, including ffp, platelets, prbcs, and cell saver blood. vascular surgery intervened, placing a stitch at the graft anastomosis, which resolved the bleeding. the impella 5.5 was repositioned after placement. the patient was extubated on (b)(6) 2025 but developed pulmonary edema and required re-intubation. crrt was later discontinued, and hemodialysis initiated. additional blood products were administered for recurrent anemia. the patient continued evaluation for durable lvad therapy and was successfully weaned from impella support on (b)(6) 2025. this case involved multiple adverse events, many attributed to the patient¿s severe cardiogenic shock and chronic renal disease. suction events were likely related to a small, hypertrophic ventricle. pulmonary edema after extubation was assessed as unrelated to the impella device. the bleeding at the impella 5.5 insertion site and surgical pocket was the adverse event most directly associated with device placement, requiring surgical repair and significant transfusion support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
313 | 1220648-2026-03628 | 24397909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397909&pc=OZD | 11/24/2025 |  |  | 02/19/2026 |  |  | 11/24/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026706531 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 102 KG | Non Hispanic | White | Death | Yes | Yes | Difficult to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received. buddy wires were used but to no affect. the patient position was also manipulated but it did not help. the device was not able to be returned. | Clinical assessment: sixty-one-year-old male admitted with acute myocardial infarction cardiogenic shock. the patient was taken to the operating room (or) for impella 5.5 placement with veno-arterial extracorporeal membrane oxygenation (va ecmo) already in place. the physician had difficulty getting the impella to pass through the innominate artery due to calcium buildup so interventional cardiology shockwave¿ d the vessel and the device was easily placed. the following day, the patient returned to the operation room for chest washout and drain replacement. ecmo configuration upgraded and impella 5.5 continued post operatively. there was ultimately no resolution of neurological status so the family opted to have care withdrawn. there is no impella-related event as the inability to pass the impella catheter was due to calcium in the patient¿s vessel. unable to equate chest washout to impella related complication, and death was the result of the family decision to withdraw care. engineering assessment: during impella 5.5 insertion, the pump was difficult to advance. clinical described that the physician had difficulty getting the impella in place as the 5.5 initially could not pass through the innominate artery due to calcium buildup. as a result, interventional cardiology shock waved the vessel, after which the pump was able to pass without issue. the difficult to advance is considered a reportable malfunction. |  |  |  |  | B5. additional information was received and added. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
314 | 1220648-2026-03630 | 24397915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397915&pc=OZD | 11/24/2025 |  |  | 02/19/2026 |  |  | 11/24/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026736378 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical assessment: a sixty-six-year-old patient was admitted with decompensated heart failure (hf). inotropes were initiated, however, the patient had ventricular tachycardic arrest and the decision was made to take the patient to the or for placement of impella 5.5. no issues during pump placement. the patient was supported with impella 5.5 while waiting for heart transplantation. during the intensive care unit (icu) stay, the placement sensor began failing with faulty aortic placement signal triggering suction alarms. the p level was reduced in attempts to correct the issue however, the placement alarms continued and were ultimately silenced. the placement signal alarm was disabled throughout the remainder of the patient¿s icu stay until (b)(6) 2025 when the patient went for heart transplant with removal of impella device. engineering assessment: during impella 5.5 support, low placement signal was noted compared against arterial line reference pressures, causing placement signal low alarm and false suction alarms. patient support continued without issue until the patient received a heart transplant. there were no adverse events related to this device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
315 | 1220648-2026-03634 | 24397924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397924&pc=OZD | 11/24/2025 |  |  | 02/19/2026 |  |  | 11/24/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026749871 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Tachycardia | Hematuria | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A sixty-four-year-old male was admitted with acute myocardial infarction related cardiogenic shock and initially supported with an impella 5.5, which was removed on (b)(6) 2025. shortly after explant, he developed ventricular tachycardia (vt) arrest requiring cpr and multiple shocks, and an impella 5.5 was reinserted for recurrent cardiogenic shock. following reinsertion, the patient experienced another vt arrest treated with a single defibrillation. echocardiography showed the impella positioned 7 cm deep, it was retracted to 4.7 cm. inotropic therapy was adjusted, and magnesium and antiarrhythmics were administered for intermittent vt. the patient also required continuous renal replacement therapy (timing relative to impella placement unknown). he subsequently underwent percutaneous coronary intervention (pci) and received 1 unit packed red blood cells (prbcs) for declining hemoglobin/hematocrit (h/h). mild hematuria and melena were noted, prompting discontinuation of heparin. with continued h/h drop, patient received an additional 3 units prbc. renal function gradually improved with return of urine output. the impella 5.5 was removed without complication on (b)(6) 2025. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
316 | 1220648-2026-03635 | 24397929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397929&pc=OZD | 11/24/2025 |  |  | 02/19/2026 |  |  | 11/24/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026768607 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A fifty-nine-year-old patient was admitted for planned heart catheterization. post procedure, the patient developed bradycardia and cardiac arrested. she was cannulated with veno-arterial extracorporeal membrane oxygenation (va ecmo) and transported to the cardiovascular operating room (cvor) for placement of an impella 5.5. the device smoothly transitioned past the anastomosis and innominate/subclavian junction into the ascending aorta without complication however, it was noted that there was resistance within the ascending aorta and the procedure was aborted. the physician opted to place an impella cp instead for left ventricular unloading on ecmo. the patient underwent embolization of the splenic artery due to a bleed and subsequent thrombus from cardiopulmonary resuscitation (cpr) injury. the patient underwent coronary artery bypass surgery, and the impella cp was removed on (b)(6) 2025. the thrombus noted in the splenic artery appears to be secondary to active cpr and not associated with the impella pump. no other clinical information was given. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
317 | 1220648-2026-03636 | 24397930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397930&pc=OZD | 11/24/2025 |  |  | 02/19/2026 |  |  | 11/24/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026710430 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A sixty-year-old male was admitted for coronary artery bypass surgery and elective impella 5.5 placement. during the procedure, it was noted that there was perforation of the left ventricular wall. the impella device was removed and the perforation was closed. once closure was confirmed, the impella 5.5 was re-inserted through the open aortotomy. the ventricle was noted to be grossly diseased and fragile per the surgeon. there was no clinical information provided and the ventricular perforation was secondary to the surgeon implant process so unable to attribute to the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
318 | 1220648-2026-03637 | 24397931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24397931&pc=OZD | 11/24/2025 |  |  | 02/19/2026 |  |  | 11/24/2025 | 02/19/2026 |  |  |  |  | 02/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745726 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A twenty-six-year-old female patient was admitted with cardiomyopathy. an impella 5.5 was placed as a bridge to durable left ventricular assist device (lvad). on (b)(6)2025, the patient cardiac arrested and was placed on veno-arterial extracorporeal membrane oxygenation (va ecmo). some suction issues with impella once ecmo was initiated. no further information was provided for patient so unsure whether patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
319 | 1220648-2026-03391 | 24376288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24376288&pc=OZD | 02/10/2026 |  |  | 02/18/2026 |  |  | 02/10/2026 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Priming Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
320 | 1220648-2026-03401 | 24376614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24376614&pc=OZD | 02/11/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1431841 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an automated impella controller generated a failure alarm upon being turned on. there was no patient involved in the reported event. | [5 empty] | There was no patient involved in the reported event. section a was left blank. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
321 | 1220648-2026-03381 | 24376648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24376648&pc=OZD | 10/16/2025 |  |  | 02/18/2026 |  |  | 10/16/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713132 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Low Blood Pressure/ Hypotension | Pulmonary Edema | Thrombocytopenia | Heart Failure/Congestive Heart Failure | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Patient presented to er with non¿st-segment elevation myocardial infarction (nstemi). taken to cath lab. during cath, patient became symptomatic, bradycardic, and atropine was given. impella cp placed. echo imaging showed cp deep and physician repositioned. dobutamine started due to worsening clinical status despite impella cp. tachycardia secondary to higher dosing of dobutamine. packed red blood cells given. (b)(6) 2025, the patient went to the or for change to impella 5.5 for additional hemodynamic support. transesophageal echocardiogram (tee) demonstrated possible clot formation around impella cp catheter in the descending aorta and was confirmed once the cp was removed. vascular was consulted performed thrombectomy prior to impella cp removal and impella 5.5 insertion. there was a brief episode of hypotension leading to asystole. cardiopulmonary resuscitation (cpr) was performed briefly and successful return of spontaneous circulation was achieved. flows and waveforms stable on impella 5.5. with improving hemodynamics. (b)(6) 2025 the patient went back to cath lab for stent placement. intermittent suction alarms occurring secondary to patient position. heparin-induced thrombocytopenia positive (hit+). for left ventricular assist device (lvad) workup. plasmapheresis being done prior to possible transplant. became acidotic and hypotensive requiring additional inotropes. worsening pulmonary edema present on chest xray. diuresing with goal of lvad surgery, which was done on (b)(6) 2025. impella removed at that time. patient was symptomatic (hypotensive and bradycardic) prior to initial impella cp placement. no definitive symptoms secondary to impella cp. impella 5.5 replaced impella cp with improvement in hemodynamics. intermittent suction alarms due to patient movement. the patient¿s symptoms of pulmonary edema are likely not related to the impella pump. hypotension is likely due to aggressive diuresis. no clinical impact of impella 5.5 pump. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
322 | 1220648-2026-03402 | 24376676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24376676&pc=OZD | 02/10/2026 |  |  | 02/18/2026 |  |  | 02/10/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026655834 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that during explantation of an impella 5.5 the inlet was broken off the cannula. explantation was completed successfully, and no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
323 | 1220648-2026-03403 | 24376696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24376696&pc=OZD | 02/11/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026757675 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
324 | 1220648-2026-03405 | 24376793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24376793&pc=OZD | 02/11/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801122 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 93 KG |  | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced a rip of the proximal end of the sterile sleeve. the sleeve was cleaned and reinforced with tegaderm. no patient harm was reported. | [5 empty] | A5 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
325 | 1220648-2026-03407 | 24376816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24376816&pc=OZD | 02/11/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026803212 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
326 | 1220648-2026-03409 | 24376896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24376896&pc=OZD | 02/06/2026 |  |  | 02/18/2026 |  |  | 02/10/2026 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Device Sensing Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
327 | 1220648-2026-03411 | 24376959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24376959&pc=OZD | 10/22/2025 |  |  | 02/18/2026 |  |  | 10/22/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026635374 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 71 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 68-year-old male with ami cardiogenic shock (scai e) underwent high-risk pci with periprocedural impella cp and ecmo support. following the procedure, ecmo and impella cp support were transitioned to impella 5.5. access-related bleeding occurred at the graft/impella insertion site. hemostasis was achieved in the icu following correction of coagulopathy and application of approximately 10 minutes of manual compression. bleeding from ecmo cannulation sites also decreased with ongoing management. limited additional clinical details were available. however, an ischemic stroke was reported on day 2 of impella support. ongoing thrombocytopenia prompted a change in anticoagulation regimen and initiation of a diagnostic workup for suspected hit. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
328 | 1220648-2026-03412 | 24376960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24376960&pc=OZD | 02/10/2026 |  |  | 02/18/2026 |  |  | 02/10/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1123074 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 82 KG | Non Hispanic | Asian | Injury | Yes | No | Application Program Problem | Device Sensing Problem | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in a 60-year-old male patient presenting with cardiomyopathy, with a past medical history significant for coronary artery disease (cad), prior cabg, and renal insufficiency. during support, the console displayed a controller error with a yellow alarm and loss of placement signal on console ic2081. the placement signal subsequently returned, the alarm resolved, and the console was exchanged for another unit without issue. the reported events include controller failure, placement signal issue, device revision or replacement, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and hemodynamic support was maintained. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
329 | 1220648-2026-03416 | 24377048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24377048&pc=OZD | 02/11/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026812357 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female | 66 KG | Non Hispanic | White | Injury | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in an 84-year-old female patient presenting for high-risk percutaneous coronary intervention (hrpci), with a history of coronary artery disease (cad). during the attempted insertion, the 5.5 catheter was unable to be advanced through the innominate artery. the reported events include failure to advance, medical device removal, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the device removal; however, the removal was performed with no consequence to the patient, and hemodynamic support was maintained. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
330 | 1220648-2026-03417 | 24377067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24377067&pc=OZD | 02/12/2026 |  |  | 02/18/2026 |  |  | 02/12/2026 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026814985 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
331 | 1220648-2026-03418 | 24377107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24377107&pc=OZD | 02/12/2026 |  |  | 02/18/2026 |  |  | 02/12/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026747714 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 87 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Increase in Pressure | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in a 71-year-old male patient presenting with ami/cgs, with a medical history significant for coronary artery disease and renal insufficiency. overnight, the purge flow decreased to <3 ml/hour, prompting administration of tpa per hospital protocol. the patient was already on systemic anticoagulation and was transitioned to systemic bivalirudin due to concern for potential hit. later that morning, a purge system blocked alarm was triggered. the local team and csc were notified, and the patient was weaned to p2 under echocardiographic guidance. it could not be determined whether thrombus was present at the inflow or outflow. left ventricular function had improved, and the patient no longer required inotropic support. a decision was made to explant the device with cardiothoracic surgery in the cvor. during induction of anesthesia, the patient required systemic epinephrine and milrinone infusions. the impella catheter was removed without complication, and no clot was visualized on the cannula or inflow/outflow segments. the patient remained hemodynamically stable throughout the procedure and was returned to the cvicu for ongoing management. the reported events include high purge pressure, medical device removal, medication administration, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the removal; however, the removal was completed without any adverse consequence to the patient, and support was appropriately maintained. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
332 | 1220648-2026-03423 | 24377222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24377222&pc=OZD | 02/11/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026753182 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 70 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Malposition of Device | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 pump was being placed via the right axillary artery graft site to support the 72 year old female patient who had been admitted in with known cardiomyopathy presenting in scai stage c shock. other underlying history, beyond the patient being on inotropes and vasopressors, was not shared. the pump was being inserted via the graft site when there was an observation made of vessel dissection. the pump was removed, however the guidewire remained within the axillary to allow for vessel treatment in the true lumen. vascular stented the dissection, however afterwards the 5.5 failed to deliver through the stented anatomy, and the pump was removed as delivery failed. the patient was noted be stable after the pump removal. vascular injury is a known risk associated with impella support as described in the instructions for use. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
333 | 1220648-2026-03429 | 24377403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24377403&pc=OZD | 02/11/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026785111 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
334 | 1220648-2026-03424 | 24377425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24377425&pc=OZD | 02/08/2026 | 02/18/2026 |  | 02/18/2026 |  |  | 02/08/2026 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026751657 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
335 | 1220648-2026-03432 | 24377556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24377556&pc=OZD | 02/11/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795369 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 66 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 68-year-old female patient presenting with cardiomyopathy and a history of coronary artery disease (cad). during attempts to advance the device, the clinical team tried to optimize the access pathway by repositioning the patient¿s arm to straighten the vasculature, utilizing an amplatz .025 wire, and manually stabilizing the graft near the anastomosis to provide additional leverage. the graft was also tunneled laterally by approximately 2 cm in an effort to improve the alignment for insertion. despite these measures, the device was unable to be advanced past the clavicle due to the patient¿s vascular anatomy. the team subsequently decided to abort the attempt and proceed with placement of an impella cp. the reported events include failure to advance, product damage, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was continued without any known adverse events. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella 5.5 (pump 3) device report is one of four pumps used during the event. only two of the four pumps generated reportable complaints. please refer to the related manufacturer report number listed in section h10 for the corresponding medical device report on the impella 5.5 (pump 2). | This report is being submitted to add an additional h6 medical device problem code that was unintentionally omitted from the initial report. h6 medical device problem code a040601 ¿ deformation due to compressive stress has been added. no new adverse event information is being provided in this submission.
336 | 1220648-2026-03439 | 24378042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24378042&pc=OZD | 02/03/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026731767 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted into the left femoral artery in a 62-year-old male patient presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), an out-of-hospital cardiac arrest requiring cardiopulmonary resuscitation (cpr), on an extracorporeal membrane oxygenation (ecmo), on a ventilator for respiratory support, and in scai stage c shock, prior to initiation of support. the impella was inserted for ventricular support for a high-risk percutaneous coronary intervention (pci). two weeks later, the impella was successfully weaned. the impella functioned as intended. the post-procedure outcome is survived at explant. one week after removal, a follow-up computed tomography (ct) scan revealed a type a dissection. the patient hemodynamics were stable. the causal relationship is unknown. | [5 empty] | A4 weight is unknown. device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
337 | 1220648-2026-03443 | 24378586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24378586&pc=OZD | 11/14/2025 |  |  | 02/18/2026 |  |  | 11/14/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745436 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 68 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 55-year-old female with history of hopkins lymphoma with radiation to chest, mitral regurgitation and aortic insufficiency admitted through clinic for hf services due to decompensated heart failure after beta blocker dose. taken to cath lab with attempted impella 5.5 placement via axillary artery. unable to get the impella to pass through the ostia of the axillary artery. multiple attempts were made with wire changes and ultimately kinked cannula. second impella 5.5 attempted with same issues to pass through the axillary artery. found to have stenosis at the ostia of the axillary artery and the aorta. due to the need for mechanical circulatory support (mcs), the care team elected to place a cp through axillary artery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
338 | 1220648-2026-03444 | 24378600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24378600&pc=OZD | 11/15/2025 |  |  | 02/18/2026 |  |  | 11/15/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745681 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 55-year-old female patient with decompensated heart failure, cardiogenic shock. the initial attempt to insert the 5.5 was unsuccessful due to axillary artery stenosis and further attempts caused catheter to kink near the motor housing. a second 5.5 was attempt but was also unsuccessful due to axillary artery stenosis. the patient ended up receiving an impella cp to accommodate the vascular anatomy. the impella cp was inserted via axillary graft. once the pump was in the left ventricle, the pump was turned on. anesthesia began seeing air bubble blouses about every 3-5 minutes. this was significant enough to suggest air introduction and cause patient to hemodynamically respond, as if air went down coronaries. the support team ran through the algorithm-settling on cavitation. this will be conservatively reported as an air embolism for the impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
339 | 1220648-2026-03447 | 24378627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24378627&pc=OZD | 12/02/2025 |  |  | 02/18/2026 |  |  | 12/02/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026768597 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Inaccurate Flow Rate | [6 empty] | Tachycardia | Hypovolemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 59-year-old male patient was admitted for acute mi cardiogenic shock. the patient was on va ecmo, and an impella cp was added. he underwent high risk pci with stent placement. had long ventricular tachycardia runs which were shocked by his implanted defibrillator. he did have some suction alarms necessitating the volume administration. an echo was done which confirmed good placement. the patient went to the or on (b)(6) 2025 and had the cp upgraded to a 5.5 with continued ecmo support. waveforms appeared unusual with lv waveforms elevated and flows falsely high. the pump had sudden higher motor current number and flows went up to 4.3 on p4. the 5.5 was removed and the surgeon removed white fibrin stretchy clot from the motor housing. the impella was replaced with a new 5.5 and flows more appropriate. no further information on this patient noted. the tachycardia could have been from the impella device; however, this patient was also admitted in cardiogenic shock and was also on concomitant ecmo so unable to confirm. the white fibrin around the motor housing could have been from old clot, lower anticoagulation (none documented) or from ecmo sluggish circulation so unable to attribute this to the impella device. during impella 5.5 support, falsely high 'saturated' pump flow was reported. the pump was removed replaced with a second impella 5.5 and support continued. on explant, fibrous clot was observed at the motor housing of the removed pump. this event is considered reportable per sop-ra04-f002. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
340 | 1220648-2026-03451 | 24378750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24378750&pc=OZD | 02/11/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796213 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinician narrative: an impella 5.5 was inserted via the right axillary artery in a 50-year-old male who presented with cardiomyopathy and acute decompensated heart failure. the patient had a chronic history of diabetes mellitus and was in cardiogenic shock, scai stage d. he required inotropic support, vasopressors, and mechanical ventilation for respiratory failure. an intra-aortic balloon pump (iabp) was also in place for circulatory support. the patient was taken to the operating room for axillary impella 5.5 placement. during advancement of the device, it was unable to pass the anastomosis. the device was removed for further evaluation. intraoperative assessment identified an axillary artery dissection. vascular surgery was consulted, and the affected artery was surgically repaired. due to inability to advance the impella 5.5 and the vascular complication encountered, the device was removed and additional circulatory support was required. the patient remained on iabp support post-procedure and was transferred to the icu. the reported patient effects included axillary artery dissection, surgical vascular intervention, medical device removal, and requirement for additional circulatory support. given the patient¿s acute decompensated cardiomyopathy, underlying diabetes, cardiogenic shock (scai stage d), and need for multiple vasoactive agents and ventilatory support, the vascular complication occurred in the setting of significant hemodynamic instability and high-risk vascular access. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
341 | 1220648-2026-03453 | 24378937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24378937&pc=OZD | 12/02/2025 |  |  | 02/18/2026 |  |  | 12/02/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026743378 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 90 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A seventy-year-old male patient with a history of coronary artery disease, metastatic renal cell cancer, and chronic renal insufficiency received an impella cp smartassist device for management of acute myocardial infarction¿related cardiogenic shock. the patient had received cardiopulmonary resuscitation prior to device placement and presented in advanced cardiogenic shock consistent with stage e classification. the patient was receiving three inotropic medications and mechanical ventilation at the time of implantation. on the first day after device insertion, the patient developed oozing and a groin hematoma at the access site, which was managed using a femstop device. the following day, the patient was escalated to an impella 5.5 device, and the impella cp device was removed. on the sixth day of support, the patient experienced mild hypotension following administration of furosemide. due to significant comorbidities, multiple risk factors, and progressive clinical deterioration, the patient¿s condition continued to worsen, and the family elected for palliative management. on the ninth day of support, the patient developed recurrent ventricular tachycardia and ventricular fibrillation that did not respond to multiple defibrillation shocks, and per the family¿s wishes, life-sustaining care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
342 | 1220648-2026-03454 | 24378976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24378976&pc=OZD | 12/01/2025 |  |  | 02/18/2026 |  |  | 12/02/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026769377 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 84 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A fifty-six-year-old male patient with acute heart failure¿related cardiogenic shock secondary to ischemic cardiomyopathy received an impella cp device. prior to insertion, the patient was receiving two inotropic medications, veno arterial extracorporeal membrane oxygenation, and mechanical ventilation, consistent with stage d cardiogenic shock. following impella placement, the patient demonstrated gradual hemodynamic improvement and was evaluated for potential heart transplantation. due to the anticipated need for prolonged support, the patient was escalated to an impella 5.5 device and transitioned from femoral extracorporeal membrane oxygenation to axillary extracorporeal membrane oxygenation on the ninth day. on the twenty fifth day of support, the patient experienced an ischemic stroke that was managed with thrombectomy. after twenty-nine days of impella support, the existing impella 5.5 was exchanged for a new device, and the patient remained ambulatory with notable clinical improvement while awaiting transplant. on the fortieth day, the patient developed a second ischemic stroke but was not a candidate for thrombectomy due to thrombocytopenia, and hemodynamics remained stable. on the forty third day, the patient experienced a third stroke; no intervention was performed, and the heart team consulted palliative care due to the patient no longer being a candidate for durable left ventricular assist device therapy without neurological recovery. the patient was intubated and reinitiated on extracorporeal membrane oxygenation. after fifty-nine total days of mechanical support and given the severity of illness, life sustaining care was withdrawn and the impella device was removed and the patient ultimately expired. cp to 5.5 escalation. 27 days on 5.5 support. purge pressure issue, (b)(4), attributable to cp prior to exchange. ppae on cp - stroke. subsequent strokes on 5.5 3x. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
343 | 1220648-2026-03457 | 24379052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24379052&pc=OZD | 12/02/2025 |  |  | 02/18/2026 |  |  | 12/02/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026725268 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 103 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Renal Failure | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A fifty three year old male patient received an impella cp device for acute myocardial infarction¿related cardiogenic shock. the patient was receiving two inotropic medications prior to impella placement. due to persistent need for inotropes and vasopressors and ongoing profound shock, the patient was escalated to an impella 5.5 device on the fifth day and continued on support in the intensive care unit. immediately after impella 5.5 insertion, the surgeon also cannulated the patient for venoarterial extracorporeal membrane oxygenation, and continuous renal replacement therapy was initiated the following day. the patient¿s clinical condition gradually improved over the next several days, and inotropic requirements decreased. on the seventeenth day of support, the patient developed leukocytosis, intermittent fever, and low systemic vascular resistance concerning for possible sepsis; the condition responded to treatment over the next few days with subsequent improvement in hemodynamics. the sepsis might have been caused by combination of therapies for a longer duration. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
344 | 1220648-2026-03458 | 24379142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24379142&pc=OZD | 02/12/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1189530 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female | 70 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted into the right femoral artery in a 51-year-old female patient with a past medical history of a prior coronary artery bypass graft (cabg), coronary artery disease (cad), and diabetes, presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), and in scai stage d shock, on multiple inotropes and vasopressors, and was on a ventilator for respiratory support, prior to initiation of support. while the patient was in the intensive care unit (icu), the impella was experiencing placement signal low alarms. an echocardiogram was performed which showed the impella measuring at 4.7cm deep in the left ventricle (lv). the impella was pulled back to 3.8cm. pink-tinged urine was noted, which cleared after repositioning. the flows improved and the alarm resolved. the following day, there was a controller error alarm. the pulmonary artery (pa) waveform disappeared intermittently, then came back and resolved. flows adjusted to p-8 at 3.0l/min. the patient had been coughing and fighting the ventilator. there was no evidence of clot. advised not to change the console due to the risk of the back-up controller not recognizing the catheter. the post-procedure outcome is unknown. poor positioning of the impella can cause hemolysis. hemolysis is listed as a potential adverse event, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this automated impella controller device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella rp flex and the impella cp.
345 | 1220648-2026-03461 | 24379156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24379156&pc=OZD | 12/02/2025 |  |  | 02/18/2026 |  |  | 12/03/2025 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026617711 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Low Blood Pressure/ Hypotension | Hypoxia | Renal Failure | Sepsis | Respiratory Failure | Heart Failure/Congestive Heart Failure | [31 empty] | Invalid/Incomplete |  |  |  | Initial submission
346 | 1220648-2026-03465 | 24379258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24379258&pc=OZD | 02/08/2026 |  |  | 02/18/2026 |  |  | 02/08/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775462 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Failure | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was placed via the axillary artery graft to support the 64 year old male patient. the patient was placed on the 5.5 to allow the high risk bypass surgery (cabg) to occur in the operating suite. the patient was in scai stage c shock, but any other underlying medical history was not shared. the pump was placed however the positioning within the left ventricle was suboptimal. the pump was repositioned and flows were low, but the pump remained in that position. the patient began to suffer from ventricular tachycardia and persistent bigeminy. again, pump position was assessed and repositioning attempts failed, so after 2 days, the pump was explanted and replace by an intra-aortic balloon pump (iabp). of note, the original positioning was done by a jr4 coronary catheter, and due to this the team was educated as this may have contributed to the positional issues. the positioning issues and pump removal and replacement by an iabp will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange, however the exchange was performed with no consequence to the patient and support. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
347 | 1220648-2026-03466 | 24379348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24379348&pc=OZD | 12/03/2025 |  |  | 02/18/2026 |  |  | 12/03/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765210 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 54-year-old male patient was admitted with acute mi, cardiogenic shock. he was taken to the cath lab and was cannulated with va ecmo and impella cp. the patient developed klebsiella bacteremia and cellulitis in bilateral groins. platelets were given for low platelet count. the patient was converted to an impella 5.5 via axillary artery on (b)(6) 2025. he developed some runs of ventricular tachycardia post procedure. echo was done that showed a small left ventricular cavity. the device was repositioned. ultimately, the patient was bridged to a durable ventricular assist device and the impella 5.5 was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
348 | 1220648-2026-03467 | 24379460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24379460&pc=OZD | 12/03/2025 |  |  | 02/18/2026 |  |  | 12/03/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026772348 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 69-year-old male patient with a history of coronary artery disease received an impella 5.5 device for management of acute myocardial infarction complicated by cardiogenic shock as a bridge to heart transplantation or durable left ventricular assist device therapy. prior to escalation to the impella 5.5 device, the patient was supported with three inotropic medications and an intra-aortic balloon pump. following impella placement, the patient demonstrated reduced inotropic requirements and showed signs of hemodynamic improvement. on the fourth day of support, the patient experienced a mild episode of epistaxis (nosebleed), which resolved after adjustment of anticoagulation dosing. after 29 days of impella 5.5 support, the patient was successfully bridged to heart transplantation as originally planned. 29 days on support. nosebleeds day 4-5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
349 | 1220648-2026-03469 | 24379524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24379524&pc=OZD | 12/04/2025 |  |  | 02/18/2026 |  |  | 12/04/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1302710 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | A 43-year-old male patient was admitted with acute decompensated heart failure and was implanted with an impella 5.5. he was also on va ecmo. the ecmo was discontinued on (b)(6) 2025 and the impella support continued. the patient was supported with multiple vasopressors and inotropes for his cardiogenic shock (cs). there was a small amount of oozing from the axillary site noted. impella support continues as of (b)(6) 2026. on (b)(6) 2026, there were purge cassette issues warranting a cassette change out and ultimately a new aic. no patient consequences were noted. oozing from the impella site from the impella device. during impella 5.5 support, blood was noted oozing was noted at the axillary access site. 42 days later, "purge disc not detected" alarms occurred when trying to replace a purge cassette. the console was exchanged, the purge disc of the cassette was detected, and support continued. the patient is still on support 30 days later as of (b)(6) 2026. these are considered reportable events per (b)(4). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
350 | 1220648-2026-03471 | 24379634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24379634&pc=OZD | 12/04/2025 |  |  | 02/18/2026 |  |  | 12/04/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026753239 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 102 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | Difficult to Advance | [5 empty] | Stroke/CVA | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 49-year-old male patient with multiple comorbidities and a prior history of heart transplantation received an impella 5.5 device for management of heart failure with cardiogenic shock following transplantation. the patient had previously been supported with an impella cp device before escalation to the impella 5.5 device. at the time of impella 5.5 implantation, the patient was classified as stage e shock according to the society for cardiovascular angiography and interventions scale, indicating severe cardiogenic shock. on the thirty-third day of impella 5.5 support, the patient developed a hemorrhagic stroke. throughout the course of care, the patient was treated with multiple medications, including inotropic agents, vasopressors, immunosuppressive therapy, and anticoagulation. the patient was also supported with multiple devices, including extracorporeal membrane oxygenation, continuous renal replacement therapy, a right ventricular assist device, and mechanical ventilation, reflecting the extreme severity of illness. the hemorrhagic stroke may have been associated with the use of anticoagulation medications. given the patient's multiple comorbidities, the absence of an available donor for repeat heart transplantation, the occurrence of stroke, and the overall complexity and severity of the clinical condition, the family elected to withdraw care on 37th day of support. death was conservatively coded to the impella 5.5 while most likely to be caused by the underlying disease and unrelated to impella. 37 days on support. difficult to advance during implant day 0, suction day 13 (likely due to ecmo), stroke day 33. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
351 | 1220648-2026-03476 | 24379866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24379866&pc=OZD | 12/04/2025 |  |  | 02/18/2026 |  |  | 12/04/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025547824 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 13 YR |  | Female | 60 KG | Non Hispanic |  | Injury | Yes | No | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A thirteen-year-old female patient received an impella 5.5 device for treatment of heart-failure¿related cardiogenic shock secondary to acute myocarditis. prior to impella placement, the patient was receiving multiple inotropic medications, vasopressor therapy and veno-arterial extracorporeal membrane oxygenation. the patient was brought to the operating room for extracorporeal membrane oxygenation decannulation, during which the impella 5.5 device was inserted. during initial setup, the aortic pressure sensor displayed values significantly higher than the patient¿s arterial line tracing; however, the device was successfully positioned and appeared to be functioning appropriately. following extracorporeal membrane oxygenation decannulation, device flow decreased. multiple troubleshooting attempts, including repositioning and imaging evaluation, were unsuccessful. because adequate flow could not be restored and the patient was no longer supported by extracorporeal membrane oxygenation, the clinical team elected to replace the device. a new impella 5.5 device was successfully implanted, and the original device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
352 | 1220648-2026-03478 | 24380029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24380029&pc=OZD | 02/09/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1429644 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that while the impella device was in use, the automated impella controller displayed a controller error. despite the error message, the controller did not require replacement, and support was successfully continued using the same device. no additional event details were provided. there were no reported signs or symptoms of patient injury, and no harm was associated with the event. | [5 empty] | A. patient information is unknown. device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
353 | 1220648-2026-03479 | 24380111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24380111&pc=OZD | 10/06/2025 |  |  | 02/18/2026 |  |  | 10/06/2025 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1556637 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
354 | 1220648-2026-03485 | 24380505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24380505&pc=OZD | 12/04/2025 |  |  | 02/18/2026 |  |  | 12/04/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026677218 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 65-year-old male patient was scheduled for a coronary artery bypass grafting surgery with impella support. the patient progressed well without clinical issues and was explanted on (b)(6) 2025 without incident. there was an issue with the aortic placement signal waveform which caused intermittent impella unknown alarms despite good pulsatility; however, this was of no consequence to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
355 | 1220648-2026-03487 | 24380577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24380577&pc=OZD | 12/05/2025 |  |  | 02/18/2026 |  |  | 12/05/2025 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026671967 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
356 | 1220648-2026-03486 | 24380636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24380636&pc=PYX | 10/04/2025 |  |  | 02/18/2026 |  |  | 10/07/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP: |  |  | IMPELLA PUMP, RP-IJ, STERILE |  |  | PYX |  |  | IMPELLA PUMP, RP-IJ, STERILE |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1928005 |  | 1000186 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male | 133 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Crack | Fluid/Blood Leak | Increase in Pressure |  |  |  |  | Stroke/CVA | Pericardial Effusion | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Medical device report (mdr). patient information: age: 38 years, sex: male. relevant history: heart failure with reduced ejection fraction, atrial fibrillation, s/p aicd, endocarditis, s/p mitral valve replacement, lvad implantation. event description: the patient underwent left ventricular assist device (lvad) implantation and developed scai stage e cardiogenic shock on postoperative day 1. a transesophageal echocardiogram (tee) revealed hypokinetic right ventricle with rv failure, prompting the heart team decision to implant impella rp flex for right-sided mechanical circulatory support. the following day, the patient returned to the operating room for sternotomy closure. the surgeon was unable to close the chest and placed the patient on extracorporeal membrane oxygenation (ecmo) support. on the next day, the impella rp flex purge system blocked alarm occurred. the patient was given tissue plasminogen activator (tpa) through the purge, which resolved the alarm. two days later, ct of the brain revealed global hypoxic-ischemic injury. the following day, the patient returned to the operating room for chest closure. a pericardial drain was placed post sternal closure with drainage of 300 cc/hour. that evening, the automated impella controller displayed high purge pressure and low purge flow alarms. inspection revealed a crack in the purge side arm. a purge side arm bypass was performed using a 20-gauge iv catheter, restoring normal purge pressure and flow. the patient remained on impella rp flex, ecmo, and lvad support, along with high-dose inotropic/pressor therapy, but expired while on full support. the patient was supported with impella rp flex for 183.76 days. impella rp flex will be conservatively reported for death; however, the device is unlikely to have contributed to the patient's death, which was most likely due to the underlying critical clinical condition. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
357 | 1220648-2026-03489 | 24380690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24380690&pc=OZD | 02/11/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026807964 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 115 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinician narrative an impella 5.5 was inserted via direct aortic access in a 48-year-old male who underwent elective high-risk aortic valve surgery. the patient had known coronary artery disease and required inotropic support, vasopressors, and mechanical ventilation for respiratory support postoperatively. he was classified as scai stage c cardiogenic shock. less than 48 hours postoperatively, while sitting in a chair and conversing, the patient experienced a sudden pulseless electrical activity (pea) arrest. resuscitative efforts were initiated; however, the patient could not be stabilized and expired. the clinical team did not believe the cardiac arrest to be related to impella function. the reported patient effects included cardiac arrest and death. the event occurred in the setting of recent high-risk cardiac surgery, underlying coronary artery disease, vasoactive support, and postoperative hemodynamic vulnerability. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's presenting native critical condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
358 | 1220648-2026-03491 | 24380750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24380750&pc=OZD | 12/08/2025 |  |  | 02/18/2026 |  |  | 12/08/2025 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1417971 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Operating System Becomes Nonfunctional | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
359 | 1220648-2026-03492 | 24380765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24380765&pc=OZD | 02/10/2026 |  |  | 02/18/2026 |  |  | 02/11/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1165516 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | Yes | Display or Visual Feedback Problem | Failure to Read Input Signal | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted via the right femoral artery in a 65-year-old male patient presenting for high-risk pci (hrpci) with a history of coronary artery disease (cad). post-procedure, staff reported that console (b)(6) was pulled for the case, but upon initiation it displayed an error related to the purge cassette/disk. the clinical team proceeded with support using a different aic console. later in the day, when the original console was powered back on, the start-up screen was described as ¿acting weird¿ and progressing ¿very slowly.¿ the site contacted us the following day, and they were instructed to remove the console from service, and a complaint was entered. the reported events include failed-to-detect purge cassette, display issue, medical device removal and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
360 | 1220648-2026-03495 | 24380916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24380916&pc=OZD | 02/12/2026 |  |  | 02/18/2026 |  |  | 02/12/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026710601 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 89 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 pump was placed via the axillary artery graft site for the 58 year old male patient who had been admitted in with acute myocardial infarction and cardiogenic shock, presenting in scai stage e shock. the patient had previously been on an intra-aortic balloon pump (iabp) and inotropes, vasopressors, and a vent for respiratory needs. any other underlying medical history and comorbidities were not shared. the pump supports was ongoing when on day 20 of the support the patient decompensated, had cpr done, but when no return of spontaneous circulation (rosc) occurred, the patient expired. the patient was not a candidate for any other advanced therapies like transplant or an lvad. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's presenting native critical condition as presented in scai stage e. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
361 | 1220648-2026-03500 | 24381264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24381264&pc=OZD | 12/03/2025 |  |  | 02/18/2026 |  |  | 12/03/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744145 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Atrial Fibrillation | Cardiac Arrest | Unspecified Infection | Tachycardia | Hematuria | [8 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 78-year-old male patient was admitted for an elective impella placement to support his coronary artery bypass grafting surgery (cabg). he received multiple blood products intraoperatively. he was hemodynamically unstable with a cardiac index of 1.3 and hypotensive with mean arterial pressures in the low 60¿s. rv function was poor on echo. the patient developed hematuria. he did cardiac arrest secondary to his ongoing hypotension. he was in cardiogenic shock (cs) and required multiple pressors to support his blood pressure and unstable hemodynamics. his impella ¿stopped working¿ during his arrest per staff, however, there was no change in motor current during this time. he again required multiple blood products. he developed a rapid atrial fibrillation (af) post op followed by ventricular tachycardia, requiring cardioversion. echo showed the impella in deep, and it was repositioned. he developed some leukocytosis and cultures were done. his af returned and was unable to convert with multiple cardioversion attempts. the family opted to withdraw care on (b)(6) 2025. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
362 | 1220648-2026-03502 | 24381402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24381402&pc=OZD | 12/04/2025 |  |  | 02/18/2026 |  |  | 12/04/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026686642 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 87 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Renal Failure | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 76-year-old male patient was scheduled for elective coronary artery bypass surgery (cabg) plus valve replacement. he was implanted with an impella 5.5. post operatively, he was found to have a non-hemorrhagic cerebral vascular accident (cva). anticoagulation was held. continuous renal replacement therapy (crrt) was started. the patient did not progress and care was withdrawn on (b)(6) 2025. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
363 | 1220648-2026-03503 | 24381506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24381506&pc=OZD | 12/03/2025 |  |  | 02/18/2026 |  |  | 12/04/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765817 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Pumping Stopped | [6 empty] | Renal Failure | Cardiogenic Shock | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 70-year-old male patient was admitted for acute decompensated heart failure. he was initially supported on both va ecmo and intra-aortic balloon pump (iabp). an impella cp was placed to vent the left ventricle. the cp was noted to be deep on echo and was repositioned. the patient was started on continuous renal replacement therapy (crrt). the patient was upgraded to an impella 5.5 with ongoing ecmo support. due to crrt and volume removal, the patient was hypovolemic requiring volume replacement. the patient developed bilateral numbness and was urgently decannulated from ecmo. due to ¿wet¿ lungs on xray, the patient was treated for pneumonia. the patient had a ct scan which showed a subdural hematoma and anticoagulation was stopped. an impella stop alarm was noted on transport but self-resolved. the patient was deemed not to be a candidate for advanced therapies and was placed on palliative care. on (b)(6) 2025, the patient expired on support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
364 | 1220648-2026-03505 | 24381602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24381602&pc=OZD | 01/20/2026 |  |  | 02/18/2026 |  |  | 01/20/2026 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766899 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Backflow | Patient Device Interaction Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
365 | 1220648-2026-03248 | 24381649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24381649&pc=OZD | 12/07/2025 | 02/16/2026 |  | 02/18/2026 |  |  | 12/07/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026774650 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 33-year-old female patient with a history of coronary artery disease, diabetes mellitus, renal insufficiency, and dialysis received an impella 5.5 device for treatment of heart failure with cardiogenic shock secondary to cardiomyopathy. prior to placement of the device, the patient had developed liver failure and acute kidney injury, indicating multiple organ dysfunction syndrome. during support, the patient experienced thrombocytopenia and hemolysis, most likely related to mechanical circulatory support combined with the severity of shock. the patient remained on impella support for 27 days, which was an off-label use of the device. due to the severity of the underlying condition, the family elected to withdraw support. death was conservatively coded to the impella 5.5 while most likely to be caused by the underlying disease and unrelated to impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
366 | 1220648-2026-03509 | 24384418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24384418&pc=OZD | 02/07/2026 |  |  | 02/18/2026 |  |  | 02/07/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1398229 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 39 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Fluid/Blood Leak | Device Damaged by Another Device | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | A 78-year-old male patient was supported with an impella pump (pump sn (b)(6); ip1 sn (b)(6)), placed electively on (b)(6) 2026 via direct left sided surgical aortic access. the customer reported that after changing the purge cassette, purge solution was observed leaking from beneath the console door, and the system displayed a low purge pressure alarm. a de air procedure was not performed. the purge cassette had not been removed and reinserted before the leak was noticed, and no excessive force was reported. the rn obtained a second purge cassette; however, leakage recurred. the pump was then transferred to a new aic with new purge tubing. during troubleshooting, the rn identified that a component on the aic where the purge disc seats was broken, and this damaged component appeared to puncture the purge discs and/or tubing, resulting in repeated purge fluid leakage. based on the information available, the event appears related to a damaged aic component that punctured the purge disc and/or tubing, leading to purge fluid leakage and a low purge pressure alarm. the patient experienced no harm, and the device was exchanged to resolve the issue. returned product analysis will be required to confirm the root cause. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
367 | 1220648-2026-03516 | 24385050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24385050&pc=OZD | 02/10/2026 | 02/18/2026 |  | 02/18/2026 |  |  | 02/10/2026 | 02/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026658659 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
368 | 1220648-2026-03518 | 24385104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24385104&pc=OZD | 02/10/2026 | 02/18/2026 |  | 02/18/2026 |  |  | 02/10/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2026802180 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the axillary artery in an 85-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs), scai stage e. the patient was initially supported with an impella cp but required escalation due to profound hemodynamic instability and the need for additional cardiac output. transfer for va-ecmo or impella 5.5 placement was declined by outside facilities due to the patient¿s critically ill condition. the patient was highly unstable, not transportable, and dependent on impella support and a suprarenin infusion. the clinical team obtained an impella 5.5 device from a nearby center; however, due to very complex and anatomical vessel characteristics, advancement of the pump was significantly challenged. despite attempts at placement, it was not possible to advance and position the device quickly enough. the therapy was ultimately withdrawn. the pump was implanted, but the patient remained under cpr throughout the intervention and never achieved return of spontaneous circulation (rosc). the reported events include failure to advance, medical device removal, hemodynamic instability, and death. the death is conservatively being reported to the impella 5.5 because the device was in use at the time of death; however, based on the available information, the patient¿s outcome is most plausibly attributed to the underlying scai stage e cardiogenic shock and profound instability. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
369 | 1220648-2026-03527 | 24385190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24385190&pc=OZD | 12/04/2025 |  |  | 02/18/2026 |  |  | 12/04/2025 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026741989 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | False Alarm | [6 empty] | Fever | Renal Failure | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 50-year-old male developed post cardiotomy cardiogenic shock after valve surgery. an impella 5.5 was placed for support in addition to va ecmo and intra-aortic balloon pump (iabp). impella in aorta alarms present post operatively despite impella position confirmed on echo to be satisfactory. the iabp was removed followed by va ecmo. continuous renal replacement therapy (crrt) was initiated post operatively while listing the patient for transplant. the patient developed a persistent fever; cultures were negative. he was trach for ongoing mechanical ventilation support. despite ongoing support, care was withdrawn on (b)(6) 2025. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
370 | 1220648-2026-03529 | 24385208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24385208&pc=OZD | 02/10/2026 | 02/18/2026 |  | 02/18/2026 |  |  | 02/10/2026 | 02/18/2026 |  |  |  |  | 02/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026730007 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 80 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 65-year-old male with cardiomyopathy and a pre support clinical state consistent with scai cardiogenic shock stage d remains on impella support via a right axillary/subclavian surgical arterial access. on (b)(6) 2026, the patient received 1 unit of prbcs for a hemoglobin of 5.4 g/dl and hematocrit of 17.3%. platelets were 204. the rn reported dark, tarry stool, suggesting a possible gastrointestinal source; no ct scan was reported as completed. on (b)(6) 2026, the patient received 2 additional units of prbcs for continued low hemoglobin. the heparin infusion was stopped due to hemoglobin decline. based on the available clinical information, the patient experienced anemia requiring multiple transfusions while on impella support. clinical assessment did not identify bleeding or hematoma at the impella access site, and urine remained clear without evidence of hemolysis. ldh levels were not significantly elevated, and there was no indication of device malfunction. the presence of dark, tarry stool suggests a likely gastrointestinal bleeding source unrelated to the impella device. the heparin infusion was appropriately paused in response to declining hemoglobin. at this time, there is no evidence that the impella device contributed to the patient¿s anemia or suspected bleeding, and device involvement is not suspected. bleeding was noted; this is a known risk per our ifu (instructions for use) because of our anticoagulation and purge requirements. the patient remains on support at the time of reporting. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
371 | 1220648-2026-03334 | 24363462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24363462&pc=OZD | 02/06/2026 |  |  | 02/17/2026 |  |  | 02/06/2026 | 02/17/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026614158 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Failure to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
372 | 1220648-2026-03339 | 24363656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24363656&pc=OZD | 02/06/2026 |  |  | 02/17/2026 |  |  | 02/07/2026 | 02/17/2026 |  |  |  |  | 02/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806198 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: a 70-year-old male in scai cardiogenic shock stage e underwent impella 5.5 implantation following vt ablation and post procedural neurological decline, with ct confirming an ischemic (embolic) stroke. the patient had been supported with va ecmo and iabp prior to escalation to impella therapy due to concerns for left ventricular thrombus, which was not visualized on later imaging. thrombus formation was subsequently noted on the impella inlet cage while anticoagulation was reportedly therapeutic. based on the available information, the patient¿s hemodynamic instability, the presence of multiple mechanical circulatory support devices, a history of lv thrombus, and an underlying pro thrombotic state may have contributed to the event. at this time, no evidence confirms a device malfunction. product return is not expected and no device data download was required. the patient remained on impella support at the time of reporting. thrombus and embolic complications are known risks associated with cardiogenic shock, severe cardiac dysfunction, and mechanical circulatory support, as outlined in the impella instructions for use. further analysis may be performed if additional information becomes available. without product return, no evaluation or analysis can be performed. (b)(6) 2026. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
373 | 1220648-2026-03341 | 24363886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24363886&pc=OZD | 02/04/2026 |  |  | 02/17/2026 |  |  | 02/04/2026 | 02/17/2026 |  |  |  |  | 02/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026800113 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Female |  |  |  | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinician narrative an impella cp was utilized in a 44-year-old female who presented with acute myocardial infarction complicated by cardiogenic shock (ami-cgs), scai stage d. past medical history was not provided. the patient required inotropic and vasopressor support, mechanical ventilation for respiratory failure, and extracorporeal membrane oxygenation (ecmo) due to the severity of her hemodynamic compromise. right femoral arterial access was obtained. the left main (lm) coronary artery was treated with coronary angioplasty, ptca, and stent placement. the complaint states that the case changed from continuation of therapy to patient complication; however, it was reported that the event did not result in patient harm. a planned escalation from impella cp to impella 5.5 was attempted; however, advancement of the 5.5 device was unsuccessful due to small vessel diameter. as a result, the 5.5 could not be advanced and an impella cp was implanted instead. the patient experienced device revision/replacement and transient hemodynamic instability related to failure to advance the larger device. given the patient¿s critical presentation with ami-cgs, scai stage d shock, need for ecmo, vasopressors, inotropes, and ventilatory support, episodes of hemodynamic instability are consistent with the underlying cardiogenic shock state and high-risk coronary intervention. inability to advance the impella 5.5 due to vessel size is a recognized anatomical limitation and not unexpected in certain patients. the reported events appear consistent with patient anatomy and the severity of cardiogenic shock requiring mechanical circulatory support. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
374 | 1220648-2026-03343 | 24363957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24363957&pc=OZD | 02/09/2026 |  |  | 02/17/2026 |  |  | 02/09/2026 | 02/17/2026 |  |  |  |  | 02/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026703333 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 74 KG | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Cardiac Arrest | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Pericardial Effusion | Thrombocytopenia | [8 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted into the right axillary/subclavian artery in a 74-year-old female presenting with an out-of-hospital cardiac arrest requiring cardiopulmonary resuscitation (cpr), in acute myocardial infarction (ami) and cardiogenic shock (cgs), in scai stage e shock, prior to initiation of support. the impella 5.5 was inserted as an escalation of therapy from an impella cp. while the patient was in the intensive care unit (icu), the impella was on p-4 support. the physician stated that he was unable to run the flows above p-4 without suction alarms and ectopy. the patient was on inotrope/vasoactive support. the lactate dehydrogenase (ldh) was elevated, so the physician inquired about a plasma free hemoglobin level, but the nurse stated that the lab was unable to obtain at that facility. lactic acid resulted in 4.2 and 4.4, so additional vasoactive support was added. the physician noticed discoloration of the toes to the right lower extremity. an echocardiogram was done and showed the pump inlet was pointed towards the mitral valve. a moderate pericardial effusion was also noted. the physician attempted to reposition the catheter, but was unable to, and expressed concern for clotting in the graft making the pump immovable. the patient was off of heparin due to reported bleeding and low platelets. the decision was made to bring the patient back to the cardiovascular operating room (cvor) for a pericardial window, graft clean out, and pump repositioning. upon arrival to the cvor, the patient's heart rate was in the 40's and mean arterial pressure (map) 40's with no pulsatility. the staff were unable to obtain a pulse. cpr was initiated. the patient expired on support. the device is unlikely to have contributed to the patient's death, which was most likely due to the clinical condition of a critically ill patient with acute myocardial infarction complicated by cardiogenic shock, along with the complications of stage e shock. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
375 | 1220648-2026-03116 | 24364230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24364230&pc=OZD | 02/04/2026 |  |  | 02/17/2026 |  |  | 02/05/2026 | 02/17/2026 |  |  |  |  | 02/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745453 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 68-year-old female patient presenting with ami/cgs with a history of diabetes mellitus. during insertion, the md noted that the axillary artery was tight, and the pump was damaged upon advancement. the front looked kinked. the decision was made to exchange for a new impella 5.5. the reported events are failure to advance, device revision or replacement, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
376 | 1220648-2026-03350 | 24364291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24364291&pc=OZD | 02/07/2026 |  |  | 02/17/2026 |  |  | 02/07/2026 | 02/17/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026803390 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
377 | 1220648-2026-03351 | 24364443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24364443&pc=OZD | 12/22/2025 |  |  | 02/17/2026 |  |  | 12/22/2025 | 02/17/2026 |  |  |  |  | 02/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 118 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | During impella 5.5 support, automated impella controller (aic) experienced a placement signal issue. team stated that placement signal reading lower than map. no other information was provided. this is considered a reportable malfunction. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
378 | 1220648-2026-03352 | 24364784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24364784&pc=OZD | 02/07/2026 |  |  | 02/17/2026 |  |  | 02/07/2026 | 02/17/2026 |  |  |  |  | 02/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1158010 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: customer reported a loss of placement signal during support. standard troubleshooting was initiated, and the clinical team proceeded with an automated impella controller (aic) exchange, after which the issue was resolved. no additional device abnormalities were reported, and support was maintained following the exchange. the reported events are controller failure, medical device removal, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and full support was resumed without any known adverse events. | [5 empty] | A. patient information is unknown. device status. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed.
379 | 1220648-2026-03355 | 24364961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24364961&pc=OZD | 02/07/2026 |  |  | 02/17/2026 |  |  | 02/07/2026 | 02/17/2026 |  |  |  |  | 02/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801125 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 93 KG |  | Black Or African American | Injury | Yes | Yes | Deformation Due to Compressive Stress | Device Damaged by Another Device | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 39-year-old male patient presenting with cardiomyopathy, with a history of renal insufficiency. after successful device placement, the physician began closing the axillary access site while another physician was placing a perclose suture at the previously removed iabp femoral access site. during this process, the wire used for perclose deployment interacted with the impella catheter and became wrapped around it. attempts to advance a catheter over the wire to free it were unsuccessful, and both the wire and catheter remained immobile. when the wire and catheter were pulled back, the impella was inadvertently pulled out of the left ventricle. the device was removed and noted to have a kink and an indentation on the catheter near the motor housing. a new impella 5.5 was placed successfully. the reported events are product damage, mechanical issue, medical device removal, device revision or replacement, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and full support was resumed without any known adverse events. | [5 empty] | A5, a6 is unknown. device status. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed.
380 | 1220648-2026-03359 | 24365921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24365921&pc=OZD | 12/22/2025 |  |  | 02/17/2026 |  |  | 12/22/2025 | 02/17/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1302104 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
381 | 1220648-2026-03363 | 24366682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24366682&pc=OZD | 02/07/2026 |  |  | 02/17/2026 |  |  | 02/08/2026 | 02/17/2026 |  |  |  |  | 02/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026811404 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 109 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was placed via the axillary graft site to support the 72-year-old male patient who was admitted in with cardiomyopathy, presenting in scai stage d shock. prior to the 5.5 being placed the patient was known to be on an intra-aortic balloon pump along with inotropes and vasopressors for support. after 2 days of support the patient's heparin dose was increased. after this there was a hematoma that developed at the pump insertion site which prompted the team to stop the heparin and apply a pressure dressing. the pump remains on for support despite the bleed. anticoagulation necessary for the pump placement and support can contribute to access site bleeding. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
382 | 1220648-2026-03365 | 24367448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24367448&pc=OZD | 02/02/2026 | 02/17/2026 |  | 02/17/2026 |  |  | 02/02/2026 | 02/17/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1620355 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
383 | 1220648-2026-03372 | 24368050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24368050&pc=OZD | 12/27/2025 |  |  | 02/17/2026 |  |  | 12/28/2025 | 02/17/2026 |  |  |  |  | 02/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET APAC |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET APAC |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737673 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Anemia | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp was placed via an unknown/unreported access site to support the 59 year old female patient who had been admitted in with acute myocardial infarction and cardiogenic shock, presenting in scai stage d shock. the patient's medical history was not shared beyond the prior cardiac arrest with cpr and being supported by inotropes and vasopressors. on the day of the access and pump placement there was an access site ooze that did not prompt intervention. on day 10 of the support the patient had a drop in hemoglobin with prompted the infusion of blood cells (rbc). the team assessed for a possible gastrointestinal bleed, though none was found. after the rbc was infused the hemoglobin rose by 12g/dl. the pump continued to support the patient til the pump was weaned and explanted. at the time of explant the team observed significant thrombus to be present and adhered to the pump. there was no treatment or intervention noted to treat the thrombosis observed on day 17 of the support. anticoagulation necessary for the pump placement and support can contribute to access site bleeding. the patient had both the heparin bolus and systemic heparin via iv delivery. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
384 | 1220648-2026-03374 | 24368527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24368527&pc=OZD | 12/29/2025 |  |  | 02/17/2026 |  |  | 12/29/2025 | 02/17/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1301514 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
385 | 1220648-2026-03373 | 24368816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24368816&pc=OZD | 02/04/2026 | 02/17/2026 |  | 02/17/2026 |  |  | 02/08/2026 | 02/17/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026720053 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Decrease in Pressure | Material Split, Cut or Torn | [5 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
386 | 1220648-2026-03375 | 24369003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24369003&pc=OZD | 12/27/2025 |  |  | 02/17/2026 |  |  | 12/29/2025 | 02/17/2026 |  |  |  |  | 02/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1636031 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 94 KG | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was on impella 5.5 support. left-lung bleeding and position alarms were reported. the patient remains on impella support. additional information indicated that the lung bleeding was not related to the impella device. a ¿position unknown¿ alarm was displayed on the console. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
387 | 1220648-2026-03376 | 24369176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24369176&pc=OZD | 12/30/2025 |  |  | 02/17/2026 |  |  | 12/30/2025 | 02/17/2026 |  |  |  |  | 02/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026652239 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | During impella 5.5 support, the patient experienced a pump stop. a note was made from csc that a red "impella stopped. retrograde flow" alarm appeared after 5 days on support. the impella was removed as a result and continued with ecmo. the pump stop is considered a reportable malfunction. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
388 | 1220648-2026-03382 | 24370438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24370438&pc=OZD | 11/11/2025 |  |  | 02/17/2026 |  |  | 12/31/2025 | 02/17/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1882177 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission
389 | 1220648-2026-03385 | 24373987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24373987&pc=OZD | 11/24/2025 |  |  | 02/17/2026 |  |  | 11/24/2025 | 02/17/2026 |  |  |  |  | 02/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2026670348 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiogenic shock secondary to an acute myocardial infarction was implanted with an impella cp for temporary mechanical circulatory support. this first impella cp device had complications and was replaced with another impella cp device pump 2, which was successfully inserted and ran well, presenting no issues per the report. as the patient's condition worsened, both the physician and the family decided to withdraw care leading to the patient's demise. | [5 empty] | Patient information age, weight, ethnicity and race are unknown. the exact date of death is unknown but has been captured as the date of explant. if actual information becomes known, a supplemental report will be submitted. medical safety review: medical safety/clinical reviewed the event an noted the patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp for mechanical circulatory support. the event describes a crack in the purge arm resulting in a leak. the pump ran on performance level p-9 with motor current in range (986/823/888mmhg). a different impella cp device (pump 2) was successfully inserted. patient's pulmonary condition worsened and withdrawal of care was decided with the physician and family. in this case, the patient's underlying condition most likely contributed to the demise outcome. regarding the impella cp pump 2, the physician reported that the impella cp pump 2 was running well, presenting no issues. however, conservatively, the death will be reported as the impella cp pump 2 was in use at the time of expiration. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
390 | 1220648-2026-03195 | 24344029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344029&pc=OZD | 02/08/2026 | 02/14/2026 |  | 02/16/2026 |  |  | 02/08/2026 | 02/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026811402 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
391 | 1220648-2026-03197 | 24344090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344090&pc=OZD | 02/08/2026 | 02/14/2026 |  | 02/16/2026 |  |  | 02/08/2026 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026710458 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp was inserted via right femoral artery in a 53 year old male for acute myocardial infarction with cardiogenic shock. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was identified as scai shock stage d. the patient received support from the cp for the coronary angioplasty procedure. on day 2 of support, hematuria was observed, which appeared to be clearing slowly. the p-level had been decreased from p-9 to p-5. the device was explanted. hemolysis is a known risk associated with impella cp as described in the instructions for use. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
392 | 1220648-2026-03198 | 24344117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344117&pc=OZD | 02/08/2026 | 02/14/2026 |  | 02/16/2026 |  |  | 02/08/2026 | 02/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798044 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
393 | 1220648-2026-03235 | 24344174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344174&pc=OZD | 11/06/2025 |  |  | 02/16/2026 |  |  | 11/06/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026749972 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Death | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 78-year-old male patient who had undergone redo aortic valve replacement and aortic root replacement required re-opening due to postoperative bleeding and hemodynamic instability, during which a single-vessel cabg (rca) and intra-aortic balloon pump (iabp) placement were performed. within the following hours the patient developed worsening low-cardiac-output syndrome (lcos), necessitating bailout escalation to impella 5.5 and rp flex support. adequate hemodynamic support was achieved with both devices under tee guidance. during mechanical circulatory support, hemolysis was observed. despite these interventions, the patient subsequently expired. the event is being conservatively reported; based on the available information, no direct causal relationship with the impella 5.5 has been established. | [5 empty] | Section d catalog number was corrected. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
394 | 1220648-2026-03236 | 24344209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344209&pc=OZD | 11/10/2025 |  |  | 02/16/2026 |  |  | 11/10/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026729379 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Protected cardiac surgery with postoperative bleeding of the surgical intervention field. no relationship to the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
395 | 1220648-2026-03238 | 24344258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344258&pc=OZD | 10/31/2025 | 02/19/2026 |  | 02/16/2026 |  |  | 10/31/2025 | 02/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1338086 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
396 | 1220648-2026-03241 | 24344333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344333&pc=OZD | 12/05/2025 |  |  | 02/16/2026 |  |  | 12/05/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026728824 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 88 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A fifty-five-year-old male patient with a history of type 2 diabetes mellitus, renal insufficiency, and coronary artery disease received an impella 5.5 device for management of heart failure cardiogenic shock secondary to ischemic cardiomyopathy. the patient presented in stage e of the society for cardiovascular angiography and interventions shock classification and was supported with an intra-aortic balloon pump, mechanical ventilation, and three inotropic medications prior to impella implantation, indicating profound shock and severe underlying illness. after one week of impella 5.5 support, the patient demonstrated some hemodynamic improvement and was successfully extubated. on day ten of support, gastrointestinal bleeding was suspected but not confirmed, and the patient was transfused with two units of packed red blood cells. after twenty-eight days of impella 5.5 support, in light of the patient¿s severe disease burden and overall clinical status, the family elected to withdraw care and the patient expired. death was conservatively coded to the impella 5.5 while most likely to be caused by the underlying disease and unrelated to impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
397 | 1220648-2026-03243 | 24344401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344401&pc=OZD | 12/15/2024 |  |  | 02/16/2026 |  |  | 12/05/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025564031 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A sixty-six-year-old male patient with a medical history of type 2 diabetes mellitus, hypertension, atrioventricular nodal reentry tachycardia, and cardiomyopathy with an ejection fraction of ten to fifteen percent presented with three weeks of chest pressure and persistent episodes of ventricular tachycardia. the patient was placed on left atrial veno-arterial extracorporeal membrane oxygenation for a planned ventricular tachycardia ablation; however, the clinical team elected not to proceed with the ablation and instead decannulated and implanted an impella 5.5 device via axillary artery surgical cutdown with the eventual goal of evaluation for heart transplantation. following implantation, the patient demonstrated significant hemodynamic improvement, though intermittent ventricular tachycardia episodes persisted and were controlled with anti-arrhythmic medications. on the ninth day of impella support, the patient experienced a massive stroke resulting in severe neurological deficits and impaired respiratory function. in light of the grave prognosis, the family elected to withdraw care, and the patient expired while on support. the impella 5.5 will be coded for stroke and death as a contribution of the pump cannot be ruled out. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
398 | 1220648-2026-03242 | 24344408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344408&pc=OZD | 02/08/2026 |  |  | 02/16/2026 |  |  | 02/08/2026 | 02/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773857 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
399 | 1220648-2026-03245 | 24344476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344476&pc=OZD | 02/07/2026 |  |  | 02/16/2026 |  |  | 02/09/2026 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026615106 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 116 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the nurse stated that the s retrograde flow alarm has appeared. also stated that they lost placement signal 2 days ago. at that time the device was at p5, flow 2.8 with 4.2/0 for max/min. mc 411/303, pf 8.0, pp 574. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
400 | 1220648-2026-03249 | 24344614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344614&pc=OZD | 12/07/2025 |  |  | 02/16/2026 |  |  | 12/07/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775991 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 94 KG | Non Hispanic | White | Death | Yes | No | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 78-year-old male patient with a history of coronary artery disease received an impella 5.5 device for management of heart failure with cardiogenic shock due to acute decompensation. the patient presented in stage e shock as per the society for cardiovascular angiography and interventions classification and had undergone multiple rounds of cardiopulmonary resuscitation prior to impella implantation. before placement of the impella device, the patient was supported with extracorporeal membrane oxygenation and three inotropic medications. given the patient¿s advanced age, multiple comorbidities, and severe disease state, the family elected to withdraw care after 13 days of support. death was conservatively coded to the impella 5.5 while most likely to be caused by the underlying disease and unrelated to impella. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
401 | 1220648-2026-03250 | 24344671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344671&pc=OZD | 11/29/2025 |  |  | 02/16/2026 |  |  | 11/29/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026736482 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Device Sensing Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 61-year-old male patient was admitted for a diagnostic catheterization. the cath showed clean coronaries but a severely reduced ef of 10% so an impella cp was placed for support and for further work up and evaluation. there were initial issues with femoral artery flow after impella placement, with a hematoma noted, so a new impella cp was placed later that evening. a covered stent was placed at the initial impella insertion site. the team was unable to up-titrate the inotropes as the patient would have increased ectopy. since the patient remained hemodynamically unstable, he was transferred for higher level of care. at the hub hospital, an impella 5.5 was placed and the cp was removed. while on impella 5.5 support, the patient had multiple episodes of supraventricular tachycardia but was asymptomatic. he required additional medications for the arrhythmia. an echo was done which showed impella to be in good position. placement signal and lv signal became inaccurate as the sensor appeared to be failing. care staff disabled placement monitoring due to the failing sensor. false suction alarms were also noted due to the failing sensor and unreliable lv waveform; therefore, the suction alarm audio was disabled. the patient was ultimately transplanted without incident on (b)(60 2025. the obvious complication was with the impella cp on initial implant. the second impella cp was not enough to provide enough cardiac support, warranting an upgrade to a 5.5. while on the 5.5, the patient had episodes of svt, however, this cannot be directly attributed to the pump. sensor failures were obviously pump related. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
402 | 1220648-2026-03255 | 24344858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24344858&pc=OZD | 02/09/2026 |  |  | 02/16/2026 |  |  | 02/09/2026 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802332 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | False Alarm | Malposition of Device | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted into the right subclavian artery in a 72-year-old male patient with a past medical history of coronary artery disease (cad), diabetes, renal insufficiency, and dialysis. the patient presented in scai stage a shock prior to initiation of support. the impella was inserted for ventricular support for a high-risk percutaneous coronary intervention (pci). during the procedure, the pump migrated more than 5cm from the aortic valve (av), prior to leaving the operating room (or). the physician decided not to advance the pump. on arrival to the cardiovascular intensive care unit (cvicu), an "impella position in aorta" alarm was on. the physician attempted to reposition the impella without success; therefore, the impella was replaced with a new impella 5.5. the post-procedure outcome is unknown. impella migration is a common, often movement-related, complication where the device moves out of its optimal position across the aortic or pulmonic valve, leading to alarms on the console. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
403 | 1220648-2026-03258 | 24345040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24345040&pc=OZD | 02/08/2026 |  |  | 02/16/2026 |  |  | 02/08/2026 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1390968 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during the preparation of a case an automated impella controller (aic) (loaner) would not reminisce the purge cassette. a different aic was used. no patient was involved. | [5 empty] | No patient is involved in this issue. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
404 | 1220648-2026-03259 | 24345205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24345205&pc=OZD | 02/08/2026 |  |  | 02/16/2026 |  |  | 02/08/2026 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1350280 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during the preparation of a case an automated impella controller (aic) console had a battery failure alarm. no further information regarding the case was provided. no patient was involved. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
405 | 1220648-2026-03262 | 24345294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24345294&pc=OZD | 09/26/2025 |  |  | 02/16/2026 |  |  | 09/26/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1461907 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 60 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | Impella cp was placed in a 70-year-old female for high-risk percutaneous coronary intervention (hrpci) through the left axillary artery. impella was placed without incident and shockwave was performed on left main, left anterior descending artery (lad), and circumferential (cx). the placement signals on the automated impella controller (aic) were absent for most of the case with no consequence to hemodynamic support or risk to the patient. this issue is referenced in the optimal device performance (odp) guidance regarding use of shockwave and impella. | [5 empty] | Date of manufacture is unknown at the time of this mdr writing. the automated impella controller device was not received from the customer; therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
406 | 1220648-2026-03268 | 24345488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24345488&pc=OZD | 02/10/2026 |  |  | 02/16/2026 |  |  | 02/10/2026 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026738753 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 115 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the impella was positioned deep in the left ventricle under echo. the measurement was7.5 cm. there is no alarms and flows are appropriate. urine is still clear. the medical doctor does not want to reposition despite the echo imaging. the patient stayed on support for 3 additional days until successful weaning and explant. | The pump was ultimately repositioned. this patient experienced excellent heart recovery. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated with additional information.
407 | 1220648-2026-03271 | 24345585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24345585&pc=OZD | 02/10/2026 |  |  | 02/16/2026 |  |  | 02/10/2026 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1599846 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a that when they turned on automated impella controller (aic) at (b)(6) and received error message "system self check failed change console immediately". shut down console by holding power button for 30 seconds and plugged console in and did try to restart one more time but received same message. this was only for a cath lab inservice training, so no patient delay whatsoever and no pump was attached. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
408 | 1220648-2026-03276 | 24346965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24346965&pc=OZD | 12/08/2025 |  |  | 02/16/2026 |  |  | 12/08/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026770412 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 93 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Unspecified Infection | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A seventy-year-old male patient with a history of type 2 diabetes mellitus, coronary artery disease, and prior coronary artery bypass grafting received an impella 5.5 device for management of acute myocardial infarction cardiogenic shock. before impella implantation, the patient required more than three inotropic and vasopressor medications, veno-arterial extracorporeal membrane oxygenation, and mechanical ventilation, consistent with stage e of the society for cardiovascular angiography and interventions shock classification, indicating severe disease. on the third day of impella support while still on combined impella and extracorporeal membrane oxygenation therapy, the patient developed fever concerning for blood infection; cultures were obtained, although results were not available. the following day, veno-arterial extracorporeal membrane oxygenation was decannulated, after which the patient experienced ventricular fibrillation requiring direct-current cardioversion with return to normal rhythm. on the eighth day of impella support, the patient developed recurrent episodes of ventricular tachycardia and ventricular fibrillation requiring cardiopulmonary resuscitation; despite intervention, the patient ultimately succumbed to the severity of illness. death was conservatively coded to the impella 5.5 while most likely to be caused by the underlying disease and unrelated to impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
409 | 1220648-2026-03277 | 24346993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24346993&pc=OZD | 12/03/2025 |  |  | 02/16/2026 |  |  | 12/03/2025 | 02/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026774635 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Cardiac Arrest | Low Blood Pressure/ Hypotension | Cardiogenic Shock | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission
410 | 1220648-2026-03280 | 24348308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24348308&pc=OZD | 04/17/2025 |  |  | 02/16/2026 |  |  | 12/09/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025585330 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 106 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A sixty-seven-year-old male patient with a history of coronary artery disease and valvular dysfunction received an impella cp device along with veno-arterial extracorporeal membrane oxygenation for management of post-cardiotomy cardiogenic shock consistent with stage d classification. prior to impella cp placement, the patient experienced multiple episodes of ventricular tachycardia and ventricular fibrillation that required defibrillation, and he was receiving multiple inotropic medications and mechanical ventilation. during insertion of the initial impella cp device, the pump was found to be malfunctioning when connected to the white connector cable; the device was not implanted and was immediately replaced with a second impella cp device, which was successfully inserted. two days after impella cp placement, the patient was escalated to an impella 5.5 device to provide increased hemodynamic support and facilitate weaning from extracorporeal membrane oxygenation. the patient¿s clinical condition subsequently continued to deteriorate, and the family elected to withdraw care after a day of 5.5 and ecmo support. death was conservatively coded to the impella 5.5 while most likely to be caused by the underlying disease and unrelated to impella. death was conservatively coded to the impella 5.5 while most likely to be caused by the underlying disease and unrelated to impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
411 | 1220648-2026-03281 | 24348620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24348620&pc=OZD | 12/09/2025 |  |  | 02/16/2026 |  |  | 12/10/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026760783 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A sixty-two-year-old male patient received an impella 5.5 device for acute decompensated heart failure shock secondary to non-ischemic cardiomyopathy. the patient had been receiving multiple doses of inotropic medications, an intra-aortic balloon pump, and mechanical ventilation for more than seventy-two hours before impella implantation, indicating severe underlying disease. hemodynamic measurements were consistent with stage e cardiogenic shock classification. the patient remained on impella support for thirty days. on the twenty-eighth day of support, during placement of a vascular line through the jugular vein, the patient experienced excessive bleeding. anticoagulation therapy with heparin was stopped, and the insertion site was compressed. after heparin therapy was restarted, the patient sustained a catastrophic intracranial hemorrhage. the surgeon reported that the patient likely experienced an ischemic stroke during the period when heparin was discontinued and compression was applied, which then converted to hemorrhage once anticoagulation resumed. while most probably caused by the line insertion event, stroke and major bleed will conservatively be coded to the impella device, as the required anticoagulation might have aggravated the bleeding and the mechanical support during anticoagulation pause might have contributed to the stroke. supportive care was withdrawn two days later due to the patient¿s overall clinical condition. death was conservatively coded to the impella 5.5 while most likely to be caused by the underlying disease and unrelated to impella. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the root cause of the injuries was not determined due to insufficient clinical details.
412 | 1220648-2026-03282 | 24348636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24348636&pc=OZD | 02/09/2026 |  |  | 02/16/2026 |  |  | 02/10/2026 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1347548 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the loaner automated impella controller will not turn on. no further information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
413 | 1220648-2026-03285 | 24349487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24349487&pc=OZD | 05/31/2025 |  |  | 02/16/2026 |  |  | 05/31/2025 | 02/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026676225 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
414 | 1220648-2026-03287 | 24349613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24349613&pc=OZD | 12/17/2025 |  |  | 02/16/2026 |  |  | 12/17/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775779 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A fifty nine year old male patient with a history of coronary artery disease and diabetes mellitus received an impella cp device for acute myocardial infarction¿related cardiogenic shock. prior to device placement, the patient was supported with three inotropic medications and mechanical ventilation, consistent with severe illness corresponding to society for cardiovascular angiography and interventions stage e shock. following impella insertion, the patient¿s hemodynamics showed slight improvement within one day. the next day, continuous renal replacement therapy was initiated, and two days later the patient was escalated to an impella 5.5 device. percutaneous coronary intervention was performed on the fifth day. on the thirteenth day of support, the patient¿s hemoglobin had been gradually decreasing, and one unit of packed red blood cells was administered; mild oozing was noted at multiple invasive sites, with no evidence of hemolysis. after approximately one month of total support, the impella device was successfully removed due to recovery of native heart function. h&h assumed to mean hemoglobin and hematocrit. blood given. oozing throughout, not just impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
415 | 1220648-2026-03289 | 24349630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24349630&pc=OZD | 12/04/2025 |  |  | 02/16/2026 |  |  | 12/04/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737494 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Anemia | Tachycardia | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 73-year-old male patient was admitted with acute mi, cardiogenic shock. he was initially implanted with an impella cp. he became oliguric and was noted to have cold extremities with decreased capillary refill. the team decided to escalate to an impella 5.5. urine output increased post 5.5 implant. vasopressors were able to be quickly weaned. the patient did develop some ventricular tachycardia and showed signs of hematuria. pump position was checked via echo. his hemoglobin dropped requiring blood transfusion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
416 | 1220648-2026-03290 | 24349734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24349734&pc=OZD | 12/04/2025 |  |  | 02/16/2026 |  |  | 12/04/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765218 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 73-year-old patient was scheduled for elective coronary artery bypass surgery (cabg) with impella 5.5 support. he initially had cardiac tamponade requiring surgical take-back for washout. received blood post operatively. support during his icu was without incident and he was explanted on (b)(6) 2025. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
417 | 1220648-2026-03292 | 24349812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24349812&pc=OZD | 12/05/2025 |  |  | 02/16/2026 |  |  | 12/05/2025 | 02/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746977 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
418 | 1220648-2026-03296 | 24349984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24349984&pc=OZD | 12/06/2025 |  |  | 02/16/2026 |  |  | 12/06/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779706 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 77 KG | Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 59-year-old male patient with a history of diabetes mellitus, hypertension, and coronary artery disease underwent high-risk coronary artery bypass surgery (cabg) and received impella 5.5 for mechanical circulatory support. after 12 days of support, the device was exchanged for another impella 5.5, which provided support for an additional 14 days. on 12/19, the patient experienced several episodes of ventricular tachycardia that resolved without intervention and did not result in hemodynamic compromise. following a total support duration of 26 days, the impella device was successfully weaned and explanted. 26 days total support (5.5 to 5.5 exchange at day 12), ao offset day 1, ali in upper arm day 3, vt events on day 14 (second pump). during first impella 5.5 support, care team contacted the call center overnight about impella position unknown alarms triggered on automated impella controller. placement signal was not correlating with the patient¿s radial arterial line, indicating a probable issue with the fibrotic sensor. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
419 | 1220648-2026-03297 | 24350201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24350201&pc=OZD | 12/17/2025 |  |  | 02/16/2026 |  |  | 12/17/2025 | 02/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745766 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
420 | 1220648-2026-03300 | 24350559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24350559&pc=OZD | 12/17/2025 |  |  | 02/16/2026 |  |  | 12/17/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775971 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A twenty-five-year-old male patient received an impella 5.5 device as preoperative support for high-risk cardiac surgery. the patient tolerated the device well, and hemodynamics improved after surgery. on the third day of support, the patient developed ectopy, and echocardiography showed the impella positioned close to the interventricular septum; the device was repositioned at the bedside, amiodarone therapy was restarted, and the impella support level was reduced from performance level seven to performance level five. after nine days of support, the patient demonstrated recovery of native heart function and was taken to the operating room for impella explantation. during removal, difficulty was encountered, and the surgeon suspected that the motor housing was catching at the angle of the patient¿s innominate artery takeoff. a sternotomy was performed, and manual pressure was applied to the innominate artery to straighten the vessel, after which the device was successfully removed. pec of device explant issue - triggers chest opening to reposition device via manual pressure. (ppae, excluded because separate pec). vt caused by device position. 5.5 noted as malposition on echo, triggered repositioning. repositioning done per standard ifu, so not noted as separate pec. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
421 | 1220648-2026-03301 | 24350772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24350772&pc=OZD | 12/20/2025 |  |  | 02/16/2026 |  |  | 12/20/2025 | 02/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1078703 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
422 | 1220648-2026-03303 | 24351044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24351044&pc=OZD | 12/06/2025 |  |  | 02/16/2026 |  |  | 12/06/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775821 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 93 KG | Hispanic |  | Injury | Yes | No | Device Sensing Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | 67 year old gentleman presenting with myocardial infarction and cardiogenic shock coming in on intra-aortic balloon pump and extracorporeal membrane oxygenation. on day 3 of circulatory support, the balloon pump was removed and impella cp placed for left ventricular unloading post cardiac arrest. in the evening, impella cp was escalated to 5.5 for better support and potential weaning of the extracorporeal membrane oxygenation. decision for impella 5.5 was apparently not driven by any device complaint but rather as a therapeutic strategy towards a permanent outcome. the aic showed intermittent placement signal alarms, however echo confirmed optimal position of 5.5 and an increase in p-level solved the alarm. on the same day, a large left sided stroke was diagnosed that was later confirmed by computer tomography scan and resulted in loss of eligibility for permanent assist device placement. after 12 days of support, the impella 5.5 was successfully weaned and removed. the stroke was identified one day after impella 5.5 implantation and will conservatively be coded to the this device, while it remains unclear whether it already occured during the pump exchange. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
423 | 1220648-2026-03309 | 24352483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24352483&pc=OZD | 02/03/2026 | 02/16/2026 |  | 02/16/2026 |  |  | 02/04/2026 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1154178 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that the console started with a yellow battery comm failure. there was no patient involvement. no additional information was provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
424 | 1220648-2026-03317 | 24354689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24354689&pc=PYX | 10/21/2025 |  |  | 02/16/2026 |  |  | 10/21/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PUMP, RP-IJ, STERILE |  |  | PYX | [5 empty] | ABIOMED, INC. |  |  | US |  | 1922345 |  | 000000000001000186 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 155 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiogenic Shock | Respiratory Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | 73-year-old male with a history of prostate cancer treated with radiation therapy. the patient presented to an outside hospital with dizziness and nausea and was found to be in atrial fibrillation with rapid ventricular response and acute kidney injury. he was transferred to a higher-level facility after an elevated troponin was identified. a urinalysis revealed a urinary tract infection, which was treated with antibiotics. the following day, chest ct angiography confirmed bilateral pulmonary emboli with right ventricular strain. he underwent successful thrombectomy but continued to decompensate, requiring increasing pressors, inotropes, pulmonary vasodilators, and oxygen therapy. a heart team determined the need for impella rp flex for right-sided circulatory support. use of impella rp flex in pulmonary embolism is off-label due to the risk of clot ingestion. right heart catheterization confirmed severe right ventricular dysfunction with the following hemodynamics: right atrium (ra): 18 mmhg. right ventricle (rv): 49/17 mmhg. pulmonary artery (pa): 56/38 (mean 42) mmhg. pulmonary capillary wedge pressure (pcwp): 16 mmhg. pulmonary artery saturation: 45%. cardiac index: 1.2 l/min/m². pulmonary artery pulsatility index: 0.9. during implantation, the patient required dobutamine, vasopressin, levophed, and intermittent epinephrine boluses. hemostasis was achieved with a hemostatic suture, manual pressure, and a fem-stop device applied to the access site. hemodynamics improved with impella rp flex at p-8, and transthoracic echocardiography demonstrated improved right ventricular function. the following day, the fem-stop was removed, and the patient was placed on systemic heparin and heparin purge, later transitioned to argatroban systemically and sodium bicarbonate in the purge. a venous duplex revealed residual thrombus in the right femoral vein. despite improved rv performance, the patient developed increasing oxygen requirements and did not tolerate impella rp flex weaning. he was transferred to a higher level of care and escalated to va-ecmo and protek duo for respiratory failure and right-sided support. detailed event chronology initial presentation. presented with dizziness and nausea. found in atrial fibrillation with rvr; developed acute kidney injury. elevated troponin prompted transfer to tertiary center. uti identified and treated with antibiotics. diagnosis of pulmonary embolism. chest cta confirmed bilateral pulmonary emboli with rv strain. thrombectomy and clinical decompensation. underwent successful pulmonary embolism thrombectomy. continued to decompensate with rising vasopressor and inotrope requirements. pulmonary vasodilators and oxygen therapy initiated. impella rp flex implantation (off-label) heart team elected to implant impella rp flex for severe rv failure. right heart catheterization showed severe hemodynamic compromise. patient supported with dobutamine, vasopressin, levophed, and intermittent epinephrine during procedure. access site secured with hemostatic suture, manual pressure, and fem-stop. post-implant course. improved hemodynamics on impella rp flex at p-8. improved rv function demonstrated on transthoracic echo. day 1 post-implant. fem-stop removed. anticoagulation transitioned from systemic heparin + heparin purge to argatroban + sodium bicarbonate purge. duplex showed residual right femoral vein thrombus. clinical decline and escalation patient developed increasing oxygen requirements. unable to tolerate impella rp flex weaning. transferred to higher acuity center. escalated to va-ecmo + protek duo for respiratory failure and right-sided circulatory support. outcome patient could not be weaned from impella rp flex. escalated to va-ecmo and protek duo. impella rp flex support was discontinued as part of transition to alternative mechanical support. | [5 empty] | The impella device was not received from the customer; therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
425 | 1220648-2026-03319 | 24355127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24355127&pc=OZD | 02/08/2026 | 02/25/2026 |  | 02/16/2026 |  |  | 02/08/2026 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1276707 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 85 KG | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | An impella cp was inserted via right femoral artery in a 82 year old male for acute myocardial infarction with cardiogenic shock. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage e. during procedure, an ¿optical sensor system error¿ message displayed on the automated impella controller. this occurred after the impella was connected to the controller and the purge solution was confirmed. the consoles were switched and no further error displayed, and the cp successfully implanted. the device has been explanted. the patient has expired. the signal issue will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange, however the exchange was performed with no consequence to the patient and support. death is a known potential adverse event of the impella cp per the instructions for use. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Updated information: d1 brand name updated. d4 catalog, lot, and serial numbers added. h4 mfr date.
426 | 1220648-2026-03323 | 24356368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24356368&pc=OZD | 02/09/2026 | 02/16/2026 |  | 02/16/2026 |  |  | 02/09/2026 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. |  |  | US |  | 1648111 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | An automated impella controller was used in a patient of unknown age or gender. patient's comorbidities and clinical state were unknown. during preparation for a case, it was reported the message "yellow controller error" displayed once the console was turned on. a new console was used prior to implanting a device. the controller error will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange, however the exchange was performed with no consequence to the patient and support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
427 | 1220648-2026-03327 | 24357172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24357172&pc=OZD | 02/09/2026 | 02/16/2026 |  | 02/16/2026 |  |  | 02/09/2026 | 02/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. |  |  | US |  | 1997618 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Decrease in Pressure | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
428 | 1220648-2026-03332 | 24358143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24358143&pc=PYX | 10/22/2025 |  |  | 02/16/2026 |  |  | 10/22/2025 | 02/16/2026 |  |  |  |  | 02/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PUMP, RP-IJ, STERILE |  |  | PYX | [5 empty] | ABIOMED, INC. |  |  | US |  | 1922354 |  | 000000000001000186 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 113 KG | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Cardiac Arrest | Cardiac Tamponade | [10 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | 70-year-old male with a history of hypertension, type 2 diabetes mellitus, hyperlipidemia, coronary artery disease, and dementia. presented with chest pain and syncope. event summary: the patient was transported to the cath lab and had a temporary pacing wire placed. he underwent emergent coronary angiography that demonstrated multivessel coronary artery disease with occluded proximal right coronary artery (rca). angioplasty to the rca was performed. the temporary pacing wire was removed and a pulmonary artery catheter was placed and revealed right ventricular (rv) cardiogenic shock with papi 0.67 and cardiac output 2.2 l/min. a decision was made to implant impella rp flex for right-sided mechanical circulatory support. following implantation, the patient tolerated weaning of inotropes and vasopressors with an associated decrease in lactate from 2.9 to 1.5. the next day, he returned to the cardiac catheterization laboratory for temporary pacing wire placement and replacement of a pulmonary artery catheter. approximately three hours later, the patient experienced cardiac arrest; cpr was initiated. bedside echocardiography demonstrated cardiac tamponade. pericardiocentesis was attempted with minimal blood aspirated. the patient expired while on impella rp flex support. the impella rp flex will be conservatively reported for death; however, the device is unlikely to have contributed to the patient's death, which was most liely due to the underlying critical clinical condition. detailed event chronology: day 0 ¿ presentation and index procedure. presentation: chest pain and syncope. cardiac cath findings: multivessel cad, proximal rca occlusion. intervention: rca angioplasty performed. rhc hemodynamics consistent with rv cardiogenic shock: papi: 0.67. cardiac output: 2.2 l/min. decision: impella rp flex implantation for right-sided mcs. immediate post-implant course: inotropes/pressors weaned successfully. lactate improved from 2.9 1.5. patient hemodynamically improved on rp flex support. day 1 ¿ additional procedures: returned to cath lab for temporary pacing wire placement and replacement of the pa catheter. post-procedure: returned to unit. ~3 hours post-procedure ¿ arrest and deterioration patient coded; cpr initiated. bedside echo revealed cardiac tamponade. pericardiocentesis performed with minimal blood return. despite resuscitative efforts, the patient expired on rp flex support. | [5 empty] | The exact date of death is unknown but has been captured as the date of explant. if actual information becomes known, a supplemental report will be submitted. device status. the impella device was not received from the customer; therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
429 | 1220648-2026-03201 | 24341398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341398&pc=OZD | 12/29/2025 |  |  | 02/15/2026 |  |  | 12/29/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026780217 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 79 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemoptysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 43 year old male with cardiomyopathy and a pre support clinical state consistent with scai shock stage c remains on impella 5.5 support (sn (b)(6)), implanted on (b)(6) 2025 via right axillary/subclavian surgical arterial access. the patient is currently being evaluated for bivad versus transplant candidacy. the reported controller error was transient, self resolving, and did not affect pump performance, flow, or hemodynamic stability. there is no evidence at this time that the impella 5.5 contributed to the patient¿s hemoptysis, which is more likely related to pneumonia and recent adjustments to anticoagulation therapy. the device remains in place and functioning as intended while the patient continues evaluation for advanced heart failure therapies. | [5 empty] | Ppae (hemoptysis): the cause of the injury was not determined due to insufficient clinical details. the pump passed all the post sterile inspection checks.
430 | 1220648-2026-03202 | 24341528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341528&pc=OZD | 02/10/2026 | 02/15/2026 |  | 02/15/2026 |  |  | 02/10/2026 | 02/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026802423 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
431 | 1220648-2026-03203 | 24341530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341530&pc=OZD | 10/30/2025 | 02/20/2026 |  | 02/15/2026 |  |  | 11/06/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026688943 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Respiratory Failure | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 65-year-old male underwent impella 5.5 placement for bridge to heart-transplant. following device insertion, the patient reported intermittent numbness, tingling, and mild pain in the right upper extremity. symptoms transiently worsened two days prior in association with access-site oozing and finally improved. in addition to these symptoms, the patient initially experienced an otherwise uneventful course of bridge-to-transplant support. this course was later interrupted by pulmonary failure requiring intubation and differential oxygen therapy. after approximately three days of partial pulmonary stabilization, the family and healthcare team elected to withdraw care. the death is being conservatively reported; however, based on the clinical circumstances, a causal relationship to the impella¿5.5 is considered highly unlikely. | The complainant reported additional information upon request: "the issue self-resolved for numbness and tingling in the extremity without any intervention, the patient outcome was successful without any further incidents related to the device. i do not have the device in my possession to return for investigation as this account will explant the device and not retain it for us to return to headquarters for investigation." |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated. the investigation is still ongoing.
432 | 1220648-2026-03204 | 24341533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341533&pc=OZD | 11/06/2025 | 02/15/2026 |  | 02/15/2026 |  |  | 11/06/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026745075 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hypovolemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 67-year-old female patient with postcardiotomy cardiogenic shock (pccs) was supported with an impella 5.5 device following cardiac surgery. during support, the patient experienced ectopy and arrhythmias occurring after suction alarms. these events resolved following the recommendation in the ifu: reduction of the performance level and administration of intravenous volume. additionally, a lidocaine infusion at 1 mg/min was initiated, with subsequent stabilization. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. these events resolved following the recommendation in the ifu: reduction of the performance level and administration of intravenous volume. additionally, a lidocaine infusion at 1 mg/min was initiated, with subsequent stabilization.
433 | 1220648-2026-03206 | 24341536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341536&pc=OZD | 11/07/2025 | 02/15/2026 |  | 02/15/2026 |  |  | 11/07/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026655240 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 72 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Ischemia | Hematuria | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 66 year-old male with a past medical history significant for coronary artery disease with prior percutaneous coronary intervention, severe mitral regurgitation, hypertension, hyperlipidemia, and type ii diabetes mellitus presented to the emergency department with complaints of shortness of breath and lethargy. on arrival, the patient appeared shocky and was cold, clammy, and diaphoretic. he was admitted to the medical intensive care unit (micu), where his clinical condition rapidly deteriorated. the medical team decided to place an impella cp for hemodynamic support. the femoral artery was free of disease but noted to be slightly narrow; however, the team proceeded with insertion and the impella device was successfully implanted. the physician then performed a left heart catheterization via single arterial access, which demonstrated that all coronary vessels were patent. an angiogram performed through the repositioning sheath subsequently demonstrated slow blood flow to the access-site leg. a trialysis catheter was placed, and the patient was transferred back to the micu for continued monitoring. upon arrival in the intensive care unit, the patient continued to have neither a palpable nor doppler-detectable pulse in the access-site leg. the treating physicians anticipated improvement as vasopressor medications were reduced. the patient¿s urine output was initially clear but later became bloody. the physician attributed this finding to the device operating at maximum flow speed and suspected it represented older hemolysis. due to the patient¿s ongoing need for increased hemodynamic support and the continued absence of pulses in the affected leg, the decision was made to escalate support to an impella 5.5 device. the impella 5.5 was successfully implanted, the impella cp device was removed, and the patient was returned to the intensive care unit for further management. no device malfunction was reported, and all device-related actions were taken based on the patient¿s clinical condition and procedural findings. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
434 | 1220648-2026-03207 | 24341537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341537&pc=OZD | 11/09/2025 | 02/15/2026 |  | 02/15/2026 |  |  | 11/09/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026748344 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Renal Failure | Hypovolemia | Thrombocytopenia | Thrombosis/Thrombus | Heart Failure/Congestive Heart Failure | [6 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 78-year-old male in cardiogenic shock (scai d) with severe three-vessel coronary artery disease, pre-existing acute kidney injury (creatinine 3.4 mg/dl), and anemia (hemoglobin 10.2 g/dl) was supported with an impella 5.5 device as a bridge to high-risk cabg (type i). during mechanical circulatory support, the patient required renal-replacement therapy for fluid management. ultrasound imaging identified a thrombus in the deep vein of the right upper extremity at the access site, without clinical signs of harm. subsequent diagnostic testing was positive for hit, and anticoagulation was switched to argatroban. on day 14 of support, a four-vessel cabg was performed. the postoperative course was complicated by coagulopathy requiring substantial transfusion support, including blood products and factor vii, as well as multiple re-operations for arterial bleeding at the ima. fluid management remained challenging; episodes of suction events related to fluid requirements and right-heart failure during volume replacement occurred, prompting escalation to additional right-sided support with the rp flex device. during the prolonged and highly complex icu stay, the patient developed an infection of unknown origin and received empiric antibiotic therapy. the combination of infection, need for ongoing mechanical circulatory support, and continued renal-replacement therapy might have contributed to the family¿s decision to withdraw care. the patient subsequently died. this case is reported conservatively; however, given the extensive comorbidities, surgical complexity, and multifactorial clinical deterioration, a causal relationship to the device is considered highly unlikely. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
435 | 1220648-2026-03209 | 24341538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341538&pc=OZD | 11/09/2025 | 02/15/2026 |  | 02/15/2026 |  |  | 11/09/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026713097 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Low Blood Pressure/ Hypotension | Sepsis | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | "a 78-year-old male with a history of diabetes, prior stroke, and acute kidney injury presented on (b)(6) 2025 with stemi and cardiogenic shock. an impella cp was attempted but showed persistent ¿air in purge line,¿ prompting replacement with a new cp. the patient required prbcs and albumin for suction alarms. echo revealed the pump had migrated deep and it was repositioned to 91 cm. additional blood was given for ecmo circuit needs. by (b)(6) 2025, the patient developed labored breathing on high-flow oxygen with gi bleeding and hematuria. on (b)(6) 2025, support was upgraded from impella cp to impella 5.5 to facilitate ecmo removal. following ecmo decannulation on (b)(6) 2025, the left leg was pulseless, and vascular surgery performed thrombectomy; the patient received further transfusion. post-procedure, the impella orientation flipped and was corrected under echocardiographic guidance. the patient remained on crrt with bloody output, minimal urine, and became hypotensive during impella weaning attempts. by (b)(6) 2025, the patient was suspected to be in dic with ongoing gi bleeding requiring multiple transfusions and had positive blood cultures. mechanical ventilation and crrt continued, with vasopressors gradually weaned. a tracheostomy was performed on (b)(6) 2025, and the impella 5.5 was discontinued on (b)(6) 2025. given the patient¿s extensive cardiac deterioration and cardiogenic shock, most complications¿including respiratory failure, dic, gi bleeding, hematuria, renal failure, vasopressor dependence, and need for crrt¿are attributable to the underlying clinical condition. ecmo support likely contributed to major bleeding, hypovolemia, and limb ischemia with thrombosis. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. given the patient¿s extensive cardiac deterioration and cardiogenic shock, most complications¿including respiratory failure, dic, gi bleeding, hematuria, renal failure, vasopressor dependence, and need for crrt¿are attributable to the underlying clinical condition. ecmo support likely contributed to major bleeding, hypovolemia, and limb ischemia with thrombosis. ".
436 | 1220648-2026-03211 | 24341540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341540&pc=OZD | 11/09/2025 | 02/15/2026 |  | 02/15/2026 |  |  | 11/09/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026736942 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Renal Failure | Tachycardia | Cardiogenic Shock | [9 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 60-year-old male with nstemi was admitted for ongoing chest pain and evaluated for cardiac surgery. during an mri viability study, he experienced a cardiac arrest and was successfully resuscitated. an impella cp was placed for urgent transfer to a higher-level care center. on arrival, echocardiography showed an ef of 5%, prompting upgrade to an impella 5.5 with cp removal. initial positioning was shallow and subsequently advanced under echo guidance. the patient was extubated and started on argatroban, and he received a blood transfusion. on (b)(6) 2025, he underwent off-pump cabg ×3. he later developed right-sided failure and returned to the or for protek placement. he was also on vv ecmo (cannulation date unknown). with escalating milrinone, he experienced increased ventricular tachycardia. ongoing volume overload required initiation of continuous renal replacement therapy. he was tracheostomized on (b)(6) 2025 and started on lidocaine for persistent arrhythmias. social work was engaged to identify a healthcare proxy. on (b)(6) 2025, care was withdrawn and the impella 5.5 was explanted. the initial impella cp was unable to provide adequate support after cardiac arrest, necessitating upgrade to the impella 5.5. subsequent multiorgan failure, prolonged ventilation, and need for dialysis were most consistent with cardiogenic shock following arrest rather than device malfunction. tachyarrhythmias may occur with mechanical circulatory support but were more likely secondary to the patient¿s profound shock state. hemoglobin decline could be attributed to cabg or protek placement; no clear association with the impella device was identified. death occurred after withdrawal of care. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. subsequent multiorgan failure, prolonged ventilation, and need for dialysis were most consistent with cardiogenic shock following arrest rather than device malfunction. tachyarrhythmias may occur with mechanical circulatory support but were more likely secondary to the patient¿s profound shock state. hemoglobin decline could be attributed to cabg or protek placement; no clear association with the impella device was identified. death occurred after withdrawal of care.
437 | 1220648-2026-03212 | 24341541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341541&pc=OZD | 11/09/2025 |  |  | 02/15/2026 |  |  | 11/09/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026731295 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Heart Failure/Congestive Heart Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 60 y/o patient was admitted with a non-st segment myocardial infarction. he experienced a cardiac arrest and had an impella cp placed, which was replaced with an impella 5.5 the following day. the patient underwent a high-risk pci with drug eluting stents placed. he had some ventricular tachycardia after the procedure which warranted the initiation of amiodarone. he was then transferred on (b)(6) 2025. he was ultimately extubated and rehabbing with pt while on the 5.5 which led to his weaning and subsequent removal without incident. the arrhythmias were likely due to the impella; however, it was found to be in good position. difficult to say whether it was the pump or the reperfusion post pci. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the arrhythmias were likely due to the impella; however, it was found to be in good position. difficult to say whether it was the pump or the reperfusion post pci.
438 | 1220648-2026-03213 | 24341545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341545&pc=OZD | 11/10/2025 | 02/15/2026 |  | 02/15/2026 |  |  | 11/10/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026745090 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An 82-year-old male patient underwent high-risk pci (scai b) after being declined as a candidate for cabg. an impella 5.5 was placed for hemodynamic support prior to the procedure. the patient continued on ticagrelor therapy and subsequently developed coagulopathy with bleeding from multiple sites, including the nares (following a prior nasogastric tube attempt), the right internal jugular cordis/swan site (removed during transfer), the right axillary surgical pocket, and the right femoral access site from the intervention. hemostasis was eventually achieved, and the patient was successfully supported and later weaned from the impella 5.5. given the concurrent antiplatelet therapy, coagulopathy, and multiple non-device-related bleeding sources, the relationship of these bleeding events to the impella 5.5 is considered highly unlikely. the case is being reported out of an abundance of caution. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. given the concurrent antiplatelet therapy, coagulopathy, and multiple non-device-related bleeding sources, the relationship of these bleeding events to the impella 5.5 is considered highly unlikely. the case is being reported out of an abundance of caution.
439 | 1220648-2026-03214 | 24341551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341551&pc=OZD | 11/11/2025 | 02/15/2026 |  | 02/15/2026 |  |  | 11/11/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026757838 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 66-year-old male with underlying cardiomyopathy received an impella 5.5 device as a bridge-to-transplant therapy. on postoperative day two, the patient developed transient, stroke-like neurological symptoms, prompting and two days after diagnostic evaluation with cranial ct and eeg. both examinations showed no evidence of acute cerebral injury. the case is being reported conservatively, as a device-related contribution to the temporary neurological symptoms cannot be definitively ruled out. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
440 | 1220648-2026-03215 | 24341555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341555&pc=OZD | 11/11/2025 | 02/15/2026 |  | 02/15/2026 |  |  | 11/11/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026677111 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Stroke/CVA | Ventricular Fibrillation | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Patient: 53m with acute myocardial infarction (ami) and cardiogenic shock. initial intervention: emergent pci of left main; lm dissection occurred during attempt to wire lad. ventricular tachycardia cardioverted. impella cp placed via rfa; procedural difficulties with access and positioning. device use: on (b)(6) 2025: impella cp inserted for lv unloading; persistent suction alarms, unreliable position signals, right groin hematoma, severe hemolysis (ldh >2000). va ecmo initiated after cardiac arrest. console fault controller swap. continued ecmo + impella cp; anticoagulation held due to bleeding. persistent thrombocytopenia, multiple hematomas, crrt initiated, ongoing suction alarms, ldh later up to 10k. repeated transfusions for anemia/bleeding. escalation: on (b)(6) 2025: upgraded to impella 5.5 for lv unloading while on va ecmo due to cp age and ongoing support needs. post-implant: episodes of vf, concern for north¿south syndrome, continued high vasopressor use, pulmonary edema, suspected neurologic injury. complications: recurrent suction alarms, severe hemolysis, multiple bleeding sites (access hematomas, bloody ett secretions), console malfunction, thrombocytopenia, coagulopathy, pulmonary edema, skin ulcers, suspected vasospasm with compartment syndrome workup, arrhythmias (vt/vf), persistent organ dysfunction (renal failure requiring crrt, shock liver suspicion). north¿south syndrome (hypoxia) secondary to va ecmo outcome: on (b)(6) 2025: family requested withdrawal of support. patient expired on full mcs (va ecmo + impella 5.5), ventilator, and crrt. summary statement: impella support in combination with va ecmo was complicated by repeated malposition signals, suction events despite volume optimization, device-related bleeding, severe hemolysis, access site hematomas, console fault, and escalation from cp to 5.5. patient remained critically unstable with refractory shock, multiorgan failure, and ultimately died secondary to family decision to withdraw care. impella cp devices most likely contributed to hemolysis and hematoma. cardiogenic shock secondary to patient clinical status was the cause of other multicystem organ failure symptoms, especially given that they were in scai stage e prior to any impella device placement. this was also true for the second cp device and ultimately, the impella 5.5, as the patient appeared to have a life-threatening stroke with fixed and dilated pupils and care was discontinued. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp devices most likely contributed to hemolysis and hematoma. cardiogenic shock secondary to patient clinical status was the cause of other multicystem organ failure symptoms, especially given that they were in scai stage e prior to any impella device placement. this was also true for the second cp device and ultimately, the impella 5.5, as the patient appeared to have a life threatening stroke with fixed and dilated pupils and care was discontinued.
441 | 1220648-2026-03217 | 24341566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341566&pc=OZD | 11/12/2025 | 02/15/2026 |  | 02/15/2026 |  |  | 11/12/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026658687 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Cardiac Tamponade | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Post cardiotomy cardiogenic shock after coronary artery bypass surgery and mitral valve replacement. ef 10%. patient placed on ecmo immediately post op and taken back to or for impella 5.5 insertion for lv unloading. continuous renal replacement therapy initiated on post op day 1. returned to or for tamponade on post op day 2 with washout. post op day 3 patient returned to or for ecmo decannulation, requiring multiple units of blood. immediately post ecmo decannulation, patient went into acute right heart failure, and va ecmo re-initiated with new cannulation. unable to wean from ventilator so patient underwent tracheostomy. ultimately decannulated from ecmo on (b)(6). impella 5.5 with slow wean over next few days. impella therapy discontinued without issue. patient symptoms most likely due to post cardiotomy shock and not the impella pump. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
442 | 1220648-2026-03219 | 24341570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341570&pc=OZD | 11/12/2025 | 02/15/2026 |  | 02/15/2026 |  |  | 11/13/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026731216 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Increase in Pressure | Device Sensing Problem | [5 empty] | Anemia | Arrhythmia | Hemolysis | Low Blood Pressure/ Hypotension | Cardiogenic Shock | Respiratory Failure | Heart Failure/Congestive Heart Failure | [6 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | 73-year-old presented in cardiogenic shock(cs), scai stage d, positive for st segment elevation mi. patient taken to cath lab and impella cp placed prior to stent placement and percutaneous coronary intervention. ventricular fibrillation with successful shocks to sinus rhythm. afterwards, patient was still hemodynamically unstable requiring multiple vasopressors, and shock team was activated. on (b)(6) 2025- cardiogenic shock with ongoing lactate elevation and corresponding hemodynamic compromise. patient continues on vasopressors and antiarrhythmic iv medications. patient escalated to impella 5.5 due to ongoing cs and hemolysis. placed on nitric oxide due to right heart failure. patient received a unit of prbcs overnight to maintain a hemoglobin of 8 or greater. cvp trending up with need for diuresis. hypotensive and started on vasopressin. on (b)(6) 2025, care team determined that patient would need either a durable vad or palliative care. ultimately, the patient decided to decline durable vad and go on palliative care. the plan was to discharge to home with hospice on iv milrinone. on (b)(6) 2025- purge flow diminished to < 1 liter. tpa given through purge. purge pressure high and blocked. md determined there was a clot in the device and tpa again attempted via purge. on (b)(6) 2025- the patient was placed on comfort care and all iv medications and lab draws discontinued. on (b)(6) 2025-care withdrawn and patient expired. patient was admitted with advanced cardiogenic shock, scai stage d with hemodynamic instability requiring vasopressors. impella cp most likely caused arrhythmia and hemolysis. with ongoing cs, patient was escalated to impella 5.5. with worsening right heart failure, most likely due to cs state and not the pump. hypotension also most likely due to cs but hard to confirm there was no impact from the 5.5. hypervolemia due to right heart failure and not pump. impella 5.5 with diminished purge flows and high purge pressures requiring tpa administration. death incurred after withdraw of care and impella so most likely due to patient¿s acuity and ongoing cs. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
443 | 1220648-2026-03220 | 24341572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341572&pc=OZD | 11/12/2025 | 02/27/2026 |  | 02/15/2026 |  |  | 11/13/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026717596 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Heart Failure/Congestive Heart Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | 68 year old male with previous history of ischemic cardiomyopathy, cad s/p pci, atrial fib s/p aicd, ef 20%. presented to er with worsening fatigue, shortness of breath and abdominal pain. impella 5.5 placed via right subclavian artery. patient was previously on eliquis, notable oozing present. a jp drain was placed. 10/30/2025-some suction events and 5.5 repositioned. patient listed for heart transplant. device repositioned multiple times throughout support. increased ventricular ectopy intermittently. (b)(6) 2025- upper gib. transfused with 4 units prbc¿s. to ir for clipping for ongoing gib. multiple blood products given. (b)(6) 2025 patient to ir for coiling for ongoing upper gib. started on dobutamine for cardiac support. transplanted on (b)(6) 2025. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
444 | 1220648-2026-03210 | 24341576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341576&pc=OZD | 11/09/2025 | 02/23/2026 |  | 02/15/2026 |  |  | 11/09/2025 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026743287 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Patient admitted for stemi and underwent coronary artery bypass grafting (cabg) x 4. difficult to get patient off bypass and patient ultimately arrested requiring subsequent va ecmo placement. impella 5.5 implanted when patient returned to or for chest washout. chest was left open and the patient was sent to the icu. required initiation of milrinone for borderline hemodynamics. patient bleeding and his anticoagulation was held. ecmo was ultimately removed on (b)(6) 2025 and the impella was removed on (b)(6) 2025 without incident. clinical symptoms occurred prior to impella 5.5 placement, other than hemodynamic instability and bleeding, which cannot be directly attributed to the impella device, and more likely due to post cardiac arrest and post-operative post cardiotomy cardiogenic shock. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. clinical symptoms occurred prior to impella 5.5 placement, other than hemodynamic instability and bleeding, which cannot be directly attributed to the impella device, and more likely due to post cardiac arrest and post-operative post cardiotomy cardiogenic shock.
445 | 1220648-2026-03216 | 24341586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341586&pc=OZD | 02/06/2026 | 02/15/2026 |  | 02/15/2026 |  |  | 02/06/2026 | 02/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025589143 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
446 | 1220648-2026-03221 | 24341595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341595&pc=OZD | 02/07/2026 | 02/15/2026 |  | 02/15/2026 |  |  | 02/09/2026 | 02/15/2026 |  |  |  |  | 02/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781918 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp was inserted via the left femoral artery in a 72-year-old male with acute myocardial infarction (ami) and cardiogenic shock (cgs). the patient¿s clinical state prior to initiation of support was unspecified. the procedure was aborted within 30 minutes due to an inability to advance the impella catheter. according to the clinical team, double stents had been placed in both inguinal arterial channels, and the impella delivery catheter could not be passed through the vasculature. no excessive force was applied, a guidewire was used, and the approximate insertion angle was 45°. resistance was encountered at the femoral artery, and the vessel characteristics were described with an estimated vessel diameter of <4 mm, limiting compatibility for a cp device. no deterioration in the patient¿s condition occurred, and no additional mcs was required. the device was removed, and the patient survived. the delivery failure was caused by patient specific vascular anatomy, particularly the presence of bilateral inguinal stents and small vessel diameter, which prevented advancement of the impella catheter. the implantation attempt was appropriately aborted without patient harm. the impella cp will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the removal; however the removal was performed with no consequence to the patient and support. | [5 empty] | Patient's height and weight are unknown. the impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
447 | 1220648-2026-03186 | 24340861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24340861&pc=OZD | 11/10/2025 | 02/26/2026 |  | 02/14/2026 |  |  | 11/10/2025 | 02/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026733495 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
448 | 1220648-2026-03187 | 24340862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24340862&pc=OZD | 11/11/2025 |  |  | 02/14/2026 |  |  | 11/11/2025 | 02/14/2026 |  |  |  |  | 02/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026738697 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A forty-four-year-old female patient underwent escalation of temporary mechanical circulatory support from impella cp to impella 5.5 after a complex clinical course. the patient had a long-standing medical history significant for prior strokes, thrombotic events, and bleeding episodes approximately ten years before the initiation of temporary mechanical circulatory support. patient support pathway included ecmo, subsequently ecmo + impella cp, and ultimately ecmo + impella 5.5. during the course of support, a cranial ct scan revealed a subdural hematoma (sdh). based on the severity of findings and overall prognosis, care was withdrawn. the patient later became an organ donor. given the clinical complexity, the extended use of multiple mechanical circulatory support devices, and the presence of significant pre-existing risk factors, we are conservatively reporting the patient¿s death. at this time, based on the available information, the impella 5.5 device is considered unlikely to be associated with the fatal outcome. the patient subsequently expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
449 | 1220648-2026-03189 | 24340863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24340863&pc=OZD | 11/11/2025 |  |  | 02/14/2026 |  |  | 11/11/2025 | 02/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026739095 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
450 | 1220648-2026-03190 | 24340864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24340864&pc=OZD | 11/11/2025 | 02/20/2026 |  | 02/14/2026 |  |  | 11/11/2025 | 02/14/2026 |  |  |  |  | 02/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025486107 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported additional information: "pt was transferred to ucsf for higher level of care post op day 0. chest left open for +- 6 days and closed at ucsf. recovered completely and discharged from hospital per ucsf team. device is not available for return." | A high-risk, type ii protected coronary artery bypass grafting (cabg) procedure was performed in a 60-year-old male patient. the patient had pre-existing thrombocytopenia with a baseline platelet count of 86×10 /l, which further declined during the procedure to 36×10 /l. significant postoperative bleeding occurred and required extensive transfusion support, including multiple blood products, factor vii, and kcentra, after which hemostasis was achieved. this case of multifactorial bleeding is being reported conservatively, as a device-related contribution cannot be fully excluded. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated. the investigation is still ongoing.
451 | 1220648-2026-03191 | 24340881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24340881&pc=OZD | 02/09/2026 |  |  | 02/14/2026 |  |  | 02/09/2026 | 02/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798642 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Patient Device Interaction Problem | Material Split, Cut or Torn | [5 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
452 | 1220648-2026-03193 | 24340900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24340900&pc=OZD | 02/09/2026 |  |  | 02/14/2026 |  |  | 02/09/2026 | 02/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801040 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
453 | 1220648-2026-03196 | 24341231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24341231&pc=OZD | 02/08/2026 |  |  | 02/14/2026 |  |  | 02/08/2026 | 02/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026641096 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Failure to Pump | [6 empty] | Thrombosis/Thrombus | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
454 | 1220648-2026-03113 | 24329788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24329788&pc=OZD | 12/24/2025 |  |  | 02/13/2026 |  |  | 12/24/2025 | 02/13/2026 |  |  | 01/31/2026 |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025461072 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiogenic Shock | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A seventy-one-year-old male with acute myocardial infarction complicated by cardiogenic shock (ami/cgs) and prior coronary artery bypass graft surgery required temporary mechanical circulatory support. a pulmonary artery catheter was in place prior to device escalation, and the patient was receiving inotropic medications, vasopressors, and ventilatory support. pump 1 ¿ impella cp (day 1 day 3): day 1 ((b)(6) 2025): right femoral arterial access was obtained by percutaneous approach for impella cp support (pre procedural stabilization for revascularization). day 3 ((b)(6) 2025): the patient was successfully weaned and impella cp was explanted. no impella cp complaints were documented during this interval. pump 2 ¿ impella 5.5 (escalation of therapy; right axillary/subclavian graft): day 10 ((b)(6) 2025): due to ongoing cardiogenic shock, the patient was escalated to impella 5.5 via a right axillary/subclavian surgical graft (10 mm vascutek gelweave). purge solution was dextrose 5% in water plus 25 meq/l sodium bicarbonate. the clinical course continued in the coronary care unit with inotropic and ventilatory support. during ongoing impella 5.5 support ¿ actual complaint (purge flow issue): clinical staff documented persistently low purge flow < 3.5 ml/hour despite routine troubleshooting. the medical team elected to administer tissue plasminogen activator into the purge line at physician discretion to improve purge performance and notified the local support team. the pump remained in use, and the patient continued on advanced supportive therapy. status at last documentation: the patient remained on impella 5.5 (survival outcome at explant: to be determined on support). the purge flow disturbance occurred during impella 5.5 support and was therefore device associated by timing. low purge flow is a known potential event and may reflect thrombotic or proteinaceous burden in the motor compartment, purge viscosity effects, catheter pathway constraints, or physiologic milieu in critically ill patients. in this case, the team administered tissue plasminogen activator into the purge system to restore flow and initiated a clinical investigation owing to the purge fluid change. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
455 | 1220648-2026-03114 | 24331513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24331513&pc=OZD | 02/04/2026 |  |  | 02/13/2026 |  |  | 02/04/2026 | 02/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026767252 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
456 | 1220648-2026-03115 | 24331525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24331525&pc=OZD | 02/02/2026 |  |  | 02/13/2026 |  |  | 02/03/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1455140 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Shipping Damage or Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a sprinkler went off above the automated impella controller. questionable amount of damage was done. no patient was involved. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
457 | 1220648-2026-03117 | 24331576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24331576&pc=OZD | 02/04/2026 |  |  | 02/13/2026 |  |  | 02/05/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026648400 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  |  | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: a 48 year old male patient underwent placement of an impella cp device for hemodynamic support due to acute myocardial infarction (ami) with cardiogenic shock (cgs). based on the available information, the impella cp device initially demonstrated appropriate placement signals; however, the placement signal subsequently rose above 250 mmhg without indication of device migration, as confirmed by fluoroscopy. the cause of the transient inaccurate placement signal could not be determined from the information provided. after several minutes, the signal returned to normal without repositioning or corrective action. the patient experienced significant femoral access site bleeding that necessitated surgical intervention and device removal. the relationship between the bleeding event and the impella cp device cannot be conclusively determined based on current data. access site bleeding is a known potential complication of large bore percutaneous access. bleeding is consistent with known device-related and patient-related factors documented in the instructions for use (ifu). a device related malfunction has not been confirmed. further investigation, including device return and analysis, is required to determine whether any device related factors contributed to the reported events. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
458 | 1220648-2026-03125 | 24331931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24331931&pc=OZD | 02/06/2026 |  |  | 02/13/2026 |  |  | 02/06/2026 | 02/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782830 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
459 | 1220648-2026-03131 | 24332040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24332040&pc=OZD | 02/06/2026 |  |  | 02/13/2026 |  |  | 02/06/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026761573 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in a 74-year-old male patient presenting for high-risk percutaneous coronary intervention (hrpci). the patient had a medical history of coronary artery disease and renal insufficiency, with hypertension noted as an additional contributing factor. during support, red-colored urine was observed. the patient remains on support. the reported hematuria is a known risk associated with impella support and may be due to patient-specific factors such as underlying critical illness, renal insufficiency, hemodynamic instability, and hypertension, as well as the anticoagulation and purge requirements of the device as described in the instructions for use (ifu). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
460 | 1220648-2026-03133 | 24332148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24332148&pc=OZD | 02/03/2026 |  |  | 02/13/2026 |  |  | 02/04/2026 | 02/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026804402 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
461 | 1220648-2026-03128 | 24332193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24332193&pc=OZD | 02/06/2026 | 02/13/2026 |  | 02/13/2026 |  |  | 02/06/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796220 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was placed via the right axillary artery post percutaneous coronary intervention to support a 45-year-old female patient admitted in with cardiomyopathy. the patient¿s comorbidities included lymphoma. the patient¿s clinical state prior to initiation of support was scai stage e. the pump was supporting when on day two of support there was oozing noted from the 5.5 access site and bruising on the patient's arm. a purse string suture was added to site and anticoagulation was held. the patient's partial thromboplastin time level was 30.2. flows were reported lower than expected and echocardiogram showed mosaic color flow at the outlet. pump repositioning did no improve flows. the patient received 1 unit of blood products. on day five of support, the patient received an additional 4 units of blood products for a gastrointestinal bleed and brain bleed. heparin was held, pressors were administered, and the patient received an additional two units of blood products. an exploratory laparotomy was performed. the patient was also hypovolemic and experienced ventricular tachycardia requiring 3 shocks. on day six of support, care was withdrawn and the pump explanted, and the patient expired. bleeding, vascular injury, cerebrovascular accident/stroke, and death are known risks associated with impella 5.5 support as described in the instructions for use. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
462 | 1220648-2026-03135 | 24332198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24332198&pc=OZD | 03/09/2020 |  |  | 02/13/2026 |  |  | 01/23/2026 | 02/13/2026 |  |  | 10/31/2021 |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP 371 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1460680 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Nerve Damage | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | In (b)(6) 2020 an impella cp was placed via the right femoral artery to support the 66 year old male patient who was admitted in with acute myocardial infarction and cardiogenic shock. the patient was known to have been cardioverted multiple times in the emergency room before arrival to the cardiac catheterization lab. when in the lab the team performed angioplasty for the coronary artery disease and the patient went then to the icu on the cp pump for continued hemodynamic support and monitoring. the pump was explanted after 2 days of successful support. later in (b)(6) 2026 there was an update forwarded to abiomed that noted that post explant the patient was noted to have a large retroperitoneal bleed with a lumbosacral plexopathy in the right lower extremity with weakness proximally at the hip and knee and numbness at the thigh and shin, pulmonary emboli with bilateral lower extremity dvt. the patient has received medical management for the limb harm and survived with the support of inpatient rehabilitation. vascular injury and bleeding are both known risks associated with impella support as described in the instructions for use. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
463 | 1220648-2026-03138 | 24332468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24332468&pc=OZD | 02/05/2026 |  |  | 02/13/2026 |  |  | 02/05/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026809035 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male | 75 KG | Non Hispanic | White | Death | Yes | Yes | Use of Device Problem | Device Sensing Problem | [5 empty] | Multiple Organ Dysfunction Syndrome | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A patient that was on impella cp was successfully escalated to an impella 5.5 that was inserted via axillary artery in a 74-year-old male patient for acute decompensated heart failure. the patient¿s comorbidities were unknown. the patient did have a massive infarct with known vsd prior to insertion. the patient¿s clinical state prior to initiation of support was identified as scai shock stage c. while on support the patient had a map around 70mmhg while on vav ecmo with 4.5 lpm. impella creates flow around 0.7 lpm with an ao placement signal 133/112 which does not fit to the patient's parameters. the p level on p7 is out of normal range, 0.7 lpm. the physician's theory was that the sensor may have been damaged which could result in a false high ao signal and false low impella flow. the issue was discussed with the perfusionist and physicians with the decision to make regularly echo control of impella and impella flows. the patient ultimately expired on support due to multi organ failure most likely due to the acuity of the clinical presentation | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
464 | 1220648-2026-03139 | 24332742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24332742&pc=OZD | 02/06/2026 |  |  | 02/13/2026 |  |  | 02/06/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026786790 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 95 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was surgically implanted via the right axillary artery in a 55-year-old male patient with cardiomyopathy presenting in scai staged shock requiring ecmo. prior to insertion, the patient experienced major bleeding from the axillary graft anastomosis site. surgeons applied hemostatic agents and re-sutured the graft, after which bleeding improved. impella was successfully inserted and ecmo was removed. the pump was not removed due to the event, and the patient remained hemodynamically supported. during repositioning under transesophageal echocardiogram (tee), the device unintentionally migrated back into the aorta despite confirmation of good pump position. device replacement was not requested; however, an impella cp was placed in the femoral artery to get the patient on support. the patient survived. the major bleed associated with the impella 5.5 appears related to surgical graft manipulation rather than device malfunction, as bleeding improved after surgical correction and the device subsequently functioned normally. the positioning issue of the device occurred during intraoperative manipulation. bleeding is a known risk associated with impella 5.5 per the instructions for use (ifu). | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
465 | 1220648-2026-03140 | 24333615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24333615&pc=OZD | 02/07/2026 |  |  | 02/13/2026 |  |  | 02/07/2026 | 02/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026756945 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
466 | 1220648-2026-03142 | 24333762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24333762&pc=OZD | 02/08/2026 |  |  | 02/13/2026 |  |  | 02/08/2026 | 02/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026811531 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
467 | 1220648-2026-03145 | 24334016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24334016&pc=OZD | 02/02/2026 | 02/17/2026 |  | 02/13/2026 |  |  | 02/09/2026 | 02/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED EUROPE GMBH - 3007904765 |  |  | GM |  | 1429248 |  | 0042-0000-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | [9 empty] | Malfunction | No | Yes | Application Network Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
468 | 1220648-2026-03141 | 24334028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24334028&pc=OZD | 02/08/2026 | 02/13/2026 |  | 02/13/2026 |  |  | 02/08/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026693704 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 69 YR |  | Male | 83 KG | Non Hispanic | White | Death | No | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was being placed via the axillary artery graft for the 69 year old male patient who had been admitted in with acute myocardial infarction and cardiogenic shock, presenting in scai stage e shock. the access was evaluated prior to the procedure and the team believed the vasculature to be of appropriate size. as the 5.5 was advanced the team encountered resistance and performed balloon angioplasty to allow for the pump delivery, however the delivery failed and was aborted. the calcification was not appreciated til after the aborted delivery. the team instead placed an impella cp as the patient hemodynamic status was declining. the patient expired on the cp pump when care was withdrawn. it was not shared if the delay caused/contributed to the eventual death. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's presenting native critical condition as presented in scai stage e. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
469 | 1220648-2026-03147 | 24334542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24334542&pc=OZD | 02/08/2026 |  |  | 02/13/2026 |  |  | 02/08/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026814136 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 72 KG | Non Hispanic | White | Injury | Yes | Yes | Crack | Fluid/Blood Leak | [5 empty] | Hemodynamic instability | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the left axillary artery in a 77-year-old female patient presenting with cardiomyopathy. patient's medical history included coronary artery disease. in the operating room, the purge reservoir at the red plug was found to be crushed. upon further assessment, the purge reservoir was noted to be cracked and broken at the yellow luer-lock connection, rendering the team unable to reconnect the purge line, resulting in pump failure. the reported events are fluid leak, product damage, medical device removal, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the device removal; however, the removal was performed with no consequence to the patient, and support was maintained without any known adverse events. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
470 | 1220648-2026-03149 | 24335104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24335104&pc=OZD | 02/06/2026 |  |  | 02/13/2026 |  |  | 02/06/2026 | 02/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026805669 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
471 | 1220648-2026-03152 | 24335215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24335215&pc=OZD | 02/04/2026 |  |  | 02/13/2026 |  |  | 02/05/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026812174 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. during support, a small amount of purge fluid was observed dripping from the pressure reservoir of the purge sidearm during use. the pump remained operational and continued to function as intended with stable performance parameters. no alarms were reported. the clinic evaluated the device and determined that pump exchange was not required. the device remained in use. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1, b2, h1: updated reportability per new information. h6: updated code to f29 per new information. clinical rationale: an impella cp device was inserted into the right femoral artery in a 72-year-old male patient presenting in an out-of-hospital cardiac arrest requiring cardiopulmonary resuscitation (cpr), in acute myocardial infarction (ami) and cardiogenic shock (cgs), and in scai stage c shock, on multiple inotropes and vasopressors, and was on bilevel positive airway pressure (bipap) continuous positive airway pressure (cpap) for respiratory support, prior to initiation of support. while the patient was in the intensive care unit (icu), a small amount of purge fluid was dropping out of the pressure reservoir of the purge sidearm. the pump was functioning as intended, so the physician didn't want to change the pump. five days later, the family elected to withdraw care, and the patient expired on support. the device is unlikely to have contributed to the patient's death, which was most likely due to the clinical condition of a critically ill patient with acute myocardial infarction complicated by cardiogenic shock, along with the complications of stage c shock. return status. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed.
472 | 1220648-2026-03158 | 24335497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24335497&pc=OZD | 12/05/2025 |  |  | 02/13/2026 |  |  | 12/05/2025 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026756985 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 41-year-old male with a history of chemo induced heart failure was admitted for acute decompensated heart failure requiring initial va ecmo for his cardiogenic shock. he was implanted with an impella 5.5 for left ventricular unloading and optimization for durable left ventricular assist device (lvad). after impella implant, the va ecmo was weaned off and the patient was decannulated. there was some high purge pressure noted which was treated with alteplase in the purse solution. no further information was provided. the patient was supported with impella 5.5 until he was transplanted on (b)(6) 2026. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
473 | 1220648-2026-03150 | 24335778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24335778&pc=OZD | 02/03/2026 |  |  | 02/13/2026 |  |  | 02/03/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026799383 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female | 70 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical narrative: an impella cp device was inserted into the right femoral artery in a 71-year-old female patient presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), in scai stage e shock, prior to initiation of support. the impella was inserted for ventricular support during a high-risk percutaneous coronary intervention (pci). pump 1: (b)(4) while the patient was in the intensive care unit (icu), the patient went into asystole and was treated with cardiopulmonary resuscitation (cpr) with return of spontaneous circulation (rosc). after cpr, there was suction in all p-levels and nearly no impella flow, with flat motor current (mc). a transthoracic echocardiogram (tte) was done, which showed the left ventricle (lv) filled and not unloaded. the impella was repositioned without success. it was noted that there was a kink in the cannula visible under fluoroscopy; therefore, the pump was replaced. pump 2: (b)(4) the patient was successfully replaced with another impella cp; however, several hours after the second device was implanted, the family elected to withdraw care, and the patient expired on support. the device is unlikely to have contributed to the patient's death, which was most likely due to the clinical condition of a critically ill patient with acute myocardial infarction complicated by cardiogenic shock as they presented in stage e shock. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
474 | 1220648-2026-03161 | 24337608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24337608&pc=OZD | 02/08/2026 |  |  | 02/13/2026 |  |  | 02/08/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026726981 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp device was inserted via the left femoral artery in a 79-year-old female patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). the patient was noted by the provider to be at high risk for bleeding. she had received eliquis, brilinta, and heparin. the impella was positioned via the left common femoral artery. the patient also had an arterial stick in the right common femoral artery, which was closed with a perclose device at the end of the procedure. at the impella access site, slight oozing was observed, and a pressure dressing was applied. at the right femoral access site, where the perclose device had been placed, a large hematoma developed, and the bedside nurse applied manual pressure. the last documented act was >400. the pump was explanted and patient survived. the reported hematoma is consistent with access-site complications and the anticoagulation/purge requirements associated with impella support, which are known risks described in the instructions for use. the patient¿s elevated bleeding risk may also have contributed to the development of the hematoma and it is unknown whether recommended best practices for access management were followed to mitigate the risk of bleeding. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
475 | 1220648-2026-03162 | 24337609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24337609&pc=OZD | 02/09/2026 | 02/25/2026 |  | 02/13/2026 |  |  | 02/09/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026709954 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 68 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | Hemodynamic instability | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 was inserted via direct aorta approach in a 67 year old male for planned surgery. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage a. during placement of the 5.5, the pump was positioned 5.5 cm from aortic valve area and the patient experienced some ectopy. it was decided to place the pump more shallow at 4.7 cm. the patient received support with the 5.5 for a bentall procedure with ascending hemiarch replacement and cabg. upon transfer from the operating room, the patient's heart rhythm was junctional. on day two of support, the patient had a run of ventricular tachycardia when moving to sitting position in bed. the arrhythmia was self limiting and additional short runs of 2 beats were observed when the patient coughed. echocardiogram confirmed the device measured to tip at 5.7 cm. it was decided not to reposition the device as the patient was hemodynamically stable. on day three of support, platelets were low at 48 and 47. aspirin and heparin were held. minimal surgical bleeding was reported. the pump repositioning will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the repositioning, however the reposition was performed with no consequence to the patient and support. thrombocytopenia and bleeding are known potential adverse events of the impella 5.5 per the instructions for use. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
476 | 1220648-2026-03164 | 24337767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24337767&pc=OZD | 02/09/2026 | 02/13/2026 |  | 02/13/2026 |  |  | 02/09/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026801848 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female | 65 KG | Non Hispanic | White | Injury | Yes | No | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in an 82-year-old female patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). patient had a history of coronary artery disease and diabetes mellitus. during the right axillary approach, the impella 5.5 was unable to be advanced. a 10 mm graft was placed, and a 0.035" wire with an al1 catheter was successfully crossed across the aortic valve. the 0.018" impella wire was then advanced across the valve; however, the impella 5.5 could not be advanced beyond the innominate artery. placement was aborted. the mid-cannula was noted to have a slight crease and prolapse at that location. the reported events include failure to advance, product damage, medical device removal, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was maintained. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the reported events include failure to advance, product damage, medical device removal, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was maintained.
477 | 1220648-2026-03168 | 24338018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24338018&pc=OZD | 01/15/2026 |  |  | 02/13/2026 |  |  | 01/17/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026759484 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 30 YR |  | Male | 80 KG |  |  | Injury | Yes | No | Malposition of Device | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp device was inserted into the left femoral artery in a 30-year-old male patient presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), and in scai stage c shock, on an extracorporeal membrane oxygenation (ecmo), prior to initiation of support. while the patient was in the intensive care unit (icu), the pump became dislocated due to the patient's anatomy and could not be repositioned. the impella was exchanged for a new impella cp. the new impella cp functioned without issues. the post-procedure outcome is survived at explant. the impella cp will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | Patient's race and ethnicity is unknown. the impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
478 | 1220648-2026-03177 | 24338644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24338644&pc=OZD | 02/08/2026 |  |  | 02/13/2026 |  |  | 02/08/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026814039 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 63 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella cp was inserted via right femoral artery post percutaneous coronary intervention in a 63-year-old male who was admitted for acute myocardial infarction with cardiogenic shock. the patient¿s comorbidities were not included. the patient¿s clinical state prior to initiation of support was scai stage c. on day three of support, the patient self extubated and then pulled the impella out. during closure there was a suspected clot in the femoral artery. thrombus is a known potential adverse event of the impella cp per the instructions for use. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
479 | 1220648-2026-03179 | 24338715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24338715&pc=OZD | 02/05/2026 | 02/13/2026 |  | 02/13/2026 |  |  | 02/09/2026 | 02/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2026797539 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
480 | 1220648-2026-03180 | 24339390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24339390&pc=OZD | 02/02/2026 | 02/18/2026 |  | 02/13/2026 |  |  | 02/09/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026815534 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male | 70 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient's renal function subsequently returned to its original level, and spontaneous urination resumed, allowing him to discontinue dialysis. impella was successfully removed on february 10th. | Clinical rationale: a 77 year old male in pre support clinical state consistent with scai cardiogenic shock stage d underwent impella 5.5 implantation for off pump coronary artery bypass support. the day after insertion, the patient developed worsening renal function with anuria. given baseline chronic renal impairment (egfr ~30, cre 1.5), the treating physician attributed the acute deterioration to intraoperative hypotension associated with suction events during rotation and intraoperative bleeding. the patient required two days of chdf (continuous hemodiafiltration). the renal deterioration was considered secondary to intraoperative hypotension in the setting of pre existing renal dysfunction and scai stage d shock rather than device malfunction. the aic shutdown anomaly was resolved with forced power off and did not impact patient outcome or pump performance. the patient remains on support. |  |  |  |  | B5 and d6b updated. the investigation is still ongoing. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
481 | 1220648-2026-03181 | 24340102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24340102&pc=OZD | 02/07/2026 |  |  | 02/13/2026 |  |  | 02/08/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797031 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 pump for mechanical circulatory support. it was reported that the purge system was blocked. no further information was provided. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
482 | 1220648-2026-03182 | 24340115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24340115&pc=OZD | 02/07/2026 |  |  | 02/13/2026 |  |  | 02/08/2026 | 02/13/2026 |  |  |  |  | 02/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798736 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 63 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Pain | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was inserted via right axillary artery to support a 67 year old male patient admitted with cardiomyopathy the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was identified as scai shock stage d. the patient was receiving support when on day 17 of support a hematoma was noted at the groin site. the patient was experiencing pain and it was reported the hematoma was growing in the last 24 hours, with pedal pulse present but diminished. heparin was stopped. vascular injury is a known potential adverse event of the impella 5.5 per the instructions for use. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
483 | 1220648-2026-02961 | 24321305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24321305&pc=OZD | 02/03/2026 |  |  | 02/12/2026 |  |  | 02/03/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026731217 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
484 | 1220648-2026-02962 | 24321315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24321315&pc=OZD | 01/29/2026 |  |  | 02/12/2026 |  |  | 02/04/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1149033 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | False Alarm | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during set up for a clinical procedure, an automated impella controller (aic) had controller alarms after being turned on. it was reported that the issue occurred during preparation for use, and was not connected to any devices at the time of occurrence. the aic was turned off, and a back up aic was used. there was no patient involvement reported, and therefore not patient harm reported. | [5 empty] | A4 weight is unknown e1. initial reporter first name is unknown device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed
485 | 1220648-2026-02698 | 24321370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24321370&pc=OZD | 02/03/2026 |  |  | 02/12/2026 |  |  | 02/03/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801549 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
486 | 1220648-2026-02968 | 24321446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24321446&pc=OZD | 12/02/2025 |  |  | 02/12/2026 |  |  | 12/02/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026754442 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A seventy-six-year-old male patient was scheduled for coronary artery bypass grafting (cabg) plus aortic root replacement with impella 5.5 support. post operatively, the chest was kept open, and the patient was sent to the intensive care unit (icu). the patient received blood products and factors for bleeding. systemic anticoagulation was held. the patient returned to the operating room (or) and had his chest closed successfully with the impella discontinued at that time. cannot attribute bleeding to the pump as the patient had undergone cabg with full heparinization and bypass. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
487 | 1220648-2026-02972 | 24321587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24321587&pc=OZD | 02/03/2026 |  |  | 02/12/2026 |  |  | 02/05/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026804032 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male | 80 KG |  |  | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 was inserted via a graft in the right axillary artery in a 57-year-old male patient with acute myocardial infarction (ami) and cardiogenic shock (cgs). the patient¿s clinical state prior to initiation of support was consistent with scai stage d shock. during support, the patient developed hemolysis, evidenced by elevated ldh and bilirubin levels, although no detailed laboratory values were available. the device was removed, after which the patient remained stable and laboratory markers returned to normal ranges. imaging did not confirm proper pump position during the hemolysis episode. blood products were given, repositioning did not resolve the issue, and no anatomic cardiac abnormalities were present. the patient survived to explant. based on the available information, the hemolysis was clinically associated with the impella 5.5 during active support, as evidenced by elevated hemolysis markers that resolved following device removal. the absence of anatomical abnormalities and the inability to confirm appropriate pump positioning suggest that pump position may have contributed to red blood cell trauma. hemolysis is a known risk associated with impella 5.5 as described in the instructions for use (ifu). | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
488 | 1220648-2026-02973 | 24321635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24321635&pc=OZD | 02/04/2026 |  |  | 02/12/2026 |  |  | 02/05/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026810000 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that implantation of an impella 5.5 was aborted due to patient anatomy. the pump could not pass the anastomosis of the ascending aorta. another device was placed to support the patient. | [5 empty] | A4 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
489 | 1220648-2026-02974 | 24321667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24321667&pc=OZD | 02/04/2026 |  |  | 02/12/2026 |  |  | 02/04/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026786084 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
490 | 1220648-2026-02979 | 24321760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24321760&pc=OZD | 02/05/2026 |  |  | 02/12/2026 |  |  | 02/05/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802284 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 49-year-old male patient presenting with cardiomyopathy, with a history of prior coronary artery bypass grafting, known coronary artery disease, and renal insufficiency. during rounding, urine was noted to be pink/red tinged. laboratory evaluation showed an elevated lactate dehydrogenase level of 598 and a haptoglobin level of 67. the most recent echocardiogram documented the pump positioned 6.3 centimeters from the aortic valve annulus. the pump had previously been shallow and was advanced. the nurse notified the provider, and an interventional decision is pending based on the clinical findings. the pump was repositioned and hemolysis resolved. the reported events are hemolysis and device in wrong position. the hemolysis is most plausibly related to patient-specific factors, including the underlying cardiomyopathy, critical illness, renal impairment, and potential hemodynamic instability, with possible contribution from device positioning. hemolysis is a known risk associated with impella support as described in the instructions for use. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
491 | 1220648-2026-02976 | 24321863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24321863&pc=OZD | 01/26/2026 |  |  | 02/12/2026 |  |  | 01/26/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1170099 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female | 43 KG |  |  | Malfunction | No | Yes | Appropriate Term/Code Not Available | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the automated impella controller failed to recognize a purge cassette and generated a purge fluid not recognized alarm. the cassette was replaced to resolve the issue. no patient harm was reported. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
492 | 1220648-2026-02987 | 24322316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24322316&pc=OZD | 01/21/2026 |  |  | 02/12/2026 |  |  | 01/21/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779525 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | Company representation | Distributor | [9 empty] | Malfunction | No | Yes | Shipping Damage or Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
493 | 1220648-2026-02988 | 24322334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24322334&pc=OZD | 02/04/2026 |  |  | 02/12/2026 |  |  | 02/05/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775284 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
494 | 1220648-2026-02989 | 24322439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24322439&pc=OZD | 02/04/2026 |  |  | 02/12/2026 |  |  | 02/04/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796491 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Unknown |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted via the right femoral artery in a 75-year-old patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). during the intervention, the impella cp became dislodged and required explant. the dislodgement occurred accidentally during staff manipulation. the patient remained clinically stable throughout the event. a new impella cp device was subsequently inserted, and hemodynamic support was successfully re-established. the reported events include device in wrong position, device revision or replacement, and hemodynamic instability. the impella cp will be conservatively reported for serious injury due to a temporary interruption in hemodynamic support during the exchange; however, the exchange was performed without consequence to the patient, and support was resumed with no known adverse patient outcomes. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
495 | 1220648-2026-02994 | 24322589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24322589&pc=OZD | 02/05/2026 |  |  | 02/12/2026 |  |  | 02/05/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026753279 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 50 KG |  |  | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was placed via direct aortic access in a 70-year-old patient per elective placement for support during a high-risk coronary artery bypass graft surgery. scai classification prior to impella placement is unknown. medical history includes known coronary artery disease. while on impella support, the patient¿s urine became red with decreasing urine output, elevated creatinine and worsening plasma free hemoglobin value (105). an echocardiogram was completed and the impella was repositioned. the following day the patient¿s urine was clear yellow with a decrease in plasma free hemoglobin (30.9). the creatinine continued to rise, and the patient was subsequently treated with dialysis. the reported hemolysis is a known risk described in the instructions for use. hemolysis may lead to renal failure requiring dialysis. the patient survived. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
496 | 1220648-2026-02995 | 24322661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24322661&pc=OZD | 02/05/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2027829719 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male | 74 KG |  |  | Injury | Yes | No | Decrease in Pressure | Increase in Pressure | Patient Device Interaction Problem |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp device was inserted into the left femoral artery in a 58-year-old male patient presenting in scai stage d shock, prior to initiation of support. the impella was inserted for ventricular support during an electrophysiology ablation (ep ablation). while the patient was in the intensive care unit (icu), there was increasing hemolysis since transfer from the cath lab. plasma free hemoglobin (pfhg) levels were 143, 376, 744, 862, and 988 mg/dl. positioning was verified several times using a transthoracic echocardiogram (tte) and repositioned. the impella was functioning at p-2 at the time of explant. the patient needed to be treated with circulatory support medication. after explant of the pump, the pfhg dropped to 3 mg/dl. in addition, there was a sudden drop in purge flow from 13ml/h to 2.8ml/h and an increase in purge pressure from 400 to 875 mmhg. the post-procedure outcome is survived at explant. poor positioning of the impella can cause hemolysis. hemolysis is listed as a potential adverse event, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
497 | 1220648-2026-03000 | 24322788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24322788&pc=OZD | 10/29/2025 |  |  | 02/12/2026 |  |  | 10/29/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745459 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 57-year-old male with severe cardiogenic shock following an anterior stemi underwent pci and was initially supported with an impella cp in the catheterization laboratory. due to subsequent oxygenation impairment, va-ecmo support was initiated. the patient continued to demonstrate hemodynamic instability, and prior to escalation of mechanical circulatory support, a pericardial window and drain were required due to tamponade occurring after pci. the patient was then escalated to impella 5.5 support via axillary access, and both va-ecmo and the impella cp were removed. during ongoing management, right-sided circulatory support became necessary, and a vp-ecmo configuration was initiated. complications related to the previously ecmo cannulated limb, including limb ischemia, required surgical intervention, multiple blood transfusions, and further clinical treatments. despite prolonged support and stabilization efforts, sedation weaning revealed absence of neurological response. no pathological diagnosis or confirmatory neurologic data were reported. on day 12, during ongoing advanced mechanical circulatory support, the patient expired. no additional clinical details were provided. the event is being conservatively reported due to the fatal outcome; however, based on the information available, a causal relationship to the impella 5.5 appears unlikely, given the overall complexity of the patient¿s condition, multiple comorbid events, and the use of several concurrent support modalities. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
498 | 1220648-2026-03001 | 24322898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24322898&pc=OZD | 12/05/2025 |  |  | 02/12/2026 |  |  | 12/05/2025 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026716259 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
499 | 1220648-2026-03004 | 24322938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24322938&pc=OZD | 10/30/2025 |  |  | 02/12/2026 |  |  | 10/30/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026733575 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Pleural Effusion | Sepsis | Hypovolemia | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 64-year-old male with acute on chronic heart failure, intolerant to gdmt, had a recent pea arrest and was supported with ecmo and iabp. impella 5.5 was implanted and ecmo/iabp were removed successfully. a pleural drain was placed. during support, the patient developed an infection of unclear origin and aki without crrt. the infection progressed to sepsis with fluid imbalance. the patient coded, was briefly resuscitated, and care was withdrawn two days later. the fatal outcome is reported conservatively; a device relationship appears unlikely in this complex case. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
500 | 1220648-2026-03003 | 24322939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24322939&pc=OZD | 02/05/2026 |  |  | 02/12/2026 |  |  | 02/05/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026763599 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted via the femoral artery in a patient presenting with ami/cgs and a history of prior cabg. the patient was supported with ecmella. following re-thoracotomy, heparin was paused, and a high purge pressure alarm was noted. the pump was subsequently explanted, and ecmo support was continued. the reported events are purge pressure high, medical device removal, hemostasis, and hemodynamic instability. the impella cp will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the device removal; however, the removal was performed with no consequence to the patient and support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
501 | 1220648-2026-03006 | 24323112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323112&pc=OZD | 11/04/2025 | 02/20/2026 |  | 02/12/2026 |  |  | 11/04/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026655572 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Death | Yes | No | Increase in Pressure | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Asystole | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | "we are conservatively reporting a highly complex case involving a 68-year-old male who presented for tavr. during initiation of the procedure, the patient developed asystole and cardiac arrest, requiring cpr. after returning to spontaneous circulation and initiation of ecmo tavr was completed. post-procedure, the patient required multiple defibrillation shocks for ventricular irritability and developed significant bleeding from the groins, rib fractures, and retroperitoneum, with ct confirming diffuse intra-abdominal bleeding. for lv-unloading he was taken to the or for impella5.5 placement, followed by ir coiling. despite these interventions, the patient expired while on support. in this complex setting of active bleeding, ecmo-related anticoagulation, and profound instability, impella5.5 implantation did not result in hemodynamic or bleeding stabilization. based on available information, no direct causal relationship to the impella5.5 has been identified. " | The complainant reported additional information: "the pump is not available for return. i cannot recall the requested information as the ca was from (b)(6)." |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 udpated. the investigation is still ongoing.
502 | 1220648-2026-03009 | 24323284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323284&pc=OZD | 12/05/2025 |  |  | 02/12/2026 |  |  | 12/05/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026770525 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 96 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A fifty eight year old male patient with a history of diabetes mellitus and coronary artery disease received an impella 5.5 device as support for high risk coronary artery bypass grafting. the patient was receiving two inotropic medications prior to device implantation. coronary artery bypass grafting was successfully completed with impella assistance, and the patient was transferred to the surgical intensive care unit for monitoring. on the following day, the patient developed moderate bleeding and was taken to the operating room, where surgical exploration identified bleeding from the right internal mammary artery, which was repaired. this bleeding was considered related to the coronary artery bypass procedure but will conservatively be coded to the impella device as the required continuous anticoagulation might have aggravated the bleeding. following surgical intervention, bleeding remained controlled, and the patient continued on impella support for eight days when the device was successfully removed. | [5 empty] | A fifty eight year old male patient with a history of diabetes mellitus and coronary artery disease received an impella 5.5 device as support for high risk coronary artery bypass grafting. the patient was receiving two inotropic medications prior to device implantation. coronary artery bypass grafting was successfully completed with impella assistance, and the patient was transferred to the surgical intensive care unit for monitoring. on the following day, the patient developed moderate bleeding and was taken to the operating room, where surgical exploration identified bleeding from the right internal mammary artery, which was repaired. this bleeding was considered related to the coronary artery bypass procedure but will conservatively be coded to the impella device as the required continuous anticoagulation might have aggravated the bleeding. following surgical intervention, bleeding remained controlled, and the patient continued on impella support for eight days when the device was successfully removed.
503 | 1220648-2026-03013 | 24323420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323420&pc=OZD | 02/06/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026742688 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted into the right subclavian artery in a 57-year-old female patient with a past medical history of renal insufficiency, presenting in cardiomyopathy, and in scai stage d shock, on an intra-aortic balloon pump (iabp), on an extracorporeal membrane oxygenation (ecmo), on multiple inotropes and vasopressors, prior to initiation of support. while the patient was in the intensive care unit (icu), the aic controller (aic) placement signal was reading out of range and not correlating with the patient's arterial line. a transthoracic echocardiogram (tte) was ordered. the impella was repositioned and moved away from the papillary muscle, sitting in the mid-ventricular space at 4cm. at this point, the placement signal was still reading high. a "placement signal not reliable" alarm was also noted on the console. the physician was aware and muted the audio on the alarm. the impella was already repositioned multiple times the previous day. the plasma free hemoglobin (pfhgb) increased to 290, and hemolysis noted in the urine post-repositioning. the pump was repositioned again at night, and the urine color started improving. pfhgb drawn again, and was down to 50 in the morning. the post-procedure outcome is survived at explant. the impella functioned at p-8 at 4.0l/min with no issues. poor positioning of the impella can cause hemolysis. hemolysis is listed as a potential adverse event, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). additional information: the pump was replaced. this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella aic controller. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
504 | 1220648-2026-03014 | 24323458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323458&pc=OZD | 10/29/2025 |  |  | 02/12/2026 |  |  | 10/29/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026699548 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  |  |  | Injury | Yes | No | Increase in Pressure | [6 empty] | Hematoma | Tachycardia | Hypovolemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 35-year-old male patient with unknown medical history presented with suspected pneumonia and subsequently decompensated into cardiogenic shock. impella 5.5 support was initiated. following the procedure, ectopic beats were observed. evaluation showed that the impella inlet appeared positioned too deep within the left ventricle. the device was repositioned, which led to partial improvement; however, the patient continued to experience intermittent episodes of ventricular tachycardia (vt). during support one episode hypovolemia was solved with albumin. inseration showed a small haematome, no treatment was necessary. the patient was subsequently weaned from impella support as planned, without further reported complications. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
505 | 1220648-2026-03015 | 24323476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323476&pc=OZD | 02/06/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026814192 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
506 | 1220648-2026-03016 | 24323497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323497&pc=OZD | 10/30/2025 |  |  | 02/12/2026 |  |  | 10/30/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026688210 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Heart Failure/Congestive Heart Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 75-year-old female patient undergoing an elective high-risk aortic valve replacement (ef 25%) received planned perioperative support with an impella 5.5 device. the patient had pre-existing anemia, and periprocedural bleeding¿including gastrointestinal bleeding¿necessitated multiple transfusions of red blood cell units and one unit of platelets. during support, the patient¿s known atrial fibrillation led to episodes of rapid ventricular response, resulting in biventricular failure and requiring escalation to ecls therapy. after stabilization, successful weaning from temporary mechanical circulatory support was achieved on day 12. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
507 | 1220648-2026-03018 | 24323655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323655&pc=OZD | 02/06/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1468305 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | [10 empty] | Injury | Yes | Yes | Failure to Read Input Signal | Priming Problem | [5 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
508 | 1220648-2026-03017 | 24323658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323658&pc=OZD | 10/31/2025 |  |  | 02/12/2026 |  |  | 10/31/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025481862 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Renal Failure | Tachycardia | Cardiac Tamponade | Cardiogenic Shock | Obstruction/Occlusion | Cardiac Perforation | Multiple Organ Dysfunction Syndrome | Thrombocytopenia | Heart Failure/Congestive Heart Failure |  | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On (b)(6) 2025, a 63-year-old male presented with acute decompensated heart failure and cardiogenic shock, found to have a flail posterior mitral valve leaflet with severe mitral regurgitation, status post mitraclip and intra-aortic balloon pump (iabp) insertion on (b)(6) 2025. the patient¿s course was complicated by worsening pulmonary edema, hemoptysis, and ventricular tachycardia, requiring amiodarone, intubation, and escalating pressor/inotrope support. iabp was removed and impella 5.5 placed, with subsequent hemolysis, left lower extremity deep venous thrombosis dvt, and inability to anticoagulate due to ongoing bleeding. the patient underwent mitral valve replacement (mvr) on (b)(6) 2025 with new impella 5.5 implantation. postoperatively, complications included two episodes of tamponade requiring surgical washout, continuous renal replacement therapy (crrt), cardiogenic shock, multiorgan failure, thrombocytopenia with suspected heparin-induced thrombocytopenia (hit), recurrent right ventricular failure, and repeated impella repositioning. device use was associated with malposition, hemolysis, and suspected vascular access-related thrombosis. ultimately, the patient was trached. they were anuric, however, and were improving hemodynamically. impella was able to be removed on (b)(6) 2025. as this patient was admitted with cardiogenic shock, scai stage e, it is expected that they have multiple comorbidities and multisystem organ failure. as there were 2 impella 5.5 pumps used to support this patient, it may be assumed that the hemolysis was a complication directly impacted by the pump. the patient also underwent mvr surgery after being septic and fungemic, so the infection and thrombosis issues were most likely secondary to that, including the major bleeding issue of cardiac tamponade with chest washout, causing anemia and the requirement of blood transfusions. cardiac perforation directly related to impella 5.5 pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
509 | 1220648-2026-03019 | 24323718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323718&pc=OZD | 11/03/2025 | 02/23/2026 |  | 02/12/2026 |  |  | 11/03/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737084 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 52-year-old female patient on hemodialysis underwent a planned high-risk coronary artery bypass graft (cabg) procedure and received impella 5.5 support for high-risk cardiac surgery (type 2 support) during the operation. post-procedure, the patient experienced an intracranial bleed and stroke, which required a craniotomy. despite these complications, the patient was successfully weaned from impella 5.5 in a timely manner and as planned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
510 | 1220648-2026-03021 | 24323826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323826&pc=OZD | 02/05/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779000 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
511 | 1220648-2026-03022 | 24323837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323837&pc=OZD | 02/05/2026 |  |  | 02/12/2026 |  |  | 02/05/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1129885 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Battery Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: it was reported that an automated impella controller (aic) console had a battery failure. no further information regarding the case was provided. the automated impella controller will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | A. patient information is unknown device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
512 | 1220648-2026-03023 | 24323884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323884&pc=OZD | 02/06/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025491318 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 117 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 device was inserted directly into the left aorta in a 58-year-old male patient with a past medical history of coronary artery disease (cad). the patient presented in scai stage d shock, and on a ventilator for respiratory support prior to initiation of support. the impella was inserted for ventricular support during a high-risk surgery: coronary artery bypass surgery (cabg). after the cabg, the patient arrived at the intensive care unit (icu), with a closed chest. the physician decided to bring the patient back to the operating room (or) to uncover and evacuate the chest. the patient returned to the icu with the chest left open. the post-procedure outcome is unknown. the impella functioned at p-7 at 4.1l/min with no issues. high-risk surgeries require intense blood thinners, increasing the risk of both active bleeding and postoperative blood loss. bleeding (hematoma) is also a known risk due to the impella's anticoagulation and purge requirements. bleeding is listed as a potential adverse event, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | D6b: explant date is unknown.
513 | 1220648-2026-03024 | 24323943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323943&pc=OZD | 02/05/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026803146 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted into the right subclavian artery in a 54-year-old female patient with a past medical history of coronary artery disease (cad) and diabetes, presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), and in scai stage c shock, on multiple inotropes and vasopressors, and was on a ventilator for respiratory support prior to initiation of support. while the patient was in the intensive care unit (icu), the patient had a suspected ischemic stroke. anticoagulation was stopped and tissue plasminogen activator (tpa) was initiated. the patient was sent for a computed tomography (ct) scan and an ischemia/symbolic stroke was ruled out. tpa treatment was continued. the post-procedure outcome is unknown. stroke is listed as a potential adverse event, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed
514 | 1220648-2026-03026 | 24323993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24323993&pc=OZD | 11/03/2025 |  |  | 02/12/2026 |  |  | 11/03/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026753236 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Ischemia | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 56-year-old male patient with acute myocardial infarction complicated by cardiogenic shock was undergoing weaning from circulatory support when he developed ventricular tachycardia (vt) and experienced cardiac arrest. the patient was successfully resuscitated, and the clinical team proceeded with immediate implantation of an impella 5.5 device. during impella support, the patient experienced intermittent episodes of vt. in addition, heparin-induced thrombocytopenia with thrombosis (hitt) was diagnosed, and absent distal lower extremity pulses were noted. no specific treatment was reported other than the administration of two units of platelets. on day three of support, blood cultures returned positive. weaning from impella was initiated but was ultimately unsuccessful, and the patient expired. this event is being reported conservatively, although a relationship to the device appears highly unlikely given the complexity of the clinical course and multiple concomitant complications. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
515 | 1220648-2026-03020 | 24324010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324010&pc=OZD | 02/03/2026 |  |  | 02/12/2026 |  |  | 02/04/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026778182 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. it was reported that no placement signal was displayed on the system. the physician confirmed that the pump position had been verified using fluoroscopy and was correctly positioned in the ventricle. at the time of the call, the pump was operating with a flow rate of 3.5 l/min. the investigation revealed that the connector had not been properly plugged in. after firmly reinserting the connector, both the placement signal and the lv (left ventricular) signal were immediately restored. the issue was completely resolved on site. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
516 | 1220648-2026-03027 | 24324073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324073&pc=OZD | 02/06/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798609 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Aortic dissection | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
517 | 1220648-2026-03028 | 24324090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324090&pc=OZD | 11/03/2025 |  |  | 02/12/2026 |  |  | 11/03/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026688240 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 54-year-old male patient underwent impella 5.5 implantation to support a protected high-risk cardiac surgery (type¿ii protected high-risk procedure). postoperatively, the patient developed ventricular tachycardia (vt), requiring re-opening of the sternotomy. the patient subsequently entered sustained vt, necessitating paddle cardioversion and administration of lidocaine boluses for arrhythmia management. due to ongoing hemodynamic instability, the patient was transitioned to vva ecmo support. with this management approach, the patient subsequently experienced a successful mechanical circulatory support course with timely, appropriate, and successful weaning. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
518 | 1220648-2026-03030 | 24324169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324169&pc=OZD | 11/04/2025 |  |  | 02/12/2026 |  |  | 11/04/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026736327 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 68-year-old female was admitted with an acute exacerbation of congestive heart failure. an intra-aortic balloon pump (iabp) was initiated, followed by escalation to impella 5.5 support as a bridge to recovery and ultimately to durable lvad therapy. during an otherwise successful course of impella support, the clinical team encountered persistently elevated purge pressures that did not resolve despite purge system exchanges. pump thrombosis was suspected, and the institution¿s standard thrombolytic (lyse) protocol was initiated. following administration, purge pressure abnormalities resolved. during the subsequent impella removal and lvad implantation procedure, the patient was found to have severe aortic insufficiency. an aortic valve replacement (avr) was performed in the same surgical session. the patient proceeded with lvad implantation as planned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
519 | 1220648-2026-03032 | 24324315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324315&pc=OZD | 12/07/2025 |  |  | 02/12/2026 |  |  | 12/07/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773832 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Tachycardia | Aortic dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 69-year-old female patient with a history of coronary artery disease underwent elective high-risk coronary artery bypass graft surgery and received an impella 5.5 device for mechanical circulatory support. on the fifth day of support, the patient experienced an episode of ventricular tachycardia, which was managed successfully with anti-arrhythmic medication. subsequently, the patient reported inability to move her left upper arm, prompting a computed tomography scan. the scan did not reveal any neurological abnormality but incidentally identified a 3.5-centimeter tubular aortic pseudoaneurysm adjacent to the impella device. the treating physician was uncertain whether the pseudoaneurysm was present prior to device placement. additionally, the patient developed hemolysis, evidenced by red-colored urine and elevated plasma free hemoglobin levels. after a minor adjustment of the impella device position, the hemolysis resolved by the following day. the device was successfully explanted after 16 days of support, and the patient demonstrated recovery of native heart function. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
520 | 1220648-2026-03033 | 24324381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324381&pc=OZD | 01/31/2026 |  |  | 02/12/2026 |  |  | 02/02/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026807744 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Female |  |  |  | Injury | Yes | No | Malposition of Device | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical narrative: a 47-year-old female with an unknown medical history received an impella cp device for temporary mechanical circulatory support due to cardiogenic shock secondary to an acute myocardial infarction. the device was implanted via the femoral artery in accordance with the impella instructions for use. at the time of device initiation, the patient was classified as society for cardiovascular angiography and interventions (scai) stage e cardiogenic shock, and the final classification remained stage e. the patient arrived at the cardiac catheterization laboratory with active cardiopulmonary resuscitation in progress. the impella cp device was initially implanted without incident. following implantation, during removal of the introducer sheath, the physician inadvertently displaced the impella pump from the left ventricle and was unable to reposition the device. the device was explanted, and a new impella cp device (sn649207) was successfully implanted. the initial impella cp device did not exhibit a device malfunction. the second impella cp, there was no report or indication of device malfunction and/or no impella related adverse event and remained in use. a little over six hours later, care was withdrawn, and the patient expired. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
521 | 1220648-2026-03034 | 24324404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324404&pc=OZD | 12/07/2025 |  |  | 02/12/2026 |  |  | 12/07/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026679257 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Increase in Pressure | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 51-year-old female patient with a history of diabetes mellitus and hypertension received an impella 5.5 device for management of acute decompensated heart failure with cardiogenic shock as a bridge to heart transplantation. the patient was also supported with an extracorporeal membrane oxygenation device. on the fourth day of treatment, systemic anticoagulation was discontinued due to a surgical intervention. as a contribution of the impella to the development of stroke cannot be ruled out, it will be coded to the device. on the sixteenth day of support, the patient developed a stroke. as of the latest update, the patient remains on mechanical circulatory support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
522 | 1220648-2026-03036 | 24324458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324458&pc=OZD | 02/06/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1249361 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
523 | 1220648-2026-03037 | 24324523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324523&pc=OZD | 12/07/2025 |  |  | 02/12/2026 |  |  | 12/07/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746974 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 67 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A seventy-one-year-old male patient with a history of coronary artery disease received an impella 5.5 device for post-cardiotomy cardiogenic shock. prior to impella implantation, the patient was receiving mechanical ventilation and inotropic support. the device was successfully implanted in the hybrid operating room. on the day after implantation, the patient developed cardiac ectopy and an elevation in plasma-free hemoglobin, indicating hemolysis. the impella position was adjusted, after which both the hemolysis and the ectopy resolved. the patient remained on impella support for a total of eleven days and demonstrated gradual clinical improvement throughout this period. after eleven days, native cardiac recovery was observed, and the impella device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
524 | 1220648-2026-03040 | 24324562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324562&pc=OZD | 12/08/2025 |  |  | 02/12/2026 |  |  | 12/08/2025 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745072 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Pumping Stopped | [6 empty] | Cardiac Arrest | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
525 | 1220648-2026-03039 | 24324585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324585&pc=OZD | 02/03/2026 |  |  | 02/12/2026 |  |  | 02/03/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782520 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiogenic Shock | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinician narrative: an impella cp was utilized in a 65-year-old patient of unspecified sex who presented with acute myocardial infarction complicated by cardiogenic shock (ami-cgs). the patient¿s past medical history was not provided, and scai stage was not specified. femoral arterial access was obtained. according to the implant and explant report, the patient was in cardiogenic shock and underwent resuscitation both prior to and following device implantation. prior to impella placement, the patient required triple inotropic support and dual vasopressor therapy. following implantation, inotropic and vasopressor requirements were reduced to two inotropes and one vasopressor, indicating transient hemodynamic support while on device therapy. despite ongoing support, the patient experienced cardiac shock, resuscitation, and subsequent clinical deterioration. the patient remained on impella cp support for approximately 20 hours and expired in the procedural area due to a cardiac cause. after a comprehensive review of the available information, the reporting event did not identify evidence suggesting a causal relationship between the impella cp and the reported event. the patient¿s death is being reported conservatively in the setting of refractory cardiogenic shock and critical illness. the patient expired. | [5 empty] | Section a is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
526 | 1220648-2026-03042 | 24324634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324634&pc=OZD | 12/08/2025 |  |  | 02/12/2026 |  |  | 12/08/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026738702 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 103 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Tamponade | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A forty-nine-year-old male patient with a history of type 2 diabetes mellitus, coronary artery disease, ventricular tachycardia, ventricular fibrillation, and renal insufficiency received an impella 5.5 device for management of acute myocardial infarction cardiogenic shock. the patient presented in stage d of the society for cardiovascular angiography and interventions shock classification and was supported with an intra-aortic balloon pump, mechanical ventilation, and two inotropic medications before device implantation. after impella support was initiated, the patient demonstrated hemodynamic improvement with reduced need for inotropic therapy. prior to implantation, the patient experienced recurrent ventricular tachycardia and ventricular fibrillation with multiple shocks delivered by the automated implantable cardioverter defibrillator. on the tenth day of support, the device again delivered a shock to terminate arrhythmia, and the electrophysiology team recommended ventricular tachycardia ablation. on the nineteenth day of support, during attempted ablation in the electrophysiology laboratory, the electrophysiology physician caused cardiac perforation with a wire, resulting in significant pericardial effusion requiring surgical intervention and subsequent transfer to the cardiovascular operating room for a pericardial window; the ablation was not performed. while not related to the impella device, the effusion will be conservatively coded to the impella device as the required continuous anticoagulation might have aggravated the bleeding. after thirty-nine days of impella support, the patient demonstrated native heart recovery, and the device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
527 | 1220648-2026-03043 | 24324646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324646&pc=OZD | 02/06/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1477997 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [12 empty] | Other | [24 empty] | R |  |  |  | Initial submission | [9 empty] | A purge disc was inserted into the purge disc receiver; however, the receiver did not detect the purge disc when inserted. the reported events are purge disc not detected, medical device removal, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was restored without any known adverse events. | [5 empty] | A: patient information is unknown. e: initial reporter information is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
528 | 1220648-2026-03045 | 24324741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324741&pc=OZD | 02/03/2026 |  |  | 02/12/2026 |  |  | 02/03/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026772528 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male | 89 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinician narrative: an impella 5.5 was inserted via the right axillary artery in an 86-year-old male presenting with acute myocardial infarction complicated by cardiogenic shock, scai stage c. the patient had a known history of coronary artery disease and required inotropic and vasopressor support for hemodynamic management. during the procedure, advancement of the impella 5.5 across the aortic valve was unsuccessful due to the patient¿s anatomy. the device was subsequently removed, and an impella cp was placed for continued mechanical circulatory support. the patient experienced hemodynamic instability temporally associated with the inability to advance the impella 5.5 and the need for device exchange. the instability occurred in the context of acute myocardial infarction, underlying cardiogenic shock, and critical illness requiring vasoactive support. after a comprehensive review of the available information, the reported hemodynamic instability is consistent with the patient¿s underlying cardiogenic shock and critical clinical condition during attempted device placement and exchange. the patient survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
529 | 1220648-2026-03046 | 24324815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24324815&pc=OZD | 12/16/2025 |  |  | 02/12/2026 |  |  | 12/16/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775688 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 164 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A thirty-four-year-old male patient with a history of renal insufficiency received an impella 5.5 device for treatment of heart failure and cardiogenic shock secondary to non-ischemic cardiomyopathy. the patient was receiving two inotropic medications and mechanical ventilation prior to device implantation, indicating severe clinical status. after impella support, the patient¿s hemodynamic condition gradually improved, and he was successfully transitioned to a durable left ventricular assist device on the ninth day of support. on the seventh day of support, the patient developed tar-colored stool with a mild decrease in hemoglobin. anticoagulation therapy with heparin was continued without interruption, and no further complications were reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
530 | 1220648-2026-03059 | 24326546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24326546&pc=OZD | 12/06/2025 |  |  | 02/12/2026 |  |  | 12/06/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026717826 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 61-year-old female patient with a history of coronary artery disease and renal insufficiency presented with acute myocardial infarction complicated by cardiogenic shock. the patient was escalated from an impella cp device to an impella 5.5 device and was in stage e shock as per the society for cardiovascular angiography and interventions classification at the time of impella 5.5 implantation. one day after implantation, the patient developed atrial fibrillation, which required direct current cardioversion and initiation of anti-arrhythmic medication. the arrhythmia was successfully corrected following these interventions. after a total of 7.5 days of impella 5.5 support, the patient achieved complete native heart recovery, and the device was successfully explanted. 7.5 days on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
531 | 1220648-2026-03066 | 24326961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24326961&pc=OZD | 12/07/2025 |  |  | 02/12/2026 |  |  | 12/07/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026688936 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hematoma | Hemolysis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 63-year-old male patient with a history of renal insufficiency and dialysis received an impella 5.5 device for treatment of acute decompensated heart failure caused by cardiomyopathy. the patient has been on impella support for 39 days as of today. on the second day of support, the patient developed a small access site hematoma, which resolved without further complications. after one week of support, the patient exhibited elevated plasma free hemoglobin levels, indicating hemolysis, which was later resolved. three weeks after initiation of support, the patient experienced an episode of arrhythmia identified as bigeminy, accompanied by chest pain. the patient was administered magnesium sulfate as per physician orders, after which both the bigeminy and chest pain resolved. ultimately, after an off-label prolonged support duration of 45 days, the patient was successfully bridged to transplantation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
532 | 1220648-2026-03070 | 24327366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24327366&pc=OZD | 02/06/2026 |  |  | 02/12/2026 |  |  | 02/07/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026756574 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male | 96 KG | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient, while partially sedated, removed the pump on his own and caused damage to the component near the blue suture hub. the patient is currently stable without mechanical circulatory support, and another device will be implanted later if clinically required. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
533 | 1220648-2026-03071 | 24327421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24327421&pc=OZD | 02/06/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2027816387 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male | 67 KG |  | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp device was inserted via the right femoral artery in an 85-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). an icu call was initiated when the impella was found positioned in the aorta. the physician reported that the patient, who has dementia, had pulled the impella catheter, resulting in migration of the device into the aorta. repositioning was not possible. the patient remained stable, and discussion was initiated regarding device explantation, with plans for the responsible cardiology provider to proceed with removal. the reported events are device in wrong position, medical device removal, device revision or replacement, and hemodynamic instability. the impella cp will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and full support was maintained without any known adverse events. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
534 | 1220648-2026-03072 | 24327473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24327473&pc=OZD | 10/01/2025 |  |  | 02/12/2026 |  |  | 10/01/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1281602 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | During impella cp support, users reported a controller error and placement signal issue. no other information was provided. based on reported controller error, this is a reportable malfunction. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
535 | 1220648-2026-03074 | 24327611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24327611&pc=OZD | 02/05/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026804473 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 101 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Inaccurate Flow Rate | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella cp was placed via the right femoral artery to support a 69-year-old male patient admitted in with acute myocardial infarction and cardiogenic shock, in scai stage d shock. there was no other medical history shared. on day three of cp support, the patient was bridged to impella 5.5 via right axillary artery for bridge to heart transplant. on day two of 5.5 support, the patient experienced neurological changes and head computed tomography (ct) revealed no findings. on day four of 5.5 support, the patient's urine was red. the patient was also on venoarterial extracorporeal membrane oxygenation (va-ecmo) and there was concern this could be a contributing factor. an unknown amount of blood products were transfused. the pump flow increased from 2.5l to 3l. on day eight of 5.5 support, care was withdrawn and the pump explanted, and the patient expired. hemolysis and death are known risks associated with impella cp and impella 5.5 as described in the instructions for use. this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the original impella 5.5. | [5 empty] | Section e initial reporter information was added since it was omitted from the initial. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
536 | 1220648-2026-03076 | 24327680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24327680&pc=OZD | 02/06/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796345 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
537 | 1220648-2026-03080 | 24327927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24327927&pc=OZD | 02/07/2026 |  |  | 02/12/2026 |  |  | 02/07/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775777 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
538 | 1220648-2026-03082 | 24328043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24328043&pc=OZD | 02/06/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801781 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was surgically implanted via direct left sided arterial access in a 72-year-old female patient with postcardiotomy cardiogenic shock (pccs) who presented in scai stage c. during support, the patient experienced a neurological event consistent with embolic stroke. the patient had facial drooping and was found to have an m1 clot and required neuro interventional thrombectomy for the occlusion. following intervention, the patient¿s neurological status returned to baseline. the patient also received one unit of packed red blood cells. per clinical staff, the patient had been off systemic anticoagulation in anticipation of potential explant. the patient was on heparin since device implant but was discontinued for operative washout. dual antiplatelet therapy was started the following day but was subsequently held overnight due to chest tube bleeding. the device remained in place during the stroke, and no device replacement was requested. the patient ultimately survived to explant and was successfully weaned from support. while on support, the device functioned within expected flow range at p-4 (2.2 ¿ 3.4l/min). based on the available information, the patient¿s embolic stroke occurred during impella 5.5 support; however, the absence of anticoagulation at the time of each neurological event represents a significant contributing factor and increases the likelihood of thromboembolic complications independent of device malfunction. the impella continued to function appropriately throughout support, and the patient recovered neurologically following intervention. stroke is a known risk associated with impella 5.5 per the instructions for use (ifu). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
539 | 1220648-2026-03084 | 24328364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24328364&pc=OZD | 02/07/2026 |  |  | 02/12/2026 |  |  | 02/07/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026730378 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 77 KG | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that high purge pressure occurred. the purge cassette had been discarded during a system switch, and the patient remained on support. a tissue plasminogen activator protocol was initiated and appeared to resolve the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
540 | 1220648-2026-03090 | 24328534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24328534&pc=OZD | 02/08/2026 |  |  | 02/12/2026 |  |  | 02/08/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026809821 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
541 | 1220648-2026-03091 | 24328573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24328573&pc=OZD | 02/06/2026 |  |  | 02/12/2026 |  |  | 02/06/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794830 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was inserted via direct aorta approach for elective placement in a 65-year-old male who was admitted for cabg and valve surgery. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage b. during placement, repositioning of the pump was attempted. the impella was not moving appropriately with the repositioning sheath and felt stuck. as the patient's chest was open, the pump was manually maneuvered into place. the pump repositioning will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the repositioning, however the reposition was performed with no consequence to the patient and support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
542 | 1220648-2026-03092 | 24328581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24328581&pc=OZD | 02/08/2026 | 02/18/2026 |  | 02/12/2026 |  |  | 02/08/2026 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026786889 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
543 | 1220648-2026-03095 | 24329255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24329255&pc=OZD | 12/19/2025 |  |  | 02/12/2026 |  |  | 01/22/2026 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1122790 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 62 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Pumping Stopped | Application Program Problem | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 device was inserted via the right axillary artery in a 39-year-old male patient presenting with cardiomyopathy. during support, the device experienced a pump stop associated with a controller error. the pump was successfully restarted on the first attempt, and the team proceeded with a controller exchange. no further issues were reported following the controller swap. the reported events include a temporary pump stop, controller failure, and device revision or replacement, along with hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
544 | 1220648-2026-03096 | 24329602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24329602&pc=OZD | 12/16/2025 |  |  | 02/12/2026 |  |  | 12/16/2025 | 02/12/2026 |  |  |  |  | 02/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026731161 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 95 KG |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a seventy-eight-year-old male patient with a history of prior coronary artery bypass grafting, renal insufficiency, diabetes mellitus, and coronary artery disease received an impella cp device for acute myocardial infarction¿related cardiogenic shock. prior to impella cp placement, the patient was receiving two inotropic medications and was supported with mechanical ventilation, and his society for cardiovascular angiography and interventions shock classification was stage d, indicating severe disease. due to worsening shock, the patient was escalated to an impella 5.5 device the following day. immediately after impella 5.5 insertion, the patient developed suction alarms, and fluoroscopic imaging identified a left sided hemothorax with an estimated blood loss of approximately one thousand seven hundred milliliters. the treating physician reported that the bleeding originated from left subclavian vessel access for swan ganz catheter placement performed prior to the procedure. the patient received multiple units of blood products and full reversal of heparin before transferring to the unit. while caused by the swan ganz catheter placement, major bleed will conservatively be coded to the impella 5.5 as it might have been aggravated by the required continous anticoagulation. despite interventions, the patient ultimately succumbed to his underlying disease. death was conservatively coded to the impella 5.5 while most likely to be caused by the underlying disease and unrelated to impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
545 | 1220648-2026-03100 | 24329658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24329658&pc=OZD | 12/20/2025 | 02/12/2026 |  | 02/12/2026 |  |  | 12/21/2025 | 02/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT, PKGD, JP |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT, PKGD, JP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0040-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
546 | 1220648-2026-02827 | 24310972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24310972&pc=OZD | 12/11/2025 |  |  | 02/11/2026 |  |  | 12/11/2025 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026679795 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 79 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The team elected to escalate mechanical circulatory support from impella cp to impella 5.5, which was successfully performed. approximately three days later the patient expired. | [5 empty] | Medical safety/clinical reviewed the event. the patient presented with cardiogenic shock secondary to an acute myocardial infarction and underwent placement of impella cp support. on the same day as implantation, the patient developed clinical findings concerning for hemolysis, including dark/discolored urine and increasing plasma-free hemoglobin levels. the team elected to escalate mechanical circulatory support from impella cp to impella 5.5, which was successfully performed. on the following day, the patient remained intubated and hemodynamically stable while receiving epinephrine and vasopressor support. urine output was clear and yellow; however, plasma-free hemoglobin levels remained elevated. a procedure to further treat the patient was anticipated later in the week or early the following week. device function and position, laboratory trends, hemodynamics, and overall clinical status were continuously monitored. approximately two days later, life-sustaining care was withdrawn, and the patient expired. the patients underlying critical condition most likely contributed to the death. however, conservatively, the death will be reported for the impella 5.5 as this device was in use at time of expiration. the event involving the impella cp is a serious injury. date of death: the exact date of death is unknown but has been captured as the date of explant. if actual information becomes known, a supplemental report will be submitted. note: h6: device problem/component/investigation coding remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
547 | 1220648-2026-02828 | 24312488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24312488&pc=OZD | 01/22/2026 |  |  | 02/11/2026 |  |  | 01/22/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1183021 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Application Program Problem | Device Sensing Problem |  |  |  |  | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | An impella 5.5 device was inserted into the right subclavian artery in a 67-year-old male patient with a past medical history of coronary artery disease (cad) and diabetes, presenting in cardiomyopathy, and in scai stage e shock, prior to initiation of support. while the patient was in the intensive care unit (icu), the physician was concerned with high purge pressure on the automated impella controller (aic). in addition, there was a yellow controller failure alarm. the nurse also stated that there were a couple of instances where the placement signal (ps) disappeared; however, it came back and was present when the controller failure alarm appeared. the pump and purge were successfully switched. no issues were noted after the change to the new aic. the impella functioned at p-5 at 2.5l/min. the post-procedure outcome is survived at explant with a bridge to transplant. the automated impella controller will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
548 | 1220648-2026-02820 | 24312653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24312653&pc=OZD | 02/02/2026 | 02/10/2026 |  | 02/11/2026 |  |  | 02/02/2026 | 02/11/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026811553 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Ventricular Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
549 | 1220648-2026-02831 | 24312677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24312677&pc=OZD | 02/03/2026 |  |  | 02/11/2026 |  |  | 02/03/2026 | 02/11/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1917655 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
550 | 1220648-2026-02832 | 24312695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24312695&pc=OZD | 02/03/2026 |  |  | 02/11/2026 |  |  | 02/03/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 29 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller generated a brief controller error alarm that self resolved. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
551 | 1220648-2026-02834 | 24312797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24312797&pc=OZD | 02/04/2026 |  |  | 02/11/2026 |  |  | 02/05/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1204375 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | An automated impella controller was used for a patient of unknown age or gender. patient's comorbidities and clinical state were unknown. it was reported the console did not detect the purge disc, in which there was a red alarm for purge disc not detected. the input issue will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1: product problem was inadvertently omitted on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
552 | 1220648-2026-02836 | 24312868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24312868&pc=OZD | 01/22/2026 |  |  | 02/11/2026 |  |  | 01/22/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1133728 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Priming Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | An impella cp device was inserted surgically into the right femoral artery in a 70-year-old male patient with a past medical history of coronary artery disease (cad), presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), scai stage a shock, and was on bilevel positive airway pressure (bipap)continuous positive airway pressure (cpap) for respiratory support prior to initiation of support. the impella was inserted for ventricular support during a percutaneous coronary intervention (pci) of the ramus artery. while the patient was in the intensive care unit (icu), the patient was experiencing hemolysis with frank hematuria and a plasma free hemoglobin of 1440. the device placement was confirmed to be in good position by echocardiogram. the post-procedure outcome is unknown. the patient was on a heparin drip which can cause hemolysis. hemolysis is listed as a potential adverse event and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
553 | 1220648-2026-02837 | 24312934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24312934&pc=OZD | 02/01/2026 |  |  | 02/11/2026 |  |  | 02/02/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026813503 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 47 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 47-year-old female with an unknown medical history received an impella cp device for temporary mechanical circulatory support due to cardiogenic shock secondary to an acute myocardial infarction. the device was implanted via the femoral artery in accordance with the impella instructions for use. at the time of device initiation, the patient was classified as society for cardiovascular angiography and interventions (scai) stage e cardiogenic shock, and the final classification remained stage e. the patient arrived at the cardiac catheterization laboratory with active cardiopulmonary resuscitation in progress. the impella cp device was initially implanted without incident. following implantation, during removal of the introducer sheath, the physician inadvertently displaced the impella pump from the left ventricle and was unable to reposition the device. the device was explanted, and a new impella cp device ((b)(6)) was successfully implanted. the initial impella cp device did not exhibit a device malfunction. the second impella cp, there was no report or indication of device malfunction and/or no impella related adverse event and remained in use. a little over six hours later, care was withdrawn, and the patient expired. based on the information available, there is no evidence to suggest a malfunction of either impella cp device. the patient expired while supported by the second impella cp device; however, the outcome is consistent with the patient¿s underlying condition, including cardiogenic shock classified as scai stage e. supported care was withdrawn. conservatively for usmdr, the death is reported as the second impella cp device was in use at the time of patient expiration. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
554 | 1220648-2026-02841 | 24313083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24313083&pc=OZD | 02/04/2026 |  |  | 02/11/2026 |  |  | 02/04/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026789851 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Unknown | 69 KG |  |  | Injury | Yes | No | Fluid/Blood Leak | Decrease in Pressure | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Impella cp was inserted via right femoral artery in a male of unknown age who was admitted for acute myocardial infarction with cardiogenic shock. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage e. on day 2 of support, "pp low" alarmed. leakage was at the yellow luer on the purge side arm, not at the side of the purge cassette purge line. the device was explanted and pump exchanged. the purge issue will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange, however the exchange was performed with no consequence to the patient and support. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
555 | 1220648-2026-02842 | 24313086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24313086&pc=OZD | 02/05/2026 |  |  | 02/11/2026 |  |  | 02/05/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775866 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 87 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematuria | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella5.5 was surgically implanted via the right axillary/subclavian artery in a 74-year-old male patient with acute myocardial infarction (ami) and cardiogenic shock (cgs). the patient¿s clinical state prior to initiation of support was consistent with scai stage d shock. during ongoing mechanical circulatory support, the clinical team reported that the patient experienced a brain bleed, characterized as a hemorrhagic stroke. a second event of hematuria was also documented later the same day, with no associated suction alarms from the impella5.5. the patient¿s overall course resulted in withdrawal of care, and the patient ultimately expired during unit support. the hemorrhagic stroke occurred during support; however, no evidence has been presented to demonstrate a device related malfunction. hemorrhagic stroke is a known clinical risk in critically ill patients with cardiogenic shock and is often multifactorial, involving patient comorbidities, anticoagulation status, and overall hemodynamic instability. the subsequent hematuria also showed no indicators of device malfunction. as the impella5.5 remained functional and no device issues were reported, the event appears more consistent with patient specific and clinical factors rather than a device related cause. the death is conservatively being reported on the impella 5.5 but is unlikely a contributing factor to cause of death and is more likely due to the patient¿s declining clinical condition. death is also known risk associated with impella 5.5 per the instructions for use (ifu). | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
556 | 1220648-2026-02845 | 24313156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24313156&pc=OZD | 01/27/2026 |  |  | 02/11/2026 |  |  | 01/27/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026804481 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 80 KG | Non Hispanic | Asian | Injury | Yes | Yes | Backflow | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 was placed emergently along with va-ecmo for the 52 year old male patient who had deteriorated with the acute myocardial infarction and after the coronary angioplasty. the patient was seen to have deteriorated due to having wide open mitral regurgitation and a chordae rupture. the pump and va-ecmo were supporting when there was suction and gave blood products to remedy and have volume resuscitation. the suction resolved with the blood products and pump remained on for support. the underlying right ventricle weakness and surgery for the mitral and chordal rupture could have contributed to the suction. suction is a known issue with impella support as described in the instructions for use and may be influenced by patient-specific factors. | It was reported the pump was explanted. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | It was reported the pump was explanted, however the explant date was not provided. b1 product problem was inadvertently omitted on the initial manufacturer device report.
557 | 1220648-2026-02846 | 24313211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24313211&pc=OZD | 01/27/2026 |  |  | 02/11/2026 |  |  | 01/27/2026 | 02/11/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779510 |  | 000000000001000100 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
558 | 1220648-2026-02853 | 24313294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24313294&pc=OZD | 10/28/2025 |  |  | 02/11/2026 |  |  | 10/28/2025 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662754 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 71 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 45-year-old male with no significant past medical history presented with worsening shortness of breath, chest pressure, and fatigue several days after returning from international travel. he was initially evaluated at an urgent care center and diagnosed with pneumonia. due to worsening symptoms, he presented to a local emergency department, where he was found to be hypoxemic with st-segment elevations and was transferred for suspected stemi. upon arrival to the coronary care unit, the patient developed ventricular tachycardia requiring multiple cardioversions and endotracheal intubation for worsening hypoxia. due to progressive hemodynamic instability, emergent va-ecmo was initiated at the bedside approximately 15 minutes after ccu arrival. after stabilization on ecmo, the patient was transported to the catheterization laboratory, where an impella cp device was inserted via the left common femoral artery for left ventricular unloading. coronary angiography demonstrated clean coronary arteries. the impella cp was functioning at performance level p7. after returning to the cardiothoracic icu on combined ecmo and impella support (ecpella), the patient developed recurrent ventricular tachycardia and ventricular fibrillation refractory to multiple cardioversion attempts. a repeat shock evaluation was performed, and the clinical team determined that escalation of support was required. the patient was taken to the operating room, where centrimag support was initiated. the impella cp device was removed during this procedure. vascular surgery assisted with ecmo decannulation and surgical closure. no impella device alarms or malfunctions were reported during use. the device was removed intact. at the time of reporting, the patient remained on centrimag support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
559 | 1220648-2026-02854 | 24313300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24313300&pc=OZD | 02/03/2026 |  |  | 02/11/2026 |  |  | 02/04/2026 | 02/11/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026742081 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Decrease in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
560 | 1220648-2026-02855 | 24313308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24313308&pc=OZD | 10/28/2025 |  |  | 02/11/2026 |  |  | 10/28/2025 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699669 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Difficult to Remove | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The patient was a 45-year-old male undergoing impella cp placement for hemodynamic support. the patient was unable to lie flat and was positioned with the head of bed elevated to approximately 15 degrees. the right common femoral artery was accessed, a 5f sheath was placed, pre-closure sutures were deployed, and the access site was upsized with impella dilators over a stiff wire. the device was primed, the left ventricle was accessed, and the impella cp was advanced into the lv without initial difficulty. the device was started in auto mode and ramped up. shortly after activation, the impella displayed superimposed waveforms suggesting suboptimal positioning. alarms appeared after a delay. the device was visualized and retracted slightly, with subsequent improvement in waveforms. the impella sheath was noted to have migrated outward and was repositioned. the patient subsequently became pulseless and entered pulseless electrical activity. cardiopulmonary resuscitation was initiated. during chest compressions, reevaluation of the impella revealed the catheter had been displaced into the ascending aorta and was kinked in half. the device could not be straightened, advanced, or explanted. cpr continued for approximately 30 minutes. after discussion with the family, resuscitation efforts were stopped, and the patient expired in the catheterization laboratory. prior to the displacement event, no device alarms or malfunctions had been reported. the inability to straighten, advance, or remove the device occurred after the catheter became kinked during chest compressions. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
561 | 1220648-2026-02859 | 24313338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24313338&pc=OZD | 10/28/2025 |  |  | 02/11/2026 |  |  | 10/28/2025 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026689285 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 61-year-old male who previously had an impella cp placed for treatment of stemi and had remained stable without vasoactive support. following heart-team evaluation, the plan was to upgrade support to an impella 5.5 and maintain the patient on mechanical support with planned cabg the following week. for the impella 5.5 procedure, the right axillary artery was accessed, and a 10×30 graft was anastomosed without tunneling. the left ventricle was crossed using a 5 fr × 65 cm pigtail catheter and 0.035 wire. a 0.018 wire was then positioned, and the impella 5.5 was advanced after priming. the device was double-barreled with the existing cp, and the cp catheter was withdrawn into the descending aorta to allow ramp-up of the 5.5. a 14 fr manta device was used to close the groin access site from the cp. the axillary graft was shortened to approximately 10 cm, and the blue suture ring was placed and secured. the repositioning sleeve and locking mechanism were advanced, accordioning the remaining graft into the pocket. the impella 5.5 was secured, repositioned under transesophageal echocardiography guidance, and locked at 43 cm. echocardiography showed an inlet-to-aortic valve distance of 4.75 cm. three-point securement was applied, and the patient was returned to room 3001 in stable condition. no device-related malfunctions were reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
562 | 1220648-2026-02860 | 24313363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24313363&pc=OZD | 02/04/2026 |  |  | 02/11/2026 |  |  | 02/04/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026783742 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 109 KG | Non Hispanic | White | Injury | Yes | No | Pumping Stopped | Material Split, Cut or Torn | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Impella 5.5 was inserted via right axillary artery in a 68 year old male who was admitted for cardiomyopathy the patient¿s comorbidities were unknown the patient¿s clinical state prior to initiation of support was scai stage e. on day 5 of support, the device required repositioning. the device was rotated, which created torque at the red impella plug so that the white impella cable broke off. the fiber optic wires snapped and led to a pump stop. the device was explanted and another 5.5 implanted. the pump stop will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the replacement; however, the replacement was performed with no consequence to the patient and support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
563 | 1220648-2026-02866 | 24313724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24313724&pc=OZD | 02/03/2026 |  |  | 02/11/2026 |  |  | 02/03/2026 | 02/11/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026811391 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
564 | 1220648-2026-02874 | 24314064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24314064&pc=OZD | 10/22/2025 |  |  | 02/11/2026 |  |  | 10/22/2025 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1157455 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | Application Network Problem | Difficult to Open or Close |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the automated impella controller (aic) not connecting to impella connect. it presented purge pressure alarms even with brand new purge cassette. additionally, the cassette door was sticking with button either not releasing to open or not releasing to close. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
565 | 1035166-2026-00007 | 24314174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24314174&pc=DYB | 01/12/2026 | 02/11/2026 |  | 02/11/2026 |  |  | 01/14/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | INTEGER HOLDINGS CORPORATION |  |  | US |  | S9649279 |  | 0052-3006 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 55 YR |  | Male |  |  |  | Injury | Yes | No | Break | Fluid/Blood Leak | [5 empty] | Cardiogenic Shock | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | We have been notified of a complaint involving an introducer (material #0052-3009, lot #1993040). it was reported that the surgeon placed 23fr introducer into graft. when clamp removed, there was a significant amount of blood leaking through what he felt were cracks in the introducer. attempted to tie black tie x 2 around introducer and graft and was unsuccessful. it appeared the blood was coming through the side of the sheath. sheath removed and new axillary insertion kit used. the introducer was split not at the tip. what is the name of the facility where event occurred? (b)(6) hospital- (b)(6). what is the outcome of the patient? (e.g. patient stable, patient alive, no injury, serious injury, death) patient stable. what procedure was the product used for? insertion of impella 5.5. what was the condition of the patient when presented to the physician? unknown. was there an issue with the product? if yes, then provide details of the issue and answer the questions a) and b). sheath split at the tip. when was the issue identified? (before, during or after the procedure) during use. what was the user (physician/ doctor/ nurse) doing when the issue occurred? inserting the introducer. please provide the contact information of the physician who used the product to obtain additional clinical information of the event in question. n/a. there were no other devices used during the procedure. customer stated the introducer will be returning. customer reported no picture taken of the leaking. the increased amount of bleeding was first noticed coming from the sides of the sheath and when it was removed it was noted that the tip was split. the patient had just been transported in from an outside facility on an intra-aortic balloon pump. the patient was in cardiogenic shock and in need of increased support with a 5.5. the patient stabilized and now has a durable left ventricular assist device. the patient stabilized quickly with the 5.5 and is now recovering from placement of the durable left ventricular assist device. the introducer will be returned. the pump was discarded by the hospital during the subsequent surgery. while there was significant enough blood leaking from the introducer, the patient had sufficient hemoglobin levels to not require a transfusion. the leaking introducer was replaced quickly by a new introducer for impella insertion. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. integer is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by integer which the company may have not been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, integer, or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, integer, or its employees, that the report constitutes an admission that the device, integer, or its employees caused or contributed to the event described in this report.
566 | 1220648-2026-02876 | 24314380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24314380&pc=OZD | 01/29/2026 |  |  | 02/11/2026 |  |  | 01/30/2026 | 02/11/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798721 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
567 | 1220648-2026-02878 | 24314480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24314480&pc=OZD | 02/05/2026 |  |  | 02/11/2026 |  |  | 02/05/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1639305 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 90 KG |  | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during transport of the patient and aic to the intensive care unit (icu) from the cardiac catheterization laboratory (cath lab), it was noted that the placement signal had disappeared, but impella flows continued. shortly after this occurred, the aic then gave a controller failure error that spontaneously resolved. it was advised that the controller be switched to a different aic but the staff decided to wait to see if it occurred again as it appeared to be working without problems. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
568 | 1220648-2026-02884 | 24315040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24315040&pc=OZD | 02/05/2026 |  |  | 02/11/2026 |  |  | 02/05/2026 | 02/11/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1663431 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
569 | 1220648-2026-02905 | 24315743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24315743&pc=OZD | 02/02/2026 |  |  | 02/11/2026 |  |  | 02/03/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026721793 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  |  |  | Malfunction | No | Yes | Shipping Damage or Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | Clinician narrative: an impella cp was involved in a reported event for which no patient age or sex was provided. during preparation of the impella cp, purge fluid was not observed at the outlet of the pump, and the automated impella controller (aic) did not recognize the yellow-to-yellow luer connection. in response, the customer used a new impella cp box. the reported event was experienced as hemodynamic instability and required device revision or replacement. no additional patient, procedural, or outcome information was provided. after a comprehensive review of the available information, there was insufficient information to establish a causal relationship between the impella cp and the reported event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
570 | 1220648-2026-02907 | 24315798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24315798&pc=OZD | 02/06/2026 |  |  | 02/11/2026 |  |  | 02/06/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1649540 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The console came in for annual preventative maintenance. upon initial startup, the device exhibited a red battery failure alarm. | [5 empty] | A1: no patient involvement. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed.
571 | 1220648-2026-02908 | 24315808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24315808&pc=OZD | 02/03/2026 |  |  | 02/11/2026 |  |  | 02/03/2026 | 02/11/2026 |  |  | 01/31/2026 |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025453876 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 90 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. nurse notified of a high purge pressure alarm. appropriate troubleshooting steps were performed. tubing was inspected with no kinks identified. the purge cassette was replaced; however, the alarm persisted. purge flow was noted to be <1.0 ml/hr. the alarm subsequently progressed to ¿purge system blocked.¿ tpa was ordered to be administered through the purge system in an attempt to restore adequate purge flow. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
572 | 1220648-2026-02913 | 24315947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24315947&pc=OZD | 02/06/2026 |  |  | 02/11/2026 |  |  | 02/06/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1763601 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the bedside nurse notified the impella team that the placement signal was reading out of range, not correlating with the patient¿s a-line. a transthoracic echocardiogram (tte) was ordered, the impella 5.5 was repositioned by the advanced heart failure physician. the pump was moved away from the papillary muscle to the mid ventricular space, measuring 4 centimeters. at this point, the placement signal was still reading high, the physician was aware. during rounds, ¿placement signal not reliable¿ noted on the impella console. the physician elected to mute the audio for the alarm. it was reported that the procedure was successful and the patient was stable. this impella aic controller report is one of two devices associated with the event. the medical device report on the impella 5.5 is in process. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
573 | 1220648-2026-02912 | 24315948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24315948&pc=OZD | 02/03/2026 |  |  | 02/11/2026 |  |  | 02/03/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1701427 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinician narrative: an event was reported involving an impella system optical automated impella controller (aic). no patient age or demographic information was provided. the reported event involved a console error, followed by a switch to the backup controller. a yellow alarm was noted upon boot-up of the aic. the event was experienced as hemodynamic instability and required device revision or replacement associated with the optical aic. no additional clinical, procedural, or patient outcome information was provided. after a comprehensive review of the available information, there was insufficient information to establish a causal relationship between the optical aic and the reported event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section a is unknown. e initial reporter is unknown.
574 | 1220648-2026-02917 | 24316301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24316301&pc=OZD | 02/06/2026 |  |  | 02/11/2026 |  |  | 02/06/2026 | 02/11/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026760974 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
575 | 1220648-2026-02918 | 24316480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24316480&pc=OZD | 02/07/2026 |  |  | 02/11/2026 |  |  | 02/07/2026 | 02/11/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779283 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
576 | 1220648-2026-02920 | 24316615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24316615&pc=OZD | 02/12/2026 |  |  | 02/11/2026 |  |  | 02/08/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775912 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Difficult to Advance | Patient Device Interaction Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during implant, the 5.5 device would not make the turn from the axillary artery into the subclavian. the physician decided to use 035 glide advantage to place device. the device was advanced without issue over the new wire. | [5 empty] | D6b: the explant date is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
577 | 1220648-2026-02921 | 24316693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24316693&pc=OZD | 02/06/2026 |  |  | 02/11/2026 |  |  | 02/06/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1210713 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that they were unable to power off aic even after unplugging the white cable. a hard reset was completed at the bottom of aic. | [5 empty] | A: patient information is unknown. d4: the udi and manufacturer date is unknown, not found with the serial number and lot number provided. e: initial reporter information is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
578 | 1220648-2026-02927 | 24317654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24317654&pc=OZD | 12/22/2025 |  |  | 02/11/2026 |  |  | 12/22/2025 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026750035 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 73 year old male patient treated with impella cp due to acute myocardial infarction complicated by cardiogenic shock. access via left femoral artery and tamponade had to be used during impella insertion and wound had to be surgically closed. complication of a left ventricle perforation. patient was later escalated to an impella 5.5 support. very limited information available in the case description. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
579 | 1220648-2026-02942 | 24318362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24318362&pc=OZD | 12/15/2025 | 02/11/2026 |  | 02/11/2026 |  |  | 12/15/2025 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781968 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 77 YR |  | Male | 84 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A seventy-seven-year-old male patient with a history of coronary artery disease received an impella 5.5 device for hemodynamic support during an elective high risk coronary artery bypass graft surgery. prior to device implantation, the patient was receiving two inotropic medications and mechanical ventilation. on the first postoperative day in the surgical intensive care unit, the patient developed a small hematoma at the axillary insertion site, which resolved with manual pressure. the patient¿s condition gradually improved over the following days, and after eleven days of support, the impella device was successfully weaned and removed. | [5 empty] | Investigation summary. no product return. (asae/hematoma): the cause of the access site adverse event was not determined due to insufficient clinical details.
580 | 1220648-2026-02948 | 24319202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24319202&pc=OZD | 02/05/2026 | 02/11/2026 |  | 02/11/2026 |  |  | 02/05/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026801796 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Male | 62 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Increase in Pressure | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiogenic Shock | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient is a 15-year-old male supported with an impella 5.5 device (sn (b)(6)) for the indication of cardiomyopathy. his pre support clinical state was consistent with scai cardiogenic shock stage c. at approximately 2300 on 2/4, the patient developed acute onset crushing chest pain. electrocardiogram demonstrated st segment elevation. based on the available information, no allegation or deficiencies noted by customer, the impella maintained hemodynamic support with stable flows and purge parameters. a transient elevation in plasma free hemoglobin resolved with standard device management interventions, and no sustained hemolysis or device malfunction was identified. no evidence suggests that the device contributed to the patient¿s reported chest pain or transient st elevation. clinical findings and device performance indicate that the impella system continued to operate within expected parameters. the patient remains on support at the time of reporting. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
581 | 1220648-2026-02950 | 24319235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24319235&pc=OZD | 02/05/2026 | 02/24/2026 |  | 02/11/2026 |  |  | 02/05/2026 | 02/11/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1379222 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
582 | 1220648-2026-02952 | 24319301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24319301&pc=OZD | 02/05/2026 | 02/11/2026 |  | 02/11/2026 |  |  | 02/05/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1102528 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 80 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 and automated impella controller (aic) were being prepped for support of a 63 year old male patient who had been admitted in with cardiogenic shock and cardiomyopathy. during the prep of the aic there was a controller error that prompted the team to replace the aic and proceed with a new aic for the support initiation. the console exchange will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
583 | 1220648-2026-02954 | 24319372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24319372&pc=OZD | 02/05/2026 | 02/24/2026 |  | 02/11/2026 |  |  | 02/05/2026 | 02/11/2026 |  |  |  |  | 02/11/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2025602801 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male | 85 KG |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57 year old male on an impella 5.5 due to cardiomyopathy. during support, a red alarm ¿purge pressure high¿ occurred and the purge flow (pf ) was less than 1ml/h and pp 1000mmhg. a recommendation was given to replace the pump and start tissue plasminogen activator (tpa) lysis. patient outcome was unknown. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
584 | 1220648-2026-02675 | 24302461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24302461&pc=OZD | 02/04/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026804270 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
585 | 1220648-2026-02678 | 24302590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24302590&pc=OZD | 02/04/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802395 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
586 | 1220648-2026-02680 | 24302856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24302856&pc=OZD | 01/31/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1438724 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Read Input Signal | Priming Problem | [5 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
587 | 1220648-2026-02682 | 24302929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24302929&pc=OZD | 02/04/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737113 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | Company representation | Distributor | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
588 | 1220648-2026-02683 | 24302973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24302973&pc=OZD | 02/04/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026805672 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 59-year-old female patient presenting with cardiomyopathy. during hospitalization, there was concern for cerebral edema, and an external ventricular drain (evd) was placed due to elevated intracranial pressures. per bedside nursing report, it was unclear whether the patient had an intracranial bleed. the most recent ct scan demonstrated a possible small subdural hematoma. the reported event involves a neurologic event (stroke). based on the available information, the intracranial findings are more plausibly related to the patient¿s underlying critical condition and neurologic pathology. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
589 | 1220648-2026-02685 | 24303052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303052&pc=OZD | 01/30/2026 |  |  | 02/10/2026 |  |  | 02/01/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1200543 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
590 | 1220648-2026-02684 | 24303071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303071&pc=OZD | 02/04/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026805575 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. the impella experienced intermittent ¿in aorta¿ alarms. multiple transthoracic echocardiograms demonstrated the device measuring 5.1¿5.3 cm across the aortic valve toward the apex of the left ventricle. the alarms resolved at pump performance level 7 during weaning from extracorporeal membrane oxygenation and cessation of sedation with suctioning. the device subsequently returned to pump performance level 4, with recurrent ¿in aorta¿ alarms. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
591 | 1220648-2026-02686 | 24303072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303072&pc=OZD | 01/30/2026 |  |  | 02/10/2026 |  |  | 01/30/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026731324 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced low pump flow and suction. the pump was repositioned. the patient was in stable condition. | [5 empty] | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
592 | 1220648-2026-02687 | 24303097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303097&pc=OZD | 01/29/2026 |  |  | 02/10/2026 |  |  | 01/29/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1366755 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella pump experienced a purge disc not detected alarm on the supporting automated impella controller. no patient harm was reported. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
593 | 1220648-2026-02690 | 24303207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303207&pc=OZD | 02/04/2026 | 02/10/2026 |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026691065 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 64 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via a right axillary artery graft in a 64-year-old male patient with a history of diabetes mellitus who underwent pre-operative mechanical circulatory support placement as part of a planned surgical prehabilitation strategy. pre-procedural computed tomography of the chest was obtained to evaluate vessel suitability. during the procedure, multiple attempts were made to advance the impella 5.5, including the use of a buddy wire to facilitate delivery into the innominate artery; however, the device could not be advanced as intended. the impella 5.5 was subsequently removed from the axillary graft, and an impella cp device was placed via the right axillary graft access to provide continued hemodynamic support. the reported events include failure to advance, device revision or replacement, and hemodynamic instability. the impella 5.5 is conservatively reported for serious injury due to a temporary interruption of hemodynamic support during the device exchange; however, the exchange was performed without known clinical consequence to the patient, and hemodynamic support was promptly resumed. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
594 | 1220648-2026-02689 | 24303210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303210&pc=OZD | 01/28/2026 |  |  | 02/10/2026 |  |  | 01/28/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775981 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
595 | 1220648-2026-02692 | 24303245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303245&pc=OZD | 10/31/2025 |  |  | 02/10/2026 |  |  | 01/21/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1849119 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the cord of an automated impella controller was frayed and only charged in certain positions. the pump was removed from service. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
596 | 1220648-2026-02693 | 24303284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303284&pc=OZD | 02/04/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1263482 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that after setting the impella device to pump performance level 0 for routine explantation, the white cable was removed from the automated impella controller (aic). the power button was then pressed for 2¿3 seconds. immediately, a high continuous signal tone was emitted with flashing green lights, and it was not possible to shut down the aic. the perfusionist was advised to press and hold the power button for a longer duration, which successfully powered down the aic. the aic was subsequently restarted, and the system booted up, displaying the message: ¿unexpected controller shutdown.¿ no further information is available. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section a is unknown. e initial reporter is unknown.
597 | 1220648-2026-02697 | 24303360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303360&pc=OZD | 02/02/2026 |  |  | 02/10/2026 |  |  | 02/02/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1226867 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Unintended Application Program Shut Down | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | An unexpected impella controller shutdown occurred during preparation for automated impella console (aic) use. clinical liaison (cl) reported that attempts to restart the controller were unsuccessful. replacement of the aic was recommended. the console was subsequently replaced, and the replacement aic arrived and was exchanged. the defective aic was returned for evaluation and repair. the reported events include shutdown or restart problem, medical device removal, device revision or replacement, and hemodynamic instability. the device will be conservatively reported for serious injury due to a temporary interruption in hemodynamic support during the exchange; however, the exchange was performed without consequence to the patient, and support was resumed with no known adverse patient outcomes. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
598 | 1220648-2026-02699 | 24303398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303398&pc=OZD | 01/27/2026 |  |  | 02/10/2026 |  |  | 02/01/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1129860 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Application Program Problem | Device Sensing Problem | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | An impella support was ongoing with a pump and automated impella controller (aic). the icu staff observed the aic to have alarmed for a "controller error" and "placement signal not reliable". to remedy the issue the team replaced the aic and support proceeded with no harm. the console replacement will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however the exchange was performed with no consequence to the patient and support. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
599 | 1220648-2026-02705 | 24303508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303508&pc=OZD | 02/02/2026 |  |  | 02/10/2026 |  |  | 02/02/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026783085 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  |  | Death | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp device was inserted into the left femoral artery in a 85-year-old male patient with a past medical history of coronary artery disease (cad). the patient presented in acute myocardial infarction (ami) and cardiogenic shock (cgs), in scai stage c shock, prior to initiation of support. the impella was inserted for ventricular support during a high-risk percutaneous coronary intervention (pci). the physician attempted to insert the impella; however, was unsuccessful due to the patient's anatomy. the impella insertion was aborted. the physician continued with the pci; however, the patient expired during the procedure. the physician did not attribute the death to the impella. the death is conservatively being reported to the impella cp but is not a contributing factor as the device insertion was aborted, and is most likely attributed to patients clinical condition of a critically ill patient with acute myocardial infarction complicated by cardiogenic shock. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted.
600 | 1220648-2026-02709 | 24303534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303534&pc=OZD | 02/01/2026 |  |  | 02/10/2026 |  |  | 02/02/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026764485 |  | 1000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | False Alarm | Pumping Stopped | [5 empty] | Sepsis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
601 | 1220648-2026-02717 | 24303648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303648&pc=OZD | 02/02/2026 |  |  | 02/10/2026 |  |  | 02/02/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026701085 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Device Sensing Problem | [6 empty] | Cardiac Tamponade | [37 empty] | R |  |  |  | Initial submission | Followup
602 | 1220648-2026-02722 | 24303728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303728&pc=OZD | 01/31/2026 |  |  | 02/10/2026 |  |  | 02/01/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026729306 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Insufficient Information | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
603 | 1220648-2026-02725 | 24303891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303891&pc=OZD | 01/30/2026 |  |  | 02/10/2026 |  |  | 01/30/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766873 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
604 | 1220648-2026-02727 | 24303903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303903&pc=OZD | 01/27/2026 |  |  | 02/10/2026 |  |  | 01/28/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026783688 |  | 000000000001000100 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
605 | 1220648-2026-02729 | 24303916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24303916&pc=OZD | 02/04/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026809807 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via a direct surgical aortic approach in a 58-year-old female patient undergoing elective placement. during intraoperative implantation, oozing was noted from the graft used for device insertion. the patient received hemostatic management including kcentra, 4 units of platelets, and 5 units of cryoprecipitate, along with cautery and topical hemostatic agents. the clinical team attributed the bleeding to underlying coagulopathy. hemostasis was achieved with standard intraoperative interventions, and troubleshooting was successful. the patient was transferred to the icu without further issue, and the impella 5.5 device remained in use and functioned as intended. the reported event involves a major bleed requiring blood transfusion and surgical intervention. bleeding associated with surgical graft placement and anticoagulation requirements is a known risk described in the impella 5.5 instructions for use (ifu). based on the available information, the bleeding was effectively managed, and most likely attributed to underlying coagulopathy. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
606 | 1220648-2026-02737 | 24304259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24304259&pc=OZD | 02/04/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806949 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission
607 | 1220648-2026-02745 | 24304526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24304526&pc=OZD | 10/30/2025 |  |  | 02/10/2026 |  |  | 10/30/2025 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026728914 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | On (b)(6) 2025, a 58-year-old male patient with a past medical history of hypertension, hyperlipidemia, diabetes mellitus, and non-ischemic cardiomyopathy with a reported left ventricular ejection fraction of 15 percent presented with shortness of breath and acute decompensated heart failure. a right heart catheterization was performed, and the decision was made to place an impella cp. due to the patient¿s critical condition, veno-arterial extracorporeal membrane oxygenation was emergently initiated as well. on (b)(6) 2025, the patient was upgraded to an impella 5.5 device. during this period, the patient was evaluated for advanced heart failure therapies but was unable to be listed due to concurrent treatment for pneumonia. on (b)(6) 2025, a sensor drift was noted, with an aortic pressure signal discrepancy greater than fifty millimeters of mercury compared with systemic mean arterial pressures. the system was manually re-zeroed. on (b)(6) 2025, the decision was made to remove the impella device. the patient was noted to have recovery of left ventricular ejection fraction, and no additional mechanical circulatory support device was placed. during impella 5.5 pump support, the patient experienced placement signal issues and mechanical interaction with blood. clinical stated placement signal unreliable alarms, placement signal disabled. pump continued to provide support as intended, no patient consequence. additionally, (b)(6) red urine noted in foley. registered nurse (rn) said acute change over last hour. surgical team notified. echo done by fellow. noted that the inlet of impella is interacting with the mitral valve. standard repositioning done for impella. this is considered a reportable malfunction. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
608 | 1220648-2026-02746 | 24304541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24304541&pc=OZD | 02/04/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1105809 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | Yes | False Alarm | Pumping Stopped | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Clinical review/rationale: an impella 5.5 was placed via the direct aortic approach to support the 77 year old male patient who had been admitted in for surgical support for the bypass/cabg. the patient had presented in scai stage d shock. the pump was supporting when on day 2 of the support there was a pump stop alarm of 2 seconds that self-resolved. the surgical team made the decision to replace the console. the console was replaced and the pump support did continue. the console exchange will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | H6 code was added as it was omitted from the initial report that was submitted. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
609 | 1220648-2026-02749 | 24304754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24304754&pc=OZD | 01/26/2026 |  |  | 02/10/2026 |  |  | 01/28/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1432278 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female | 60 KG |  |  | Injury | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Hemodynamic instability | [12 empty] | Other | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted in a 67-year-old female patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). during the evening prior to reporting, the customer noted multiple controller failure alarms occurring approximately three to four times overnight, which appeared immediately upon activation. concurrently, the placement signal and left ventricular (lv) waveform were noted to intermittently disappear on the controller display. the clinical team exchanged the controller. the patient subsequently survived to explant. the reported events include controller failure, placement signal issues, medical device revision or replacement, and hemodynamic instability. the impella cp is being conservatively reported for serious injury due to a temporary interruption of hemodynamic support during the controller exchange; however, the exchange was performed without consequence to the patient, and support was resumed without any known adverse events. | [5 empty] | A5. ethnicity is unknown. a6. race is unknown. device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
610 | 1220648-2026-02752 | 24304804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24304804&pc=OZD | 10/24/2025 |  |  | 02/10/2026 |  |  | 10/24/2025 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744171 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | Black Or African American | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Inflammation | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 61-year-old male with severe non-ischemic cardiomyopathy (ef 10¿15%) presented in cardiogenic shock and underwent impella 5.5 implantation via the right axillary artery without immediate procedural complications. within 24 hours of implantation, the patient developed significant bleeding at the insertion site, requiring return to the operating room on (b)(6) 2025 for surgical exploration and transfusion. bleeding improved, but recurrent oozing occurred on (b)(6) 2025, prompting temporary cessation of bivalirudin. during his icu course, the patient developed frequent premature ventricular contractions. echocardiography showed the impella positioned deeply in the left ventricle (~6 cm). the device was repositioned, resulting in resolution of arrhythmias and suction events. subsequently, brown drainage and local inflammatory changes were noted at the axillary site. the patient underwent surgical washout on (b)(6) 2025, and a wound vac was applied. empiric antibiotics were initiated for suspected infection. all blood and wound cultures remained negative. the patient was later determined not to be a candidate for advanced therapies. his condition progressively declined despite ongoing device support. after goals-of-care discussions, life-sustaining treatments were withdrawn, and the patient died with the impella 5.5 in place. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
611 | 1220648-2025-48599 | 24304975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24304975&pc=OZD | 10/26/2025 | 11/24/2025 |  | 02/10/2026 |  |  | 10/26/2025 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722236 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | Heart Block | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 59-year-old male with out-of-hospital cardiac arrest was transported to the catheterization lab on va-ecmo after multiple shocks and bystander cpr. pre-procedure imaging demonstrated apical dyskinesia and mitral regurgitation, and vascular access was obtained using ultrasound guidance and micropuncture technique. an impella device was inserted for mechanical circulatory support, after which the patient immediately developed heart block requiring placement of a temporary pacemaker. during subsequent swan-ganz catheter insertion, the right ventricle was inadvertently perforated, resulting in cardiac perforation and immediate impella restricted flow rate/suction events accompanied by bleeding. vascular surgery was consulted emergently, leading to surgical intervention and placement of an additional device, specifically a reperfusion catheter. the patient was transferred to the icu, where further evaluation revealed a renal mass and gastrointestinal bleeding. despite ongoing management, the patient showed no signs of cardiac recovery, remained completely laminar on support, and ultimately experienced death. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
612 | 1220648-2026-02756 | 24304985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24304985&pc=OZD | 02/03/2026 |  |  | 02/10/2026 |  |  | 02/03/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748981 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Ventricular Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
613 | 1220648-2026-02759 | 24305114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24305114&pc=OZD | 02/02/2026 |  |  | 02/10/2026 |  |  | 02/05/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026780314 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
614 | 1220648-2026-02760 | 24305250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24305250&pc=OZD | 02/05/2026 |  |  | 02/10/2026 |  |  | 02/05/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED EUROPE GMBH - 3007904765 |  |  | GM |  | 1383039 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Other | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, when the control unit was unplugged for rehabilitation purposes, the charge dropped from 100 percent to 50 percent, and a yellow "half battery capacity" alarm sounded. the staff immediately plugged it back into the outlet and returned to 100% charging. it was reported that the procedure was successful and the patient stable. | [5 empty] | D6b: explant date is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
615 | 1220648-2026-02762 | 24305393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24305393&pc=OZD | 02/02/2026 |  |  | 02/10/2026 |  |  | 02/05/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026805259 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Malposition of Device | [6 empty] | Ischemia | Renal Failure | Thrombosis/Thrombus | [10 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | An impella 5.5 device was inserted into the right subclavian artery in a 61-year-old male patient presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), prior to initiation of support. the impella was inserted for ventricular support during an electrophysiology ablation (ep ablation) and post-high-risk percutaneous coronary intervention (pci). during the impella 5.5 insertion, the brachiocephalic artery was highly tortuous. a 0.35 stiff guidewire was substituted for the 0.18 guidewire, and access to the left ventricle was achieved. however, torque was not transmitted, and the tip was unable to be directed toward the apex. while the patient was in the intensive care unit (icu), a thrombus was noted that was not seen before the impella insertion. the tip of the impella was pointing toward the posterior wall, causing blood stagnation at the apex, leading to thrombus formation. the left ventricular thrombus caused intestinal ischemia, leading to the decision to withdraw the catheter. the post-procedure outcome is unknown. thrombus (thrombosis) can occur due to inadequate anticoagulation. thrombosis is an adverse event associated with both impella's mechanical support and percutaneous coronary intervention. | Clinician narrative updated: an impella cp was inserted via the femoral artery in a 61-year-old male who was transported as a stemi. the patient underwent pci without mechanical circulatory support, followed by ventricular tachycardia ablation the next day. impella cp support was subsequently initiated to facilitate continued management and eligibility for ablation therapy. at the time of impella insertion, the patient was not reported to be in cardiogenic shock. past medical history was not provided. during hospitalization, the patient required ventilatory support and was treated with inotropes and vasopressors. an intra-aortic balloon pump had been used prior to impella support. pump: 1 reported on report number 1220648-2026-02280. while supported with the impella cp via femoral arterial access, the patient developed progressive lower limb ischemia. the decision was made to upgrade support to an impella 5.5. an impella 5.5 was subsequently inserted. during insertion, the brachiocephalic artery was described as highly tortuous. a 35 stiff guidewire was substituted for the 18 guidewire, and access to the left ventricle was achieved. it was reported that torque was not transmitted effectively, and the device tip was unable to be directed toward the left ventricular apex. the physician later reported that the tip of the impella 5.5 was positioned toward the posterior wall, resulting in blood stagnation at the apex. no left ventricular thrombus was present prior to 5.5 insertion; however, a left ventricular thrombus. the thrombus subsequently led to intestinal ischemia. the patient developed circulatory failure requiring dialysis following the progression of ischemia. clinical status continued to deteriorate due to intestinal necrosis and multiorgan dysfunction. care was ultimately withdrawn, and the patient expired. thrombus, limb ischemia, renal failure and intestinal necrosis are recognized complications in critically ill patients requiring large-bore arterial access and mechanical circulatory support, particularly in the setting of vascular tortuosity, hemodynamic instability, vasopressor use, and systemic hypoperfusion. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's presenting native critical condition. |  |  |  |  | A4: patient's height and weight are unknown. d6b: the explant date is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1 and b2: updated reportability based on new information from the customer. d6b: updated explant date. h1: updated reportability. h6: updated code. h11: updated clinical rationale. updated clinical rationale: an impella 5.5 device was inserted into the right subclavian artery in a 61-year-old male patient presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), prior to initiation of support. the impella was inserted for ventricular support post-high-risk percutaneous coronary intervention (pci). the following day, the patient had a ventricular tachycardia (vt) ablation performed. during the impella 5.5 insertion, the brachiocephalic artery was highly tortuous. a 0.35 stiff guidewire was substituted for the 0.18 guidewire, and access to the left ventricle was achieved. however, torque was not transmitted, and the tip was unable to be directed toward the apex. while the patient was in the intensive care unit (icu), a thrombus was noted, that was not seen before the impella insertion. the tip of the impella was pointing toward the posterior wall, causing blood stagnation at the apex, leading to thrombus formation. the left ventricular thrombus caused intestinal ischemia, leading to the decision to withdraw the catheter. the patient's condition worsened due to intestinal necrosis, and the decision was made to withdraw care. the patient expired on support. thrombus (thrombosis) can occur due to inadequate anticoagulation. thrombosis is an adverse event associated with both impella's mechanical support and percutaneous coronary intervention. hemolysis and acute renal dysfunction are listed as potential adverse events and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). the device is unlikely to have contributed to the patient's death, which was most likely due to the clinical condition of a critically ill patient with acute myocardial infarction complicated by cardiogenic shock. | B5 was updated and h10 updated with the related report number.
616 | 1220648-2026-02763 | 24305397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24305397&pc=OZD | 10/27/2025 |  |  | 02/10/2026 |  |  | 10/27/2025 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744977 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 87 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The patient undergoing high-risk pci with impella 5.5 support experienced an acute overnight decrease in hemoglobin. this prompted transfusion and a full gastrointestinal evaluation. the suspected gi bleed remained clinically stable, and diagnostic studies including capsule endoscopy and colonoscopy were performed without identifying any contraindication to continuation of cardiac care. medical therapy, including anticoagulation, was temporarily adjusted while the gi workup was underway. throughout the hospitalization, the impella 5.5 functioned as expected with no alarms, complications, or evidence of device malfunction or use issue. staged pci procedures to the rca, lcx, and lad were successfully completed with stable hemodynamic support. the clinical events observed were attributed to underlying patient condition and medical management rather than to the device. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
617 | 1220648-2026-02764 | 24305404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24305404&pc=OZD | 10/28/2025 |  |  | 02/10/2026 |  |  | 10/28/2025 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026756399 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 69 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The patient, with a history of esrd on dialysis, coronary artery disease, atrial fibrillation, and hypertension, was transferred in cardiogenic shock after an anterior septal stemi with prior angioplasty of the ramus and lad and support with an intra-aortic balloon pump. the patient was brought for impella placement, swan-ganz catheter placement, and coronary angiography. following angiography, the existing iabp access was used for impella insertion, and a 14 fr sheath was placed. the device was positioned and initiated at p8 per ifu. hemodynamic data were obtained, and although rv support was considered, it was not pursued based on bedside echocardiographic assessment. after lv unloading, periods of reduced flow with suction were observed and improved with volume administration. a right ij hematoma and oozing at the right femoral access site were noted in the setting of elevated anticoagulation, and both sites were monitored and redressed. during the icu course, the rn reported a decline in pulse pressure on the arterial interface along with continued bleeding at the femoral access where pressure was being held. the patient¿s hemoglobin decreased, and bleeding from multiple access sites was documented. ct was performed due to concern for bleeding around the swan-ganz catheter. upon returning from ct, the patient experienced a cardiac arrest that required cpr and resuscitative measures. following return of circulation, the impella displayed continuous suction and appeared shallow, consistent with malposition and sensing difficulty; the device was repositioned at bedside with subsequent resolution of alarms and restoration of appropriate flow signals. the patient was ultimately determined not to be a surgical candidate. goals of care were transitioned to dnr and comfort-measures-only. the patient¿s condition continued to decline, and life-sustaining therapies were withdrawn on (b)(6), after which the patient expired. throughout the course, recurrent bleeding, suction episodes, flow restriction, device sensing issues, and the need for repositioning occurred in the context of cardiogenic shock, severe comorbidities, anticoagulation, and hemodynamic instability. no device malfunction or use error was identified. the clinical deterioration and death were attributed to the patient¿s critical underlying condition and bleeding complications rather than to a device-related problem. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
618 | 1220648-2026-02765 | 24305498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24305498&pc=OZD | 10/28/2025 |  |  | 02/10/2026 |  |  | 10/28/2025 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733805 |  | 000000000001000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 19-year-old female patient with myocarditis-induced cardiogenic shock underwent placement of an impella 5.5 on (b)(6) via an innominate artery graft, with simultaneous removal of an intra-aortic balloon pump (iabp). following implantation, initial pump performance was stable, with flows ranging from 2.9¿3.6 l/min on p4¿p6 support settings. during routine monitoring on the automated impella controller (aic), the clinical team noted placement signal drift, reflected by impella aortic pressures appearing lower than non-invasive cuff pressures. despite this discrepancy, the device maintained a consistent and interpretable placement waveform, and no alarms or malfunctions were reported. initial transthoracic echocardiography (tte) suggested possible device malposition with the inlet appearing directed toward the mitral valve; however, repeat imaging shortly thereafter confirmed appropriate mid-left-ventricular positioning, and the initial finding was determined to be an imaging artifact rather than true migration. the patient demonstrated no clinical symptoms associated with the pressure discrepancy, with no evidence of mitral valve injury, hemolysis, suction events, arrhythmias, decreased flow, or hemodynamic instability. pump performance remained stable, and support continued without interruption. the pump was not exchanged or removed, and although initially expected to be returned for evaluation, it was ultimately not retained by the facility. the patient remained hemodynamically stable and listed status 2e for heart transplantation throughout support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
619 | 1220648-2026-02783 | 24306178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24306178&pc=OZD | 02/04/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765805 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. the device exhibited a low placement signal with associated suction alarms. the placement signal was no longer correlating with the hemodynamic monitor. recalibration of the aortic pressure was performed. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
620 | 1220648-2026-02786 | 24306552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24306552&pc=OZD | 02/04/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1914207 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 85 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. a yellow controller error alarm occurred twice within the past several hours. each occurrence was transient, lasting approximately one minute. all other automated impella controller (aic) values remained within normal limits. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
621 | 1220648-2026-02787 | 24306553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24306553&pc=OZD | 02/04/2026 |  |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782188 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Clinician narrative an impella 5.5 was utilized in a 67-year-old male who presented with acute myocardial infarction complicated by cardiogenic shock (ami-cgs). the patient had a known history of coronary artery disease. at presentation, the patient required extracorporeal membrane oxygenation (ecmo), inotropic and vasopressor support, and was placed on mechanical ventilation for respiratory support. the patient was classified as scai stage d cardiogenic shock. right surgical axillary subclavian arterial access was obtained. during the course of support, the patient developed a pocket bleed. additional information received states that the patient had received antiplatelet loading with brilinta and plavix due to an occluded coronary stent that was approximately one week old, and the antiplatelet therapy was not related to the impella device or access site issues. the patient experienced a major bleed and required surgical intervention in the setting of antiplatelet therapy and critical illness. additional information received further states that the axillary access site bleeding was completely resolved following intervention, with no ongoing bleeding identified after measures were taken to control the site. after a comprehensive review of the available information, the reporting event did not identify evidence suggesting a causal relationship between the impella 5.5 and the reported event. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D6b: updated explant date.
622 | 1220648-2026-02788 | 24306562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24306562&pc=OZD | 02/05/2026 |  |  | 02/10/2026 |  |  | 02/05/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1455140 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
623 | 1220648-2026-02789 | 24306974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24306974&pc=OZD | 01/30/2026 |  |  | 02/10/2026 |  |  | 02/02/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026810030 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Difficult to Advance | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: a 74-year-old patient female with a history of cardiomyopathy and diabetes mellitus. the patient¿s clinical state prior to initiation of support was identified as scai shock stage d. the case was managed as an extracorporeal cardiopulmonary resuscitation (ecpr) scenario. due to narrow femoral vessels, femoral impella placement was not feasible; therefore, the patient was initially supported with an intra aortic balloon pump (iabp) and extracorporeal membrane oxygenation (ecmo). in an effort to wean the patient from ecmo, placement of an impella 5.5 via the right surgical axillary/subclavian artery approach was planned. subclavian passage was achieved without issue; however, resistance was encountered when advancing the catheter tip at the junction of the brachiocephalic artery and ascending aorta, possibly related to vascular sclerosis. after the application of torque and forward pressure, the catheter advanced and the impella 5.5 was successfully positioned. on the following day, the patient developed anisocoria, and computed tomography imaging demonstrated a cerebral infarction. the patient expired while on support in the setting of underlying critical clinical condition. the direct cause of the cerebral infarction could not be confirmed due to limited post event evaluation. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
624 | 1220648-2026-02795 | 24307350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24307350&pc=OZD | 02/05/2026 |  |  | 02/10/2026 |  |  | 02/05/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1124582 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
625 | 1220648-2026-02796 | 24307369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24307369&pc=OZD | 02/04/2026 | 02/10/2026 |  | 02/10/2026 |  |  | 02/04/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806220 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 80 YR |  | Male | 72 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella support with a 5.5 pump. the complainant reported that the placement signal was higher than the left radial line. the mean arterial pressure showed double on the console than on the arterial line. a second arterial line was placed. the placement signal was still reading higher. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
626 | 1220648-2026-02799 | 24307551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24307551&pc=OZD | 02/02/2026 |  |  | 02/10/2026 |  |  | 02/02/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026803425 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 74 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Embolism | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombocytopenia | [8 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Clinical review/rationale: an impella 5.5 was inserted via right axillary artery in a 66 year old male post aortic and tricuspid valve surgery. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage d. the patient received support with the 5.5 for post-cardiotomy cardiogenic shock with low cardiac output. the patient had pre-existing ecmo support and dual prosthetic valves. on day 3 of support, platelets were down to 23,000 and the patient was evaluated for hitt. on day 5 of support, the patient had a ventricular tachycardia event, which was treated with antiarrhythmics. echocardiogram confirmed position of the 5.5, and other devices concurrently providing support. thrombocytopenia is a known potential adverse events of the impella 5.5 per the instructions for use and may be influenced by patient-specific factors such as underlying cardiac dysfunction and hemodynamic conditions. | Clinical review/rationale updated: an impella 5.5 was inserted via right axillary artery in a 66-year-old male post aortic and tricuspid valve surgery. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage d. the patient received support with the 5.5 for post-cardiotomy cardiogenic shock with low cardiac output. the patient had pre-existing ecmo support and dual prosthetic valves. on day 3 of support, platelets were down to 23,000 and the patient was evaluated for hitt. oropharyngeal bleeding was noted due with pressure related to the endotracheal tube. anticoagulation, including heparin, was held and dissolvable sutures were placed in the posterior oropharynx. a bronchoscopy was also performed for the pulmonary hemorrhage. the patient also demonstrated a change in neurologic status, in which ct revealed right cerebellar infarction. on day 5 of support, the patient had a ventricular tachycardia event, which was treated with antiarrhythmics. echocardiogram confirmed position of the 5.5, and other devices concurrently providing support. on day 18 of support, care was withdrawn and the patient expired. thrombocytopenia, bleeding, stroke and death are known potential adverse events of the impella 5.5 per the instructions for use and may be influenced by patient-specific factors such as underlying cardiac dysfunction and hemodynamic conditions. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received therefore b2, b5, d6b, h1, and h6 updated.
627 | 1220648-2026-02804 | 24307748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24307748&pc=OZD | 01/13/2026 | 02/10/2026 |  | 02/10/2026 |  |  | 01/13/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026764628 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
628 | 1220648-2026-02803 | 24307795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24307795&pc=OZD | 01/30/2026 |  |  | 02/10/2026 |  |  | 02/02/2026 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1839845 |  | 0042-0040-EU |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Unknown |  |  |  | Death | Yes | No | Application Program Problem | Device Sensing Problem | [5 empty] | Hemodynamic instability | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical narrative: clinical narrative: a 75-year-old male with an unknown medical history underwent implantation of an impella cp device for temporary mechanical circulatory support due to cardiogenic shock secondary to an acute myocardial infarction. the device was percutaneously implanted via the femoral artery (side unknown). at the time of impella initiation, the patient was classified as society for cardiovascular angiography and interventions (scai) stage e cardiogenic shock. however, during device preparation, upon connection of the pump to the automated impella controller (aic), an alarm displayed indicating a ¿system error in the optical sensor,¿ with instructions to switch to a backup controller that supports optical sensing or to continue without placement and suction control. the clinical team connected a backup aic, and the alarm resolved. the implant procedure proceeded without reported delay. approximately 16 hours after initiation of impella support, care was withdrawn, and the patient expired. this event involved an aic malfunction identified during preparation, characterized by an optical sensor system error. in accordance with the impella instructions for use, a backup automated impella controller was available and successfully utilized, allowing continued set up for therapy without interruption. there was no reported adverse delay in therapy delivery. the patient¿s underlying condition contributed to the patient¿s demise. the death is consistent with the severity of underlying cardiogenic shock (scai stage e) and acute myocardial infarction. based on available information, the aic malfunction did not materially contribute to the patient¿s clinical deterioration or death. conservatively for usmdr, however, the death is being reported in association with the automated impella controller malfunction, although no causal relationship between the controller issue and the patient outcome has been identified. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section a is unknown.
629 | 1220648-2026-02808 | 24308106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24308106&pc=OZD | 02/03/2026 |  |  | 02/10/2026 |  |  | 02/03/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802379 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Increase in Pressure | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
630 | 1220648-2026-02817 | 24308496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24308496&pc=OZD | 12/22/2025 |  |  | 02/10/2026 |  |  | 12/22/2025 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026733576 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 100 KG | Non Hispanic | White | Death | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 67-year-old male with a medical history significant for ischemic cardiomyopathy with heart failure and reduced ejection fraction of approximately 10 percent, coronary artery disease status post percutaneous coronary intervention in 2022 and 2024, mitral valve repair with mitraclip in 2023, biventricular implantable cardioverter defibrillator, hypertension, hyperlipidemia, atrial fibrillation, diabetes mellitus, and obstructive sleep apnea received an impella 5.5 pump 1 device (serial number: (b)(6) for temporary mechanical circulatory support. the indication for use was cardiogenic shock secondary to an acute myocardial infarction. prior to impella implantation, the patient underwent a scheduled right heart catheterization on (b)(6), which demonstrated elevated filling pressures and low cardiac output. a pulmonary artery catheter was left in place, and the patient was admitted for diuresis. the patient was being evaluated for heart and kidney transplantation and was initiated on milrinone and dopamine for a cardiac index of 1.8 liters per minute per square meter. at the time of impella 5.5 pump 1 device initiation, the patient was classified as society for cardiovascular angiography and interventions stage d cardiogenic shock. on post-implant day nine, the patient developed bacteremia. blood cultures were obtained and were pending, with subsequent concern for an infected cardiac pacemaker lead. the patient was placed on contact precautions, and plans were made to replace vascular lines and consider device exchange if the infection persisted. the physician documented understanding of the impella instructions for use, including indications and recommended duration of support. the long-term treatment goal remained heart transplantation. the patient continued on impella 5.5 pump 1 support for an additional six days before device exchange. the patient was successfully transitioned to impella 5.5 pump 2 (serial number: (b)(6). approximately 18 days later, a placement signal not reliable alarm was reported. the care team was informed that sensor-related alarms do not impact pump function, as described in the impella instructions for use. alarm functionality and monitoring recommendations were reviewed, and no patient harm was reported for event. the care team was instructed on disabling the audible alarm. nine days later, an elevated purge pressure alarm greater than 1100 mmhg with low purge flow was reported. no kinks were observed in the purge line. tissue plasminogen activator protocol was recommended and followed, with resolution of the issue, it appears as no further alarms were reported. the patient remained on impella 5.5 pump 2 support for an additional 41 days. care was subsequently withdrawn, and the patient expired. infection is a known risk associated with impella support, vascular access, and prolonged mechanical circulatory support. this event represents a serious injury associated with impella 5.5 pump 1. regarding impella 5.5 pump 2 and the high purge pressure and placement signal not reliable alarm, these incidents were managed per instructions and appear to have been successfully resolved based on the information available. these device malfunction events are unlikely associated with the patient's outcome. however, for reporting purposes, the death is attributed to impella 5.5 pump 2, as this device was in use at the time of patient expiration. in this case, the patient¿s underlying cardiac and systemic conditions contributed most to the demise outcome. this even story involves two product complaints. please relate manufacturing report numbers.: impella 5.5 pump 1 (b)(4): serious injury. impella 5.5 pump 2 (b)(4): death. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
631 | 1220648-2026-02819 | 24309149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24309149&pc=OZD | 02/02/2026 |  |  | 02/10/2026 |  |  | 02/02/2026 | 02/10/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1133367 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
632 | 1220648-2026-02826 | 24310464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24310464&pc=OZD | 11/03/2025 |  |  | 02/10/2026 |  |  | 11/04/2025 | 02/10/2026 |  |  |  |  | 02/10/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2026705343 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male | 110 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock secondary to an acute myocardial infarction was initially implanted with an impella cp device for temporary mechanical circulatory support with extracorporeal membrane oxygenation (ecmo). subsequently, the patient was escalated to impella 5.5 for left ventricular unloading. after multiple days on combined ecmo and impella 5.5 support there were concerns of hypoxic brain damage and care was withdrawn and the patient passed away. additional information received from the after reveal of the hypoxic brain damage, therapy was withdrawn; there were no complications associated with the impella therapy. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: (stroke): the root cause of the injury was unable to be determined due to insufficient clinical details. | Patient information a4 weight is unknown. the exact date of death is unknown but has been captured as the date of explant. if actual information becomes known, a supplemental report will be submitted. medical safety review. medical safety/clinical reviewed the event. the patient had impella cp placed with extracorporeal membrane oxygenation (ecmo) that was later escalated to impella 5.5 for left ventricular unloading. after multiple days on combined ecmo and impella 5.5 support there were concerns of hypoxic brain damage and care was withdrawn and the patient passed away. brain hypoxia is an increased risk with impella 5.5 and ecmo. this would be an adverse event but not a definite contributing factor to the death of the patient as this is a risk of critically ill patients. device status: the impella device was not received from the customer; therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
633 | 1220648-2026-02531 | 24290241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24290241&pc=OZD | 02/02/2026 |  |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1095486 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 96 KG | Non Hispanic | White | Injury | Yes | No | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Perfusion called to notify us of a controller error alarm. aic was working perfectly fine prior, but when the patient was in the or for a cabg the alarm came on. alarm resolved but controllers were exchanged. aic was taken to biomed to be evaluated and reach out if aic needs to be sent in for any reason. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. sections d9 and h6 have been updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
634 | 1220648-2026-02533 | 24290316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24290316&pc=OZD | 02/01/2026 |  |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 |  |  | 01/31/2026 |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025461506 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 118 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Clinician narrative. an impella 5.5 device was inserted via the right axillary subclavian artery in a 53-year-old male who presented with acute myocardial infarction complicated by cardiogenic shock. the patient had a known history of coronary artery disease and renal insufficiency and required intra-aortic balloon pump support, inotropic therapy, vasopressors, and mechanical respiratory support. the patient was classified as scai stage d. during support with the impella 5.5 device, the patient developed thrombocytopenia and required blood product transfusion. additional information received indicated that the patient was diagnosed with thrombotic thrombocytopenic purpura/immune thrombocytopenic purpura (ttp/itp) attributed to cefepime exposure. a heparin-induced thrombocytopenia evaluation was performed and returned negative. no active bleeding was identified. the patient experienced thrombocytopenia and blood transfusion during impella 5.5 support. after a comprehensive review of the available information, the reported event did not identify evidence suggesting a causal relationship between the impella 5.5 device and the thrombocytopenia. the thrombocytopenia was determined to be medication-induced and associated with cefepime-related ttp/itp in the setting of critical illness and cardiogenic shock. the patient survived. | [5 empty] | B1 and b2: updated reportability per new information. h1: updated reportability. h6: updated codes per new information. updated clinical review: an impella 5.5 device was inserted via the right axillary subclavian artery in a 53-year-old male who presented with acute myocardial infarction complicated by cardiogenic shock. the patient had a known history of coronary artery disease and renal insufficiency and required intra-aortic balloon pump support, inotropic therapy, vasopressors, and mechanical respiratory support. the patient was classified as scai stage d. during support with the impella 5.5 device, the patient developed thrombocytopenia and required blood product transfusion. additional information received indicated that the patient was diagnosed with thrombotic thrombocytopenic purpura/immune thrombocytopenic purpura (ttp/itp) attributed to cefepime exposure. a heparin-induced thrombocytopenia evaluation was performed and returned negative. no active bleeding was identified. the patient experienced thrombocytopenia and blood transfusion during impella 5.5 support. after a comprehensive review of the available information, the reported event did not identify evidence suggesting a causal relationship between the impella 5.5 device and the thrombocytopenia. the thrombocytopenia was determined to be medication-induced and associated with cefepime-related ttp/itp in the setting of critical illness and cardiogenic shock. additional information received reported that the patient expired. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's presenting native critical condition. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
635 | 1220648-2026-02534 | 24290320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24290320&pc=OZD | 01/31/2026 |  |  | 02/09/2026 |  |  | 01/31/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802655 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
636 | 1220648-2026-02535 | 24290334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24290334&pc=OZD | 01/23/2026 |  |  | 02/09/2026 |  |  | 01/30/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798483 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | No Apparent Adverse Event | [5 empty] | Heart Failure/Congestive Heart Failure | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient with an unknown demographic background was supported with an impella 5.5 for the indication of cardiomyopathy. the first impella 5.5 device (sn (b)(6)) was surgically implanted via the right axillary/subclavian artery on (b)(6) 2026 at 6:30 pm eastern (cet (b)(6) 2026 12:30 am) and explanted 15 minutes later due to pump stop. the patient survived to explant, and the device was replaced with a second impella 5.5 (sn (b)(6)), which was implanted at 6:45 pm eastern (cet (b)(6) 2026 12:45 am). the second device remained in use until (b)(6) 2026 at 7:30 pm eastern (cet (b)(6) 2026 1:30 am), at which time care was withdrawn and the patient subsequently expired. based on the available information, the patient¿s death occurred following withdrawal of care and cannot be directly attributed to device malfunction. no evidence has been provided to suggest that either impella 5.5 device contributed to the fatal outcome. the impella 5.5will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange, however the exchange was performed with no consequence to the patient and support. the death is conservatively being reported on the impella 5.5 but is unlikely a contributing factor to cause of death and is more likely due to the patient condition. no evidence has been provided to suggest that either impella 5.5 device contributed to the fatal outcome. | [5 empty] | Clinical rationale update: upon further review the death will be coded to pump 1 (mfr 1220648-2026-02256) and withdrawn from pump 2 (mfr 1220648-2026-02535). clinical rationale: a patient with an unknown demographic background was supported with an impella 5.5 for the indication of cardiomyopathy. the first impella 5.5 device (sn (b)(6)) was surgically implanted via the right axillary/subclavian artery and explanted 15 minutes later due to pump stop. the device was replaced with 5.5 (sn (b)(6)), which was implanted at. the second device remained in use until at which time care was withdrawn and the patient subsequently expired. based on the available information the death will be coded to the initial impella 5.5 but is unlikely a contributing factor to the patients death. any additional information received will be filed with mfr 1220648-2026-02256. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
637 | 1220648-2026-02536 | 24290370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24290370&pc=OZD | 01/29/2026 |  |  | 02/09/2026 |  |  | 01/29/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746109 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
638 | 1220648-2026-02537 | 24290386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24290386&pc=OZD | 01/25/2026 |  |  | 02/09/2026 |  |  | 01/25/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026778651 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
639 | 1220648-2026-02538 | 24290396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24290396&pc=OZD | 02/02/2026 |  |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745082 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female | 95 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was surgically inserted via direct aortic access in a 46-year-old female who presented with acute myocardial infarction complicated by cardiogenic shock. the patient had a known history of coronary artery disease and diabetes mellitus and was receiving inotropic support, vasopressors, and mechanical respiratory support. the patient was classified as scai stage e. the device was placed electively in preparation for planned coronary artery bypass graft surgery. during support with the impella 5.5 device, the patient experienced a neurologic event consistent with stroke. additional information reported significant left internal carotid artery stenosis of approximately 75 percent. the patient experienced stroke during impella 5.5 support. after a comprehensive review of the available information, the reported event did not identify evidence suggesting a causal relationship between the impella 5.5 device and the stroke. the event was more likely attributable to the patient¿s underlying cerebrovascular disease, significant carotid artery stenosis, severe cardiogenic shock, and critical illness in the perioperative setting. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
640 | 1220648-2026-02542 | 24290538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24290538&pc=OZD | 01/28/2026 |  |  | 02/09/2026 |  |  | 01/28/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026789751 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
641 | 1220648-2026-02544 | 24290590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24290590&pc=OZD | 01/31/2026 |  |  | 02/09/2026 |  |  | 01/31/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026799966 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
642 | 1220648-2026-02545 | 24290616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24290616&pc=OZD | 01/31/2026 |  |  | 02/09/2026 |  |  | 02/01/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782789 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella cp device was inserted into the left femoral artery in a 41-year-old male patient presenting with cardiomyopathy and in scai stage d shock on an intra-aortic balloon pump (iabp) prior to initiation of support. while in the intensive care unit (icu), hemolysis was observed with pink/red urine on level p-7. the impella was lowered to p-3. the lactate dehydrogenase (ldh) was 1.8 (previously 2.7 while on iabp pre-impella). the patient subsequently experienced ventricular tachycardia (vt) and extracorporeal membrane oxygenation (ecmo) was initiated for additional support. the patient¿s vital signs stabilized after ecmo insertion. the physician indicated vt was caused by myocarditis, not impella. the device functioned at p-3 at 2.0 l/min with no issues. the post-procedure outcome is unknown. hemolysis and arrhythmias are listed as potential adverse events and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | Explant date provided therefore d6b updated. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
643 | 1220648-2026-02547 | 24290824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24290824&pc=OZD | 02/02/2026 |  |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798738 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
644 | 1220648-2026-02551 | 24290982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24290982&pc=OZD | 02/02/2026 |  |  | 02/09/2026 |  |  | 02/03/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802441 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 65 KG |  |  | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was placed via the axillary graft site for the 71 year old female patient admitted in with acute myocardial infarction and cardiogenic shock, presenting in scai stage d shock. there were no known cardiac or systemic comorbidities shared. the pump was being placed however the subclavian artery was found to be dissected post the insertion of the 5.5 pump through the introducer. the pump placement was aborted and patient has survived the harm. vascular injury is a known risk associated with impella support as described in the instructions for use. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
645 | 1220648-2026-02557 | 24291576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24291576&pc=OZD | 01/29/2026 |  |  | 02/09/2026 |  |  | 01/30/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025571895 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Inflammation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
646 | 1220648-2026-02558 | 24291589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24291589&pc=OZD | 02/02/2026 |  |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795750 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was placed via the axillary artery graft to support a 54-year-old male patient who had been admitted in for surgery, the pump was placed pre-operatively. there was no medical history shared for the patient. the pump was supporting when after admitting to the intensive care unit (icu) the team noted a hematoma at the axillary site. the patient returned to the operating room and had the pocket re-opened, and the team observed a small arterial muscular bleed that was surgically repaired. hemostasis was achieved. the pump remains on for support and the anticoagulation necessary for the pump placement and support can contribute to bleeding. vascular injury is a known risk associated with impella support as described in the instructions for use. additional information the customer reported that the bleed resolved. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
647 | 1220648-2026-02563 | 24291746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24291746&pc=OZD | 02/02/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806916 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Use of Device Problem | Failure to Advance | [5 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
648 | 1220648-2026-02565 | 24291897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24291897&pc=OZD | 02/03/2026 |  |  | 02/09/2026 |  |  | 02/03/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797226 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female | 56 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received indicating that the impella device was explanted. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the nurse noticed that the distal end of the anticontamination sleeve no longer attached to touy-borst lock. tegaderm applied for additional securement. it was reported that the device position was stable. |  |  |  |  | Additional information received for d6 explant. section d catalog number was corrected section e initial reporter information was added since it was omitted from the initial. | D6b: explant date is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
649 | 1220648-2026-02567 | 24291975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24291975&pc=OZD | 02/03/2026 |  |  | 02/09/2026 |  |  | 02/03/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797225 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
650 | 1220648-2026-02575 | 24292157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24292157&pc=PYX | 02/01/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP PUMP SET, EU |  |  | PYX |  |  | IMPELLA RP PUMP SET, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025593625 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 70 YR |  | Male | 95 KG |  |  | Injury | Yes | No | Difficult or Delayed Positioning | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella rp was placed to support the 70 year old male patient admitted in with right heart failure, post operative for cardiothoracic surgery, presenting in scai stage e shock. the pump was supporting for 5 days when the pump dislocated out of appropriate position and the console was alerting for suction. the pump was successfully repositioned however the pump had persistent decreasing in pump flow. despite the drop in flows the pump was supporting effectively per the team. the pump was explanted due to the positional issues and flow, and at the time of explant the team observed thrombus in the right atrium which prompted the team to use the inari device for thrombectomy. a new rp was placed and support continues. the pump replacement due to flow and positional issues will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange, however the exchange was performed with no consequence to the patient and support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
651 | 1220648-2026-02578 | 24292314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24292314&pc=OZD | 01/30/2026 |  |  | 02/09/2026 |  |  | 01/30/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775744 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 140 KG | Hispanic |  | Injury | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the left axillary artery in a 78-year-old female patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). the patient¿s medical history was significant for prior coronary artery bypass grafting (cabg) and diabetes mellitus. during the procedure, the left axillary artery was determined to be too small to accommodate the device, and the impella 5.5 could not be advanced despite standard procedural attempts. as a result, a device revision or replacement was required. the reported events are failure to advance, medical device revision or replacement, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
652 | 1220648-2026-02585 | 24292594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24292594&pc=OZD | 01/30/2026 |  |  | 02/09/2026 |  |  | 01/30/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795404 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
653 | 1220648-2026-02582 | 24292597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24292597&pc=OZD | 01/30/2026 |  |  | 02/09/2026 |  |  | 01/30/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796313 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 89 KG | Non Hispanic | White | Injury | Yes | No | Failure to Pump | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in a 63-year-old female patient presenting with cardiomyopathy and a medical history significant for renal insufficiency. the patient was in cardiogenic shock, scai stage, and was taken for escalation of mechanical circulatory support from an impella cp to an impella 5.5. the surgeon successfully obtained axillary access and advanced the impella 5.5 into the left ventricle without difficulty. the impella cp was withdrawn from the left ventricle immediately prior to placement of the impella 5.5, and a double-barrel technique was not utilized. when the impella 5.5 was advanced from p-0 to p-1, an "impella stopped" alarm was noted. troubleshooting and restart measures were attempted without success, and the pump was unable to start. the decision was made to prime and implant another impella 5.5 device. the reported events are pump unable to start, medical device revision or replacement, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however, the event occurred in the setting of scai stage e shock, where the patient had limited physiologic reserve and was undergoing high-risk escalation of support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
654 | 1220648-2026-02586 | 24292605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24292605&pc=OZD | 02/01/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026807801 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 71 YR |  | Unknown |  |  |  | Injury | Yes | No | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp was placed via the femoral access to support the patient of unknown age or gender who had been admitted in with acute myocardial infarction and cardiogenic shock, presenting in scai stage d shock. the pump was placed but, there were low flows on the console. the patient had been effectively anticoagulated with heparin and the clotting time (act) was out of range, and so the team made decision to explant and replace the pump due to the flow issue and concerns of clot ingestion. the pump was explanted and replaced and support proceeded. the low flow will be reported for the support interruption during the exchange, however the exchange was performed with no consequence to the patient and support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
655 | 1220648-2026-02590 | 24292716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24292716&pc=OZD | 02/02/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026774605 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
656 | 1220648-2026-02594 | 24292832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24292832&pc=OZD | 02/01/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796297 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
657 | 1220648-2026-02595 | 24292853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24292853&pc=OZD | 02/03/2026 |  |  | 02/09/2026 |  |  | 02/03/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026809838 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male | 101 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted surgically into the left aorta in a 33-year-old male patient presenting in cardiomyopathy, in scai stage e shock, prior to initiation of support. the patient was in the intensive care unit (icu), recovering from a heart and liver transplant from the previous week. the patient had worsening post-transplant shock and placed on venous-arterial (va) extracorporeal membrane oxygenation (ecmo) and impella 5.5. the patient developed a slow heart rate that was reading 0 on the monitor but was more likely ventricular escape beats and not capturing with a temporary pacemaker. the physician attributed the patient's critical condition and bradycardia to the heart transplant and metabolic disorders from being in multi-organ failure. two days later, the family elected to withdraw care, and the patient expired on support. the device is unlikely to have contributed to the patient's death, which was most likely due to the patient's existing comorbidities, along with the complications of stage e shock. | [5 empty] | Product complaint # (b)(4) the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
658 | 1220648-2026-02597 | 24292950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24292950&pc=OZD | 02/02/2026 |  |  | 02/09/2026 |  |  | 02/03/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026812155 |  | 1000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female | 70 KG |  |  | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | Asystole | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella cp device was inserted into the right femoral artery in a 71-year-old female patient presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), in scai stage e shock, prior to initiation of support. the impella was inserted for ventricular support during a high-risk percutaneous coronary intervention (pci). while the patient was in the intensive care unit (icu), the patient went into asystole and was treated with cardiopulmonary resuscitation (cpr) with return of spontaneous circulation (rosc). after cpr, there was suction in all p-levels and nearly no impella flow, with flat motor current (mc). a transthoracic echocardiogram (tte) was done, which showed the left ventricle (lv) filled and not unloaded. the impella was repositioned without success. it was noted that there was a kink in the cannula visible under fluoroscopy; therefore, the pump was replaced. the patient was successfully replaced with another impella cp; however, several hours after the second device was implanted, the family elected to withdraw care, and the patient expired on support. the device is unlikely to have contributed to the patient's death, which was most likely due to the clinical condition of a critically ill patient with acute myocardial infarction complicated by cardiogenic shock as they presented in stage e shock. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
659 | 1220648-2026-02562 | 24292974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24292974&pc=OZD | 02/01/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
660 | 1220648-2026-02599 | 24293085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24293085&pc=OZD | 10/31/2025 | 02/27/2026 |  | 02/09/2026 |  |  | 10/31/2025 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026741023 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 65-year-old male with no significant known past medical history presented with chest pain and shortness of breath. the patient was taken to the cardiac catheterization laboratory and underwent percutaneous coronary intervention. following the procedure, the patient experienced worsening shortness of breath and pulmonary edema with an elevated end-diastolic pressure, and an impella cp device was placed for hemodynamic support. on the same day, due to worsening hemodynamics, the decision was made to escalate the patient to an impella 5.5. following placement of the impella 5.5 via a vascular graft to the right axillary artery, oozing at the access site occurred requiring abortion of placement. vascular surgery treated the access site and a mini thoracotomy was carried out to access the anterior aorta for direct placement of another pump. during this procedure, the patient was still supported with the previous impella 5.5 and a direct exchange of both pumps was carried out. following implantation, in the intensive care unit, the patient developed supraventricular tachycardia. he was treated with magnesium and dobutamine rate was decreased. the patient¿s heart rate returned to normal. after two days, the patient could be successfully weaned and the impella 5.5 was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
661 | 1220648-2026-02601 | 24293121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24293121&pc=OZD | 02/02/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026689082 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
662 | 1220648-2026-02615 | 24295172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24295172&pc=OZD | 02/02/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026655577 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 58 YR |  | Female | 67 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received indicating that the impella device was explanted. | A 58 year old female undergoing pre operative placement for impella 5.5 support, with pre support status consistent with scai stage d, was noted to be thrombocytopenic prior to implantation with a platelet count of 64. the impella 5.5 was surgically placed via the right axillary/subclavian artery. following transfer to the icu, the patient developed access site bleeding characterized as oozing from both the cutdown and insertion sites. the patient had received 6,000 units of heparin intraoperatively; however, due to bleeding, heparin administration was held overnight. hemostasis was attempted using sandbag pressure and 4x4 gauze. two units of ffp were administered, and estimated blood loss was approximately 0.25 units. based on available information, the access site bleeding appears most likely related to the patient¿s underlying thrombocytopenia and anticoagulation exposure rather than a device malfunction. the patient continues with impella 5.5 and no device related injury can be confirmed at this time while the patient remains on support. |  |  |  |  | Additional information received for d6 explant. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
663 | 1220648-2026-02618 | 24295559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24295559&pc=OZD | 01/31/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 01/31/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026705948 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
664 | 1220648-2026-02621 | 24295736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24295736&pc=OZD | 01/22/2026 |  |  | 02/09/2026 |  |  | 01/22/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026775869 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 84 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An 81-year-old male with a medical history significant for prior coronary artery bypass graft (cabg) surgery, coronary artery disease, peripheral vascular disease, and diabetes mellitus underwent implantation of an impella cp device for temporary mechanical circulatory support. the indication for use was cardiogenic shock secondary to an acute myocardial infarction. at the time of impella cp initiation, the patient was classified as society for cardiovascular angiography and interventions (scai) stage b cardiogenic shock, with clinical deterioration to scai stage e during the course of hospitalization. the impella cp device was placed; however, it was explanted on the day of insertion due to limb ischemia. documentation indicated loss of distal pulses in the affected limb, and imaging demonstrated absence of flow. no other intervention beyond the pump explant is known. following explant of the impella cp, the patient was transitioned to an axillary impella 5.5 device for continued mechanical circulatory support. the impella 5.5 was successfully implanted and remained in place for approximately seven days, after which the patient expired. the cause of death was not further detailed in the available information. in this case, based on the available information, the patient¿s underlying condition contributed most to the demise outcome (progression of acute myocardial infarction¿related cardiogenic shock [scai stage b to stage e] in the setting of advanced coronary artery disease with prior cabg, peripheral vascular disease, diabetes mellitus). there was no report or indication of a device malfunction or device-related adverse event; however, as the impella 5.5 device was in place at the time of the patient¿s demise, the event is conservatively reported as a death-type reportable event. | [5 empty] | Patient information: ethnicity is unknown. device status: the impella device was not received from the customer; therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
665 | 1220648-2026-02622 | 24295879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24295879&pc=OZD | 10/24/2025 |  |  | 02/09/2026 |  |  | 10/24/2025 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703539 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during support of impella cp the patient had bleeding around the impella site. impella redressed. protamine given and bleeding resolved. the day after there was pink tinged urine. elevated ldh. impella repositioned with echo small amount of volume given. remains on low dose dobutamine. plan to continue to rest lv and possibly wean impella successfully weaned and explanted. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
666 | 1220648-2026-02623 | 24295922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24295922&pc=OZD | 02/02/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796263 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An (b)(6) year old female underwent a high-risk percutaneous coronary intervention (hrpci) on (b)(6) 2026. a 5.5 impella device was selected for placement via the right axillary/subclavian artery using a surgical approach. implantation was attempted; however, the patient¿s anatomy, including vessel stenosis and the presence of resistance within the axillary artery, prevented advancement of the device. the event was attributed to patient specific anatomical factors, including stenosis within the axillary artery, which prevented successful delivery of the impella 5.5 device. there was no evidence of device malfunction or damage, and the complaint outcome was procedural abortion due to anatomical constraints rather than device performance. the impella 5.5 will be conservatively reported for hemodynamic instability due to temporary hemodynamic. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
667 | 1220648-2026-02637 | 24297566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24297566&pc=OZD | 02/02/2026 |  |  | 02/09/2026 |  |  | 02/02/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1186531 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | During setup for an emergent impella cp case, the automated impella controller (aic) failed to recognize the purge disc. a second aic was obtained, and the same purge cassette was successfully recognized and functioned as intended. | [5 empty] | E4 was inadvertently omitted on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
668 | 1220648-2026-02648 | 24298311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24298311&pc=OZD | 02/01/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 02/01/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026795340 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Heart Block | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
669 | 1220648-2026-02651 | 24298426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24298426&pc=OZD | 02/01/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 02/01/2026 | 02/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026756820 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
670 | 1220648-2026-02656 | 24298587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24298587&pc=OZD | 02/01/2026 | 02/09/2026 |  | 02/09/2026 |  |  | 02/01/2026 | 02/09/2026 |  |  |  |  | 02/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026775831 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was surgically inserted via direct aortic access in a 68-year-old male who presented for elective mechanical circulatory support placement in the setting of aortic valve disease and cardiogenic shock, scai stage d. the patient¿s medical history was notable for renal insufficiency. the patient required inotropes, vasopressors, and ventilator support. postoperatively, the patient was noted to have chest tube bleeding following chest closure. evaluation identified bleeding along the right sternal edge, with laboratory values consistent with coagulopathy. the patient received blood products, and hemostatic management was initiated. heparin was held at that time. no bleeding was identified at the impella insertion site or at the aortic graft anastomosis, and the bleeding was localized to the sternum following chest closure. the patient experienced a major bleed requiring blood transfusion and hemostatic intervention. after a comprehensive review of the available information, the reported event is consistent with postoperative surgical bleeding and coagulopathy. comprehensive review of the event did not identify evidence suggesting a causal relationship between the impella and the reported event. the patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. comprehensive review of the event did not identify evidence suggesting a causal relationship between the impella and the reported event. the patient survived.
671 | 1220648-2026-02518 | 24288288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24288288&pc=OZD | 02/02/2026 |  |  | 02/08/2026 |  |  | 02/02/2026 | 02/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026785129 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
672 | 1220648-2026-02519 | 24288292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24288292&pc=OZD | 01/31/2026 |  |  | 02/08/2026 |  |  | 02/02/2026 | 02/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806203 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
673 | 1220648-2026-02521 | 24288325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24288325&pc=OZD | 02/01/2026 |  |  | 02/08/2026 |  |  | 02/02/2026 | 02/08/2026 |  |  |  |  | 02/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1288546 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 122 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Display or Visual Feedback Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | An impella cp device was inserted via the right femoral artery in a 54-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs) and a known history of coronary artery disease. during support, the customer reported disappearance of the white waveform on the automated impella conroller(aic) . impella flows remained adequate throughout the event. out of caution and to ensure continuity of care, the clinical team elected to exchange the controller. the exchange was performed without complication, and no adverse patient effects were noted. following successful percutaneous coronary intervention (ppci), the patient no longer required mechanical circulatory support. the reported events include display issue, device revision or replacement, medical device removal, and hemodynamic instability. the impella cp is being conservatively reported for serious injury due to a temporary interruption of hemodynamic support during the controller exchange; however, the exchange was performed without consequence to the patient, and support was maintained without any known adverse events.t6 | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
674 | 1220648-2026-02522 | 24288329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24288329&pc=OZD | 02/02/2026 |  |  | 02/08/2026 |  |  | 02/02/2026 | 02/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802406 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
675 | 1220648-2026-02525 | 24288346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24288346&pc=OZD | 02/02/2026 |  |  | 02/08/2026 |  |  | 02/02/2026 | 02/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1138943 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
676 | 1220648-2026-02527 | 24288420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24288420&pc=OZD | 01/13/2026 |  |  | 02/08/2026 |  |  | 01/13/2026 | 02/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026795549 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission
677 | 1220648-2026-02511 | 24285688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24285688&pc=OZD | 02/01/2026 | 02/07/2026 |  | 02/07/2026 |  |  | 02/01/2026 | 02/07/2026 |  |  |  |  | 02/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744142 |  | 1000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 37 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | No | Backflow | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 was placed via the axillary graft access to support the 37 year old male patient. he was admitted in with cardiomyopathy, presenting in scai stage e shock on inotropes and vasopressors. on thee 4th day of support the patient had a postoperative bleed in the chest. the pump was having intermittent suction alarms and this prompted the team to evaluate pump position. the axillary graft site was also observed to be bleeding at the anastamosis. there was no need for blood product replacement as the hemoglobin level was satisfactory. when the p-level was adjusted the suction resolved. the pump remains on for support and the anticoagulation necessary for the pump placement and support can contribute to bleeding. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
678 | 1220648-2026-02512 | 24285693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24285693&pc=OZD | 02/01/2026 | 02/07/2026 |  | 02/07/2026 |  |  | 02/01/2026 | 02/07/2026 |  |  |  |  | 02/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1914207 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 72 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | An impella cp device was inserted via the right femoral artery in a 72-year-old male patient presenting with high-risk percutaneous coronary intervention (hrpci) and a known history of coronary artery disease, diabetes mellitus, and renal insufficiency requiring dialysis. during support, the customer reported a purge disc not detected issue while using the automated impella controller (aic). in response, a backup aic was utilized, and the affected controller (b)(6) was removed from circulation. the patient remained supported, and subsequently survived to explant. the reported events include purge disc not detected, medical device revision or replacement, and hemodynamic instability. the impella cp is being conservatively reported for serious injury due to a temporary interruption of hemodynamic support during the controller exchange; however, the exchange was performed without consequence to the patient, and support was maintained without any known adverse events. | [5 empty] | The impella device has not been received from the customer; therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | B1 product problem and e4 was inadvertently omitted on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated.
679 | 1220648-2026-02513 | 24285694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24285694&pc=OZD | 02/01/2026 | 02/07/2026 |  | 02/07/2026 |  |  | 02/01/2026 | 02/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026799246 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | [9 empty] | Injury | Yes | No | Pumping Stopped | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
680 | 1220648-2026-02514 | 24285756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24285756&pc=OZD | 01/31/2026 | 02/07/2026 |  | 02/07/2026 |  |  | 02/01/2026 | 02/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026735734 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | [9 empty] | Malfunction | No | Yes | Difficult to Open or Close | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
681 | 1220648-2026-02516 | 24285760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24285760&pc=OZD | 02/02/2026 | 02/07/2026 |  | 02/07/2026 |  |  | 02/02/2026 | 02/07/2026 |  |  |  |  | 02/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, CA |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, CA |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1096255 |  | 0042-0000-CA |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 68 YR |  | Female | 63 KG |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported their automated impella controller (aic) had a brief loss of waveforms with a controller error message. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
682 | 1220648-2026-02517 | 24285763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24285763&pc=OZD | 02/02/2026 | 02/07/2026 |  | 02/07/2026 |  |  | 02/02/2026 | 02/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026747670 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
683 | 1220648-2026-02416 | 24277028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24277028&pc=OZD | 11/05/2025 |  |  | 02/06/2026 |  |  | 11/05/2025 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026755775 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 91 KG | Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Renal Failure | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical narrative: a 59-year-old male with a medical history significant for nonischemic cardiomyopathy status post left ventricular assist device (heartmate ii) and subsequent orthotopic heart transplant, hypertension, hyperlipidemia, type 2 diabetes mellitus, and obstructive sleep apnea on home continuous positive airway pressure received an impella 5.5® device for temporary mechanical circulatory support due to cardiogenic shock secondary to heart failure. the device was surgically implanted via the right axillary artery. at the time of device initiation, the patient was classified as society for cardiovascular angiography and interventions (scai) stage c cardiogenic shock, with subsequent progression to scai stage d cardiogenic shock. on the day following implantation, the patient required and received hemodialysis with a net ultrafiltration goal of 3 liters. the treatment plan included continuation of current therapies with intent to bridge the patient to orthotopic heart and kidney transplantation. the patient received continuous intravenous infusions of bivalirudin for anticoagulation and dobutamine for inotropic support. eight days following implantation, the patient received an additional hemodialysis treatment with a net ultrafiltration goal of 2 liters. several days later, nursing staff reported a placement signal not reliable (psnr) alarm. the alarm function and recommended monitoring were reviewed with the care team. as described in the impella instructions for use, sensor-related alarms do not impact pump function, and no patient harm occurred because of the psnr alarm. the care team was instructed on disabling the audible alarm. some days following, the patient was bridged to transplant. the impella device was weaned, explanted with a survived patient outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
684 | 1220648-2026-02417 | 24277031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24277031&pc=OZD | 01/27/2026 |  |  | 02/06/2026 |  |  | 01/27/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1085765 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 102 KG | Non Hispanic | White | Injury | Yes | No | Decrease in Pressure | Pumping Stopped | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | An impella cp device was inserted into the left femoral artery in a 66-year-old male patient with a past medical history of coronary artery disease (cad), a prior coronary artery bypass graft (cabg), and renal insufficiency. the patient presented in cardiomyopathy and in scai stage c shock, prior to initiation of support. while the patient was in the intensive care unit (icu), the nurse called to report that there was increased purge flow (pf) and decreased purge pressure (pp) so the automated impella controller (aic) console, was switched to a new one. however, now the left ventricle (lv) is about the aortic pressure (ao), and the motor current (mc) is above 1000. while on the call, the pump stopped. the abiomed clinical representative advised the nurse to restart with success on p-2, but the mc was still over 700. the device functioned at p-7 at 3.4l/min. the post-procedure outcome is survived at explant. the automated impella controller will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
685 | 1220648-2026-02420 | 24277364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24277364&pc=OZD | 01/29/2026 |  |  | 02/06/2026 |  |  | 01/29/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1337720 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced brief controller alarms on the supporting automated impella controller. during each event, the placement signal was lost, while motor current remained unchanged. in both instances, the alarm self resolved without intervention. additionally, optical sensor alerts occurred which were resolved by unplugging and reseating the white optical cable. support continued and the patient was in stable condition. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
686 | 1220648-2026-02422 | 24277419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24277419&pc=OZD | 01/27/2026 |  |  | 02/06/2026 |  |  | 01/27/2026 | 02/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779521 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
687 | 1220648-2026-02423 | 24277465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24277465&pc=OZD | 01/26/2026 |  |  | 02/06/2026 |  |  | 01/28/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1626081 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male | 70 KG |  | White | Malfunction | No | Yes | Degraded | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the cover of the white connector cable connection on an automated impella controller was damaged and the cover will not close. no patient harm was reported. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
688 | 1220648-2026-02428 | 24277597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24277597&pc=OZD | 01/28/2026 |  |  | 02/06/2026 |  |  | 01/28/2026 | 02/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026738973 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hematoma | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
689 | 1220648-2026-02433 | 24277741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24277741&pc=OZD | 01/30/2026 |  |  | 02/06/2026 |  |  | 01/30/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1407754 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 103 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Clinical rationale: an impella 5.5 device was inserted surgically into the right aorta in a 64-year-old male patient with a past medical history of coronary artery disease (cad), presenting in scai stage b shock, prior to initiation of support. the impella was inserted for ventricular support for a high-risk surgery (coronary artery bypass graft). while the patient was in the intensive care unit (icu), the automated impella controller (aic) console had a "controller error" alarm, the consoles were exchanged, no further issues were noted. the nurse reported that the impella never stopped working. a few days later, the impella was successfully weaned, and the post-procedure outcome is survived at explant. the device functioned at p-5 at 3.3l/min with no issues. the automated impella controller will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
690 | 24277811 | 24277811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24277811&pc=OZD | 01/08/2026 | 01/26/2026 | 02/02/2026 | 02/06/2026 | 02/02/2026 | 02/06/2026 |  | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 4000021 |  | No |  |  | Not Available |  |  | No |  |  |  | User Facility report |  | Non-Healthcare Professional | [6 empty] | 81 YR |  | Male | 77 KG |  |  | Malfunction | No | Yes | Break | Disconnection | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | During transfer of patient from ct table back to bed, the repositioning sheath on the impella broke (the plastic cover became disconnected from the leur lock), and the impella was removed by 5cm because the repositioning sheath was no longer keeping the impella in place). it was noted at the time that the rn mentioned no pulling on the line.
691 | 1220648-2026-02435 | 24277829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24277829&pc=OZD | 01/29/2026 |  |  | 02/06/2026 |  |  | 01/30/2026 | 02/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781120 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
692 | 1220648-2026-02441 | 24277938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24277938&pc=OZD | 01/29/2026 |  |  | 02/06/2026 |  |  | 01/30/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781170 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female | 61 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Clinical rationale: a 44-year-old female with cardiomyopathy underwent implantation of an impella 5.5 device via the right axillary/subclavian artery using a surgical approach on (b)(6) 2026 at 2:55 pm. after transfer to the icu, a small, non-expanding hematoma was noted near the impella insertion site. based on the available information, the small access site hematoma did not expand, required no intervention, and the patient remained clinically stable. at this time, there is no evidence that the impella device directly contributed to the event; however, a definitive attribution cannot be made due to several unknown factors. the patient remains on support at the time of reporting. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
693 | 1220648-2026-02444 | 24278031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24278031&pc=OZD | 01/31/2026 |  |  | 02/06/2026 |  |  | 01/31/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798742 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 73 KG | Non Hispanic | White | Injury | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was attempted via the right axillary artery in a 63-year-old female patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs) and diabetes mellitus. during the procedure, the right axillary artery was determined to be too small in caliber to accommodate the impella 5.5. after unsuccessful advancement of a .035-inch wire and a 6 fr pigtail catheter, an angiogram was performed. the implanting physician estimated the vessel diameter to be approximately 5 mm and noted significant vessel disease at the planned anastomosis site. alternative access options, including the left axillary artery and an innominate approach, were discussed. however, consent was not available to proceed with a direct chest approach. as a result, the impella 5.5 procedure was aborted, and the device was not implanted. the existing impella cp device was left in place to continue hemodynamic support. the reported events include failure to advance, medical device removal, and hemodynamic instability. the impella 5.5 is being conservatively reported for serious injury due to a temporary interruption of hemodynamic support associated with procedural manipulation; however, the procedure was stopped without consequence to the patient, and hemodynamic support was maintained without any known adverse events. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
694 | 1220648-2026-02446 | 24278084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24278084&pc=OZD | 01/31/2026 |  |  | 02/06/2026 |  |  | 01/31/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773826 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Increase in Pressure | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right innominate artery in a 62-year-old female patient presenting with post-cardiotomy cardiogenic shock (pccs) and a known history of coronary artery disease. during support, the bedside nurse reported a high purge pressure alarm followed by a purge system blocked alarm. the purge cassette was exchanged, and the purge solution was changed from 1 l d5w with 25 meq of sodium bicarbonate to 1 l d5w with 25,000 units of heparin. the intensivist contacted the cardiothoracic surgeon, who came to the bedside. a bedside echocardiogram was performed, demonstrating recovery of left ventricular ejection fraction to approximately 50%, and the patient was successfully weaned down on inotropic and vasopressor support. based on improved hemodynamics and ventricular recovery, the treating physician elected to take the patient to the operating room for impella 5.5 removal. the device was removed due to hemodynamic stability and no longer requiring mechanical circulatory support. the reported events include purge pressure high, medical device removal, and hemodynamic instability. the impella 5.5 is being conservatively reported for serious injury due to a temporary interruption of hemodynamic support associated with purge system management and purge cassette exchange; however, the intervention and removal were performed without consequence to the patient, and the patient demonstrated clinical improvement. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
695 | 24278163 | 24278163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24278163&pc=OZD | 07/01/2025 | 02/02/2026 | 02/02/2026 | 02/06/2026 | 02/02/2026 | 02/06/2026 |  | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC |  |  | US |  |  |  | 4000021 |  | No |  |  | Not Available |  |  | No |  |  |  | User Facility report |  | Non-Healthcare Professional | [6 empty] | NA |  | Male | 122 KG |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient arrived at the micu [medical intensive care unit] from the cath lab with the distal end of the impella sheath exposed. the clear sheath that comes connected to the distal lock and covers/maintains the sterility of the impella was not secure within the distal locking mechanism. the sheath is supposed to be directly connected to the distal locking device.
696 | 1220648-2026-02453 | 24279268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24279268&pc=OZD | 12/04/2025 |  |  | 02/06/2026 |  |  | 12/04/2025 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025571160 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 78-year-old male patient was admitted to the hospital with acute myocardial infarction cardiogenic shock under cardiopulmonary resuscitation with a mean atrial pressure of 62 mmhg (under more than 3 catecholamines and respiratory support) with a left ventricular ejection fraction of 20 mmhg, in society for cardiovascular angiography and interventions shock stage e. he had a known history of coronary artery disease and renal insufficiency and was on intra-aortic balloon pump prior to impella cp support. activated clotting time prior to the impella cp device was 298 seconds. impella cp was placed for hemodynamic support and patient transferred to the intensive care unit on highest impella performance level 9. weaning of medication was ongoing and repositioning of the device by best practices was performed. continuous renal replacement therapy was re-initiated and well tolerated. surgery of mitral, tricuspid valve and coronary artery bypass grafting in addition to a maze procedure were still pending. on the second day of impella cp support, the left main and left anterior descending were stented as planned and the left circumflex artery ballooned. after that, patient suffered from atrial fibrillation, vasopressors had to be increased and lactate trended up. afterwards, an escalation to impella 5.5 was prepared in the operation room and a right axillary artery dissection was seen under fluoroscopic guidance before impella 5.5 insertion. the dissection could be repaired with a balloon and impella 5.5 advanced without complications thereafter. on the second day of impella 5.5 support, the patient was transferred to another hospital and finally successfully weaned on day 3. the impella cp is coded not for renal failure or atrial fibrillation as these were pre-existent. the impella 5.5 guide wire is coded for vascular dissection, which is related to the planned device insertion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
697 | 1220648-2026-02454 | 24279291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24279291&pc=OZD | 01/30/2026 |  |  | 02/06/2026 |  |  | 01/30/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026771309 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received indicating that the impella device was explanted. | The complainant reported that a patient implanted with an impella 5.5 experienced alarms for placement signal not reliable and that the placement signal waveforms were not present during support. it was reported that there were no values present for the placement signal; however, pump flows and motor current were unaffected. a point of care ultrasound was performed to verify the pump was properly positioned. it was noted that the pump was 4.7 cm across the aortic valve toward the left ventricle apex. it was noted that the pump's function remained as expected with only the sensor malfunctioning. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. at the time of writing this report, the patient continues to be supported by impella 5.5. |  |  |  |  | Additional information received for d6 explant. | A4 weight is unknown. device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
698 | 1220648-2026-02456 | 24279456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24279456&pc=OZD | 01/29/2026 | 02/06/2026 |  | 02/06/2026 |  |  | 01/29/2026 | 02/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781119 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
699 | 1220648-2026-02460 | 24279824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24279824&pc=OZD | 01/28/2026 |  |  | 02/06/2026 |  |  | 01/28/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1117560 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Unknown |  |  |  | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale an impella cp device was inserted via the left femoral artery in a 36-year-old patient. during hospitalization, automated impella control (aic) was turned on, at which time a controller failure was noted. in response, the care team performed a controller exchange. the reported events are controller failure, device revision or replacement, and hemodynamic instability. the impella cp is being conservatively reported for hemodynamic instability due to a temporary interruption of support during the controller exchange; however, the exchange was performed without known patient consequence, and support was resumed. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
700 | 1220648-2026-02461 | 24279858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24279858&pc=OZD | 02/01/2026 | 02/06/2026 |  | 02/06/2026 |  |  | 02/01/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025563969 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 69 YR |  | Male | 86 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Hematuria | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 device was inserted via the right axillary artery in a 69-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs), with a medical history significant for coronary artery disease and renal insufficiency. during support, hematuria was observed in the foley catheter. hematuria is a known clinical finding in critically ill patients receiving mechanical circulatory support and may be influenced by underlying shock physiology, renal dysfunction, anticoagulation requirements, and urinary catheterization. it is a known risk described in the instructions for use (ifu). the patient¿s clinical condition continued to decline, care was subsequently withdrawn, and the patient expired. the death is being conservatively reported to the impella 5.5 ; however, the device is unlikely to have contributed to the patient¿s death, which is most plausibly attributed to the underlying critical condition and refractory cardiogenic shock. | [5 empty] | Section g has been corrected since it was omitted from the initial report. h6 health effect - clinical code has been updated. | The impella device has not been received from the customer; therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
701 | 1220648-2026-02463 | 24279914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24279914&pc=OZD | 01/30/2026 |  |  | 02/06/2026 |  |  | 01/30/2026 | 02/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1254075 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Application Network Problem | Battery Problem | [5 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
702 | 1220648-2026-02464 | 24280059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24280059&pc=OZD | 02/01/2026 | 02/06/2026 |  | 02/06/2026 |  |  | 02/01/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746068 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 57 YR |  | Male | 86 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was inserted via the left axillary subclavian access site in a 57-year-old male. the patient presented with postcardiotomy cardiogenic shock and low cardiac output syndrome, skystage d, in the setting of known coronary artery disease and prior coronary artery bypass grafting with planned valve replacement. the patient required support with inotropes, vasopressors, and mechanical ventilation for respiratory support. during support with the impella 5.5, nursing staff notified the care team of active bleeding from the insertion site. the reported events included a major bleed and blood transfusion related to the access site, with an estimated 1,155 ml of output since insertion. the patient received five units of packed red blood cells and two units of platelets. vascular surgery and critical care medicine teams were aware of the bleeding. the patient was receiving a continuous heparin infusion at 22 units per hour. based on the available information, the reported access site bleeding and need for transfusion are consistent with the patient¿s critical clinical condition, including postcardiotomy cardiogenic shock and low cardiac output syndrome, scai stage d, the requirement for therapeutic anticoagulation with heparin in the setting of valve replacement, and known risks associated with large-bore axillary subclavian vascular access in a severely ill patient. comprehensive review of the event did not identify evidence suggesting a causal relationship between the impella and the reported event. the patient survived. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
703 | 1220648-2026-02466 | 24280143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24280143&pc=OZD | 01/28/2026 |  |  | 02/06/2026 |  |  | 01/29/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797927 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Difficult to Advance | Patient Device Interaction Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Intraoperative Cardiac Injury | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was being placed in a 72 year old male via the axillary artery graft to support the year old male patient who had been admitted into the medical center for a planned surgery. the patient was to have an aortic valve replacement performed for known valvular disease. during the delivery of the 5.5 pump there was difficulty traversing anatomy and getting across the valve, as such they troubleshot with catheter and wire exchanges. there was damage to the aortic root during the procedure, the cause of which could not be fully determined to be either the wire or catheter touching the root. surgical repair was successful, blood was given to the patient, and the pump was weaned and explanted with patient survival. the cardiac damage will be conservatively reported, despite evidence of causality. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
704 | 1220648-2026-02467 | 24280154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24280154&pc=OZD | 02/01/2026 | 02/06/2026 |  | 02/06/2026 |  |  | 02/01/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1129859 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 74 YR |  | Female | 74 KG | Non Hispanic | White | Injury | Yes | No | False Alarm | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp device was inserted into the right femoral artery in a 74-year-old female patient presenting with an out-of-hospital cardiac arrest requiring cardiopulmonary resuscitation (cpr), in acute myocardial infarction (ami) and cardiogenic shock (cgs), in scai stage e shock, prior to initiation of support. the impella was inserted for ventricular support for a high-risk percutaneous coronary intervention (pci). while the patient was in the intensive care unit (icu), there was access site bleeding requiring six units of packed red blood cells (prbcs), 3 units of platelets, and albumin. the patient needed to be paralyzed due to lower extremity movement leading to the bleeding. it was noted that the peel-away sheath was left in place with an antegrade sheath. the nurse noted a lack of distal pulses to the extremity, which also appeared to be mottled/pale, and cool in temperature. the automated impella controller (aic) had an ¿impella in aorta¿ alarm. the physician stated that he followed directions to decrease flows to p-2 and confirm placement. he stated that when he did this, the patient decompensated, requiring additional medications for support. upon checking placement, the impella was noted to be in good position. the impella was successfully weaned the following day, with an escalation of therapy to an impella 5.5. it was noted that the popliteal and tibial pulses were present on doppler when the impella cp was removed. the device functioned at p-9 at 3.3l/min with no issues. the reported event involves access site limb ischemia with device removal and bleeding. bleeding is a known risk due to the impella's anticoagulation and purge requirements. limb ischemia and bleeding are listed as potential adverse events, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
705 | 1220648-2026-02468 | 24280308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24280308&pc=OZD | 01/25/2026 |  |  | 02/06/2026 |  |  | 01/27/2026 | 02/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806256 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
706 | 1220648-2026-02479 | 24282586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24282586&pc=OZD | 01/29/2026 |  |  | 02/06/2026 |  |  | 01/29/2026 | 02/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1980194 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
707 | 1220648-2026-02486 | 24283254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24283254&pc=OZD | 01/30/2026 |  |  | 02/06/2026 |  |  | 01/30/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026804225 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female | 113 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via the left axillary artery in a 58-year-old female patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs) and a history of known coronary artery disease (cad). during the procedure, the impella 5.5 was unable to advance past the surgical anastomosis while being delivered over an abiomed 0.018-inch guidewire, requiring significant manipulation. the operator then attempted advancement over a 0.035-inch guidewire; however, the device remained unable to pass. the impella 5.5 was subsequently removed, and fibers/debris (sediment) were noted on the pump. out of an abundance of caution for the patient, the decision was made to exchange the device for a new impella 5.5. the original device was never activated, and there were no additional allegations regarding device performance. the reported events are failure to advance, medical device revision or replacement, and hemodynamic instability. the impella 5.5 will be conservatively reported for serious injury due to temporary hemodynamic support interruption during the exchange; however the exchange was performed with no consequence to the patient and support. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
708 | 1220648-2026-02487 | 24283260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24283260&pc=OZD | 01/27/2026 |  |  | 02/06/2026 |  |  | 01/27/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1261102 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted via the right femoral artery in a (b)(6) male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs) and a history of renal insufficiency. during support, a controller failure was reported. the impella pump continued to run, and flows did not stop. the alarm subsequently resolved; however, out of caution, an automated impella controller (aic)¿to¿aic exchange was performed, and a new controller was placed. the reported events are controller failure, device revision or replacement, and hemodynamic instability. although the case is coded as product malfunction, the impella cp is being conservatively reported for serious injury due to a temporary interruption of hemodynamic support during the console exchange; however, the exchange was performed with no consequence to the patient, and hemodynamic support was resumed. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
709 | 1220648-2026-02492 | 24283915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24283915&pc=OZD | 01/29/2026 | 02/06/2026 |  | 02/06/2026 |  |  | 01/29/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026744157 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65 year old male with an impella 5.5 due to cardiomyopathy. during support, there was an arterial placement signal drifting/low. there was no patient harm. the patient remains on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received for d6 explant.
710 | 1220648-2026-02495 | 24285262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24285262&pc=OZD | 01/30/2026 | 02/06/2026 |  | 02/06/2026 |  |  | 01/30/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026799999 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 92 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 69 year old male with cardiomyopathy and a pre support presentation consistent with scai stage d remained on impella 5.5 support via the left axillary/subclavian surgical access site. during routine patient rounding, the heart failure physician and care team reported evidence of hemolysis suspected to be related to impella positioning. preliminary review indicates a positioning related event in which the pump inflow was oriented toward the mitral apparatus, contributing to hemolysis. the device was repositioned, and no alarms or mechanical malfunctions were reported. further evaluation will be conducted upon product return. hemolysis is consistent with known device-related and patient-related factors documented in the instructions for use (ifu). the patient remains on support at the time of reporting. | [5 empty] | D6b: updated the explant date. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
711 | 1220648-2026-02496 | 24285263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24285263&pc=OZD | 01/30/2026 | 02/06/2026 |  | 02/06/2026 |  |  | 01/30/2026 | 02/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026805594 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
712 | 1220648-2026-02497 | 24285272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24285272&pc=OZD | 01/31/2026 | 02/06/2026 |  | 02/06/2026 |  |  | 01/31/2026 | 02/06/2026 |  |  |  |  | 02/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026797973 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 128 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Pain | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Impella 5.5 was inserted via right axillary artery in a 63-year-old male for pre-operative placement. the patient¿s comorbidities included coronary artery disease with previous cabg, and dialysis. the patient¿s clinical state prior to initiation of support was scai stage e. on day 9 of support, it was reported the patient was reaching above their head and experienced pain under the right arm, below the 5.5 axillary access site. a golf ball sized hematoma was noted below and lateral to the 5.5 axillary access site. manual pressure was applied for approximately 20 minutes. the hematoma and pain resolved. on day 13 of support, the impella 5.5 was explanted and the patient expired. death and vascular injury are known potential adverse events of the impella 5.5 per the instructions for use. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. death and vascular injury are known potential adverse events of the impella 5.5 per the instructions for use.
713 | 1220648-2026-02290 | 24266940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24266940&pc=OZD | 01/30/2026 |  |  | 02/05/2026 |  |  | 01/31/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026703406 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
714 | 1220648-2026-02291 | 24266953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24266953&pc=OZD | 01/19/2026 |  |  | 02/05/2026 |  |  | 01/21/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026707304 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
715 | 1220648-2026-02292 | 24267009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24267009&pc=OZD | 01/28/2026 |  |  | 02/05/2026 |  |  | 01/28/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798804 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
716 | 1220648-2026-02293 | 24267164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24267164&pc=OZD | 01/28/2026 |  |  | 02/05/2026 |  |  | 01/28/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026677013 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
717 | 1220648-2026-02295 | 24267242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24267242&pc=OZD | 01/28/2026 |  |  | 02/05/2026 |  |  | 01/28/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1547850 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 93 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced discrepant ao signals on the supporting automated impella controller. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
718 | 1220648-2026-02300 | 24267447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24267447&pc=OZD | 01/26/2026 |  |  | 02/05/2026 |  |  | 01/29/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1147695 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | [10 empty] | Injury | Yes | No | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
719 | 1220648-2026-02303 | 24267540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24267540&pc=OZD | 01/14/2026 |  |  | 02/05/2026 |  |  | 01/28/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026762849 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was inserted via direct aorta approach in a 59 year old male who was admitted for an aortic valve surgery. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage c. the patient received support with the 5.5 for the valve procedure. one day post-procedure, the patient experienced bleeding and underwent re-thoracotomy to remove the thrombus and hematoma. hypoxic encephalopathy was a result on unknown date. the source of bleeding was unknown. bleeding and thrombosis are known potential adverse events of the impella 5.5. per the instructions for use. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
720 | 1220648-2026-02305 | 24267552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24267552&pc=OZD | 01/29/2026 |  |  | 02/05/2026 |  |  | 01/29/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026705915 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information received that the pump got exchanged because the placement signal was not reliable and had an issue. | A 41 year old male with an indication for use of impella cp for acute myocardial infarction (ami) and cardiogenic shock (cgs), with pre support presentation consistent with scai stage c, underwent support with two impella cp pumps. the reported events were reviewed, including the loss of placement signal alarm observed with pump 1. despite intermittent loss of placement signal, motor current remained pulsatile, echocardiography consistently confirmed appropriate pump position, and the device continued to provide effective support. the customer acknowledged this as a known issue and managed the situation with increased imaging surveillance. pump 2 functioned without complaint. there is no indication that either impella device malfunctioned or contributed to the patient¿s subsequent decline. the patient¿s death followed a clinical decision to withdraw care and was determined by the treating team to be unrelated to the impella devices. the death will be conservatively reported, however it is more likely to the underlying critical condition of the patient. |  |  |  |  | Section b1, b5, d9, and h6 were updated based on additional information. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
721 | 1220648-2026-02309 | 24267602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24267602&pc=OZD | 01/29/2026 |  |  | 02/05/2026 |  |  | 01/29/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2027816370 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male | 63 KG | Non Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Cardiac Tamponade | Cardiac Perforation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp device was inserted into the right subclavian artery in a 78-year-old male patient with a past medical history of coronary artery disease (cad), presenting in scai stage a shock, prior to initiation of support. the impella was inserted for ventricular support for a high-risk percutaneous coronary intervention (pci). during the pci, after a rotablation of the right circumflex (rcx) was performed, a perforation was noted, along with a pericardial tamponade. cardiopulmonary resuscitation (cpr) was performed for over 60 minutes. the patient expired on support in the procedure room. the impella cp will conservatively be reported for death; however, the physician stated that there was no causal relationship between the device and the patient's death. the impella cp functioned as intended with no issues. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
722 | 1220648-2026-02316 | 24267783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24267783&pc=OZD | 01/29/2026 |  |  | 02/05/2026 |  |  | 01/29/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026726684 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient is a 70 year old male with cardiomyopathy whose pre support condition was consistent with scai stage d. an impella 5.5 (sn (b)(6)) was implanted via the right axillary/subclavian artery using a surgical approach on (b)(6) 2026 at 08:50 am. shortly after implantation, a purge cassette leak occurred. the defective purge cassette triggered an alarm, and the cassette was replaced, which fully resolved the purge issue. the physician reported hypoperfusion of the right arm. after evaluation with vascular surgery, the team determined that the ischemia was not impella induced, but rather the result of insufficient vessel diameter in combination with catecholamine use. the reported purge cassette leak was resolved by replacing the cassette, and no additional device-related malfunction was identified. the right arm hypoperfusion was clinically attributed to patient specific vascular limitations and catecholamine effects rather than to impella performance. explanting the device successfully resolved the ischemia, and the patient survived the procedure. ischemia and hemolysis are consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
723 | 1220648-2026-02318 | 24267836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24267836&pc=OZD | 11/07/2025 |  |  | 02/05/2026 |  |  | 11/10/2025 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026739125 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
724 | 1220648-2026-02323 | 24267896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24267896&pc=OZD | 01/29/2026 |  |  | 02/05/2026 |  |  | 01/29/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1481065 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
725 | 1220648-2026-02327 | 24267986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24267986&pc=OZD | 01/27/2026 |  |  | 02/05/2026 |  |  | 01/27/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1127438 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Battery Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Clinical rationale: an impella cp device was inserted into the right femoral artery in a 69-year-old male patient. the impella was inserted for ventricular support during a high-risk percutaneous coronary intervention (pci). during the procedure, the automated impella controller (aic) console had a "battery failure" alarm, followed by an unexpected controller shut down that resolved three seconds later. the physician unplugged the impella catheter and plugged the catheter into a new aic. a "purge cassette failure" alarm sounded. to troubleshoot, the cassette was replaced, which resolved the issue. the impella was weaned an hour later, and hemostasis was achieved with a perclose device. the post-procedure outcome is survived at explant. the automated impella controller will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | D6a and d6b should have been left blank on the initial manufacturer device report. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
726 | 1220648-2026-02329 | 24268030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24268030&pc=OZD | 01/28/2026 |  |  | 02/05/2026 |  |  | 01/29/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794637 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
727 | 1220648-2026-02337 | 24268307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24268307&pc=OZD | 01/24/2026 | 02/05/2026 |  | 02/05/2026 |  |  | 01/24/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026775963 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 58 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted into a 68-year-old female patient via right axillary¿subclavian access for hemodynamic support. the patient initially presented with cardiomyopathy and acute decompensated chronic ischemic heart disease with known coronary artery disease. the patient was receiving inotropic support, vasopressors, and respiratory ventilator support. during the course of therapy, nursing staff reported a hematoma at the insertion site. the managing physician was notified and stated that computed tomography imaging did not demonstrate active bleeding and suspected the patient¿s underlying coagulopathy as a contributing factor. the patient experienced a hematoma and required blood transfusion. hematoma and bleeding events are known to occur in the setting of large-bore vascular access and are described in the instructions for use. additional information received stated that the patient subsequently expired. the death is being reported conservatively. comprehensive review suggests the hematoma was more likely attributed to patient-specific factors, including underlying coagulopathy, comorbid conditions, and the use of large-bore vascular access. no evidence was identified to suggest a causal relationship between the impella 5.5 device and the reported events. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
728 | 1220648-2026-02342 | 24268499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24268499&pc=OZD | 01/25/2026 | 02/05/2026 |  | 02/05/2026 |  |  | 01/26/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2026716760 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female | 105 KG |  |  | Death | Yes | No | Malposition of Device | [6 empty] | Hemodynamic instability | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp was placed via the femoral artery to support the 61 year old female patient who had been admitted in with acute myocardial infarction and cardiogenic shock. no other underlying cardiac or systemic comorbidities are known. the pump was supporting when the pump came out of position while in the icu unit. the team was unable to reposition the pump back across the valve, and the patient had hemodynamic status deteriorate with high medication and poor laboratory results. the team would not escalate to extracorporeal life support (ecls) or place another cp and the patient expired. the death is being reported as a contributing factor to the death, and patient's presenting native critical condition as presented in scai stage d and decision made not to replace or escalate care. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
729 | 1220648-2026-02346 | 24269190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24269190&pc=OZD | 01/27/2026 |  |  | 02/05/2026 |  |  | 01/27/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1991105 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
730 | 1220648-2026-02347 | 24269336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24269336&pc=OZD | 01/30/2026 |  |  | 02/05/2026 |  |  | 01/30/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025479706 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
731 | 1220648-2026-02350 | 24269452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24269452&pc=OZD | 05/13/2025 |  |  | 02/05/2026 |  |  | 01/27/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026612639 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Study | Health Professional | User facility | Company representation |  | 68 YR |  | Male | 70 KG |  | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: an impella 5.5 was placed via the axillary graft site to support the 68-year-old male who was admitted in with known coronary artery disease and plan for a bypass (cabg) surgery, presenting in scai stage d shock. this took place in (b)(6)/2025 and the pump supported for the patient's needs and was weaned and explanted after (b)(6) weeks. later in 1/2026 the impact eu study documented there were two harms attributed to the pump support. there was melena/gastrointestinal bleeding and a nosebleed with relationship noted as definite to the impella. there additionally was an elevated plasma free hemoglobin, indicative of hemolysis, with relationship to the impella listed as probable. of note, during the support the team had adjusted p levels as standard troubleshooting for hemolysis. hemolysis is a known risk associated with impella support as described in the instructions for use and may be influenced by patient-specific factors such as underlying cardiac dysfunction and hemodynamic conditions. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
732 | 1220648-2026-02352 | 24269875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24269875&pc=OZD | 01/29/2026 |  |  | 02/05/2026 |  |  | 01/29/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798077 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
733 | 1220648-2026-02211 | 24270624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24270624&pc=OZD | 11/19/2025 |  |  | 02/05/2026 |  |  | 11/19/2025 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1090628 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | No Audible Alarm | Application Program Problem | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical assessment by (b)(6): a 71-year-old male was admitted for acute myocardial infarction and cardiogenic shock. before percutaneous coronary intervention, an impella cp was placed. the patient suffered a cardiac arrest and was escalated to a veno-arterial extracorporeal membrane oxygenation before the coronary procedure was carried out. following several error messages and loss of placement signals, the automated impella controller was switched to a new controller. the new controller showed the impella as defective. impella was removed and switched out for new and second impella cp. the patient was supported for an off-label prolonged duration of six days. during this time, he was treated for a growing pulmonary edema, also requiring continous renal replacement therapy for both volume management and temporary renal failure. due to the poor prognosis, care was withdrawn and the patient expired. death was conservatively coded to the first impella cp while most likely to be caused by the underlying disease and unrelated to impella as the replacement device functioned as expected in this critically ill patient. also, pulmonary edema and renal failure are more probable to be related to the underlying disease. | [5 empty] | A4 weight is unknown. device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this automated impella controller device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device reports on the impella cp pump 1 and impella cp pump 2.
734 | 1220648-2026-02368 | 24270802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24270802&pc=OZD | 01/30/2026 |  |  | 02/05/2026 |  |  | 01/30/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026756749 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
735 | 1220648-2026-02370 | 24271259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24271259&pc=OZD | 01/31/2026 | 02/25/2026 |  | 02/05/2026 |  |  | 01/31/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737066 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 140 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | False Alarm | No Apparent Adverse Event | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported multiple suction alarms, with a possible sensor malfunction suspected. patient outcome has not yet been determined, as the patient remains on device support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
736 | 1220648-2026-02378 | 24271963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24271963&pc=OZD | 01/28/2026 |  |  | 02/05/2026 |  |  | 01/28/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025536168 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was implanted with an impella 5.5 via percutaneous right axillary/subclavian artery for mechanical circulatory support. a purge fluid leak was reported. it was found that the purge fluid was leaking from the housing that covers the pressure reservoir and purge filter during use. no purge-related alarms were noted. there was no change in purge pressure/purge fluid. the purge set was not replaced by this date. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
737 | 1220648-2026-02374 | 24271984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24271984&pc=OZD | 01/31/2026 |  |  | 02/05/2026 |  |  | 01/31/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1082388 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 68 KG | Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a nurse observed two brief controller error alarms that self-resolved. at the time of reporting, the alarms had resolved and all waveforms were within normal limits. intermittent loss of aortic (ao) and left ventricular (lv) placement signals was noted. motor current (mc) and infusion flow (if) were appropriate for the programmed p-level. patient outcome has not yet been determined, as the patient remains on device support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
738 | 1220648-2026-02376 | 24272038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24272038&pc=OZD | 01/31/2026 |  |  | 02/05/2026 |  |  | 01/31/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801031 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
739 | 1220648-2026-02381 | 24272368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24272368&pc=OZD | 02/02/2026 |  |  | 02/05/2026 |  |  | 02/02/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781192 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 107 KG | Non Hispanic | Black Or African American | Injury | No | Yes | Difficult or Delayed Positioning | Deformation Due to Compressive Stress | Device Sensing Problem | Difficult to Advance | Adverse Event Without Identified Device or Use Problem |  |  | Hematoma | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during a routine echocardiogram to assess left ventricular (lv) function, the impella catheter depth was noted to be 2.1 cm. an attempt by the heart failure team to advance the catheter was unsuccessful. the physician reported difficulty advancing the catheter through the locking mechanism using the blue button and described resistance at the hub. the patient was taken to the operating room, where the axillary pocket was opened and the graft position was revised. the catheter was clamped at the graft anastomosis using a soft clamp, and the suture hub and graft end were revised. a kink was observed in the catheter shaft near the repositioning unit, which was noted as a potential contributing factor to the limited advancement. following revision, the catheter was successfully advanced to an appropriate position in the lv at a depth of 5 cm. the patient tolerated the procedure without reported complications. the procedure was performed due to concern for potential catheter migration, as the patient is ambulatory and awaiting transplant. the physician plans to replace the pump if further issues occur. patient outcome has not yet been determined, as the patient remains on device support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
740 | 1220648-2026-02386 | 24272965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24272965&pc=OZD | 01/25/2026 | 02/05/2026 |  | 02/05/2026 |  |  | 01/26/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. |  |  | US |  | 1870587 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Based on the limited available information, the complaint involves a purge system priming issue with no patient harm reported. a device evaluation may be necessary to determine the root cause. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
741 | 1220648-2026-02390 | 24273184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24273184&pc=OZD | 01/29/2026 | 02/05/2026 |  | 02/05/2026 |  |  | 01/29/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026739585 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 37 YR |  | Male | 79 KG | Non Hispanic | White | Injury | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 37-year-old male with cardiomyopathy presented in pre support scai stage e. the patient underwent a planned right axillary/subclavian surgical insertion of an impella 5.5 (sn(b)(6) on (b)(6) 2026. during the procedure, the team encountered significant resistance due to a small subclavian artery and pronounced tortuosity at the subclavian¿innominate turn. the pump was initially introduced through a 23f sheath over a 0.018" wire but was unable to navigate the turn. multiple subsequent attempts were made using three different guidewires, a 16f dilator to assess passage, and finally a 0.025" wire; however, the device continued to fail to advance through the tortuous segment. excessive force was applied in attempts to overcome resistance, but the pump could not be advanced past the axillary artery. no product damage was noted, and no device return is expected. due to the patient¿s vascular anatomy¿including vessel diameter <7mm and tortuosity¿the procedure was aborted. the device was explanted at 4:00 pm the same day. the patient remained stable, with no device involvement in the outcome. the complaint outcome was classified as unsuccessful due to failure to advance. | [5 empty] | The impella device has not been received from the customer; therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
742 | 1220648-2026-02395 | 24273339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24273339&pc=OZD | 01/28/2026 | 02/05/2026 |  | 02/05/2026 |  |  | 01/28/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2025575156 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Restricted Flow rate | Malposition of Device | [5 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A case study reported that an impella cp was inserted into the right femoral artery of a 58-year-old male who presented with exertional fatigue two weeks earlier and was diagnosed with ischemic cardiomyopathy associated with three-vessel disease. he had persistent low cardiac function (left ventricular ejection fraction 12%) and ventricular tachycardia, and was referred to our hospital for revascularization. the impella was inserted for ventricular support for an off-pump coronary artery bypass graft (opcab). during the surgery, he experienced episodes of hypotension. transesophageal echocardiogram (tee) revealed that the impella inlet was located at the left ventricular apex, and the outlet was visualized entering and exiting the left ventricular cavity just above the aortic valve, suggesting a poorly positioned impella. frequent positioning adjustments were required to maintain flow. because the impella was already fixed to the thigh, adjusting its position was difficult, and circulatory management was necessary without relying on the impella. after anastomosis of the left circumflex artery and release of the device from rotation, hemodynamics remained stable, and the procedure was completed. the post-procedure outcome was survived at explant. correct positioning of the impella cp is directly related to hemodynamics; improper placement commonly causes hypotension, low flow, or suction alarms. when the device moves too far into the left ventricle or too far out into the aorta, it cannot effectively move blood, leading to ineffective support and low blood pressure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
743 | 1220648-2026-02398 | 24273418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24273418&pc=OZD | 01/23/2026 | 02/05/2026 |  | 02/05/2026 |  |  | 01/23/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026701740 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
744 | 1220648-2026-02400 | 24273450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24273450&pc=OZD | 01/24/2026 | 02/05/2026 |  | 02/05/2026 |  |  | 01/24/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1263045 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a 65 year old male on an impella 5.5 due to cardiomyopathy. during support, the aic (console) battery immediately dropped to 50% when unplugged to ambulate the patient. there are no other active alarms and the battery returns to 100% when plugged back in. there was no patient harm and the patient remains on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
745 | 1220648-2026-02403 | 24273552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24273552&pc=OZD | 01/30/2026 |  |  | 02/05/2026 |  |  | 01/30/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798113 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
746 | 1220648-2026-02406 | 24273640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24273640&pc=OZD | 02/03/2026 |  |  | 02/05/2026 |  |  | 01/30/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026806125 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
747 | 1220648-2026-02407 | 24273686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24273686&pc=OZD | 01/30/2026 |  |  | 02/05/2026 |  |  | 01/30/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781062 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a sixty-six-year-old male patient with acute myocardial infarction and cardiogenic shock was supported with an impella 5.5 device placed through the axillary-subclavian approach. during support, the patient developed a significant bleeding event at the insertion site. nursing staff were unable to determine whether the bleeding originated from the sutures or the surrounding wound vacuum dressing, which had been placed prior to transfer. the wound care team planned to exchange the dressing to assist with management. the patient was receiving heparin therapy and dual antiplatelet medication at therapeutic levels. the patient also experienced arrhythmias, including bigeminal premature ventricular contractions and atrial flutter. electrolyte replacement was administered, and an echocardiogram confirmed appropriate impella positioning. the arrhythmias resolved following electrolyte correction. an additional medical device was required to assist with management of the access site. the reported events¿major bleeding, arrhythmia, and the need for adjunct device support¿were considered consistent with the patient¿s underlying condition, cardiogenic shock, anticoagulation therapy, and the known risks associated with large-bore arterial access. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b: updated explant date.
748 | 1220648-2026-02402 | 24273764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24273764&pc=OZD | 01/28/2026 | 02/05/2026 |  | 02/05/2026 |  |  | 01/28/2026 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1324936 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 60 KG | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a 60 year old male on impella support for an elective procedure. during support, a ¿controller error¿ occurred. as a result, the controller was swapped. only one back up controller connected waveform came back. the patient remains on support. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
749 | 1220648-2026-02409 | 24273801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24273801&pc=OZD | 12/15/2025 | 02/25/2026 |  | 02/05/2026 |  |  | 12/15/2025 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775983 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
750 | 1220648-2026-02410 | 24273879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24273879&pc=OZD | 10/21/2025 |  |  | 02/05/2026 |  |  | 10/21/2025 | 02/05/2026 |  |  |  |  | 02/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026718433 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella 5.5 support, the placement signal was lost while the fixation pad was being secured. device position was reassessed with imaging and corrected, which restored the placement signal with stable flow throughout. later, intermittent loss of the placement signal recurred along with a yellow controller alarm, although device flow, motor current, and position remained stable. the clinical team switched to a different console and completed an automated controller transfer, after which the alarms resolved and no further issues occurred. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
751 | 1220648-2026-02414 | 24274107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24274107&pc=OZD | 01/30/2026 |  |  | 02/05/2026 |  |  | 01/30/2026 | 02/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026678491 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
752 | 1220648-2026-02200 | 24256821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24256821&pc=OZD | 01/26/2026 |  |  | 02/04/2026 |  |  | 01/26/2026 | 02/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794579 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
753 | 1220648-2026-02203 | 24256866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24256866&pc=OZD | 01/08/2026 |  |  | 02/04/2026 |  |  | 01/08/2026 | 02/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795704 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
754 | 1220648-2026-02207 | 24257043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24257043&pc=OZD | 01/26/2026 |  |  | 02/04/2026 |  |  | 01/26/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719710 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a purge system blocked alarm. no kink or blockage of the purge system could be seen. the purge cassette was replaced but the alarm did not resolve. the patient was weaned from the pump. | [5 empty] | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
755 | 1220648-2026-02208 | 24257066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24257066&pc=OZD | 01/22/2026 |  |  | 02/04/2026 |  |  | 01/22/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794397 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female | 61 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | Hemodynamic instability | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was placed to support the 70year old female via the axillary artery graft access site. the patient had been admitted in with known acute myocardial infarction and cardiogenic shock. any underlying cardiac or systemic medical history and comorbidities are unknown, though the patient was intubated and on a vent for respiratory needs. the patient was critical with norepinephrine and "blood bags" infusing. the patient had signs of anemia. the medical team made decision to withdraw care. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's presenting native critical condition. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted.
756 | 1220648-2026-02209 | 24257094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24257094&pc=OZD | 01/27/2026 |  |  | 02/04/2026 |  |  | 01/27/2026 | 02/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026803192 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Cardiac Arrest | Vascular Dissection | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
757 | 1220648-2026-02212 | 24257389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24257389&pc=OZD | 01/21/2026 | 02/05/2026 |  | 02/04/2026 |  |  | 01/21/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026763779 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 was inserted via the right axillary subclavian artery in a 65-year-old female with postcardiotomy cardiogenic shock and low cardiac output syndrome following coronary artery bypass graft surgery. the patient had a history of prior cabg and known coronary artery disease and was reported as scai shock stage e. the patient was receiving inotropes, vasopressors, and ventilator support. the patient was noted to have hemolysis, and imaging identified a structure near the inflow region that was determined by the provider to be native cardiac tissue rather than thrombus. device repositioning was performed, followed by weaning and removal of the impella 5.5. no additional mechanical circulatory support was deployed. the reported patient experiences included hemolysis, device repositioning, and medical device removal. based on the available information, the reported events appear consistent with the patient¿s underlying disease process and the known risks associated with advanced mechanical circulatory support. the patient survived. | An impella 5.5 was inserted via the right axillary subclavian artery in a 65-year-old female with postcardiotomy cardiogenic shock and low cardiac output syndrome following coronary artery bypass graft surgery. the patient had a history of prior coronary artery bypass graft surgery and known coronary artery disease and was reported as scai shock stage e. the patient was receiving inotropes, vasopressors, and ventilator support. the patient was noted to have laboratory findings concerning for hemolysis. imaging identified a structure near the inflow region that was determined by the provider to be native cardiac tissue rather than thrombus. based on the available information, the hemolysis may be attributable to the patient¿s underlying critical illness and hemodynamic instability, which are known risk factors in the setting of advanced mechanical circulatory support. device repositioning was performed, followed by weaning and removal of the impella 5.5. no additional mechanical circulatory support was deployed. the reported patient experiences included hemolysis, device repositioning, and medical device removal. additional information received states that, due to the patient¿s deteriorating clinical status, the family withdrew care and the patient expired. the death is being reported conservatively. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 and h6 updated. the investigation is still ongoing.
758 | 1220648-2026-02215 | 24257461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24257461&pc=OZD | 01/22/2026 |  |  | 02/04/2026 |  |  | 01/23/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, CANADA |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026709200 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male | 71 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted surgically into the left axillary artery in a 57-year-old male patient with a past medical history of coronary artery disease (cad), presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), scai stage c shock, on multiple inotropes and vasopressors, on a ventilator for respiratory support prior to initiation of support. the impella was inserted for ventricular support during a percutaneous coronary intervention (pci) of the left anterior descending artery (lad), and the obtuse marginal artery (om). it was reported that during preparation, the physician attempted to deploy a perclose device, but experienced difficulty withdrawing, which resulted in a laceration of the artery. the injury led to prolonged bleeding which was managed and resolved by a cardiothoracic/vascular surgeon by suturing a hemashield graft. the impella cp was subsequently implanted. the pci was performed the following day, and the impella was successfully explanted a few hours later. while on support the device functioned at p-7 at 3.1l/min. the post-procedure outcome is survived at explant. using excessive force during perclose deployment or applying too much tension on the knot can cause the suture to tear the artery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
759 | 1220648-2026-02218 | 24257578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24257578&pc=OZD | 01/21/2026 | 02/16/2026 |  | 02/04/2026 |  |  | 01/21/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796338 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 86 YR |  | Female | 80 KG | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Bradycardia | Cardiac Arrest | Low Blood Pressure/ Hypotension | Tachycardia | Hemodynamic instability | [8 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An 86-year-old female with a medical history significant for coronary artery disease underwent implantation of an impella 5.5 device for temporary mechanical circulatory support. the indication for use was postcardiotomy cardiogenic shock (pccs). the device was surgically implanted via the right axillary artery. at the time of impella initiation, the patient was classified as society for cardiovascular angiography and interventions (scai) stage e cardiogenic shock. on post-implant day 0, while being prepared for transfer from the cardiovascular operating room to the intensive care unit, the patient was noted to have a heart rate in the 150¿170 beats per minute range. the physician ordered an amiodarone bolus, which was administered intravenously by anesthesia. following administration, the patient¿s heart rate decreased to approximately 50¿70 beats per minute. subsequently, the impella device generated suction alarms. transesophageal echocardiography confirmed appropriate device positioning. the patient then developed hypotension with mean arterial pressures noted in the 30¿40 mmhg range. cardiopulmonary resuscitation was initiated, and intravenous epinephrine and calcium were administered. return of spontaneous circulation was achieved, and the patient¿s mean arterial pressure improved to the 90 mmhg range. the patient was transferred to the intensive care unit. while in the intensive care unit, the patient experienced recurrent hypotension (blood pressure began slowing dropping). an epinephrine infusion was increased without sustained improvement. the impella device again generated suction alarms. the patient subsequently experienced cardiac arrest. cardiopulmonary resuscitation was initiated, and the chest was opened at the bedside. the patient was placed on veno-arterial extracorporeal membrane oxygenation (va-ecmo). the impella support level was reduced to p-4, after which the patient¿s hemodynamic status stabilized. as of post-implant day 13, the patient remained on mechanical circulatory support. the suction alarms and hemodynamic instability occurred in the setting of cardiogenic shock (scai stage e) and acute changes in heart rate, preload, and blood pressure. administration of intravenous antiarrhythmic therapy was followed by significant bradycardia and hypotension. suction events may occur when ventricular preload is inadequate. the patient is elderly and critically ill with heart failure and shock, requiring increasing levels of life-support, including va-ecmo. | An impella 5.5 device was inserted into the right subclavian artery in a 86-year-old female patient with a past medical history of coronary artery disease (cad). the patient presented in an out-of-hospital cardiac arrest with cardiopulmonary resuscitation (cpr), and in scai stage e shock, prior to initiation of support. the impella was inserted for ventricular support during a high-risk surgery: coronary artery bypass graft (cabg). the patient was being prepared to transfer to the cardiovascular intensive care unit (cvicu) from the cardiovascular operating room (cvor). the patient's heart rate (hr) was 150-170's. the physician ordered an amiodorone bolus. the patient's hr dropped to 50-70. the impella began experiencing suction alarms. a transesophageal echocardiogram (tee) showed no issues with position. the blood pressure (bp) then dropped to a mean arterial pressure (map) of 30-40. cardiopulmonary resuscitation (cpr) was administered. epinephrine and calcium iv were given. return of spontaneous circulation (rosc) achieved. bp increased to 90's map. the patient was transferred to cvicu. while in the cvicu, the bp slowly began dropping. an epinephrine drip was increased without success. the impella began having suction alarms. the physician was at bedside and the patient coded. cpr started while the chest was opened, and the patient was placed on venous-arterial (va) extracorporeal membrane oxygenation (ecmo). impella decreased to p-4. the patient began to stabilize. two weeks after the impella was implanted, the family elected to withdraw care, and the patient expired on support. the device is unlikely to have contributed to the patient's death, which was most likely due to clinical condition of a critically ill patient, along with the complications of stage e shock. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | B2, b5, d4, d6b and h6 updated. the investigation is still ongoing.
760 | 1220648-2026-02219 | 24257601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24257601&pc=OZD | 01/27/2026 |  |  | 02/04/2026 |  |  | 01/27/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026804168 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female | 86 KG | Hispanic |  | Injury | Yes | No | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 device was inserted into the right axillary artery in a 54-year-old female patient presenting in cardiomyopathy, and in scai stage c shock, prior to initiation of support. while the patient was in the intensive care unit (icu), the patient was found to be unresponsive in the chair, in and out of ventricular tachycardia (vt). the patient was transferred to the bed, and amiodarone bolus was given. a transthoracic echocardiogram (tte) showed the tip of the catheter adjacent to the mitral valve (mv). the physician torqued the catheter towards the left ventricle (lv) apex. an echocardiogram showed the pump at 4.9cm across the valve in good position. the patient was in normal sinus rhythm (nsr) and supported at p-6 at 3.2l/min with no issues. the post-procedure outcome is unknown. improper positioning of an impella 5.5 device can directly lead to arrhythmias. mispositioning, such as when the device is too deep in the left ventricle or too high in the aorta, can cause mechanical irritation of the cardiac structures, inducing conduction changes. however, the patient's underlying condition of cardiomyopathy can also cause arrythmias, because when the heart becomes enlarged, thick, or rigid, the electrical system is damaged, causing the heart to beat too fast, too slow, or irregularly. | [5 empty] | D6b explant date provided. | A6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
761 | 1220648-2026-02229 | 24257905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24257905&pc=OZD | 01/27/2026 |  |  | 02/04/2026 |  |  | 01/27/2026 | 02/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1299286 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Device Sensing Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
762 | 1220648-2026-02232 | 24258000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24258000&pc=OZD | 01/27/2026 |  |  | 02/04/2026 |  |  | 01/28/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802396 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 63 KG | Non Hispanic | White | Injury | Yes | No | Increase in Pressure | [6 empty] | Cardiac Tamponade | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 pump was placed via the direct surgical access to support the 71-year-old male who had been admitted in for elective valve surgery. the patient presented in scai stage e shock. any other cardiac or systemic comorbidities are unknown. the pump was supporting when the console alarmed for high purge pressure and purge blockage. the team believed the alarms to be due to underlying patient condition and tamponade/compressed heart. no intervention or troubleshooting measures were taken. the pump remained on for support til the next day and was explanted. upon explant the team observed thrombosis had occurred, as there was thrombus on the pump's outflow window and sensor. there was no noted harm and with explant and new 2nd 5.5 pump was placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
763 | 1220648-2026-02233 | 24258038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24258038&pc=OZD | 01/27/2026 |  |  | 02/04/2026 |  |  | 01/28/2026 | 02/04/2026 |  |  | 05/31/2025 |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2024350342 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Anemia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A publication reported that on 10/2023 an impella cp device was inserted into the femoral artery of a 35-year-old male patient presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), prior to initiation of support. additionally, the publication noted there was myocarditis. . the impella was inserted for ventricular support for a high-risk percutaneous coronary intervention (pci). the abiomed team in japan on 1/2026 located the publication of this patient support. while the patient was in the intensive care unit (icu), hemolysis was noted, and the renal function gradually deteriorated. the physician decided an escalation of therapy was necessary and an impella 5.5 was placed via right axillary artery. upon removal of the impella cp, access site bleeding was noted. anemia gradually developed and the wound swelled. the post-procedure outcome is survived at explant per the publication. bleeding is a known risk due to the impella's anticoagulation and purge requirements. hemolysis and bleeding are listed as potential adverse events, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
764 | 1220648-2026-02235 | 24258057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24258057&pc=OZD | 01/27/2026 |  |  | 02/04/2026 |  |  | 01/28/2026 | 02/04/2026 |  |  | 08/31/2024 |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2023259572 |  | 000000000001000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Mitral Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | On (b)(6) 2023 and impella 5.5 was placed as care escalation from a prior placed impella cp for the 73-year-old male patient. he had been admitted in with drug resistant cardiogenic shock, coronary artery disease, and with impaired consciousness and shock. the pump supported the patient for 23 days and was then explanted as hemodynamically the patient had recovered. on (b)(6) 2026, the abiomed team in japan located a publication from the medical team outlining this support. there is an allegation of renal failure the prompted dialysis and mitral valve regurgitation (mr). the patient was medically managed for the mr and had a mitraclip performed. the patient did survive the events and harms. cardiac injury is a known risk associated with impella support as described in the instructions for use. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
765 | 1220648-2026-02236 | 24258098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24258098&pc=OZD | 01/27/2026 |  |  | 02/04/2026 |  |  | 01/27/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026727667 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | No | Malposition of Device | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp was placed via the femoral artery to support the 56 year old male patient who had been admitted in with acute myocardial infarction and cardiogenic shock. no other cardiac or systemic comorbidities were shared. the pump was supporting when after 9 days the team made decision to explant and replaced the pump with escalated care, of the impella 5.5 pump. the pump had been malpositioned at the mitral valve and when explanted the cp had thrombus adhered to the pump. the material was sent to pathology for review, but to date the report is not known. the 5.5 was successfully placed and supports the patient to date. | [5 empty] | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
766 | 1220648-2026-02237 | 24258108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24258108&pc=OZD | 01/27/2026 |  |  | 02/04/2026 |  |  | 01/27/2026 | 02/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026621535 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
767 | 1220648-2026-02234 | 24258111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24258111&pc=OZD | 11/17/2025 | 02/17/2026 |  | 02/04/2026 |  |  | 11/17/2025 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744131 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 86 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 62 year old male with a history of myocardial infarction presented with chest pain and suffered an episode of ventricular fibrillation in the emergency department, requiring resuscitation. patient declined surgery, but had to be readmitted to the intensive care unit due to myocardial infarction and cardiogenic shock. percutaneous coronary intervention was performed and after that, an impella cp was placed. during therapy, "placement signal low" and subsequent "impella in ventricle alarm" occurred that were most probably associated with poor ventricular function and lacking pulsatility. subsequently, the patient was escalated to impella 5.5 on the same day. during support with the 5.5, the patient remained highly unstable and in critical condition, suffering episodes of ventricular tachycardia. due to the poor prognosis, care was withdrawn and the patient expired after three days of support. death was most likely contributed to the patients underlying critical condition. | [5 empty] | Investigation summary: ppae (arrhythmia/tachycardia): the root cause of the injury was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
768 | 1220648-2026-02246 | 24259042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24259042&pc=OZD | 10/21/2025 |  |  | 02/04/2026 |  |  | 10/21/2025 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1390619 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | An impella 5.5 was inserted via the right axillary subclavian artery in a 74-year-old male who presented for high-risk percutaneous coronary intervention. the patient had known coronary artery disease and renal insufficiency and was supported with inotropes, vasopressors, and mechanical ventilation for respiratory support. the patient was classified as scai stage c. during support, a controller-related event was reported, characterized by a controller error and intermittent placement signal issue. the event was associated with hemodynamic instability, prompting clinical evaluation and management. device revision or replacement was performed at the system level to address the reported controller-related issue and restore stable support. following intervention, device function and hemodynamic parameters stabilized. review of the event indicates that the controller-related error and placement signal issue were managed clinically and are consistent with recognized management scenarios during mechanical circulatory support. comprehensive review did not identify evidence suggesting a causal relationship between the impella controller and the reported event. the patient survived. | [5 empty] | D4: the device serial number was not provided, udi is unavailable. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed.
769 | 1220648-2026-02248 | 24259327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24259327&pc=OZD | 01/22/2026 |  |  | 02/04/2026 |  |  | 01/22/2026 | 02/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026742655 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | [6 empty] | Stroke/CVA | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
770 | 1220648-2026-02250 | 24259891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24259891&pc=OZD | 01/17/2026 |  |  | 02/04/2026 |  |  | 01/17/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1446476 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 197 KG | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support it was reported that the placement left ventricle signal was not there. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
771 | 1220648-2026-02252 | 24260369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24260369&pc=OZD | 01/29/2026 |  |  | 02/04/2026 |  |  | 01/29/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1746366 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during annual preventive maintenance after initial startup, the console displayed a red battery failure alarm. no patient was involved. | [5 empty] | The controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
772 | 1220648-2026-02256 | 24261072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24261072&pc=OZD | 01/22/2026 |  |  | 02/04/2026 |  |  | 01/22/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798481 |  | 1000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient with an unknown demographic background was supported with an impella 5.5 for the indication of cardiomyopathy. the first impella 5.5 device (sn (b)(6)) was surgically implanted via the right axillary/subclavian artery and explanted 15 minutes later due to pump stop. the device was replaced with 5.5 (sn (b)(6)), which was implanted at. the second device remained in use until at which time care was withdrawn and the patient subsequently expired. based on the available information the death will be coded to the initial impella 5.5 but is unlikely a contributing factor to the patients death. | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a patient presenting with cardiomyopathy. shortly after implantation, the pump stopped. the device was removed and replaced with a new impella 5.5. no issues were reported with the current pump following the exchange, and device function has remained stable. the reported events are pump stop, medical device removal, and hemodynamic instability. the impella 5.5 will be conservatively reported for hemodynamic instability due to temporary interruption of hemodynamic support during the pump exchange; however, the exchange was performed with no consequence to the patient, and support was resumed without any known adverse events. |  |  |  |  | After further review case 1220648-2026-02256 has been consolidated with mfr 1220648-2026-02535. mfr 1220648-2026-02535 was found to be a duplicate. the following was updated. b2 death and date of death. b5 clinical rational. h1 type of report: death. h6 health effect - impact code: death. the investigation is still ongoing. the exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. | This is one of two related reports. this report represents the first pump and another report will be submitted to represent the second pump. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
773 | 1220648-2026-02260 | 24261918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24261918&pc=OZD | 01/29/2026 | 02/04/2026 |  | 02/04/2026 |  |  | 01/29/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1729761 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 70 YR |  | Male | 52 KG | Non Hispanic | Asian | Death | Yes | No | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [12 empty] | Death | Required Intervention | [23 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | An impella cp device was inserted via the right femoral artery in a 70-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs, scai stage d). during the hospitalization, the care team initiated a console exchange from a lowell general console to a tufts console. during the exchange process, when connecting the purge disc, the system displayed repeated error messages stating, ¿purge disc not detected.¿ multiple purge cassettes were attempted; however, the error persisted. subsequently, the patient expired. the reported events are purge disc not detected, medical device revision or replacement, hemodynamic instability, and death. the death is being conservatively reported to the impella cp due to a temporary interruption of hemodynamic support during the console exchange. however, the patient¿s death is more plausibly attributed to the severity of the underlying ami/cgs (scai stage d) and critical condition. | [5 empty] | The impella device has not been received from the customer; therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | E4 was inadvertently omitted on the initial manufacturer device report. h5 and h8 was erroneously entered on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
774 | 1220648-2026-02261 | 24261938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24261938&pc=OZD | 01/26/2026 |  |  | 02/04/2026 |  |  | 01/27/2026 | 02/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026784002 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
775 | 1220648-2026-02265 | 24262086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24262086&pc=OZD | 01/28/2026 | 02/04/2026 |  | 02/04/2026 |  |  | 01/29/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026727408 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 61 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported their impella cp pump stopped. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
776 | 1220648-2026-02278 | 24264655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24264655&pc=OZD | 01/19/2026 | 02/04/2026 |  | 02/04/2026 |  |  | 01/20/2026 | 02/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026773801 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
777 | 1220648-2026-02279 | 24264669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24264669&pc=OZD | 01/28/2026 | 02/04/2026 |  | 02/04/2026 |  |  | 01/29/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | ABIOMED, INC. |  |  | US |  | 1745371 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that during annual pm after initial startup, the console displayed a red battery failure alarm. there was no patient involvement. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
778 | 1220648-2026-02280 | 24264718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24264718&pc=OZD | 01/27/2026 | 02/04/2026 |  | 02/04/2026 |  |  | 01/29/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2026793957 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella cp was inserted via the femoral artery in a 61-year-old male who presented for an electrophysiology ablation procedure. the patient¿s clinical status prior to initiation of support was reported as scai shock stage d. past medical history was not reported. the patient underwent coronary intervention involving the left anterior descending, left circumflex, and right coronary arteries. the patient was receiving inotropes, vasopressors, and ventilator support for respiratory support. during the procedure, angiography of the lower limb revealed stenosis, and balloon angioplasty was performed, followed by placement of a 16 fr medikit sheath. an antegrade sheath was subsequently inserted. adjustment of the sheath position in the icu resulted in improvement in limb coloration. due to ongoing clinical concerns, including laboratory findings, a plan was made to upgrade support at an early stage. the patient experienced ischemia. based on the available information, the reported event appears consistent with the patient¿s underlying clinical condition, vascular anatomy, and procedural factors associated with large-bore femoral access and vasopressor use. no issues with pump operation were reported. the patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. based on the available information, the reported event appears consistent with the patient¿s underlying clinical condition, vascular anatomy, and procedural factors associated with large-bore femoral access and vasopressor use.
779 | 1220648-2026-02282 | 24264854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24264854&pc=OZD | 01/29/2026 | 02/04/2026 |  | 02/04/2026 |  |  | 01/29/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. |  |  | US |  | 1630913 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 59 KG |  |  | Injury | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | An impella cp device was inserted via the right femoral artery in an 82-year-old male patient undergoing high-risk percutaneous coronary intervention (hrpci). the patient¿s medical history was significant for coronary artery disease, diabetes mellitus, and renal insufficiency. after the pump was successfully started, waveforms intermittently disappeared from the console display, while flows and motor current remained within acceptable ranges. this intermittent issue occurred multiple times during the procedure. after the second or third occurrence, a ¿controller error¿ message was displayed on the console, and the decision was made to exchange the controller. at the end of the case, the patient was connected to a different console, after which no further issues were observed, and pump function remained stable. the reported events are controller failure, placement signal issue, medical device revision or replacement, and hemodynamic instability. the impella cp is being conservatively reported for hemodynamic instability due to a temporary interruption of support during the controller exchange; however, the exchange was performed without known patient consequence, and support was maintained. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
780 | 1220648-2026-02281 | 24264864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24264864&pc=OZD | 01/29/2026 | 02/10/2026 |  | 02/04/2026 |  |  | 01/29/2026 | 02/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026786861 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
781 | 1220648-2026-02283 | 24264881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24264881&pc=OZD | 01/29/2026 | 02/04/2026 |  | 02/04/2026 |  |  | 01/29/2026 | 02/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026803100 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
782 | 1220648-2026-02284 | 24264912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24264912&pc=OZD | 01/29/2026 | 02/04/2026 |  | 02/04/2026 |  |  | 01/29/2026 | 02/04/2026 |  |  |  |  | 02/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1186526 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 70 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 82-year-old male on an impella for support during a high-risk cardiac procedure. twice during the case, the placement signal disappeared and a yellow alarm stating "controller error" popped up on the controller (aic). it said to switch out controller, however the motor current and flows were normal. both times the aic corrected itself and the placement signal started working again and the "controller error" alarm resolved. the case was finished without any other issues and weaned and removed at the end of the case. | [5 empty] | D6a and d6 b should have been left blank on the initial manufacturer device report. e4 was inadvertently omitted on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
783 | 1220648-2026-02285 | 24264954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24264954&pc=OZD | 01/29/2026 | 02/04/2026 |  | 02/04/2026 |  |  | 01/29/2026 | 02/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026801827 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
784 | 1220648-2026-02286 | 24264966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24264966&pc=OZD | 01/29/2026 | 02/04/2026 |  | 02/04/2026 |  |  | 01/29/2026 | 02/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026803412 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Unspecified Infection | Sepsis | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
785 | 1220648-2026-02287 | 24265022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24265022&pc=OZD | 01/28/2026 | 02/04/2026 |  | 02/04/2026 |  |  | 01/29/2026 | 02/04/2026 |  |  | 03/31/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024319973 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
786 | 1220648-2026-02017 | 24245178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24245178&pc=OZD | 11/20/2025 | 01/31/2026 |  | 02/03/2026 |  |  | 11/20/2025 | 02/03/2026 |  |  | 12/31/2025 |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1390619 |  | 0042-0000-US |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 70 KG |  | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. after rp flex insertion automated impella controller had placement signal issue where the pa signal elevated and flows dropped below 1. swan pa in 30s and not matching. p8 flows less than 1. purge stable and pump positioning confirmed via echo. patient support continued on the same device. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
787 | 1220648-2026-02130 | 24245402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24245402&pc=OZD | 01/22/2026 |  |  | 02/03/2026 |  |  | 01/22/2026 | 02/03/2026 |  |  |  |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1257681 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during a training event an automated impella controller generated an alarm that could not be resolved. a new controller was used for the training. no patient use was reported. | [5 empty] | The reported event did not involve patient use. section a was left blank. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
788 | 1220648-2026-02132 | 24245489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24245489&pc=OZD | 01/24/2026 |  |  | 02/03/2026 |  |  | 01/26/2026 | 02/03/2026 |  |  |  |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026803700 |  | 000000000001000459 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Pumping Stopped | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | An impella 5.5 device was inserted into the right subclavian artery in a 62-year-old male patient presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), from an out-of-hospital cardiac arrest with cardiopulmonary resuscitation (cpr), in scai stage e shock, prior to initiation of support. the impella was inserted for ventricular support during a percutaneous coronary intervention (pci). the impella 5.5 was inserted for escalation of therapy from an impella cp device. while the patient was in the intensive care unit (icu), the pump stopped suddenly and would not restart, so it was removed. an intra-aortic balloon pump (iabp) was inserted; however, the patient expired several hours later. the death could be attributed to the sudden pump start on a patient who was highly dependent on the impella 5.5 for circulation, combined with the clinical condition of a critically ill patient with acute myocardial infarction complicated by cardiogenic shock as they presented in stage e shock. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B2: corrected date of death based on new information. d6b: corrected explant date based on new information. h6: updated codes based on the results of the investigation. h11: updated based on additional information and the results of the investigation. additional information ·the reason for replacing the cp was to change devices in order to strengthen support and with a view to long-term success. ·the patient passed away on the same day after removal. . when the pump stop alarm sounded and the doctor entered the room, he discovered that the tape securing the device had come off and the cord was tangled around the patient's leg. the pump had come loose by several centimeters, and its position was adjusted under ultrasound guidance. the patient was restless, and it is possible that the pump had shifted position due to movement. the timeline of the shift in position due to the patient's movement and the pump stoppage is unknown investigation summary the device was returned for evaluation. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. pump stop: it was reported that the pump suddenly stopped and did not restart. the lt logs confirmed impella stopped alarms and show stable motor current with purge values stable but with periodic pulsatility. suddenly there was alarm #115 impella stopped alarm with no motor current spike but with a drop in placement signal to about 30mmhg. logs showed failed restart attempts with an eventual motor current high pump stop on a restart attempt about 10 minutes later with placement signal going out of range. the returned product was inspected and there were minor kinks along the catheter. the pump stop reproduced in the lab. the pump was electrical tested and there were open lines on all 3 motor cable lines. the pump was evaluated for the source of the open lines and and all 3 cables were found to be stretched/necked as a result of tensile force aligning with clinical detail of patient cable being wrapped around patient's leg and possibility of patient pulling cable. the cause of the pump stop was determined to be due to an electrical open issue likely as a result of unintended use error. device history the pump passed all post sterile inspection checks | Updated clinical rationale: an impella 5.5 device was inserted into the right subclavian artery in a 62-year-old male patient presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), from an out-of-hospital cardiac arrest with cardiopulmonary resuscitation (cpr), in scai stage e shock, prior to initiation of support. the impella was inserted for ventricular support during a percutaneous coronary intervention (pci). the impella 5.5 was inserted for escalation of therapy from an impella cp device, in order to strengthen support and with a view to long-term success. while the patient was in the intensive care unit (icu), the pump stop alarm sounded and the physician entered the room. he discovered that the tape securing the device had come off and the cord was tangled around the patient's leg. the pump had come loose by several centimeters, and its position was adjusted under ultrasound guidance. the patient was restless, and it is possible that the pump had shifted position due to movement. the timeline of the shift in position due to the patient's movement and the pump stoppage is unknown. the impella was removed, and an intra-aortic balloon pump (iabp) was inserted; however, due to a do-not-attempt-resuscitation (dnar), the patient expired several hours later. the death could be attributed to the sudden pump stop on a patient who was highly dependent on the impella 5.5 for circulation, combined with the clinical condition of a critically ill patient with acute myocardial infarction complicated by cardiogenic shock as they presented in stage e shock.
789 | 1220648-2026-02134 | 24245536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24245536&pc=OZD | 01/24/2026 |  |  | 02/03/2026 |  |  | 01/27/2026 | 02/03/2026 |  |  |  |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1121932 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Clinical rationale: an impella cp device was inserted via the right femoral artery in a 64-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). during support, the automated impella controller (aic) required replacement. the reported events are controller failure, device revision or replacement, medical device removal, and hemodynamic instability. although the case is coded as product malfunction, the impella cp is being conservatively reported for serious injury due to a temporary interruption of hemodynamic support during the console exchange; however, the exchange was performed with no consequence to the patient, and hemodynamic support was resumed. | [5 empty] | A. patient information is unknown. device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated.
790 | 1220648-2026-02136 | 24245594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24245594&pc=OZD | 01/15/2026 |  |  | 02/03/2026 |  |  | 01/15/2026 | 02/03/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1338418 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Application Program Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
791 | 1220648-2026-02137 | 24245650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24245650&pc=OZD | 01/15/2026 |  |  | 02/03/2026 |  |  | 01/15/2026 | 02/03/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026671914 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
792 | 1220648-2026-02146 | 24245884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24245884&pc=OZD | 01/25/2026 | 02/18/2026 |  | 02/03/2026 |  |  | 01/26/2026 | 02/03/2026 |  |  |  |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026742004 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 93 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | An impella 5.5 was placed via an axillary graft access to support the 49 year old male admitted in with cardiogenic shock and cardiomyopathy, in scai stage c shock. no other clinical details or medical history were shared. the 5.5 was supporting when on day 10 of the support there was ventricular tachycardia that was treated by a defibrillation and support proceeded with a stable patient result. arrythmia is a known risk associated with impella support as described in the instructions for use and may be influenced by patient-specific factors such as underlying cardiac dysfunction and hemodynamic conditions | [5 empty] | The root cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
793 | 1220648-2026-02148 | 24245944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24245944&pc=OZD | 01/26/2026 |  |  | 02/03/2026 |  |  | 01/26/2026 | 02/03/2026 |  |  |  |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025601462 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male | 70 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted surgically into the aorta in a 30-year-old male patient with a past medical history of coronary artery disease (cad), multi-vessel disease, an ejection fraction of 15%, presenting in scai stage d shock and ventilated for respiratory support prior to initiation of support. the patient presented a few days prior with fatigue, shortness of breath, 100% complete total occlusion (cto) of left anterior descending artery (lad) with some collaterals. worsening symptoms associated with hypotension with diuresis and guideline-directed medical therapy (gdmt) initiation for heart failure. the impella 5.5 was inserted for ventricular support during an elective high-risk coronary artery bypass graft (cabg) surgery. during the procedure, there was some leaking from the anastomosis, along with needle sites of the vein graft. the patient was coagulopathic and received dozens of blood products. the chest was packed, reexplored, and eventually closed. the impella was in an ideal position with excellent flows at p-6 at 4.1l/min. there were periods of low pulsatility that were self-resolving. the post-procedure outcome is survived at explant. bleeding is a known risk due to the impella's anticoagulation and purge requirements. bleeding is listed as a potential adverse event, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | A6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
794 | 1220648-2026-02151 | 24246042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24246042&pc=OZD | 01/27/2026 |  |  | 02/03/2026 |  |  | 01/27/2026 | 02/03/2026 |  |  |  |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026731238 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 62-year-old female patient presenting with postcardiotomy cardiogenic shock / low cardiac output syndrome (pccs/lcos) and a medical history significant for coronary artery disease (cad). surgical access was obtained to the right axillary artery, and a 10-mm vascutek gelweave graft was placed. a peel-away sheath was inserted, and a pigtail catheter with a 0.035 wire was used to cross the aortic valve. the 0.035 wire was removed, an impella 0.018 wire was placed, the pigtail catheter was removed, and the impella 5.5 device was loaded onto the wire. upon advancement of the device into the right axillary artery, extensive resistance was encountered per the cardiothoracic surgeon. after multiple attempts to advance the impella 5.5, the surgeon elected to remove the device and a second device was placed. the surgeon noted that the vessel size was small, though it was initially not anticipated to impede device advancement. the reported events are failure to advance, device revision or replacement, medical device removal, and hemodynamic instability. although the case is coded for product malfunction, the impella 5.5 is being conservatively reported for serious injury due to a temporary interruption of hemodynamic support during the device exchange; however, the exchange was performed with no consequence to the patient, and hemodynamic support was maintained. | [5 empty] | This follow-up report is being submitted to correctly add an h6 code that was omitted in the initial report. h6. medical device problem code a040601 (deformation due to compressive stress) was not included in the initial mdr and is now correctly added. | A4 weight is unknown device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
795 | 1220648-2026-02158 | 24246775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24246775&pc=OZD | 01/21/2026 | 02/03/2026 |  | 02/03/2026 |  |  | 01/21/2026 | 02/03/2026 |  |  |  |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026734167 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 67 YR |  | Male | 73 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 device was inserted via the right axillary artery in a 67-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs), scai shock stage c. the patient had a history of diabetes mellitus. the patient was transferred from an outside facility where the impella 5.5 was placed. upon arrival to the receiving unit early the following morning, a large hematoma was noted at the axillary access site. during patient rounding, the bedside nurse notified the impella representative of the hematoma. the patient was returned to the operating room for hematoma evacuation. the patient required transfusion of two units of packed red blood cells intraoperatively. the evacuation was successful, with no evidence of ongoing bleeding. the surgical site was closed, and the patient was returned to the intensive care unit for recovery. the reported event is an access-site hematoma requiring blood transfusion, surgical intervention, and hemostasis. access-site bleeding and hematoma formation are known risks associated with impella implantation, particularly with surgical axillary access and anticoagulation and purge requirements, as described in the instructions for use. based on the available information, the hematoma was managed with standard surgical intervention, hemostasis was achieved, and no further bleeding was reported. | [5 empty] | D6b explant date provided. e4 was inadvertently omitted on the initial manufacturer device report. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
796 | 1220648-2026-02162 | 24247370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24247370&pc=OZD | 11/18/2025 |  |  | 02/03/2026 |  |  | 11/18/2025 | 02/03/2026 |  |  |  |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026754266 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 62 year old male was treated for acute myocardial infarction with cardiogenic shock. an impella cp was placed and shortly after, urine color was suggestive of hemolysis but shown to be macrohematuria due to a traumatic foley insertion, which was a preexisting condition. the patient remained inotrope and vasopressor dependant and was subsequently escalated to an impella 5.5. on 5.5, hematuria again occurred and required a blood transfusion. a search for other bleeding sources did not uncover other sources of blood loss. after an off-label prolonged 28 days of support on the impella 5.5, the patient was successfully bridged to an left ventricular assist device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
797 | 1220648-2026-02170 | 24248203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24248203&pc=OZD | 01/27/2026 |  |  | 02/03/2026 |  |  | 01/27/2026 | 02/03/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798739 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
798 | 1220648-2026-02171 | 24248320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24248320&pc=OZD | 02/01/2026 |  |  | 02/03/2026 |  |  | 01/27/2026 | 02/03/2026 |  |  |  |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801818 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 103 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinician narrative an impella 5.5 was inserted via the right axillary subclavian artery in a 70-year-old male with cardiogenic shock (cgs), scai stage d, due to acute decompensated chronic heart failure and non-ischemic cardiomyopathy, with a medical history significant for coronary artery disease, diabetes mellitus, renal insufficiency, and pre-existing arrhythmias. during support, the patient experienced recurrent ventricular tachycardia requiring resuscitation and defibrillation. the device was repositioned under echocardiographic guidance, with imaging confirming appropriate position and no further adjustments made. the reported events occurred in the setting of advanced cardiogenic shock, severe underlying structural heart disease, metabolic derangements associated with renal insufficiency, and a documented history of frequent ventricular ectopy and arrhythmias prior to impella implantation, all of which are recognized contributors to malignant ventricular arrhythmias and hemodynamic instability in critically ill patients. comprehensive review supports these patient- and disease-related factors as plausible contributors to the reported event and did not identify evidence suggesting a causal relationship between the impella 5.5 device and the reported event. the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
799 | 1220648-2026-02173 | 24248611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24248611&pc=PYX | 01/27/2026 |  |  | 02/03/2026 |  |  | 01/27/2026 | 02/03/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PUMP, RP-IJ, STERILE |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1958493 |  | 1000186 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
800 | 1220648-2026-02178 | 24250029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24250029&pc=OZD | 01/26/2026 |  |  | 02/03/2026 |  |  | 01/27/2026 | 02/03/2026 |  |  |  |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796396 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinician narrative: an impella cp was in place via the right femoral artery in a 61-year-old male who presented as scai stage d with acute myocardial infarction complicated by cardiogenic shock. the patient was intubated and receiving vasopressors and mechanical ventilation for respiratory support. after transfer to the receiving facility, bleeding was observed at the sheath outlet site. due to visible sheath bleeding, the clinical team decided to replace the pump. the patient experienced a major bleeding event and medical device removal. the use of large-bore femoral arterial access, anticoagulation, critical illness, and hemodynamic instability may have contributed to the reported bleeding event. comprehensive review of the event did not identify evidence suggesting device contribution. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
801 | 1220648-2026-02177 | 24250030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24250030&pc=OZD | 01/27/2026 |  |  | 02/03/2026 |  |  | 01/27/2026 | 02/03/2026 |  |  |  |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1546473 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: during impella support, a controller error alarm was present on the console. the automated impella controller (aic) was exchanged appropriately and successfully in accordance with standard troubleshooting guidance. the reported events are controller failure, device revision or replacement, and hemodynamic instability. although the case is coded as product malfunction, the device is being conservatively reported for serious injury due to a temporary interruption of hemodynamic support during the console exchange; however, the exchange was performed with no consequence to the patient, and hemodynamic support was resumed without any known adverse events. | [5 empty] | A. patient information is unknown. device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
802 | 1220648-2026-02181 | 24250600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24250600&pc=OZD | 01/27/2026 |  |  | 02/03/2026 |  |  | 01/27/2026 | 02/03/2026 |  |  | 08/31/2025 |  | 02/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2024391391 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Anemia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: in 10/2023 an impella 5.5 was placed to escalate care from an impella cp and ecmo that had been previously supporting for under a week. the 5.5 supported via the axillary graft site for the patient with known diagnosis of cardiogenic shock and acute myocardial infarction. the abiomed team in japan in 1/2026 located the publication of this patient support. the publication noted the patient to have suffered harms with a relationship to the pump use/procedure. additionally the publication noted there was myocarditis. the publication speaks to hemolysis, access site bleeding at the axillary site, and anemia. hemolysis had begun on the first impella pump and renal function deteriorated and this caused the coagulation issues to worsen for the patient. the medical team treated the patient with blood product infusion and suture. the hemostasis was achieved by the placement of the viabahn stent after balloon angioplasty and surgical intervention at the graft anastamosis. the patient survived the event per the publication. anticoagulation necessary for the pump placement and support can contribute to access site bleeding, hemolysis is a known risk associated with impella support as described in the instructions for use and may be influenced by patient-specific factors such as underlying cardiac dysfunction and hemodynamic conditions | [5 empty] | A4 weight is unknown device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. a medical device report on the impella cp (pump 1) with serial number 443784 will be filed.
803 | 1220648-2026-02193 | 24255143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24255143&pc=OZD | 01/27/2026 | 02/03/2026 |  | 02/03/2026 |  |  | 01/28/2026 | 02/03/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1213649 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Application Program Problem | Unintended Application Program Shut Down | [5 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
804 | 1220648-2026-02020 | 24230758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24230758&pc=OZD | 11/25/2025 |  |  | 02/02/2026 |  |  | 11/25/2025 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026755472 |  | 1000080 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | 55 year old female patient with acute myocardial infarction / cardiogenic shock was implanted with impella cp. on treatment day 2, physician suspected hemolysis based on unspecified changing lab values. patient recovered, and was successfully weaned and pump explanted. hemolysis is being conservatively coded for impella cp, although hemolysis was not formally diagnosed, and there were no patient consequences. | [5 empty] | Product complaint # (b)(4). investigation summary: no product return. no data logs available. hemolysis/mechanical interaction with blood: the cause of the hemolysis was not determined due to lack of product, data logs, and sufficient clinical information. device history review : the pump sn (b)(6) passed all the post sterile inspection checks.
805 | 1220648-2026-02019 | 24231022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24231022&pc=OZD | 11/24/2025 |  |  | 02/02/2026 |  |  | 11/24/2025 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026739491 |  | 0048-0045 |  | No |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Restricted Flow rate | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
806 | 1220648-2026-02025 | 24231112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24231112&pc=OZD | 01/23/2026 |  |  | 02/02/2026 |  |  | 01/23/2026 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781197 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Other | Health Professional | Company representation |  |  |  | 44 YR |  | Female | 77 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | [8 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 44-year-old female patient presenting with cardiomyopathy and classified as scai shock staged. the device initially demonstrated purge flows of approximately 9¿10 ml/hr. subsequently, purge flow decreased to 2.1 ml/hr with purge pressures rising to 900 mmhg. the icu initiated their in-house tpa protocol, resulting in purge flow improvement to 7 ml/hr with pressures in the 500s. the purge solution was then changed back to d5w with sodium bicarbonate. throughout this period, pump support flows and motor current remained stable. multiple tpa attempts produced minimal success, and following a purge system blocked alarm, the clinical team elected to replace the impella 5.5. an additional contributing factor was the use of d5w with sodium bicarbonate as the original purge solution. the reported events are high purge pressure, medical device removal, and hemodynamic instability. high purge pressure is a known risk identified in the impella instructions for use, associated with obstruction of the purge pathway, or purge solution¿related factors. the impella 5.5 will be conservatively reported for serious injury due to temporary interruption of hemodynamic support during the exchange; however, the exchange was completed without consequence to the patient, and hemodynamic support remained stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
807 | 1220648-2026-02034 | 24231676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24231676&pc=OZD | 11/10/2025 |  |  | 02/02/2026 |  |  | 11/10/2025 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748976 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
808 | 1220648-2026-02043 | 24232673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24232673&pc=OZD | 01/28/2026 |  |  | 02/02/2026 |  |  | 01/28/2026 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795406 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 91 KG | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 via right surgical direct artery for mechanical circulatory support. there was oozing between the locking mechanism and sterile sheath cover connection. no patient harm was reported. the pump is still being used in the patient. | [5 empty] | D6b explant date was added. | The impella device was not received from the customer as it is still being used by the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
809 | 1220648-2026-02053 | 24232887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24232887&pc=OZD | 01/29/2026 |  |  | 02/02/2026 |  |  | 01/29/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1218791 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
810 | 1220648-2026-02047 | 24232889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24232889&pc=OZD | 01/23/2026 |  |  | 02/02/2026 |  |  | 01/23/2026 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026778550 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced an impella defective alarm and the pump stopped. the pump was explanted. the patient was in stable condition. | [5 empty] | Patient information is not available. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
811 | 1220648-2026-02046 | 24232902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24232902&pc=OZD | 01/28/2026 |  |  | 02/02/2026 |  |  | 01/28/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802378 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
812 | 1220648-2026-02050 | 24232978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24232978&pc=OZD | 01/28/2026 |  |  | 02/02/2026 |  |  | 01/28/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795458 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
813 | 1220648-2026-02052 | 24233008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24233008&pc=OZD | 01/16/2026 |  |  | 02/02/2026 |  |  | 01/19/2026 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737002 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female | 70 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | An impella 5.5 was placed to support the 64 year old female admitted in with cardiogenic shock and cardiomyopathy, in scai stage e shock. prior to the 5.5 being placed via direct surgical access the patient was known to be and out of hospital cardiac arrest and on ecmo, inotropes, and vasopressors. the distributor in the eu did not furnish any other comorbidities or medical history. the pump was supporting in the icu when on day 2 the team observed with ct imaging there to be both old and new thrombosis. the patient was not responsive and the team and family discussed options, and the patient had care withdrawn. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's presenting native critical condition as presented in scai stage e. | [5 empty] | Additional information: they did a sedation hold to verify patient was unresponsive. they decided to withdraw care so the flow pump was reduced until the patient expired and explanted after she passed. the pump was discarded so is not available for return. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed and no product was returned. ppae/thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details.
814 | 1220648-2026-02057 | 24233456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24233456&pc=OZD | 01/22/2026 |  |  | 02/02/2026 |  |  | 01/22/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026757857 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
815 | 1220648-2026-02056 | 24233458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24233458&pc=OZD | 01/14/2026 |  |  | 02/02/2026 |  |  | 01/14/2026 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1170094 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Display or Visual Feedback Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | [5 empty] | Patient information was not provided. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
816 | 1220648-2026-02023 | 24233471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24233471&pc=OZD | 01/23/2026 |  |  | 02/02/2026 |  |  | 01/23/2026 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026757375 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | Device Sensing Problem |  |  |  |  | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted via the femoral artery in a patient presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), with scai stage c shock, on multiple inotropes and vasopressors, prior to initiation of support. the impella was inserted for ventricular support during a high-risk percutaneous coronary intervention (pci). it was reported that during the procedure, there were various purge alarms, the purge disc was not recognized, and the system was blocked. reinstalling the purge system did not resolve the issue. the purge system and the automated impella controller (aic) were replaced which resolved the purge issue; however, the ps signal was displayed incorrectly (1/-46), and there was a suction alarm, which was not present previously. the pump's position was confirmed under fluoroscopy, along with the volume status. the physician decided to replace the pump. after replacing the pump, there were no further issues. the pump was running at 3.5 l/min. the pump was explanted post-intervention and hemostasis was achieved with a pressure bandage. the post-procedure outcome is survived at explant. the automated impella controller and the impella pump will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | This is one of two related reports. this report represents the impella cp and another report was submitted to represent the automated impella controller. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
817 | 1220648-2026-02059 | 24233538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24233538&pc=OZD | 12/22/2025 |  |  | 02/02/2026 |  |  | 01/23/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026665307 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
818 | 1220648-2026-02074 | 24233578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24233578&pc=OZD | 01/21/2026 |  |  | 02/02/2026 |  |  | 01/21/2026 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798001 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: a 68-year-old female with postcardiotomy cardiogenic shock and a pre support clinical presentation consistent with scai stage e underwent attempted placement of an impella 5.5 device via a left axillary surgical cutdown. the impella encountered resistance preventing smooth passage and eventually was aborted. based on the information provided, the impella 5.5 was not successfully delivered due to anatomical limitations of the patient¿s axillary artery, which prevented the device from navigating the required turn. further evaluation will be performed if the product is returned for analysis. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
819 | 1220648-2026-02060 | 24233674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24233674&pc=OZD | 01/24/2026 | 02/16/2026 |  | 02/02/2026 |  |  | 01/24/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1556637 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
820 | 1220648-2026-02066 | 24233721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24233721&pc=OZD | 01/24/2026 |  |  | 02/02/2026 |  |  | 01/24/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1849082 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Material Split, Cut or Torn | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
821 | 1220648-2026-02069 | 24233810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24233810&pc=OZD | 01/28/2026 |  |  | 02/02/2026 |  |  | 01/28/2026 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794792 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 96 KG | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The registered nurse reported that while the patient was sleeping, motor current and placement signal pulsatility were reduced causing intermittent impella in aorta and position unknown alarms. an echocardiogram several hours earlier had shown proper inlet position. they were certain that the position had not changed since the echocardiogram was done several hours earlier. it was recommend that they check again when possible and to consider reducing the p-level to see if that resolves the alarms. they did not consider giving fluid bolus due to elevated filling pressures. reduction of p-levels resolved the alarms. no patient harm was reported. | [5 empty] | Explant date was provided therefore d6b was updated. | The impella device was not received from the customer as they remain on support. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
822 | 1220648-2026-02071 | 24233820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24233820&pc=OZD | 01/24/2026 |  |  | 02/02/2026 |  |  | 01/25/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026804168 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Anemia | Cardiac Arrest | Hematoma | Pain | Sepsis | Tachycardia | [32 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
823 | 1220648-2026-02082 | 24234115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24234115&pc=OZD | 10/08/2025 |  |  | 02/02/2026 |  |  | 10/08/2025 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1324018 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Other | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a bedside registered nurse caring for a patient receiving dual mechanical circulatory support noted repeated "controller error" alerts from the active controller. the alerts began approximately five minutes before the initial call and briefly resolved on their own. the issue was communicated to the clinical support center, which advised keeping a backup controller powered on and ready at the bedside while allowing the active controller to continue functioning. the active controller continued to alarm over the following hour. after further consultation with the clinical support center, staff were instructed to switch to the backup controller. the backup controller was activated, and the pump was successfully transferred without complications. clinical review: 71yo male presents with shortness of breath and fatigue and subsequently placed on impella and rp flex. a controller error alerted. the nurse switched aic to a new console. no patient consequence. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
824 | 1220648-2026-02024 | 24234241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24234241&pc=OZD | 01/23/2026 |  |  | 02/02/2026 |  |  | 01/23/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1786113 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Device Alarm System | Failure to Read Input Signal | [5 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
825 | 1220648-2026-02076 | 24234314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24234314&pc=OZD | 10/25/2025 | 02/20/2026 |  | 02/02/2026 |  |  | 10/25/2025 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026740924 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Renal Failure | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
826 | 1220648-2026-02077 | 24234318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24234318&pc=OZD | 01/29/2026 |  |  | 02/02/2026 |  |  | 01/29/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026803120 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Difficult to Open or Close | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
827 | 1220648-2026-02096 | 24235390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24235390&pc=OZD | 01/23/2026 | 02/02/2026 |  | 02/02/2026 |  |  | 01/25/2026 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT, PKGD, JP |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT, PKGD, JP |  |  | ABIOMED, INC. |  |  | US |  | 1566737 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Power Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A 72-year-old male presented in pre support clinical condition consistent with scai cardiogenic shock stage c. the impella device was indicated due to cardiac rupture, suspected to be a blowout type rupture in the lad region following myocardial infarction. on (b)(6) 2026, during the impella insertion procedure, the device was powered on; however, an initial diagnostic error occurred in the adaptive cardiac rupture control unit (b)(6). based on the available information, the reported event involved an initial diagnostic error in control unit (b)(6) during impella installation, necessitating replacement with control unit (b)(6). the issue was identified prior to full device operation, and the exchange resolved the error. the aic will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange however the exchange was performed with no consequence to the patient and support was resumed without any known adverse events. the event will continue to be monitored, and any additional information will be assessed as part of ongoing post market surveillance. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | D6b: updated explant date.
828 | 1220648-2026-02106 | 24236008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24236008&pc=OZD | 01/21/2026 | 02/02/2026 |  | 02/02/2026 |  |  | 01/21/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1400770 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
829 | 1220648-2026-02111 | 24236181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24236181&pc=OZD | 01/26/2026 | 02/02/2026 |  | 02/02/2026 |  |  | 01/26/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1349165 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Power Problem | [5 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
830 | 1220648-2026-02114 | 24236943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24236943&pc=OZD | 12/15/2025 |  |  | 02/02/2026 |  |  | 12/15/2025 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026780139 |  | 1000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 70 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. elevated purge pressures (in the 1000s) and low purge flows (<1) were observed. tissue plasminogen activator (tpa) was initiated via the purge system; however, no improvement in purge pressures was noted. the clinical team planned device weaning and explantation. the patient survived the procedure. | [5 empty] | 5.5 pump, sn: (b)(6) & data stick were returned. high purge pressure: returned product was visually inspected and no damage to pump or cannula was observed. no biomaterial was present in or around purge gap. pump was connected to lab console and purged for approximately 8 minutes with no purge alarms reproducing. pump was turned on to p-9 and again purge alarms did not reproduce. returned usb did not contain any data logs. clinical details noted purge pressure in the 1000s and purge flow of less than 1ml/hr. tpa added to purge but did not resolve issue prior to decision to explant the same day. the cause of the high purge pressure could not be determined as issue did not reproduce with returned product and no data logs were available for analysis. pump sn (b)(6) passed all post-sterile inspection checks.
831 | 1220648-2026-02115 | 24239844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24239844&pc=OZD | 01/27/2026 |  |  | 02/02/2026 |  |  | 01/27/2026 | 02/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798797 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
832 | 1220648-2026-02117 | 24240456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24240456&pc=OZD | 12/15/2025 | 02/17/2026 |  | 02/02/2026 |  |  | 12/15/2025 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782080 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female | 88 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient experienced right shoulder pain following implantation of an impella 5.5 device via the right axillary artery. on assessment, mild swelling was noted without redness or signs of infection, and the patient described soreness in the area. the clinical team was made aware of the concern and continued to monitor the site. | [5 empty] | No product was returned. pain: the cause of the injury was not determined due to insufficient clinical details. miwb/ hemolysis: patient had plasma-free hemoglobin 90, decreased from 100. the patient was subsequently administered iv fluids. no logs are available for this device. the cause of the hemolysis was not determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | Investigation summary. no product was returned. pain: the cause of the injury was not determined due to insufficient clinical details. miwb/ hemolysis: patient had plasma-free hemoglobin 90, decreased from 100. the patient was subsequently administered iv fluids. no logs are available for this device. the cause of the hemolysis was not determined due to insufficient clinical details.
833 | 1220648-2026-02119 | 24240746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24240746&pc=OZD | 12/15/2025 |  |  | 02/02/2026 |  |  | 12/15/2025 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765166 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 92 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed a hematoma around the impella 5.5 axillary insertion site. a surgical drain remained in place with minimal output, and the bedside nurse suspected that a clot was obstructing the drain. the surgeon and care team were notified. the patient remained hemodynamically stable and did not require blood transfusions, and the team chose to continue monitoring. the patient later returned to the operating room for washout and evacuation of the hematoma. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
834 | 1220648-2026-02120 | 24240898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24240898&pc=OZD | 12/16/2025 |  |  | 02/02/2026 |  |  | 12/16/2025 | 02/02/2026 |  |  |  |  | 02/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773823 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 103 KG | Non Hispanic | White | Injury | Yes | No | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an attempted right axillary impella 5.5 insertion resulted in a serious injury event. the patient had a history of a prior aortic valve replacement and newly identified severe aortic stenosis. although the risks and contraindications related to advancing the pump through a severely stenotic valve were reviewed, the surgical team proceeded. the guidewire was advanced into the left ventricle, but the pump could not be advanced across the aortic valve due to obstruction. the interventional cardiology team performed a balloon aortic valvuloplasty, but repeated attempts to advance the pump remained unsuccessful. further valvuloplasty was determined to be too risky, and the procedure was aborted. the patient required continued support with extracorporeal membrane oxygenation and an intra-aortic balloon pump, and plans were made to reassess the need for repeat valve intervention prior to any future attempt at impella placement. | [5 empty] | No product returned. failure to advance: the cause of the deliver issue was determined to be patient condition as the patient had av stenosis preventing successful delivery of impella. device passed all post sterile inspection checks.
835 | 1220648-2026-01924 | 24220160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24220160&pc=OZD | 01/21/2026 |  |  | 01/30/2026 |  |  | 01/22/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794648 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
836 | 1220648-2026-01926 | 24220435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24220435&pc=OZD | 01/22/2026 |  |  | 01/30/2026 |  |  | 01/22/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1366755 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
837 | 1220648-2026-01927 | 24220475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24220475&pc=OZD | 01/20/2026 |  |  | 01/30/2026 |  |  | 01/20/2026 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1075486 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an automated impella controller generated multiple yellow controller error alarms. pump flow remained stable and the alarms were self limiting. the patient was in stable condition. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
838 | 1220648-2026-01928 | 24220509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24220509&pc=OZD | 01/23/2026 |  |  | 01/30/2026 |  |  | 01/23/2026 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1849082 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Battery Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | While a patient was on impella support there was a low battery alarm despite the device being plugged in. they attempted to plug into an alternate wall outlet with no resolution of the alarm. the decision was made to switch to an alternate automated impella controller."the controller will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however the exchange was performed with no consequence to the patient or support. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
839 | 1220648-2026-01932 | 24220651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24220651&pc=OZD | 02/16/2026 |  |  | 01/30/2026 |  |  | 01/23/2026 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801547 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 83 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Unspecified Infection | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Impella 5.5 was inserted via axillary artery in a 49year old male with acute decompensated cardiomyopathy. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was identified as scai shock stage d. while on support the patient experienced ventricular tachycardia which required cardioversion x2. the pump functioned within expected range at p-8 4.2l/min. the patient remains on support. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia/arrhythmia; the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
840 | 1220648-2026-01933 | 24221117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24221117&pc=OZD | 01/21/2025 |  |  | 01/30/2026 |  |  | 01/21/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025592823 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemoptysis | Thrombocytopenia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
841 | 1220648-2026-01935 | 24221184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24221184&pc=OZD | 01/18/2026 |  |  | 01/30/2026 |  |  | 01/18/2026 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026769354 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female | 45 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Clinical rationale/review: an impella 5.5 was placed via the axillary graft site to support the 53 year old female patient who had been admitted in with acute myocardial infarction and cardiogenic shock, presenting in scai stage e shock. she had known history of a prior bypass/cabg for coronary artery disease and diabetes. no other history was shared. on 2nd day of support there was an observation made of the access site bleeding which required blood products to be infused back to the patient. with the bleeding there was flattening of the motor current and reduced flows were seen but pump position was appropriate. the team discussed another echo imaging to be done to assess position and volume needs. the pump remains on for support despite the bleeding and motor current and flow issues. anticoagulation necessary for the pump placement and support can contribute to access site bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b: added explant date.
842 | 1220648-2026-01937 | 24221239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24221239&pc=OZD | 01/23/2026 |  |  | 01/30/2026 |  |  | 01/23/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766844 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
843 | 1220648-2026-01942 | 24221380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24221380&pc=OZD | 01/22/2026 |  |  | 01/30/2026 |  |  | 01/23/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795714 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
844 | 1220648-2026-01943 | 24221417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24221417&pc=OZD | 01/23/2026 |  |  | 01/30/2026 |  |  | 01/23/2026 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801028 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 68 KG |  | White | Injury | Yes | No | Increase in Pressure | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Impella 5.5 was inserted via axillary artery in a patient of unknown age and gender for an unidentified 61 year old male for cardiomyopathy, acute decompensation with goal to bridge to transplant. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was documented as scai shock stage d. the impella device functioned within expect flow range p-5 2.7l/min. while on support there were purge pressure concerns so the site decided to initiate tpa protocol, which is an off label use, for the purge solution. during echocardiography assessment there was a noticeable lv thrombus which was not appreciated on previous echocardiograms. in an effort to extubate the patient they weaned sedation, the patient woke and was exqubated, however, there was unresolved ventricular tachycardia despite cardioversion. the decision was made to remove the impella device and continue support with ecmo. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
845 | 1220648-2026-01945 | 24221432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24221432&pc=OZD | 01/21/2026 |  |  | 01/30/2026 |  |  | 01/21/2026 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1112883 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 113 KG | Non Hispanic | White | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted via the right femoral artery in a 68-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). the patient had a known history of coronary artery disease. shortly after initiation of support, the customer reported a controller error alarm. the alarm subsequently resolved; however, out of caution and in accordance with guidance, the customer elected to exchange the console. the reported events are controller failure, medical device revision or replacement, and hemodynamic instability. the impella cp is being conservatively reported for serious injury due to a temporary interruption of hemodynamic support during the console exchange; however, the exchange was performed with no consequence to the patient, and support was maintained without any known adverse events. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
846 | 1220648-2026-01947 | 24221498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24221498&pc=OZD | 01/23/2026 |  |  | 01/30/2026 |  |  | 01/23/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026729870 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
847 | 1220648-2026-01951 | 24221905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24221905&pc=OZD | 01/26/2026 |  |  | 01/30/2026 |  |  | 01/26/2026 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1204376 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the purge cassette was not reading. they tried to power cycle the device, and this did not resolve the issue. they stated that there was no patient involvement and the device was replaced. | [5 empty] | A: patient demographic information is not available. e: initial reporter's information is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
848 | 1220648-2026-01952 | 24221917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24221917&pc=OZD | 01/14/2026 | 02/02/2026 |  | 01/30/2026 |  |  | 01/14/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026786804 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
849 | 1220648-2026-01953 | 24221947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24221947&pc=OZD | 01/15/2026 |  |  | 01/30/2026 |  |  | 01/15/2026 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026752000 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted via the right femoral artery in a male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs), scai shock stage c. during support, the treating physician reported bleeding at the femoral access site. best practices were reviewed, including discussion of applying a pressure bandage, with recommendation to consult vascular surgery if bleeding persisted. the patient was noted to have consumptive coagulopathy with a markedly elevated activated clotting time (act). bleeding subsequently occurred from multiple vascular access sites. following consultation with the treating physicians, it was determined that the act had been continuously maintained above 300, which was identified as a user-related management issue and the most plausible contributor to the bleeding events. no surgical intervention was required. the patient subsequently expired while on impella support. access-site bleeding is a known risk associated with impella support due to anticoagulation and purge requirements, as described in the instructions for use. based on the available information, the bleeding is most plausibly related to supratherapeutic anticoagulation in the setting of consumptive coagulopathy. the death is conservatively being reported on the impella cp; however, it is unlikely that the device contributed to the patient¿s death, which is more plausibly attributed to the underlying critical illness and cardiogenic shock. | [5 empty] | The exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. a2, a4, a5, and a6 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
850 | 1220648-2026-01959 | 24222668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24222668&pc=OZD | 01/26/2026 |  |  | 01/30/2026 |  |  | 01/26/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798648 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
851 | 1220648-2026-01962 | 24223060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24223060&pc=OZD | 01/26/2026 |  |  | 01/30/2026 |  |  | 01/26/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026632487 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
852 | 1220648-2026-01963 | 24223141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24223141&pc=OZD | 01/24/2026 |  |  | 01/30/2026 |  |  | 01/24/2026 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026767037 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 113 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted via the right subclavian artery in a 69-year-old male patient with a past medical history of coronary artery disease (cad), diabetes, renal insufficiency and gastrointestinal bleed (gi bleed), presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), with scai stage d shock, and on extracorporeal membrane oxygenation (ecmo), on multiple inotropes and vasopressors, and ventilated for respiratory support prior to initiation of support. an initial impella cp was inserted for ventricular support during a high-risk percutaneous coronary intervention (pci) of the left main artery (lm), the left anterior descending artery (lad), and the left circumflex artery (lcx). the impella cp device was removed, and an impella 5.5 device was then inserted for escalation of therapy. while in the intensive care unit (icu), the patient was on bivalirudin for which was turned off frequently for gi bleeding. there were times where the impella needed to be repositioned due to malposition towards the mitral valve. the flows were increased and decreased multiple times over the last four days. the patient had multiple episodes of ventricular tachycardia (vt) arrests and was in vt storm while on support. the plasma free hemoglobin was 133 and increased to 442. the lactate dehydrogenase (ldh) increased from 722 to 2026. four units of packed red blood cells (prbcs) were transfused for a hemoglobin of 5.7. the ventricular assist device (vad) engineer noticed a change in the left ventricle (lv) placement signal and motor current without any changes to the patient¿s flows. a bedside echo showed the impella at 3.3 cm from the aortic valve. the medical team did not see a clot on the impella or in the lv, so they started thinking hemolysis and made the decision for the impella to be exchanged. in the operating room (or), while removing the impella 5.5, access site bleeding was noted, which required blood products to be given. two pieces of tissue were also noted to be on the outflow of the impella's cannula. a new impella 5.5 was inserted into the left side. the post-procedure outcome is survived at explant. while hemolysis and bleeding are listed as potential adverse events, and consistent with known device-related and patient-related factors documented in the instructions for use (ifu), ecmo can also be attributed to these complications. however, these complications are likely due to the underlying clinical condition of a critically ill patient with acute myocardial infarction complicated by cardiogenic shock as they presented in stage d shock. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
853 | 1220648-2026-01964 | 24223150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24223150&pc=OZD | 01/26/2026 | 02/02/2026 |  | 01/30/2026 |  |  | 01/26/2026 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1109482 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [12 empty] | Other | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient was transferred to (b)(6) on impella 5.5, upon arrival the console was switched to automated impella controller (aic) 4510, 4510 was unable to detect the purge disc and an alarm continued to go off for purge disc, a new cassette was attempted however during the cassette change menu - the aic was unable to recognize the new purge cassette despite being placed correctly into the aic, the impella was then placed back onto the mon general console. another aic was obtained and this led to a successful console swap. clinical review an impella 5.5 device was inserted surgically into the aorta in a 54-year-old male patient with a past medical history of coronary artery disease (cad), presenting in society for cardiovascular angiography and interventions (scai) stage a shock and ventilated for respiratory support prior to initiation of support. the impella 5.5 was inserted for ventricular support during an elective high-risk coronary artery bypass graft (cabg) surgery. it was reported that upon transfer from another facility, the automated impella controller (aic) console was unable to detect the purge disc and an alarm sounded. while changing the cassette, the aic was unable to recognize the new purge cassette which led to a new console swap. the post-procedure outcome is unknown. the automated impella controller will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. | [5 empty] | Product complaint # (b)(4). this report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by ethicon inc, or its employees that the report constitutes an admission that the product, ethicon inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate. to date the device has not been returned. if the device or further details are received at a later date a supplemental medwatch will be sent. no further information available as reporter details have not been disclosed (confidential). | D4 udi was erroneously entered on the initial manufacturer device report number. | D4 lot number updated. the investigation is still ongoing.
854 | 1220648-2026-01975 | 24223965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24223965&pc=OZD | 01/26/2026 |  |  | 01/30/2026 |  |  | 01/26/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, CA |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026613377 |  | 0550-0004 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Contamination | Crack | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
855 | 1220648-2026-01979 | 24224111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24224111&pc=OZD | 01/19/2026 | 01/30/2026 |  | 01/30/2026 |  |  | 01/19/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798000 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
856 | 1220648-2026-01981 | 24224428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24224428&pc=OZD | 01/20/2026 | 01/30/2026 |  | 01/30/2026 |  |  | 01/20/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025482065 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
857 | 1220648-2026-01987 | 24225345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24225345&pc=OZD | 01/26/2026 |  |  | 01/30/2026 |  |  | 01/26/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026774717 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
858 | 1220648-2026-01986 | 24225353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24225353&pc=OZD | 11/12/2025 |  |  | 01/30/2026 |  |  | 11/12/2025 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026740909 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Difficult or Delayed Positioning | [6 empty] | Low Blood Pressure/ Hypotension | Ischemia | Renal Failure | Hypovolemia | Cardiogenic Shock | Respiratory Failure | Thrombosis/Thrombus | Heart Failure/Congestive Heart Failure | Hemodynamic instability |  |  |  |  | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that during an impella 5.5 insertion, there were initial attempts to cross the aortic valve using a glidewire with a diagnostic catheter were unsuccessful due to the wire repeatedly catching on the left coronary cusp. a bentson wire was subsequently utilized and successfully crossed the valve. the primed impella was then backloaded and passed through the hemostatic valve without difficulty. it was also reported that during advancement of the impella 5.5 into the left ventricle, the device became slightly oriented toward the interventricular septum. the surgeon spent several minutes applying torque in an attempt to correct its trajectory. although initially unsuccessful, he eventually achieved the appropriate combination of torque and forward advancement, allowing the pump to seat properly within the ventricle. it was reported that the patient continued to demonstrate evidence of clot formation within the extracorporeal membrane oxygenation (ecmo) circuit, even after an oxygenator changeout had been performed. in response to the patient¿s clinical condition, several interventions were undertaken. the patient remained sedated due to ventilator dyssynchrony, and anticoagulation was transitioned back to bivalirudin after an attempted switch to heparin resulted in subtherapeutic anti-xa levels and reduced antithrombin iii, findings that are commonly seen in patients with a history of hit. following the ecmo circuit changeout, the patient also received 1 unit of packed red blood cells and 1 unit of platelets. it was also mentioned that the patient was critically ill and experienced a sudden loss of extracorporeal membrane oxygenation (ecmo flow), decreasing to 1.8 l/min at 4400 rpm, following patient repositioning. the registered nurse (rn) identified visible clot within the ecmo oxygenator, raising concern for oxygenator membrane failure due to thrombosis. the patient was managed with intubation and sedation and continued v-av ecmo and impella 5.5 support. morning abg values were noted as follows 7.37 / 42 / 55 / 24 / ¿1 / 88% while on mechanical ventilation with 50% fio2 and an ecmo sweep of 2 l/min. after identifying clot burden within the oxygenator, ventilator fio2 was increased to 100%, and the sweep gas flow was increased to 3 l/min to augment oxygenation. an oxygenator changeout was performed which resulted in restored ecmo flow to 4 l/min (2.3 l/min venous / 1.7 l/min arterial). the care plan includes discontinuation of bivalirudin and initiation of heparin at 15 units/kg/hr, with continued support from both ecmo and the impella device. an ecmo reconfiguration from v-av to v-pa is planned. it was reported that the patient had concurrent emco therapy and had coagulopathy. the patient was continued with ecpella therapy which includes (extracorporeal membrane oxygenation(ecmo) and impella), and continuous renal replacement therapy (crrt)initiated for volume removal at 150 ml/hr. it was reported that patient showing new or worsening areas of mottling/discoloration of extremities or abdomen. the registered nurse noted micro clots. it was decided to transition the patient to comfort-focused care. following this decision, plans were made to withdraw life-sustaining treatments. impella support was terminated. the complaint procedure outcome was successful with this device, and the patient outcome is noted as critical. | [5 empty] | No product returned. cardiac failure, respiratory failure, hemodynamic instability, hypotension, renal failure, hypovolemia: the root cause of the injury was unable to be determined due to insufficient clinical details. heart block, ischemia, cardiogenic shock, thrombosis: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. difficult to place or position: the cause of the difficult to place or position was unable to be determined due to insufficient clinical details. device lot: 1974028. subcomponent lot: n/a. device (sn: (b)(6)) passed all post sterile inspection checks.
859 | 1220648-2026-01990 | 24225515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24225515&pc=OZD | 01/23/2026 |  |  | 01/30/2026 |  |  | 01/23/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1338086 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Power Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup
860 | 1220648-2026-01991 | 24225915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24225915&pc=OZD | 01/23/2026 | 02/26/2026 |  | 01/30/2026 |  |  | 01/24/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775144 |  | 1000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
861 | 1220648-2026-01993 | 24226024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226024&pc=OZD | 11/13/2025 |  |  | 01/30/2026 |  |  | 11/15/2025 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1471015 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | A 60-year-old female patient with cardiogenic shock had an attempted impella 5.5 insertion via the axillary artery. the patients' comorbidities were unknown. the patient¿s clinical state prior to initiation of support is/was identified as scai stage d shock and they were on an intra-aortic balloon pump, medical therapy and respiratory support. upon set up the automated impella controller the nursing staff noted that the purge disc and cassette were not recognized so an alternate automated impella controller was used without issue. due to a heavily calcified axillary artery the decision was made to abort proceeding with placing the impella 5.5 and transition to an impella cp. the impella cp was placed with no noted issues and support was initiated. the patient outcome is unknown. | [5 empty] | Data review: console logs were consistent with what was reported in the complaint. imc logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set #402) several times. (see ¿(b)(4)¿ in the attachments). device analysis: console (b)(4) was sent back to field service for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and the purge flag was confirmed to be broken. (see ¿(b)(4)¿ in the comments) root cause: the purge disc flag on ic8041 was observed to be broken and was the root cause of the inability to detect the purge disc. (b)(4) documents the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly. (see ¿(b)(4)¿ in the attachments for image of broken flag with retainer removed to inspect for fluid ingress) repair: before sending this console back to the customer, field service was instructed the following: - replace the purge disc detect flag [(b)(4)] and purge flag spring [(b)(4)] - perform sections 10-12 of pm procedure (b)(4)] - perform a full functional check on the console and optical system [(b)(4)] subcomponent name: purge flag. material number: 3001198. q1) service history review - are there any qns associated with the complaint device¿s most recent service report? no q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode? n/a q3) complaint history review - have there been any other complaints against this console? yes. three other complaints were made against this console: (b)(6): ¿black flag for the purge disc broken¿. confirmed broken purge disc flag. (b)(6): ¿purge alarms keep happening after replacement cassettes¿. the root cause of the alarm of purge fluid not detected and rfid error (invisible to user) was a defective pud (b)(6): ¿aic sn: (b)(6) lost placement signal and the console had to be switched out¿. oscillating contrast observed but the root cause could not be conclusively determined due to the inability to reproduce. phr summary: there were issues related to the complaint discovered when reviewing the product history of console ic8041 - a prior complaint with the same fm/rc (smartsolve #(b)(4)).
862 | 1220648-2026-01994 | 24226030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226030&pc=OZD | 11/13/2025 |  |  | 01/30/2026 |  |  | 11/15/2025 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748286 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
863 | 1220648-2026-01997 | 24226140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226140&pc=OZD | 01/15/2026 | 01/30/2026 |  | 01/30/2026 |  |  | 01/16/2026 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026795737 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 111 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in a 64-year-old male patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). the patient had a history of diabetes mellitus and renal insufficiency and was classified as scai shock stage¿d. during support, the patient experienced ventricular tachycardia. pump position appeared appropriate on echocardiography, and volume status was assessed as adequate. the reported ventricular tachycardia is most plausibly attributed to the patient¿s underlying cardiac rhythm disturbances and critical clinical condition. | [5 empty] | The investigation for the tachycardia has been completed. the device was not returned for evaluation. additionally, clinical details were not sufficient to determine the root cause. therefore, the root cause of the injury was unable to be determined due to insufficient clinical details.
864 | 1220648-2026-01999 | 24226218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226218&pc=OZD | 01/22/2026 | 01/30/2026 |  | 01/30/2026 |  |  | 01/22/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026798825 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Anemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
865 | 1220648-2026-02000 | 24226240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226240&pc=OZD | 01/27/2026 |  |  | 01/30/2026 |  |  | 01/27/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1338418 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
866 | 1220648-2026-02002 | 24226313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226313&pc=OZD | 01/22/2026 | 02/03/2026 |  | 01/30/2026 |  |  | 01/22/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1814730 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Application Program Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
867 | 1220648-2026-02003 | 24226366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226366&pc=OZD | 01/22/2026 | 01/30/2026 |  | 01/30/2026 |  |  | 01/22/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026780939 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
868 | 1220648-2026-02005 | 24226388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226388&pc=OZD | 01/22/2026 | 01/30/2026 |  | 01/30/2026 |  |  | 01/22/2026 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1432077 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Battery Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | An automated impella controller (aic) was found, during impella use in the icu, to have a battery charging issue. the team troubleshot by changing the plug outlet, but when this did not remedy the issue, the team replaced the aic. the aic-to-aic swap occurred successfully. support proceeds and no other intervention was shared/known. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. e4 was inadvertently omitted on the initial manufacturer device report. h5 was erroneously selected on the initial manufacturer device report. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
869 | 1220648-2026-02007 | 24226455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226455&pc=OZD | 01/22/2026 | 02/16/2026 |  | 01/30/2026 |  |  | 01/22/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026783696 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup
870 | 1220648-2026-02009 | 24226547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226547&pc=OZD | 01/23/2026 | 01/30/2026 |  | 01/30/2026 |  |  | 01/23/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026775944 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
871 | 1220648-2026-02011 | 24226562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226562&pc=OZD | 01/22/2026 | 01/30/2026 |  | 01/30/2026 |  |  | 01/23/2026 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026795532 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
872 | 1220648-2026-02010 | 24226565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226565&pc=OZD | 12/01/2025 |  |  | 01/30/2026 |  |  | 12/01/2025 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1084348 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Hemodynamic instability | [41 empty] | Initial submission
873 | 1220648-2026-02012 | 24226573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226573&pc=OZD | 10/18/2025 |  |  | 01/30/2026 |  |  | 10/18/2025 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026753466 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Difficult to Insert | Deformation Due to Compressive Stress | [5 empty] | Perforation of Vessels | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the nurse reported, that after peeling the peel away sheath, the reposheath could not be inserted inside the patient. the catheter shaft (outside the body!) has an abnormality over which the repo sheath could not be pushed. additionally, there was a puncture in the femoral artery. the puncture was fixed without any relevant complications. it was reported that the patient was stable. | [5 empty] | A: patient information is not available. e: initial reporter is unknown. investigation summary: no product was returned. product damage (catheter kink): clinical details report that there was an abnormality on the catheter shaft that prevented the repo sheath from being advanced. the cause of the catheter kink was not determined since product was not returned and insufficient clinical details were provided. difficult/unable to insert through other device: clinical details note that the repo sheath could not be advanced due to an abnormality on the catheter. ultimately, they fixed the pump with the repo sheath outside of the body. the cause of the difficulty to insert through other device was not determined since product was not returned and insufficient clinical details were provided. device history review: the pump passed all post-sterile inspection checks.
874 | 1220648-2026-02015 | 24226715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24226715&pc=OZD | 10/11/2025 |  |  | 01/30/2026 |  |  | 10/11/2025 | 01/30/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026653493 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  |  | Thrombocytopenia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
875 | 1220648-2026-02016 | 24227593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24227593&pc=OZD | 12/10/2025 |  |  | 01/30/2026 |  |  | 12/18/2025 | 01/30/2026 |  |  |  |  | 01/30/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026786130 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Heart Failure/Congestive Heart Failure | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient needing an impella 5.5 pump for mechanical circulatory support. the implant was considered a high risk percutaneous cardiac second to the patient¿s periprocedural events of stroke and congestive heart failure. while on support the patient expired. it was ruled that the cause of death was hypoxic encephalopathy and postoperative heart failure. in addition, it was reported that there was no relation to the impella device and the patient¿s death. no further information was provided. | The patient was initially implanted with impella 5.5 pump 1 (sn(b)(6)) which remained in place for approximately one day. a pump stop occurred, and both the impella 5.5 device and the automated impella controller (aic) in use were replaced. this event is documented under a related complaint file. the patient was subsequently implanted with impella 5.5 pump 2 (sn63347) (this report's device) which remained in place for approximately 11 days. while on this support, the patient expired. the documented cause of death was hypoxic encephalopathy and postoperative heart failure. per the physician¿s comments, the death was not related to the impella, and there were no issues with impella operation. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received and b5: event description was updated, corrected accordingly. medical safety/clinical review: medical safety/clinical reviewed the event. the event report, with limited available information, describes a patient with post-cardiotomy cardiogenic shock (pccs) who was implanted with an impella 5.5 device for mechanical circulatory support (mcs). the patient remained on impella support for approximately 11 days and subsequently expired. the reported cause of death was hypoxic encephalopathy and postoperative heart failure. it is unknown the order of events in relation to the impella 5.5 implant. according to the treating physician, there were no reported issues with impella device operation, and the event was assessed as unrelated to the impella device. however, the death will be reported on the impella 5.5 (pump 2) as device was in use at time of expiration. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding cardiac failure, the cause of the injury was not determined due to insufficient clinical details. for the stroke, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this is one of three devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | Medwatch follow up 1 report was erroneously submitted with a 4-feb-2026 g3 date received by manufacturer and should be disregarded. the correct g3. date for follow-up 1 is 23-dec-2025.
876 | 1220648-2026-01809 | 24206630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206630&pc=OZD | 11/05/2025 |  |  | 01/29/2026 |  |  | 11/05/2025 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713399 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
877 | 1220648-2026-01781 | 24206632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206632&pc=OZD | 10/19/2025 | 01/28/2026 |  | 01/29/2026 |  |  | 10/19/2025 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026636270 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Female | 52 KG | Hispanic |  | Injury | Yes | No | Malposition of Device | [6 empty] | Cardiac Arrest | Fever | Tachycardia | Respiratory Failure | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient had a history of intracranial bleeding and significant pulmonary compromise. despite interventions, pulmonary status remained poor, although urine output gradually improved. shortly after initiation of mechanical circulatory support, the catheter tip was found to have inverted; however, pump flow remained stable and was monitored. anticoagulation was started at reduced intensity due to a mechanical tricuspid valve. a computed tomography scan revealed clots in the medulla, prompting adjustments to anticoagulation therapy and reduction of pump support. the patient developed persistent fever and worsening pulmonary findings. episodes of ventricular tachycardia occurred, with uncertainty regarding whether the arrhythmia was related to the support device or the pulmonary artery catheter. anticoagulation was transitioned from heparin to bivalirudin, with bicarbonate added to the purge. hemodynamic assessment showed elevated pulmonary pressures. the lungs gradually improved, and sedation was maintained to support neurological recovery from the prior intraventricular hemorrhage. because the fever persisted, the pulmonary artery catheter was exchanged. electrolytes were corrected, and the complication was managed. neurological symptoms were later determined to be due to a brainstem arteriovenous malformation. cardiogenic shock began to improve, but the patient continued to lack airway protective reflexes. during tracheostomy placement, the patient developed respiratory arrest followed by cardiac arrest, requiring cardiopulmonary resuscitation. pump support was increased, and the patient was emergently placed on peripheral veno-arterial extracorporeal membrane oxygenation. suction events and differential oxygenation occurred, prompting surgical conversion to veno-venous extracorporeal membrane oxygenation, which improved hemodynamics. multiple vasopressor agents and antiarrhythmic therapy were required. subsequent evaluation in the cardiac catheterization laboratory showed sluggish extracorporeal membrane oxygenation circuit flow and clot formation within the oxygenator, raising concern for inadequate ventricular unloading. the patient remained on bivalirudin and epinephrine. the left ventricular signal on the controller was intermittently absent at low pump settings. by the end of the reported course, the patient no longer required mechanical circulatory support; however, pulmonary recovery from acute respiratory distress syndrome and neurological recovery from intracranial bleeding were ongoing. cardiogenic shock had resolved, but the patient continued to require aggressive medical management. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
878 | 1220648-2026-01813 | 24206642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206642&pc=OZD | 11/05/2025 |  |  | 01/29/2026 |  |  | 11/06/2025 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744185 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Fever | Hemolysis | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
879 | 1220648-2026-01816 | 24208326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24208326&pc=OZD | 01/20/2026 |  |  | 01/29/2026 |  |  | 01/21/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026729894 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
880 | 1220648-2026-01603 | 24208372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24208372&pc=OZD | 10/28/2025 |  |  | 01/29/2026 |  |  | 10/28/2025 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1080535 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission
881 | 1220648-2026-01818 | 24208413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24208413&pc=OZD | 01/09/2026 |  |  | 01/29/2026 |  |  | 01/09/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026785665 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a pump stop during the weaning process. the pump could not be restarted so it was explanted. the patient was stable and supported by venoarterial extracorporeal membrane oxygen. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
882 | 1220648-2026-01819 | 24208464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24208464&pc=OZD | 01/13/2026 |  |  | 01/29/2026 |  |  | 01/15/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782243 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Death | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Based on available information, the event appears related to inability to advance the impella due to patient specific vascular anatomy. the patient¿s survival outcome at explant was reported as death, though the exact time and circumstances of death were not provided. reported failure modes included failure to advance and patient death. the death is conservatively being reported to the impella cp until further information is available. no device malfunction has been confirmed at this time; final assessment will depend on product evaluation and any additional clinical information that becomes available. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the root cause of the delivery issue was determined to be patient condition as the impella cp could not be implanted in the ventricle because of the patient's vascular anatomy. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. | Patient data was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
883 | 1220648-2026-01820 | 24208492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24208492&pc=OZD | 01/14/2026 |  |  | 01/29/2026 |  |  | 01/15/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The event involved repeated controller failure alarms and a later loss of placement signal, which are malfunctions reportable under mdr requirements because they could potentially lead to serious patient injury if they were to recur. although no patient harm occurred, the alarms necessitated controller exchange and altered clinical management (echo based monitoring), was required. the impella 5.5 will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. the patient remains on support. | [5 empty] | Patient data was not provided. section a was left blank. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1 product problem was inadvertently omitted on the initial manufacturer device report. d6 should have been left blank on the initial manufacturer device report. h5 and h8 was erroneously selected on the initial manufacturer device report.
884 | 1220648-2026-01834 | 24208915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24208915&pc=OZD | 01/19/2026 |  |  | 01/29/2026 |  |  | 01/20/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775826 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Vascular Dissection | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 54 year old female with a history of cardiomyopathy underwent support with an impella 5.5 (b)(6) implanted via right axillary/subclavian surgical access on 1/2/2026 and explanted on (b)(6) 2026. the patient initially survived to explant without reported device related complications. on (b)(6) 2026, approximately one week post explant, she was readmitted for a transient ischemic event characterized primarily by dizziness upon standing. diagnostic imaging identified a small arterial dissection near the previous graft site¿where the graft had been fully removed during explant¿and a possible small thrombus within the subclavian artery. the stroke etiology was determined to be embolicbased on the available information, the patient experienced an embolic cerebrovascular event post explant, with contributing factors including thrombus formation and a localized dissection near the former graft site. the patient was initiated on a heparin infusion and the plan included ongoing anticoagulation therapy, with aspirin to be continued upon discharge, and follow up imaging in the coming weeks while the event was attributed clinically to impella support, no device was available for analysis, and no definitive device malfunction was identified. the patient survived the event and remains under continued medical management. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
885 | 1220648-2026-01836 | 24208932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24208932&pc=OZD | 01/20/2026 | 02/02/2026 |  | 01/29/2026 |  |  | 01/20/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026741805 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 95 KG | Hispanic | White | Injury | Yes | No | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 73 year old male undergoing mechanical circulatory support for acute myocardial infarction complicated by cardiogenic shock required impella therapy. the device was unable to be advanced past the innominate artery due to significant vessel tortuosity, identified as the location of resistance. no product damage was reported, no product return was indicated, and no patient harm occurred. the 5.5 was subsequently aborted and a impella cp was implanted. the reported event is consistent with a delivery challenge related to patient specific vascular tortuosity rather than device malfunction. the inability to advance the impella 5.5 appears attributable to anatomical resistance, and there is no allegation of device performance issues at this time. the impella 5.5 will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange however the exchange was performed with no consequence to the patient and support was resumed without any known adverse events. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
886 | 1220648-2026-01837 | 24208955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24208955&pc=OZD | 01/20/2026 |  |  | 01/29/2026 |  |  | 01/20/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782019 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 101 KG | Non Hispanic | White | Injury | Yes | No | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 66 year old male underwent a planned pre operative placement of an impella 5.5 via the right axillary/subclavian surgical arterial approach. during the procedure, significant difficulty was encountered advancing the device due to the patient¿s complex anatomy. initial placement was attempted over an 0.027" wire; however, the wire could not be removed secondary to a kink sustained during device advancement. the impella 5.5 could not be successfully implanted due to anatomical complexity, guidewire kinking, and suspected shaft compromise leading to inadequate torque response. the device was removed, and support was transitioned to an impella cp without reported complications. no patient harm was reported, and the patient remained hemodynamically supported following cp implantation. product return evaluation is anticipated to further assess the reported issues. the impella 5.5 will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange however the exchange was performed with no consequence to the patient and support was resumed without any known adverse events. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
887 | 1220648-2026-01840 | 24208996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24208996&pc=OZD | 01/21/2026 | 02/02/2026 |  | 01/29/2026 |  |  | 01/21/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1218086 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Priming Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup
888 | 1220648-2026-01841 | 24209177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24209177&pc=OZD | 11/18/2025 |  |  | 01/29/2026 |  |  | 11/18/2025 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766566 |  | 1000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Damaged by Another Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 59-year-old male was admitted for myocardial infarction with cardiogenic shock. an impella cp was placed and a percutaneous coronary intervention carried out. towards the end of procedure, a pulmonary catheter was placed. when manipulating the pulmonary catheter, the impella cp was accidentally pulled out of the left ventricle, requiring a replacement of the device. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
889 | 1220648-2026-01843 | 24209282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24209282&pc=OZD | 01/17/2026 |  |  | 01/29/2026 |  |  | 01/17/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026743354 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | False Alarm | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Hemodynamic instability | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
890 | 1220648-2026-01844 | 24209377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24209377&pc=OZD | 11/18/2025 |  |  | 01/29/2026 |  |  | 11/18/2025 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026769856 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 62-year-old male was treated for acute myocardial infarction with cardiogenic shock. an impella cp was placed and shortly after, urine color was suggestive of hemolysis. the impella cp was repositioned using echocardiography and the urine gradually cleared up. after two days of support, the impella cp was successfully weaned and removed. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
891 | 1220648-2026-01845 | 24209411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24209411&pc=OZD | 01/16/2026 |  |  | 01/29/2026 |  |  | 01/16/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737527 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Unspecified Infection | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
892 | 1220648-2026-01848 | 24209518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24209518&pc=OZD | 01/19/2026 |  |  | 01/29/2026 |  |  | 01/19/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797261 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
893 | 1220648-2026-01855 | 24210729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24210729&pc=PYX | 10/20/2025 |  |  | 01/29/2026 |  |  | 10/20/2025 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1668177 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the device exhibited 2 controller error alarms. the resolution for this issue is unknown. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
894 | 1220648-2026-01856 | 24210765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24210765&pc=OZD | 01/23/2026 |  |  | 01/29/2026 |  |  | 01/23/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773801 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male | 95 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Stroke/CVA | Tachycardia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella 5.5 device was inserted surgically into the aorta in a 19-year-old male patient with a past medical history of recurrent ventricular tachycardia, presenting in cardiomyopathy, with scai stage e shock, and on extracorporeal membrane oxygenation (ecmo), on multiple inotropes and vasopressors, and ventilated for respiratory support prior to initiation of support. after the impella was placed, the patient was noted to have monomorphic ventricular tachycardia (vt). the device was noted to be near the ventricular septum so the physician repositioned the device away from the septum. the patient was successfully cardioverted at 200j. it was noted that the patient had vt and was on amiodarone and lidocaine infusions pre and intra-operatively. while the patient was in the intensive care unit, the placement signal of the impella 5.5 dampened but the mean arterial pressure (map) correlated with the arterial line tracing. two hours later, "preventing retrograde flow" alarms at p-4. map on placement signal now reading 40mmhg higher than the arterial line (appropriately leveled, zeroed, and correlating with nibp). point of care ultrasound (pocus)/transthoracic echocardiogram (tte) to confirm device placement in setting of fiber optic issue. it was noted that the patient experienced vision changes. a computed tomography (ct) scan showed right ganglia ischemic stroke. a repeat ct showed hemorrhagic conversion of basal ganglia bleed and progressive intraventricular hemorrhage stroke. the patient was not a candidate for neurosurgical intervention. the family elected to withdraw care, and the patient expired on support. the device is unlikely to have contributed to the patient's death, which was most likely due to the stroke, and the clinical condition of a critically ill patient who presented in heart failure with cardiomyopathy and in stage e shock. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
895 | 1220648-2026-01857 | 24210817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24210817&pc=OZD | 11/19/2025 |  |  | 01/29/2026 |  |  | 11/19/2025 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026759575 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
896 | 1220648-2026-01858 | 24210861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24210861&pc=OZD | 10/21/2025 |  |  | 01/29/2026 |  |  | 10/21/2025 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1389732 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the controller exhibited an unknown alarm noise. the automated impella controller (aic) audible alarm notification wasn't typical ¿beep¿ sound, rather a chirping sound. the resolution for this issue is unknown. it was reported that the procedure was successful. | [5 empty] | D6a and d6b should have been left blank on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
897 | 1220648-2026-01859 | 24210967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24210967&pc=PYX | 10/21/2025 |  |  | 01/29/2026 |  |  | 10/21/2025 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1462131 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
898 | 1220648-2026-01861 | 24211249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24211249&pc=OZD | 01/21/2026 | 02/24/2026 |  | 01/29/2026 |  |  | 01/22/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1431034 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 56 KG | Non Hispanic | White | Injury | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Clinical rationale: a 57 year old male underwent elective impella placement via direct right-sided surgical access on (b)(60 2026 at 15:40. the reported issues were limited to psnr alarms and a subsequent yellow controller error, both of which resolved after controller replacement. pump performance remained stable. the impella 5.5 will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however the exchange was performed with no consequence to the patient and support. the patient remains on support. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. this is one of two related reports. this report represents the automated impella controller and another report was submitted to represent the impella 5.5.
899 | 1220648-2026-01862 | 24211264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24211264&pc=OZD | 01/25/2026 |  |  | 01/29/2026 |  |  | 01/25/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795547 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
900 | 1220648-2026-01863 | 24211469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24211469&pc=OZD | 11/08/2025 |  |  | 01/29/2026 |  |  | 11/07/2025 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026728911 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemolysis | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
901 | 1220648-2026-01865 | 24211644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24211644&pc=OZD | 11/12/2025 |  |  | 01/29/2026 |  |  | 11/12/2025 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026727096 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | Ischemia | Tachycardia | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
902 | 1220648-2026-01866 | 24211657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24211657&pc=OZD | 01/25/2026 |  |  | 01/29/2026 |  |  | 01/25/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1883624 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
903 | 1220648-2026-01867 | 24211793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24211793&pc=OZD | 01/25/2026 | 02/03/2026 |  | 01/29/2026 |  |  | 01/25/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026707304 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
904 | 1220648-2026-01869 | 24211973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24211973&pc=OZD | 01/24/2026 |  |  | 01/29/2026 |  |  | 01/25/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026805160 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the impella 5.5 pressure reservoir was leaking purge fluid. water droplets found on the pressure reservoir, crystallized material found on the external mechanism. there was no significant change in purge pressure or flow rate, no plans for replacement. continued use with gauze wrapped around the external mechanism and the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
905 | 1220648-2026-01846 | 24215140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24215140&pc=PYX | 11/04/2025 |  |  | 01/29/2026 |  |  | 11/04/2025 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1664272 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Injury | Yes | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Impella cp was placed in a 64y/o male for cardiomyopathy. hospital course complicated with suction alarms, hematuria, malpositioning. the automated impella controller (aic) will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange however the exchange was performed with no consequence to the patient and support was resumed without any known adverse events. during impella cp support, controller shutdown occurred. unexpected controller shutdown alarm appeared on screen, indicated automated impella controller (aic) restarted. automated impella controller (aic) was swapped without issue. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
906 | 1220648-2026-01875 | 24215147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24215147&pc=OZD | 11/19/2025 |  |  | 01/29/2026 |  |  | 11/19/2025 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026774575 |  | 1000080 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | [5 empty] | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 67 year old female was treated for acute myocardial infarction and cardiogenic shock. an impella cp was inserted and the patient transferred to the intensive care unit, where a retroperitoneal bleeding was diagnosed. the patient received blood transfusions. after two days of support, the impella cp could be weaned. during removal of the pump, a thrombus associated to the retroperitoneal bleeding dislodged and occluded the right femoral artery. a surgical intervention was required to restore perfusion to the limb. both major bleeding and ischemia will be coded to the impella cp. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
907 | 1220648-2026-01879 | 24217754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24217754&pc=OZD | 01/26/2026 | 01/29/2026 |  | 01/29/2026 |  |  | 01/26/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026771042 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
908 | 1220648-2026-01880 | 24217811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24217811&pc=OZD | 01/26/2026 | 01/29/2026 |  | 01/29/2026 |  |  | 01/26/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026783731 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
909 | 1220648-2026-01881 | 24217889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24217889&pc=OZD | 11/01/2025 | 01/29/2026 |  | 01/29/2026 |  |  | 01/28/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1093346 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Network Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported their automated impella controller (aic) could not update via impella connect. | [5 empty] | Investigation summary: data analysis: console logs from the reported day of event show that after handshake between aic and rlm was established, frequently occurring messages ¿icm not ready for file receive¿ were logged. ltlog and rtlog data indicates unreliable placement signal and low optical snr. rlm journal shows frequent reboots of the rlm, intermittent connectivity to rlv, rlr, and rlfs, but mostly - frequent disconnections from the wifi access point. there were also multiple rlm journal entries suggesting possible malfunction of the site network equipment (access point), e.g.: kernel: wlan0: associating with ap with corrupt probe response. there was also evidence of microsd card corruption, e.g.: kernel: ext4-fs error (device mmcblk1p11): ext4_dx_find_entry:1771: inode #2480642: block 8: comm rlc: directory block failed checksum. analysis of console ic1386 data also revealed poor optical performance with pumps (b)(6), indicated by low snr and multiple placement signal not reliable alarms - unrelated to this complaint. device analysis: console was tested at fs, and rlm connectivity failures were not reproduced. with a replaced microsd card, rl08461 operated within specification. optical parameters (5s) testing confirmed malfunction of the optical bench, evidenced by values of snr and contrast below specification. further analysis of the optical bench ccd detector output revealed noise in the signal, causing failure of 5s spec. root cause: reported connectivity issues of rl08461 on ic1386 were most likely due to site network/ap performance, but corruption of microsd card might have contributed to connectivity issues. device history lot n/a. device history batch, subcomponent name: microsd card (white), material number: 2000-0017, lot number: 2025514484, serial number: n/a. subcomponent name: optical bench, material number: 0042-1012, lot number: 1554138, serial number: (b)(6). device history review q1) service history review - are there any qns associated with the complaint device¿s most recent service report? there were no nonconformances in the most recent fsr before the complaint occurrence date which were related to the failure mode. q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode? no. q3) complaint history review - have there been any other complaints against this console? no. phr summary: there were no issues related to the complaint discovered when reviewing the product history of console (b)(6).
910 | 1220648-2026-01886 | 24218158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218158&pc=OZD | 01/27/2026 |  |  | 01/29/2026 |  |  | 01/01/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026805574 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
911 | 1220648-2026-01890 | 24218286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218286&pc=OZD | 11/19/2025 |  |  | 01/29/2026 |  |  | 11/19/2025 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026770789 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 74 year old male was admitted for an acute decompensation of a chronic heart failure. the patient arrived in a critical state and anuric. an impella cp was placed and continuous renal replacement therapy was started. on the first day of support, the patient suffered repeated episodes of ventricular tachycardia and expired. the patients underlining critical condition most likely contributed to the patients death. | [5 empty] | D3 manufacturer fax and g1 manufacturer contact fax number were added as they were omitted from the initial report that was submitted. d4 revised primary udi number as it was reported incorrectly on the initial report that was submitted. e1 added zip code extension as it was omitted from the initial report that was submitted. e4 added selection as it was omitted from the initial report that was submitted. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. peri-procedural adverse event (arrhythmia): in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
912 | 1220648-2026-01891 | 24218366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218366&pc=OZD | 01/22/2026 |  |  | 01/29/2026 |  |  | 01/26/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1174888 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male | 90 KG |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the purge cassette was inserted but not recognised. the automated impella controller (aic) displayed the following alarm: purge fluid not recognized. neither the cassette nor the purge line were flushed. the purge cassette was then replaced and the procedure could continue. the pump could then be flushed and implanted without any problems. the patient suffered no harm. the customer said that this problem has occurred several times on this aic. therefore, is sending the aic in for inspection. | [5 empty] | Correction: d6a and d6b should have been left blank on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
913 | 1220648-2026-01892 | 24218370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218370&pc=OZD | 11/19/2025 |  |  | 01/29/2026 |  |  | 11/19/2025 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026774556 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
914 | 1220648-2026-01893 | 24218389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218389&pc=OZD | 01/26/2026 |  |  | 01/29/2026 |  |  | 01/26/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1745371 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
915 | 1220648-2026-01894 | 24218494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218494&pc=OZD | 07/23/2025 |  |  | 01/29/2026 |  |  | 01/13/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025573283 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Respiratory Failure | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
916 | 1220648-2026-01897 | 24218518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218518&pc=OZD | 11/19/2025 |  |  | 01/29/2026 |  |  | 11/19/2025 | 01/29/2026 |  |  | 12/31/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025449400 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
917 | 1220648-2026-01904 | 24218585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218585&pc=OZD | 01/16/2026 | 01/29/2026 |  | 01/29/2026 |  |  | 01/16/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026798757 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 85 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Perforation of Vessels | Pericardial Effusion | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 with smartassist was inserted via the right axillary subclavian artery in a 53-year-old male patient who presented with acute decompensated chronic ischemic heart failure and cardiogenic shock. the patient was classified as scai stage d and was receiving inotropic support, vasopressors, and respiratory support via mechanical ventilation. immediately prior to leaving the procedure table, changes in waveforms were observed with an associated drop in pressure. echocardiography demonstrated a pericardial effusion. a pericardial window was performed. the patient experienced tachycardia, pericardial effusion, vessel perforation, surgical intervention, and subsequently expired. based on review of the available information, the patient¿s clinical deterioration and death were attributed to the severity of underlying cardiogenic shock, critical illness, and procedural complications associated with large-bore vascular access. the death is being reported conservatively. comprehensive review did not identify evidence suggesting a causal relationship between the impella 5.5 with smartassist device and the reported event. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. based on review of the available information, the patient¿s clinical deterioration and death were attributed to the severity of underlying cardiogenic shock, critical illness, and procedural complications associated with large-bore vascular access. the death is being reported conservatively. comprehensive review did not identify evidence suggesting a causal relationship between the impella 5.5 with smartassist device and the reported event.
918 | 1220648-2026-01906 | 24218612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218612&pc=PYX | 01/20/2026 | 01/29/2026 |  | 01/29/2026 |  |  | 01/20/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. |  |  | US |  | 1541296 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male | 62 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | A 21 year old male with unknown comorbidities and a pre support clinical state consistent with scai shock stage d was supported with an impella device for the indication of acute myocardial infarction/cardio¬genic shock. the alarm was intermittent and self clearing; however, the care team elected to exchange the aic controllers. based on the available information, the reported intermittent controller error alarm resolved following controller exchange, with no impact to patient outcome and no additional device malfunction observed. the impella cp will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange, however the exchange was performed with no consequence to the patient and support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
919 | 1220648-2026-01908 | 24218627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218627&pc=OZD | 01/16/2026 | 02/03/2026 |  | 01/29/2026 |  |  | 01/16/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026681954 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
920 | 1220648-2026-01910 | 24218688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218688&pc=OZD | 01/16/2026 | 01/29/2026 |  | 01/29/2026 |  |  | 01/16/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026801056 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
921 | 1220648-2026-01911 | 24218705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218705&pc=OZD | 01/16/2026 | 01/29/2026 |  | 01/29/2026 |  |  | 01/16/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026655826 |  | 1000100 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
922 | 1220648-2026-01914 | 24218755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218755&pc=OZD | 01/22/2026 | 01/29/2026 |  | 01/29/2026 |  |  | 01/22/2026 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026745717 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 72 KG | Non Hispanic | White | Death | Yes | No | Material Split, Cut or Torn | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 73-year-old male with cardiomyopathy underwent placement of an impella device, during which the sterile sleeve on the push button repositioning sheath was observed to be detached from the hub. the issue was identified after peel away removal on the clinical side. the push button repositioning mechanism remained intact, preventing pump movement, and a dressing was applied to avoid adherence to the sleeve. there was no harm to the patient, and no procedural delay occurred. the first pump (sn (b)(6) was implanted via the right femoral artery on (b)(6) 2026 and explanted on (b)(6) 2026, with the patient surviving to explant. a second pump (sn (b)(6) was surgically implanted via the right axillary/subclavian artery on (b)(6) 2026 and explanted the same day following withdrawal of care by the family. the patient subsequently expired, with no device related harm reported. the observed product issue consisted of a detached anti contamination sleeve; no excessive force, delivery issues, positioning issues, or interacting devices were reported. the involved pump is expected to be returned for evaluation. based on available information, the detached sleeve did not result in patient harm or contribute to the patient¿s death following withdrawal of care. the impella 5.5 will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange however the exchange was performed and no consequence to the patient was reported. the patient¿s death is conservatively reported but likely unrelated to device performance but rather the patient¿s condition and withdraw of care. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
923 | 1220648-2026-01915 | 24218769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218769&pc=OZD | 01/19/2026 | 01/29/2026 |  | 01/29/2026 |  |  | 01/22/2026 | 01/29/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026669824 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
924 | 1220648-2026-01917 | 24218849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218849&pc=OZD | 10/25/2025 |  |  | 01/29/2026 |  |  | 10/25/2025 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026736503 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 115 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult or Delayed Positioning | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. ongoing ps low alarm noted; device appears ventricularized. on (b)(6) 2025, the patient was on p4 with intermittent placement signal low alarms. echocardiography appeared to show the device positioned slightly deep. staff reported that the physician was already planning device removal later the same day and did not want to adjust the device. due to challenging anatomy during initial placement, the physician also preferred to avoid repositioning. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
925 | 1220648-2026-01919 | 24218861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24218861&pc=OZD | 08/29/2025 |  |  | 01/29/2026 |  |  | 08/29/2025 | 01/29/2026 |  |  |  |  | 01/29/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1428325 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during priming, the impella initially identified the catheter as a prior-generation device with a gray power cable, causing the prime to abort. priming was restarted and completed successfully. it was confirmed that the impella was not detected upon connection. consequently, the external power module (epm) automatically reset after 20 seconds. following the reset, the pump still was not detected. after restarting the system, the pump connected and was successfully recognized. the pump was then started at p-level 9. shortly after, the console displayed an impella flow low alarm along with a suction detection error message, which persisted throughout the procedure until the pump flow was reduced. the console was later shut down. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
926 | 1220648-2026-01676 | 24198349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24198349&pc=OZD | 01/15/2026 | 02/02/2026 |  | 01/28/2026 |  |  | 01/15/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026772521 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | Company representation | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
927 | 1220648-2026-01677 | 24198571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24198571&pc=OZD | 01/21/2026 |  |  | 01/28/2026 |  |  | 01/21/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026774696 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
928 | 1220648-2026-01678 | 24198688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24198688&pc=OZD | 01/21/2026 |  |  | 01/28/2026 |  |  | 01/21/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026705679 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A company representative reported a placement signal not reliable alarm with an impella cp. the company representative confirmed placement was checked and "ok" and the motor current (mc) was unchanged on p* with good flows and hemodynamics were stable. they have disabled pm and are monitoring mc and hemodynamics are moving forward. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
929 | 1220648-2026-01680 | 24198713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24198713&pc=OZD | 01/11/2026 |  |  | 01/28/2026 |  |  | 01/14/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026758776 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
930 | 1220648-2026-01681 | 24198737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24198737&pc=OZD | 01/13/2026 |  |  | 01/28/2026 |  |  | 01/14/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026728023 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
931 | 1220648-2026-01686 | 24198843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24198843&pc=OZD | 01/20/2026 |  |  | 01/28/2026 |  |  | 01/20/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026631953 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
932 | 1220648-2026-01689 | 24198892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24198892&pc=OZD | 01/28/2026 |  |  | 01/28/2026 |  |  | 01/17/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796289 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
933 | 1220648-2026-01688 | 24198906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24198906&pc=OZD | 01/23/2026 |  |  | 01/28/2026 |  |  | 01/23/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026700172 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 131 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was further reported that the patient was heparinized, at some point during the procedure for bypass and intervention. the bleeding was located in the chest cavity; the sternum was left open for two days post op. it has stopped now. it was not related to the device, never the less complaint warranted. | Clinician narrative an impella 5.5 device was placed in a 55-year-old male for hemodynamic support in the setting of high-risk cardiac surgery, including coronary artery bypass grafting with aortic and mitral valve intervention. the primary indication for support was postcardiotomy cardiogenic shock and low cardiac output syndrome. the patient had a history of prior coronary artery bypass grafting and known coronary artery disease and required inotropic and vasopressor support, as well as mechanical ventilation for respiratory support. the impella 5.5 was inserted via direct aortic access. during emergence, device placement was noted to be approximately 10 mean arterial pressure points discordant with the arterial line and cardiopulmonary bypass flows. post-procedure, the patient experienced a major bleeding event requiring packed red blood cell transfusion. the bleeding was assessed by the clinical team and determined not to be related to the impella 5.5 device. the care team elected to continue using the device. based on comprehensive review, the reported bleeding and blood transfusion are consistent with perioperative bleeding risks associated with complex cardiac surgery, prior surgical history, and critical illness. comprehensive review did not identify evidence suggesting a causal relationship between the impella 5.5 device and the reported patient events. the patient survived. |  |  |  |  | B5 updated with additional information. | This is one of two related reports. this report represents the impella 5.5 and another report was submitted to represent the automated impella controller. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
934 | 1220648-2026-01690 | 24198911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24198911&pc=OZD | 01/19/2026 |  |  | 01/28/2026 |  |  | 01/19/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026757892 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 224 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the purge flow was 2.3 and purge pressure was in the 900¿s. tissue plasminogen activator was initiated. there was no patient harm. | [5 empty] | The device explant date was provided therefore d6b was updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
935 | 1220648-2026-01685 | 24198915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24198915&pc=OZD | 01/26/2026 |  |  | 01/28/2026 |  |  | 01/26/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1368535 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 131 KG | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The purge cassette ended up swelling on the automated impella controller (aic) to aic transfer and popping. | The complainant reported that a patient was placed on an impella 5.5 for mechanical circulatory support. the nurse switch from one automated impella controller to another console. the nurse switched over the white impella cord first and then the purge cassette. the purge disc had started swelling and getting fat. the nurse was unable to get into the new console and it began leaking. a new cassette was recommended. the patient was off purge for 20 minutes while someone retrieved a new purge cassette tubing. the patient's outcome was stable. this is one of two related reports. this report represents the automated impella controller and another report was submitted to represent the impella 5.5. |  |  |  |  | B5 updated with additional information. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
936 | 1220648-2026-01694 | 24199031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24199031&pc=OZD | 01/08/2026 |  |  | 01/28/2026 |  |  | 01/20/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026734056 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp device was inserted via the left femoral artery in a 61-year-old female patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs). when switching from iabp to impella cp support, the iabp was removed and left coronary artery (lca) contrast was performed prior to impella cp insertion. during this transition period, the patient converted to atrial fibrillation (af). following impella cp initiation, circulatory dynamics worsened. the patient progressed to pulseless electrical activity (pea), and veno-arterial extracorporeal membrane oxygenation (va ecmo) was initiated. an impella cp smartassist (cp sa) was subsequently inserted for additional support. arrhythmias were noted to occur several times during impella cp operation. the treating physician reported no association between the impella cp and the onset of af. the patient survived following explant. the reported arrhythmia occurred in the setting of acute myocardial infarction, cardiogenic shock, hemodynamic instability, and escalation to va-ecmo, all of which are well-recognized clinical conditions associated with atrial and ventricular arrhythmias. based on the available information and provider assessment, the arrhythmia is most plausibly related to the patient¿s underlying critical condition and clinical course. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
937 | 1220648-2026-01697 | 24199127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24199127&pc=OZD | 01/18/2026 |  |  | 01/28/2026 |  |  | 01/20/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026759951 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | N impella cp device was inserted via the right femoral artery in a 55-year-old male patient presenting in cardiomyopathy and cardiogenic shock (cgs) due to myocarditis-like symptoms, with scai stage d shock, on multiple inotropes and vasopressors, and was ventilated for respiratory support prior to initiation of support. after the impella placement, the support level was set to p-9 for approximately three hours. hemolysis was detected by hematuria, so the position was adjusted, the support level was lowered, and dobutamine and pitressin were added. continuous hemodiafiltration (chdf) was initiated and haptoglobin was administered, but the hemolysis continued. the patient was managed at p-2, but the hemolysis worsened on the third day of use, so the physician removed the impella. the post-procedure outcome is unknown. hemolysis is listed as a potential adverse event and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
938 | 1220648-2026-01699 | 24199154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24199154&pc=OZD | 01/19/2026 |  |  | 01/28/2026 |  |  | 01/21/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782839 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Restricted Flow rate | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 69-year-old male with cardiomyopathy (scai stage c) underwent pci with impella cp support following ecpr on (b)(6) 2026. the patient experienced recurrent vt/vf, and vt ablation was performed under ecmo support on (b)(6). during ablation, suction events occurred when the pump level was increased from p 6 to p 4, despite no evidence of volume depletion on swan ganz parameters and no cannula kinks. suction resolved only when the pump was reduced to p 2, leading clinicians to suspect thrombus adhesion at the inlet area. systemic heparin had not been used during the preceding six days, and echocardiography did not visualize thrombus; however, thrombus was confirmed on the removed pump, with clots observed on the outlet portion. the device was not removed due to a malfunction but following successful weaning, with ecmo discontinued on (b)(6) and the impella cp explanted on (b)(6). no replacement pump was required, and the patient remained hemodynamically stable. the pump will be returned for evaluation. the thrombosis is likely related to no heparin for six days. it is unknown if the vt was pre-existing conditon however an ablation was performed. the low flow and suction events were likely related to pump associated thrombosis rather than device failure, as the device functioned at reduced support levels and was removed only after clinical weaning. the patient survived explant. the impella cp will be returned for evaluation. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
939 | 1220648-2026-01700 | 24199177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24199177&pc=OZD | 01/21/2026 |  |  | 01/28/2026 |  |  | 01/21/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798654 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
940 | 1220648-2026-01702 | 24199197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24199197&pc=OZD | 01/20/2026 |  |  | 01/28/2026 |  |  | 01/20/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026786895 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 63 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that tissue plasminogen activator was started for high purge pressure. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
941 | 1220648-2026-01703 | 24199213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24199213&pc=OZD | 01/21/2026 | 02/03/2026 |  | 01/28/2026 |  |  | 01/21/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802661 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 87 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | An impella 5.5 device was inserted surgically into the right subclavian artery in a 59-year-old male patient with a past medical history of coronary artery disease (cad). the patient presented in cardiomyopathy, scai stage c shock and was on an intra-aortic balloon pump (iabp), prior to initiation of support. while the patient was in the intensive care unit (icu), the patient was having intermittent short 5 beat runs of ventricular tachycardia (vt). the nurse stated that the patient was dry and was receiving intravenous (iv) fluids, which was decreasing the amount of vt. an echocardiogram was ordered to assess position. the post-procedure outcome is unknown. ventricular tachycardia is a potential adverse event during support with the impella 5.5 and can be associated with mechanical irritation from device malpositioning; however, cad and cardiomyopathy, are leading contributors of vt due to damage to the heart muscle that can cause the heart to beat too fast. | [5 empty] | D6b updated. the investigation is still ongoing. | A6 is unknown. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | Investigation summary: the device was not returned for evaluation. in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history: the pump passed all the post sterile inspection checks.
942 | 1220648-2026-01701 | 24199218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24199218&pc=OZD | 01/22/2026 |  |  | 01/28/2026 |  |  | 01/22/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026800145 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Priming Problem | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
943 | 1220648-2026-01698 | 24199230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24199230&pc=OZD | 01/22/2026 |  |  | 01/28/2026 |  |  | 01/22/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1484655 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | Yes | Priming Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Clinical rationale/review: ip2 aic. an impella 5.5 was placed to support a patient who had previously been supported by an intra-aortic balloon pump (iabp), ecmo, and respiratory support with a vent after being admitted with known cardiogenic shock and cardiomyopathy, scai stage d shock. the pump was placed but removed when the pump failed to prime effectively. the team troubleshot by changing the purge fluid, the cp pump, and finally the aic itself was replaced. the new aic and new cp pump support the patient. | [5 empty] | This is one of two related reports. this report represents the automated impella controller and another report was submitted to represent the impella 5.5. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
944 | 1220648-2026-01706 | 24199488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24199488&pc=OZD | 01/21/2026 | 02/24/2026 |  | 01/28/2026 |  |  | 01/22/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801516 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
945 | 1220648-2026-01707 | 24199574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24199574&pc=OZD | 11/13/2025 |  |  | 01/28/2026 |  |  | 11/13/2025 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026743362 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Malposition of Device | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient experienced hemodynamic instability following cardiac surgery, including self-limiting premature ventricular contractions and episodes of ventricular tachycardia requiring treatment. the patient could not be weaned from cardiopulmonary bypass and required an impella 5.5 for mechanical circulatory support. postoperative arrhythmias were associated with shallow device positioning and resolved after repositioning, performance-level adjustments, and volume resuscitation. the device was later removed without access complications. imaging showed significantly reduced ejection fraction, and the patient required intensive postoperative management consistent with a serious injury. | [5 empty] | No product returned. hypotension: the root cause of the injury was unable to be determined due to insufficient clinical details. tachycardia: in order to make a cause determination, all of the clinical details outlined would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position: the cause of the device in wrong position was unable to be determined due to insufficient clinical details. device passed all post sterile inspection checks.
946 | 1220648-2026-01710 | 24199677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24199677&pc=OZD | 01/16/2026 |  |  | 01/28/2026 |  |  | 01/22/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026776782 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale/review: an impella cp was placed to support the 51 year old patient admitted in with acute myocardial infarction and cardiogenic shock in scai stage c shock. the patient's underlying medical history was not shared and is unknown. the pump was placed, but the leg became ischemic on day 2 of the support. both legs had flow issues but, the leg accessed by the impella was given a perfusion catheter for flow restoration. the perfusion sheath remained until the pump was explanted on day 4. on the following day the team did perform a surgical limb amputation to leg accessed by the impella. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
947 | 1220648-2026-01711 | 24199836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24199836&pc=OZD | 01/22/2026 |  |  | 01/28/2026 |  |  | 01/22/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1186353 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella device experienced an automated impella controller (aic) battery alarm during use. it was reported that despite the aic having a fully charged battery and plugged into a working outlet, there was an alarm for the battery. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the device is reportedly available for return; however, as of the time of this report the device has not yet been returned for evaluation. | [5 empty] | Section a. patient information is unknown. device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
948 | 1220648-2026-01716 | 24200066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24200066&pc=OZD | 01/22/2026 |  |  | 01/28/2026 |  |  | 01/22/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1326728 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
949 | 1220648-2026-01718 | 24200200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24200200&pc=OZD | 01/16/2026 |  |  | 01/28/2026 |  |  | 01/16/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1080531 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 65 KG | Non Hispanic | White | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was attempted via the right femoral artery in a 77-year-old female patient presenting with acute myocardial infarction complicated by cardiogenic shock. overnight, the automated impella controller (aic 1115) demonstrated intermittent "controller failure" alarms without associated pump failure. nursing staff reported that the error would appear and disappear several times each hour. eventually, both aortic and left ventricular signals were lost despite ongoing pump function, and the "controller failure" alarm remained active. based on these findings, the clinical team elected to exchange the aic. the reported events are controller failure, medical device removal, and hemodynamic instability. the impella cp will be conservatively reported for serious injury due to hemodynamic instability associated with the temporary interruption of hemodynamic support during the attempted placement and subsequent removal; however, the removal was completed without consequence to the patient, and no known adverse events were reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
950 | 1220648-2026-01717 | 24200243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24200243&pc=OZD | 01/20/2026 |  |  | 01/28/2026 |  |  | 01/20/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1130827 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission
951 | 1220648-2026-01428 | 24200350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24200350&pc=OZD | 12/30/2025 |  |  | 01/28/2026 |  |  | 12/30/2025 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1103726 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 127 KG | Non Hispanic | White | Injury | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Other | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: an impella cp device was inserted into the right femoral artery in a 51-year-old male patient with a past medical history of coronary artery disease (cad) and renal insufficiency, presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), scai stage e shock, on multiple inotropes and vasopressors, and was on a ventilator for respiratory support prior to initiation of support. during the procedure, the automated impella controller (aic) had an error message and was successfully switched to a loaner console. the post-procedure outcome is survived at explant. the automated impella controller will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange. | [5 empty] | Device status. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. related medical device reports: this impella automated impella controller device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp.
952 | 1220648-2026-01723 | 24200715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24200715&pc=OZD | 01/07/2026 |  |  | 01/28/2026 |  |  | 01/07/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797223 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
953 | 1220648-2026-01725 | 24201069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24201069&pc=OZD | 01/17/2026 |  |  | 01/28/2026 |  |  | 01/18/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026790320 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
954 | 1220648-2026-01733 | 24202286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24202286&pc=OZD | 12/05/2025 |  |  | 01/28/2026 |  |  | 12/05/2025 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026769346 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Decrease in Pressure | Malposition of Device |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
955 | 1220648-2026-01735 | 24202511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24202511&pc=OZD | 10/26/2025 |  |  | 01/28/2026 |  |  | 10/26/2025 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744120 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the patient developed ventricular tachycardia, with the relationship to the pump described as ¿possible.¿ the patient experienced multiple episodes of polymorphic vt. during the initial episode, the impella was suspected to be positioned too deep while operating at p6, so support was reduced due to possible suction events. despite this, vt recurred. a point-of-care ultrasound (pocus) echocardiogram was performed, and the impella required repositioning approximately 1 cm. the patient was also started on amiodarone, after which no further vt episodes occurred. filling pressures improved, and impella support was increased back to p4. electrolytes were repleted as needed. the patient survived the procedure and was reported to be stable. additional information indicated that the patient remained status post mitral valve percutaneous coronary intervention (mv pci), with no acute events overnight. impella support was weaned to p2, and milrinone was initiated due to hypotension. | [5 empty] | Ppae: (hypotension): the cause of the injury was not determined due to insufficient clinical details. (tachycardia): in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. the pump sn (b)(6) passed all the post sterile inspection checks.
956 | 1220648-2026-01736 | 24202557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24202557&pc=OZD | 01/23/2026 | 02/02/2026 |  | 01/28/2026 |  |  | 01/23/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766897 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
957 | 1220648-2026-01738 | 24202870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24202870&pc=OZD | 01/23/2026 |  |  | 01/28/2026 |  |  | 01/23/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026780917 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 94 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. it was reported that upon arrival to the intensive care unit, it was noticed that the placement signal was dampened and did not correlate with the arterial blood pressure. placement signal reading 60's/50's and arterial line reading 90's/70's. no alarms present. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
958 | 1220648-2026-01740 | 24203140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24203140&pc=OZD | 01/23/2026 |  |  | 01/28/2026 |  |  | 01/23/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, CA |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1595731 |  | 0042-0000-CA |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Open or Close | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a purge release button got stuck and the door would not latch closed. there was some purge solution in the cassette space and was wiped away. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
959 | 1220648-2026-01746 | 24204738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24204738&pc=OZD | 10/03/2025 |  |  | 01/28/2026 |  |  | 10/03/2025 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1080191 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
960 | 1220648-2026-01751 | 24205987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24205987&pc=OZD | 01/19/2026 | 01/28/2026 |  | 01/28/2026 |  |  | 01/20/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026772397 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Decrease in Pressure | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
961 | 1220648-2026-01756 | 24206095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206095&pc=OZD | 01/14/2026 | 02/03/2026 |  | 01/28/2026 |  |  | 01/14/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779564 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
962 | 1220648-2026-01758 | 24206113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206113&pc=OZD | 01/14/2026 |  |  | 01/28/2026 |  |  | 01/14/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026757824 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
963 | 1220648-2026-01763 | 24206351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206351&pc=OZD | 01/21/2026 | 01/28/2026 |  | 01/28/2026 |  |  | 01/21/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026795503 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
964 | 1220648-2026-01765 | 24206372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206372&pc=OZD | 01/19/2026 | 01/28/2026 |  | 01/28/2026 |  |  | 01/21/2026 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026795477 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
965 | 1220648-2026-01766 | 24206384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206384&pc=OZD | 01/21/2026 | 01/28/2026 |  | 01/28/2026 |  |  | 01/22/2026 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | ABIOMED, INC. |  |  | US |  | 2026763287 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female | 61 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Ischemia | Hemodynamic instability | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 70 year old female with an impella cp placed on (b)(6) 2026 for support during acute myocardial infarction complicated by cardiogenic shock developed right leg ischemia at the femoral access site. on (b)(6) 2026 at approximately 21:42, the patient returned to the operating room due to the cold, ischemic limb. the impella cp (sn (b)(6)) was explanted and replaced with an impella 5.5 (sn (b)(6)) via the left axillary artery. following the procedure, the patient returned to the icu intubated and in critical hemodynamic condition, requiring high dose norepinephrine support (1.0 g/kg/min). laboratory findings included significant anemia requiring multiple blood transfusions and an elevated lactate level of 8 mmol/l, indicating ongoing circulatory insufficiency. despite imaging and bedside assessments confirming that the impella device remained in correct position, the patient¿s mean arterial pressure declined to approximately 18 mmhg. given the worsening multiorgan instability, the clinical team initiated discussions to withdraw care. the reported events¿ischemia, anemia, hemodynamic instability, and the required device replacement¿appear to be related to the patient¿s underlying severe cardiogenic shock and vascular compromise rather than a device malfunction. the patient remains on support at the time of reporting | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Device was discarded. e4 was inadvertently omitted on the initial manufacturer device report.
966 | 1220648-2026-01770 | 24206479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206479&pc=OZD | 08/23/2025 |  |  | 01/28/2026 |  |  | 08/23/2025 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during preparation for impella support, the automated impella controller (aic) failed to start up. when powered on, the system displayed the error message, ¿system self check failed. change console immediately.¿ the unit was removed from service and not used. a replacement aic was used instead. | [5 empty] | E1. initial reporter first and last name is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
967 | 1220648-2026-01771 | 24206488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206488&pc=OZD | 09/14/2025 |  |  | 01/28/2026 |  |  | 09/16/2025 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026711663 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Appropriate Term/Code Not Available | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that extracorporeal membrane oxygenation (ecmo) was initiated due to hemodynamic collapse. a second impella insertion was considered however, cardiac rupture occurred, preventing placement. as the rupture happened before any impella related steps such as guidewire or pigtail insertion was performed. the death was determined to be unrelated to the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
968 | 1220648-2026-01780 | 24206517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206517&pc=OZD | 10/17/2025 |  |  | 01/28/2026 |  |  | 10/17/2025 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026721276 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female | 81 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was supported with a five-and-a-half mechanical circulatory support device when a peri-procedural adverse event occurred. during the procedure, blood was observed leaking around the graft locks despite the locks being applied in accordance with the instructions for use. the patient subsequently developed renal failure, and an additional device was required to maintain adequate hemodynamic support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
969 | 1220648-2026-01788 | 24206540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206540&pc=OZD | 10/22/2025 |  |  | 01/28/2026 |  |  | 10/22/2025 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026705469 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 119 KG | Non Hispanic | White | Injury | Yes | Yes | Device Damaged by Another Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Hemodynamic instability | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient was implanted with an impella 5.5 device for mechanical circulatory support during an ablation procedure. significant oozing was noted from the right axillary impella access site following heparin administration, with activated clotting times ranging from 260 to 375 seconds. the patient required crystalloid, 5% albumin, and one unit of packed red blood cells, along with a pressure dressing and hemostatic agents. intermittent hemodynamic instability occurred, requiring norepinephrine, phenylephrine, and fluid boluses, along with adjustments in impella support. during the procedure at approximately 12:15, a tacticath ablation catheter became lodged in the inlet area of the impella. device support was reduced, and multiple attempts to remove the catheter were unsuccessful until it dislodged shortly before surgical assistance arrived. inspection revealed a portion of the catheter appeared to be missing, with several loose fragments identified, though it could not be confirmed whether all pieces were accounted for. a second catheter was used and briefly became stuck in the inlet but was removed intact with less difficulty. procedure time totaled 50 minutes with 49 ablation lesions. support was returned to p7, and the procedure was completed. the patient developed a new left bundle branch block and may require future biventricular pacemaker placement. additional patient outcome information is not available in the provided report. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
970 | 1220648-2026-01793 | 24206552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206552&pc=OZD | 10/22/2025 |  |  | 01/28/2026 |  |  | 10/22/2025 | 01/28/2026 |  |  |  |  | 01/28/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1849082 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 137 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the aic generated a ¿controller error¿ alarm, and the ao signal was temporarily absent. the event persisted for approximately one minute, after which the system self-resolved without requiring intervention. no recurrence of the alarm or signal loss has been observed since. as a precautionary measure, a backup console was placed in the room to ensure immediate availability of support in the event of unexpected pump stoppage. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
971 | 1220648-2026-01795 | 24206560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24206560&pc=OZD | 10/24/2025 |  |  | 01/28/2026 |  |  | 10/24/2025 | 01/28/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1631552 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Alarm System | Pumping Stopped | Failure to Read Input Signal | Application Program Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
972 | 1220648-2026-01540 | 24187593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24187593&pc=OZD | 01/06/2026 |  |  | 01/27/2026 |  |  | 01/14/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026760833 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male | 66 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Multiple Organ Dysfunction Syndrome | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Impella 5.5 implanted in a 72-year-old male for acute myocardial infarction and cardiogenic shock. there were no reported events during device support. the patient subsequently died due to multiple organ failure. review of the event did not identify a device anomaly. the device is being conservatively reported as a death; however, the device is unlikely to have contributed to the patient¿s death, which was most likely due to the underlying critical clinical condition. | [5 empty] | Date of death the exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. investigation summary ==> ppae (organ failure): the cause of the injury was not determined due to insufficient clinical details. device history lot ==> device lot: 1977557 device history review ==> the pump sn 630483 passed all the post sterile inspection checks.
973 | 1220648-2026-01541 | 24187860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24187860&pc=OZD | 08/06/2025 |  |  | 01/27/2026 |  |  | 01/14/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025586684 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
974 | 1220648-2026-01546 | 24188007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24188007&pc=PYX | 01/06/2026 |  |  | 01/27/2026 |  |  | 01/07/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2019074 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the automated impella controller exhibited a warning that it should be exchanged for a new controller. no patient harm was reported. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
975 | 1220648-2026-01547 | 24188021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24188021&pc=OZD | 01/21/2026 |  |  | 01/27/2026 |  |  | 01/21/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1116070 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 86 KG |  |  | Injury | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Clinical rationale: a 67-year-old patient with a past medical history of coronary artery disease was supported with an impella cp device for the indication of acute myocardial infarction and cardiogenic shock. on (b)(6) 2026, during use, the aic (sn (b)(6)) generated a yellow controller error alarm. shortly afterward, both the aortic placement (ao) signal and left ventricular (lv) signal went blank, while the display continued to show a flow of 3.1 l/min and a present motor current waveform. urine output at the time was greater than 30 cc/hr and pink red in color. the clinical team exchanged the aic for another controller. based on the available information, the reported controller error and loss of ao/lv waveforms were resolved by exchanging the aic, with no confirmed device malfunction of the pump itself at this time. the patient remains on impella support, and no long-term clinical impact related to the reported event has been determined. product return and data analysis are pending evaluation. the outcome to hematuria was unknown at the time of reporting. | [5 empty] | This is one of two related reports. this report represents the automated impella controller and another report was submitted to represent the pump. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
976 | 1220648-2026-01549 | 24188081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24188081&pc=OZD | 01/14/2026 |  |  | 01/27/2026 |  |  | 01/14/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026772340 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
977 | 1220648-2026-01553 | 24188176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24188176&pc=OZD | 01/15/2026 |  |  | 01/27/2026 |  |  | 01/15/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796232 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Decrease in Pressure | Increase in Pressure | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
978 | 1220648-2026-01555 | 24188189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24188189&pc=OZD | 01/15/2026 |  |  | 01/27/2026 |  |  | 01/15/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779707 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
979 | 1220648-2026-01560 | 24188354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24188354&pc=OZD | 01/14/2026 |  |  | 01/27/2026 |  |  | 01/14/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026728706 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that during implantation of an impella cp there was a kink on the 14 fr introducer and when the physician tried to pull the impella the cannula broke. a second impella was implanted successfully. no patient harm was reported. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
980 | 1220648-2026-01564 | 24188466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24188466&pc=OZD | 01/15/2026 |  |  | 01/27/2026 |  |  | 01/15/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801108 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
981 | 1220648-2026-01565 | 24188530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24188530&pc=OZD | 01/15/2026 |  |  | 01/27/2026 |  |  | 01/15/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026791578 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced high purge pressure. a trolley wheel was found to be on the impella cable, generating the alarm. however, after the trolley wheel was moved the purge pressure did not resolve. the patient survived. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
982 | 1220648-2026-01566 | 24188561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24188561&pc=OZD | 01/19/2026 |  |  | 01/27/2026 |  |  | 01/21/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766410 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
983 | 1220648-2026-01567 | 24188734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24188734&pc=OZD | 08/28/2025 |  |  | 01/27/2026 |  |  | 01/15/2026 | 01/27/2026 | [6 empty] | P140003 |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026693539 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Inflammation | Nerve Damage | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
984 | 1220648-2026-01569 | 24189141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24189141&pc=OZD | 01/22/2026 |  |  | 01/27/2026 |  |  | 01/22/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026770585 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Inaccurate Flow Rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Clinician narrative: an impella 5.5 was surgically inserted via direct left aortic access in a 66-year-old male who presented with acute myocardial infarction complicated by cardiogenic shock. the patient had a history of prior coronary artery bypass grafting and known coronary artery disease and was classified as scai stage e. the patient was supported with extracorporeal membrane oxygenation, inotropes, vasopressors, and mechanical ventilation for respiratory support. during support, the patient experienced major bleeding and progressive hemodynamic instability in the setting of critical illness and advanced cardiogenic shock. clinical documentation noted elevated and highly saturated pump flows during support. additional information received states that the treating provider assessed the patient¿s clinical deterioration and determined that the bleeding and subsequent decline were attributable to a retroperitoneal hemorrhage related to the ecmo cannulation site and ischemic bowel, rather than impella-related factors. despite ongoing management, the patient¿s condition continued to deteriorate. review of the event indicates that the reported major bleeding and death occurred in the context of profound cardiogenic shock, prior cardiac surgery, large-bore vascular access associated with ecmo cannulation, anticoagulation, and suspected ischemic gut. comprehensive review did not identify evidence suggesting a causal relationship between the impella 5.5 device and the reported events. the patient expired while on support, and the death is being reported conservatively. | The complainant reported a patient implanted with an impella 5.5 experienced saturated pump flows during support. it was reported that on the second day of support, the left ventricle signal was greater than the aorta placement signal. the aorta placement signal was noted to be 76/65 mmhg, and the left ventricle placement signal was 109/71 mmhg. it was observed that the flow rate at p4 was higher than expected at 4.3 l/min, with expected rates per instructions for use of 1.9-3.3 l/min. additionally, it was noted that the motor current was 278/227 ma. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. at the time of writing this report, the patient remains on impella 5.5 support. |  |  |  |  | This follow-up report is being submitted to update b1, b2, b5, d6b, and h6 codes for new information that was received on 27-jan-2026. b1 is being updated to include serious injury and death. b2 is being updated to include death and date of death. b5 is being updated to include new information that was not included in the initial report and clinical narrative. the revised b5 provides a complete event narrative. d6b is being updated to include explant date. h6. health effect - clinical code for e0506 (hemorrhage/blood loss/bleeding) has been added. h6. health effect - impact code f02 (death) has been added. | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
985 | 1220648-2026-01571 | 24189246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24189246&pc=OZD | 09/25/2025 |  |  | 01/27/2026 |  |  | 09/26/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026711402 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding complications at all puncture sites, not limited to the impella access site. according to the attending physician, the patient had poor liver function, which may have contributed to the bleeding. the patient expired. according to the attending physician there is no direct connection between the death and the impella therapy. | [5 empty] | No product was returned for investigation. major bleed: patient suffered from bleeding complications out of every single puncture site, not only from the impella puncture site. the cause of the bleeding is determined to be patient condition because the patient was bleeding from multiple locations. pump # (b)(4) passed all post sterile inspection checks.
986 | 1220648-2026-01573 | 24189605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24189605&pc=OZD | 09/24/2025 |  |  | 01/27/2026 |  |  | 09/24/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026731294 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 111 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support with the impella 5.5, the optical sensor demonstrated intermittent signal loss, with the aortic waveform decreasing in reliability over approximately one hour and eventually failing completely, resulting in persistent ¿placement signal not reliable¿ alerts. catheter inspection revealed no visible damage, and the nurse and patient confirmed the external portion of the catheter was intact. the clinical team managed the device using motor current and flow parameters. after a purge tubing change, the aortic signal intermittently returned but repeatedly disappeared, each time temporarily restored by pressing the gray button on the automated impella controller; this pattern continued despite confirmed secure connections. the device remained in use. additionally, the patient required cardioversion for persistent atrial fibrillation following implant, and later experienced a nosebleed while on bivalirudin, with anticoagulation subsequently held and an ear, nose, and throat consultation pending. | [5 empty] | Optical sensor issue: clinical narrative reports intermittent placement signal not reliable (psnr) alarms between 10/oct and 14/oct. pressing on automated impella controller (aic) buttons was noted to help resolve the issue intermittently. no damages or kinks were noted on the catheter. data log review confirmed the intermittent psnr alarms between 10/oct/2025 and 16/oct/2025. during this time, contrast would oscillate between +/- 3200, signal not reliable (snr) would drop to zero, lamp would max out, and gain/light would drop. the optical signal issues resolved for the most part on 16/oct. the pump was not returned for investigation. based on the intermittent nature of the pattern noted in data logs, the cause of the optical signal issue was determined to be related to the console. paroxysmal atrial fibrillation: clinical details report that the patient had to be cardioverted to troubleshoot a-fib shortly after pump insertion. in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. epistaxis: clinical details report that the patient was bleeding from the nose on (b)(6) 2025. the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided.
987 | 1220648-2026-01574 | 24189620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24189620&pc=OZD | 10/07/2025 |  |  | 01/27/2026 |  |  | 10/08/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026729998 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the aortic (ao) pressure sensor demonstrated progressive drift and eventually became sufficiently negative that the device generated an alarm indicating an unreliable placement signal. despite the sensor drift, device flow parameters continued to correlate appropriately with motor current. the pump was removed the previous evening without any knowledge, and retrieval for analysis will not be possible. there was no known injury or harm to the patient. additional information stated that the first pump was used and later removed without replacement. during its use, there were concerns regarding device sensing, recurrent suction events, and the need for repositioning due to suspected incorrect placement. possible thrombus formation was also reported, and the patient was simultaneously supported with extracorporeal membrane oxygenation. the patient required and received a blood transfusion during the course of treatment. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
988 | 1220648-2026-01575 | 24189677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24189677&pc=OZD | 10/08/2025 |  |  | 01/27/2026 |  |  | 10/08/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026733784 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Malposition of Device | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the lactate levels were observed to be progressively increasing, attributed to ongoing infection with mrsa, escherichia coli, and staphylococcus species additional information states that the initial pump required escalation to a higher-support device. during use of the second pump, clinical concerns were documented, including the need for blood transfusion, management of bleeding and hemostasis, evidence of infection, and concerns related to device positioning. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
989 | 1220648-2026-01578 | 24190092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190092&pc=OZD | 01/21/2026 |  |  | 01/27/2026 |  |  | 01/21/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026801043 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | Patient Device Interaction Problem |  |  |  |  | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
990 | 1220648-2026-01580 | 24190371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190371&pc=PYX | 11/05/2025 |  |  | 01/27/2026 |  |  | 11/05/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1881868 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the clinical staff noted an ongoing device position in aorta alarm. a bedside echocardiogram demonstrated that the device was appropriately positioned. the patient remained stable with normal flows, pulsatile motor current, and purge pressure correlating with the arterial line. the user¿s questions were addressed, and guidance was provided on disabling the audio alarm. no clinical impact to the patient was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
991 | 1220648-2026-01581 | 24190373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190373&pc=PYX | 11/22/2025 |  |  | 01/27/2026 |  |  | 11/23/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1577697 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. during impella cp support in the icu, the automated impella controller (aic) identified as (b)(6) displayed a ¿controller failure¿ alarm and then shut off. the impella cp was immediately connected to a different aic, and support was successfully resumed without interruption. the original aic (b)(6) was removed from the patient care area. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
992 | 1220648-2026-01582 | 24190406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190406&pc=OZD | 11/23/2025 |  |  | 01/27/2026 |  |  | 11/23/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1739661 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. during routine use of the automated impella controller, an rn attempted to load a new purge cassette and initiate priming. the console did not recognize the purge cassette, and priming could not be completed. no additional details were provided, and no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | D2a and d2b revised as they were removed incorrectly on the initial report that was submitted. e1 added zip code extension as it was omitted from the initial report that was submitted. h6 code c20 was reported incorrectly on the initial report and should of been reported as c21 as the investigation was pending. type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. data review: the console logs were partially consistent with what was reported in the complaint. although it was reported that the console could not recognize the purge cassette, the logs showed that the console actually had trouble detecting the purge disc. the logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. device analysis: the console was sent back to field service for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. root cause: the console did not have any evident purge cassette recognition issues. however, the purge disc flag on the console was observed to be broken and was the root cause of the console¿s inability to detect the purge disc. the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly. the console did not have any evident purge cassette recognition issues. product history review summary: there were no issues related to the complaint discovered when reviewing the product history of the console.
993 | 1220648-2026-01584 | 24190409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190409&pc=OZD | 01/14/2026 |  |  | 01/27/2026 |  |  | 01/15/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796200 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
994 | 1220648-2026-01585 | 24190416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190416&pc=PYX | 11/24/2025 |  |  | 01/27/2026 |  |  | 11/24/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1496398 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. biomed staff member reported that the sensor connection port on the front of the console was not functioning correctly. no additional details were provided, and no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
995 | 1220648-2026-01586 | 24190433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190433&pc=OZD | 10/20/2023 |  |  | 01/27/2026 |  |  | 11/13/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that, per an engineering request, a previous event warranted initiation of a complaint investigation for the console used during mechanical circulatory support. during support with the impella cp device, the automated controller experienced a placement signal error, and it was reported that the placement signal was lost. no additional information was provided regarding corrective actions taken at the time of the event. this has been classified as a reportable malfunction. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
996 | 1220648-2026-01587 | 24190442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190442&pc=OZD | 11/24/2025 |  |  | 01/27/2026 |  |  | 11/24/2025 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1885817 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
997 | 1220648-2026-01588 | 24190445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190445&pc=PYX | 11/13/2025 |  |  | 01/27/2026 |  |  | 11/13/2025 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1669290 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
998 | 1220648-2026-01589 | 24190532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190532&pc=OZD | 12/05/2023 |  |  | 01/27/2026 |  |  | 11/14/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1133831 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during support with the mechanical circulatory support device, the automated controller experienced a placement signal error and the placement signal was lost. no information on corrective actions was provided. per an engineering request, a previous event warranted initiation of a complaint investigation of the console related to the loss of placement signal. this has been classified as a reportable malfunction. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
999 | 1220648-2026-01591 | 24190558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190558&pc=OZD | 11/14/2025 |  |  | 01/27/2026 |  |  | 11/15/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1127440 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during support with the impella cp device, the automated controller experienced a placement signal error. it was noted that the controller was displaying a mean arterial pressure approximately ten points lower than the arterial line measurement. no additional information regarding corrective actions was provided. this has been classified as a reportable malfunction. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1000 | 1220648-2026-01590 | 24190562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190562&pc=OZD | 11/25/2025 |  |  | 01/27/2026 |  |  | 11/25/2025 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1098589 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission
1001 | 1220648-2026-01592 | 24190570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190570&pc=PYX | 11/11/2025 |  |  | 01/27/2026 |  |  | 11/16/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1768515 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the mechanical circulatory support specialist observed intermittent controller error alarms on the automated controller. the specific alarm color was not known. the alarms resolved without intervention, and the patient remained stable with consistent flows throughout. the clinical team elected to replace the controller. this event has been classified as a reportable malfunction. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1002 | 1220648-2026-01593 | 24190578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190578&pc=OZD | 11/24/2025 |  |  | 01/27/2026 |  |  | 11/25/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. during patient transport to ir, air entered the system. the rn attempted to de-air the system, but the purge was not flowing forward. after repeating the de-airing steps, an aic controller failure alarm occurred and then self-resolved after approximately 4 seconds. the rn was subsequently able to successfully de-air the system, and the patient remained on the same controller. clinical support contacted the csc, who advised not to change controllers unless another failure message occurs. a standby aic was placed outside the patient¿s room, and the rn was instructed not to switch aics during active alarms. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1003 | 1220648-2026-01594 | 24190583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190583&pc=PYX | 11/16/2025 |  |  | 01/27/2026 |  |  | 11/16/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1762746 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during support with the impella cp device, the automated controller displayed intermittent static on the screen after it was unplugged from alternating current power. the issue resembled brief visual interference similar to an older television screen. the display distortion was intermittent, did not persist, and did not recur after the initial report. no corrective actions were taken, and no impact to the patient was noted. this has been classified as a reportable malfunction. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1004 | 1220648-2026-01595 | 24190584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190584&pc=OZD | 11/26/2025 |  |  | 01/27/2026 |  |  | 11/26/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1401378 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. a ¿battery power low¿ alarm occurred when the automated impella controller (aic) was unplugged from ac power. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1005 | 1220648-2026-01596 | 24190598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190598&pc=OZD | 11/17/2025 |  |  | 01/27/2026 |  |  | 11/17/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1084303 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during support with the mechanical circulatory support device, the automated controller experienced intermittent placement signal issues and controller error alarms. the clinical team noted that the aortic and left ventricular placement signals were unreliable at times, although no alarms initially accompanied these fluctuations. the signals returned briefly and then were lost again, at which point a controller error alarm appeared. the waveforms later returned and remained stable until near the end of the procedure, when signal loss recurred and another controller error alarm was triggered. the controller was replaced, after which the signals remained stable for the remainder of the procedure. the event was attributed to a controller malfunction and has been classified as a reportable malfunction. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1006 | 1220648-2026-01597 | 24190607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190607&pc=OZD | 11/26/2025 |  |  | 01/27/2026 |  |  | 11/26/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1087127 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Hispanic |  | Malfunction | Yes | Yes | Application Program Problem | [6 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. a controller error occurred on aic ic1240, prompting an aic-to-aic transfer. the impella system was successfully transferred to aic ic12527, and support continued without any issue. | [5 empty] | B1 selection was added. b2 selection was added. h6 code c20 on the initial report should of been reported as code c21 as the investigation results were pending. type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. the momentary loss of placement signal and controller error alarm was due to an ¿optical pressure measuring unit¿s firmware communication protocol anomaly, resulting in misalignment of data communication packets received by main computer. the automated impella controller software operated as designed. product history review summary: there were no issues related to the complaint discovered when reviewing the product history of the console. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1007 | 1220648-2026-01598 | 24190632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190632&pc=OZD | 11/26/2025 |  |  | 01/27/2026 |  |  | 11/26/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1130896 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. a controller error occurred on the automated impella controller accompanied by an absent aortic placement signal (aops). the alarm and the aops issue self-resolved without intervention, and no additional problems were observed. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1008 | 1220648-2026-01599 | 24190805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190805&pc=OZD | 01/15/2026 |  |  | 01/27/2026 |  |  | 01/15/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775967 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Ventricular Fibrillation | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1009 | 1220648-2026-01583 | 24190894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24190894&pc=OZD | 11/12/2025 | 01/27/2026 |  | 01/27/2026 |  |  | 11/12/2025 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744935 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1010 | 1220648-2026-01602 | 24191135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24191135&pc=OZD | 12/18/2025 |  |  | 01/27/2026 |  |  | 01/16/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026786173 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | Pericardial Effusion | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale/review: an impella cp was placed to support the patient of unknown age or gender who had been admitted in acute myocardial infarction and cardiogenic shock. the pump was supporting the hemodynamics and allowed for the angioplasty of the coronary artery. when accessing with the guidewire, there was a coronary perforation that resulted in an effusion and tamponade. the team treated by thoracotomy. after 9 days the patient expired. anticoagulation necessary for the pump placement and support can contribute to bleeding. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's presenting native critical condition. the physician in japan did relay that the death was not due to the use of the impella cp. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1011 | 1220648-2026-01605 | 24191365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24191365&pc=PYX | 10/28/2025 |  |  | 01/27/2026 |  |  | 10/28/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1431034 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 79 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during pump preparation, an optical sensor system error occurred. despite multiple troubleshooting attempts and a console exchange, the error persisted. the device was not successfully implanted, the procedure with this device was aborted. a replacement device was prepared per instructions for use (ifu) and displayed no error messages or alarms. the device was successfully inserted without complications, and the patient was transported to the cardiovascular intensive care unit (cvicu) on support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1012 | 1220648-2026-01607 | 24191395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24191395&pc=OZD | 10/29/2025 | 02/18/2026 |  | 01/27/2026 |  |  | 10/29/2025 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1407754 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1013 | 1220648-2026-01608 | 24191405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24191405&pc=OZD | 10/29/2025 |  |  | 01/27/2026 |  |  | 10/29/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1108091 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the case was initiated with the pump functioning appropriately and no alarms or warnings noted. approximately 1.5 hours into the procedure, the console unexpectedly powered off despite being plugged in and fully charged. no high motor current events were observed. upon restarting, the console displayed an ¿unexpected controller shutdown¿ error. the pump and console were successfully restarted, and the remainder of the case was completed without further issues. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary the device was returned for evaluation. device analysis: the console was tested but the failure did not reproduce during testing. root cause: the cause of the unexpected reboot could not be determined due to the lack of indication in the logs and failure to reproduce. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1014 | 1220648-2026-01609 | 24191419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24191419&pc=PYX | 10/29/2025 |  |  | 01/27/2026 |  |  | 10/29/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the device indicated placement signal not reliable alarm occurred, with the placement signal absent from the screen for approximately 30 minutes. additional precautions were discussed with the nurse and physician, including closer monitoring of motor current and flow to assess pump function while continuing patient observation. the placement signal later returned, and no further issues were noted. | [5 empty] | D4. lot number, serial number, primary udi number are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1015 | 1220648-2026-01610 | 24191424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24191424&pc=OZD | 11/02/2025 | 01/30/2026 |  | 01/27/2026 |  |  | 11/02/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1103170 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during impella support, a controller error occurred. the automated impella controller (aic) simultaneously generated placement signal, suction, and controller error alarms. no corrective action was taken. | [5 empty] | D9 updated; the product has been returned. the investigation is still ongoing. | E1. initial reporter first and last name is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1016 | 1220648-2026-01613 | 24191426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24191426&pc=OZD | 11/03/2025 |  |  | 01/27/2026 |  |  | 11/03/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1156149 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during impella support setup, a controller error occurred. a controller error alarm was triggered upon powering on the automated impella controller (aic). a new console was subsequently used without issue. | [5 empty] | E1. initial reporter first and last name is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1017 | 1220648-2026-01612 | 24191428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24191428&pc=PYX | 11/02/2025 |  |  | 01/27/2026 |  |  | 11/03/2025 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1567887 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Overheating of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1018 | 1220648-2026-01615 | 24191445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24191445&pc=PYX | 11/03/2025 |  |  | 01/27/2026 |  |  | 11/03/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1656816 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during impella cp support, a controller error and placement signal error were observed. the impella cp pump lost aortic (ao) and left ventricular (lv) waveforms during percutaneous coronary intervention (pci) while ballooning the left main (lm) lesion. the waveforms initially returned but was lost again, triggering repeated controller error alarms until the waveforms were successfully restored. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1019 | 1220648-2026-01614 | 24191449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24191449&pc=OZD | 11/03/2025 |  |  | 01/27/2026 |  |  | 11/03/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1678477 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during impella support, a controller error occurred. the automated impella controller (aic) alerted staff that the system self-check failed during startup. power-cycling the battery did not clear the alert. | [5 empty] | D2a and d2b revised as they were entered incorrectly on the initial report that were submitted. e1 added initial reporter phone number as it was omitted from the initial report that was submitted. g4 revised as it was entered incorrectly on the initial report that was submitted. h6 code c20 was reported incorrectly on the initial report and it should of been c21. type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. data review: the console logs from the complaint date were consistent with what was reported in the complaint. the logs showed a communication issue with the power battery manager (pbm) during the initial bootup. due to the communication issue, the aic rebooted automatically (as designed) to attempt another power on. after rebooting, it displayed the "system self check failed" screen and was then forcefully shut down by the user. device analysis: the console was sent back to field service for further investigation. the system self check failure screen was reproduced upon boot up during testing. visual inspection confirmed the pbm corrosion which had impacted a neighboring rs232 communication channel between the pbm and 4-in-1. root cause: the root cause of the ¿system self check failure¿ on the console was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. product history review summary: there were no issues related to the complaint discovered when reviewing the product history of the console. | E1. initial reporter first and last name is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1020 | 1220648-2026-01618 | 24191947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24191947&pc=OZD | 09/26/2025 |  |  | 01/27/2026 |  |  | 09/26/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026718431 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | PHARMACIST | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Device Sensing Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the pump generated repeated false placement signal alarms despite the device being positioned appropriately in the aorta, as confirmed by echocardiography. at one point, the aortic waveform reading spiked into the 230s, inconsistent with the patient¿s actual hemodynamics. additionally, it was reported that the patient experienced right arm hematoma, noted near bicep. soft with no concerns from nurse and with observation only. no actions were taken. the patient outcome was reported to be stable. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed
1021 | 1220648-2026-01619 | 24192143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24192143&pc=OZD | 02/03/2026 |  |  | 01/27/2026 |  |  | 01/16/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796198 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1022 | 1220648-2026-01621 | 24192364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24192364&pc=OZD | 01/19/2026 | 02/03/2026 |  | 01/27/2026 |  |  | 01/19/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026739598 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1023 | 1220648-2026-01620 | 24192379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24192379&pc=OZD | 01/14/2026 |  |  | 01/27/2026 |  |  | 01/15/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744573 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Decrease in Pressure | Increase in Pressure | [5 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1024 | 1220648-2026-01622 | 24192382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24192382&pc=OZD | 01/11/2026 | 01/27/2026 |  | 01/27/2026 |  |  | 01/12/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2026696276 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient on an impella cp for pre-op placement for mitral valve surgery. when the pump was clamped with a soft jaw clamp to block the aorta, the alarm "pump stop current consumption increase" occurred, and the alarm went off even when the pump was turned to surgical mode. the operation was carried out to the end, and when the impeller was bled before the heart-lung machine was left, the pump would not restart, and the cp was removed at the same time as the heart-lung machine was left. the patient transition to the heart-lung machine and the intra-aortic balloon pump. the patient survived the procedure. | [5 empty] | E4 was inadvertently omitted on the initial manufacturer device report. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1025 | 1220648-2026-01623 | 24192500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24192500&pc=OZD | 09/07/2023 |  |  | 01/27/2026 |  |  | 01/20/2026 | 01/27/2026 |  |  | 03/31/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2024333669 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Study | Health Professional | User facility | Company representation | [7 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Pain | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1026 | 1220648-2026-01624 | 24192574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24192574&pc=OZD | 01/19/2026 | 01/27/2026 |  | 01/27/2026 |  |  | 01/19/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1287852 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1027 | 1220648-2026-01629 | 24192721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24192721&pc=OZD | 01/20/2026 | 01/27/2026 |  | 01/27/2026 |  |  | 01/20/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1399951 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Component Missing | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that impella stand is missing the iv pole portion. stand had to be switched out. there was no patient involvement. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1028 | 1220648-2026-01631 | 24192777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24192777&pc=OZD | 11/12/2025 |  |  | 01/27/2026 |  |  | 11/12/2025 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026743376 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1029 | 1220648-2026-01632 | 24192825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24192825&pc=OZD | 11/09/2025 |  |  | 01/27/2026 |  |  | 11/09/2025 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026743313 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Ventricular Fibrillation | Cardiac Tamponade | Respiratory Failure | Heart Failure/Congestive Heart Failure | Hemodynamic instability | [30 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1030 | 1220648-2026-01635 | 24193044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24193044&pc=OZD | 01/17/2026 |  |  | 01/27/2026 |  |  | 01/17/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794372 |  | 1000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male | 110 KG |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Clinical review/rationale: the impella cp was placed via the femoral access to support the 72-year-old male admitted in with cardiogenic shock and acute myocardial infarction with known coronary artery disease. the cp was supporting the patient along with ecmo in the icu when the team observed positional alarms on day 2 of support. the red alarm for "in aorta" resolved and support proceeded. the following day there was imaging performed and the team observed thrombus burden around the pigtail. there was no intervention or pump explant, rather the team kept the patient on the systemic heparin for anticoagulation needs. | Clinical narrative updated: clinical review/rationale: the impella cp was placed via the femoral access to support the 72-year-old male admitted in with cardiogenic shock and acute myocardial infarction with known coronary artery disease. the primary mcs was ecmo with the impella cp providing lv unloading. while in the icu the team observed position alarms on day 2 of support. the red alarm for "in aorta" resolved and support proceeded. the following day via transesophogeal echo a large thrombus was noted on the impella catheter. initially there was no intervention or pump explant, rather the team elected to keep the patient on the systemic heparin for anticoagulation needs. it was later decided to exchange the pump due to the thrombus and the patient expired later that day. the death will be conservatively reported on the initial impella cp however the death is most likely due to the patients underlying condition. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was provided b2, b5, d9, h1, h6 was updated.
1031 | 1220648-2026-01638 | 24193325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24193325&pc=OZD | 01/18/2026 | 01/27/2026 |  | 01/27/2026 |  |  | 01/18/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026795442 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1032 | 1220648-2026-01616 | 24193434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24193434&pc=OZD | 01/18/2026 | 01/27/2026 |  | 01/27/2026 |  |  | 01/19/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026804883 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male | 82 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella device for mechanical circulatory support. the patient was presented with a difficult anatomy. the team switched from the right to the left subclavian; however, delivery thru the left subclavian was not possible. after several attempts and consultation of senior interventional cardiologist, a decision was made to terminate the procedure. the patient was stable on veno-arterial extracorporeal membrane oxygenation support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1033 | 1220648-2026-01641 | 24193586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24193586&pc=OZD | 01/16/2026 | 01/27/2026 |  | 01/27/2026 |  |  | 01/19/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2026802205 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | White | Death | Yes | Yes | False Alarm | Restricted Flow rate | [5 empty] | Asystole | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 device was implanted in an adult male patient of unknown age for cardiogenic shock (scai stage d) in the setting of acute decompensated chronic ischemic cardiomyopathy and heart failure. the patient required extracorporeal membrane oxygenation, inotropes, and vasopressors. the device initially functioned as intended. while in the intensive care unit, decreased displayed flow and alarms indicating decoupled flow curves were noted, and imaging did not identify the device within the aorta at that time. the patient experienced cardiovascular collapse with asystole, required resuscitation, and subsequently expired. the patient¿s death was most likely due to the underlying critical clinical condition and severity of cardiogenic shock. device performance was as expected. the event is conservatively reported as death. | [5 empty] | D6b: added explant date. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1034 | 1220648-2026-01642 | 24196612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196612&pc=OZD | 01/11/2026 | 01/27/2026 |  | 01/27/2026 |  |  | 01/20/2026 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026658474 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 93 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted surgically into the right aorta in a 63-year-old male patient with a past medical history of coronary artery disease (cad). the impella 5.5 was inserted for ventricular support in a high-risk elective coronary artery bypass graft (cabg) surgery. the patient presented in scai stage e shock and was on an intra-aortic balloon pump (iabp), multiple inotropes and vasopressors, and was ventilated for respiratory support prior to initiation of support. eleven days post-implant, a perforation was noted to the left coronary cusp of the aortic valve. the physician decided to explant the impella. the patient was noted to have severe aortic insufficiency (ai) post-impella explant. the post-procedure outcome is survived at explant. aortic valve injury is listed as a potential adverse event and consistent with known device-related and patient-related factors documented in the instructions for use (ifu). | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1035 | 1220648-2026-01644 | 24196623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196623&pc=OZD | 10/23/2025 |  |  | 01/27/2026 |  |  | 10/23/2025 | 01/27/2026 |  |  | 11/30/2025 |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2024430562 |  | 1000100 |  | Device was returned to manufacturer |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 50 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that approximately 2000, local support was notified by the registered nurse (rn) that the automated impella controller (aic) was alarming ¿impella stopped.¿ the interventional cardiologist (ic) was unavailable, and initial troubleshooting was performed via facetime. according to the rn, there were no preceding alarms or changes in the patient¿s status or in the aic parameters prior to the pump stopping. while local support was en route to the facility, an additional aic was brought to the bedside, and an aic-to-aic transfer was attempted. the transfer resulted in the same alarm, and the impella pump remained stopped. notably, purge flow and purge pressure remained within nominal ranges throughout the event. the physician was contacted, and local support met him at the bedside. another aic-to-aic transfer was attempted, again yielding the same result with persistent pump stoppage. by this time, the pump had been nonfunctional for approximately one hour, but the patient remained hemodynamically stable. the device was fully explanted and the patient was reported to have survived the procedure. | [5 empty] | Section d catalog number was corrected. the product was returned. the cause of the pump stop was biomaterial build up based on log evidence. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1036 | 1220648-2026-01645 | 24196640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196640&pc=OZD | 10/18/2025 |  |  | 01/27/2026 |  |  | 10/23/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026733471 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | Pumping Stopped | Device Sensing Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient initially experienced high purge pressures with purge flow measuring <1 ml/hr. a tissue plasminogen activator (tpa) purge protocol was initiated through the purge system. concerns regarding the potential for pump stoppage were discussed with the care team. despite tpa administration, a ¿purge system blocked¿ alarm subsequently occurred. the pump¿s performance level then dropped from p6 to p0 and could not be restarted. the impella device was removed. the patient remained hemodynamically stable on a milrinone infusion. it was also reported that the placement signal was not reliable. intermittent suction alarms were noted, likely related to device positioning. and ongoing purge pressure issues. the patient was reported to have survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1037 | 1220648-2026-01646 | 24196646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196646&pc=OZD | 10/24/2025 |  |  | 01/27/2026 |  |  | 10/24/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1376304 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that when the automated impella controller (aic) was powered on, it displayed a self-check error and was unable to complete startup. as a result, the aic could not be turned on. | [5 empty] | Data review: the data logs were not recorded on complaint date. the logs shows - powermod_1 communication timeout error. the console logs also showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. (screen.png). device analysis: console was tested after arriving in field service. the reported failure mode was reproduced during testing at the field service (fs). upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination (pbm_corrosion1.jpg and pbm_corrosion2.jpg) which has impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. name: pbm material number: 0042-1125. lot number: 2019339988. serial number: (b)(6). are there an qns associated with the complaint device¿s most recent service report? console (b)(4) underwent rework due to a nonconformance according to the fsr (b)(4) which was unrelated failure mode. q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode? na. q3) have there been any other complaints against this console? no. phr summary: there were no issues related to the complaint discovered when reviewing the product history of console (b)(4).
1038 | 1220648-2026-01647 | 24196649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196649&pc=OZD | 10/24/2025 |  |  | 01/27/2026 |  |  | 10/24/2025 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026703546 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Fever | Ischemia | Pericardial Effusion | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
1039 | 1220648-2026-01648 | 24196655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196655&pc=OZD | 12/16/2025 | 02/24/2026 |  | 01/27/2026 |  |  | 12/16/2025 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782192 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1040 | 1220648-2026-01650 | 24196658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196658&pc=PYX | 01/20/2026 | 02/05/2026 |  | 01/27/2026 |  |  | 01/20/2026 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. |  |  | US |  | 1719481 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
1041 | 1220648-2026-01654 | 24196670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196670&pc=OZD | 12/16/2025 |  |  | 01/27/2026 |  |  | 12/16/2025 | 01/27/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744135 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1042 | 1220648-2026-01662 | 24196713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196713&pc=OZD | 12/16/2025 |  |  | 01/27/2026 |  |  | 12/16/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746993 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male | 91 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient experienced multiple episodes of ventricular fibrillation in the operating room. these events were not attributed to the circulatory support pump but were believed to be related to low potassium levels. the patient required four defibrillation attempts before stabilizing. a circulatory support pump was placed using right-side surgical access through the aorta and remained in use at the time of the report. no additional information was available regarding patient condition, device performance, or the potential return of the product. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1043 | 1220648-2026-01663 | 24196718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196718&pc=OZD | 12/16/2025 | 02/14/2026 |  | 01/27/2026 |  |  | 12/16/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026749965 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that attempts to place a circulatory support pump through right-side axillary access were unsuccessful due to extremely tortuous and atypical vascular anatomy, which prevented advancement of the guidewire and later caused kinking of the pump cannula. after prolonged efforts, including specialist consultation and a failed attempt at direct insertion through a mini-sternotomy due to an infectious pocket around the ascending aorta, the procedure was aborted and the clinical team transitioned the patient to veno-arterial extracorporeal membrane oxygenation for circulatory support. no patient symptoms were reported during the event, and the device required revision or replacement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1044 | 1220648-2026-01666 | 24196729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196729&pc=OZD | 10/28/2025 |  |  | 01/27/2026 |  |  | 10/28/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739131 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Cardiac Tamponade | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient arrived to the unit and experienced an episode of low blood pressure and appeared to be in a hypercoagulable state. the chest tube was placed on suction and then connected to a different drainage system. there was concern for possible cardiac tamponade; however, once the chest tube was fully suctioned, the low blood pressure resolved. the patient continued to bleed heavily from the mediastinal chest tube. during surgical exploration, the saphenous vein graft was found to be the source of bleeding and was sutured. multiple blood products were administered. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1045 | 1220648-2026-01669 | 24196731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196731&pc=OZD | 10/28/2025 |  |  | 01/27/2026 |  |  | 10/28/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026740943 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 92 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the patient subsequently experienced a stroke. the plan of care was reviewed with the registered nurse at the bedside. the patient was recovering from an ischemic stroke with some remaining weakness in the right upper extremity and right lower extremity. the patient was working with physical therapy and occupational therapy and had been removed from the orthotopic heart transplant list. the patient was experiencing brief episodes of ventricular tachycardia. the registered nurse reported that the clinical team was aware of these episodes and had been informed that the impella device was appropriately positioned. an official ultrasound examination of the heart was being considered. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1046 | 1220648-2026-01667 | 24196734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196734&pc=OZD | 10/28/2025 |  |  | 01/27/2026 |  |  | 10/28/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756673 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 53 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp device was used to provide support during treatment of a completely blocked right coronary artery. the patient experienced a pulseless electrical activity arrest during the right coronary artery chronic total occlusion percutaneous coronary intervention. the groin access site was not suitable for impella cp placement. the pump was inserted and initiated without difficulty, and return of spontaneous circulation was achieved. a successful percutaneous coronary intervention of the right coronary artery was subsequently completed. the pump was left in place overnight while the patient was transported to the coronary care unit in critical condition. the pump was successfully weaned and removed, and the patient continued to show daily improvement. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1047 | 1220648-2026-01668 | 24196735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196735&pc=OZD | 10/28/2025 |  |  | 01/27/2026 |  |  | 10/28/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739622 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 92 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp device was used in a male patient who presented with swelling of both lower extremities and shortness of breath during exertion. an ultrasound examination of the heart performed at an outside facility showed a reduced ejection fraction of twenty percent. the impella device was inserted using best practice techniques. the peel-away sheath was removed, and the repositioning sheath was secured with sutures and adhesive dressings. a pulmonary artery catheter was placed through the right femoral vein. a knee immobilizer was applied, and the patient was transferred to the medical intensive care unit in stable condition. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1048 | 1220648-2026-01673 | 24196759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24196759&pc=OZD | 11/04/2025 |  |  | 01/27/2026 |  |  | 10/31/2025 | 01/27/2026 |  |  |  |  | 01/27/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026711283 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced hypotension and arrhythmia but maintained stable on impella support. the patient was reported to have survived the procedure. | [5 empty] | Hypotension/ arrhythmia: the cause of the injury was not determined due to insufficient clinical details. the pump s/n: (b)(6) passed all the post sterile inspection checks.
1049 | 1220648-2026-01491 | 24176542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24176542&pc=OZD | 01/09/2026 |  |  | 01/26/2026 |  |  | 01/09/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026658720 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1050 | 1220648-2026-01500 | 24176673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24176673&pc=OZD | 12/28/2025 |  |  | 01/26/2026 |  |  | 12/29/2025 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026742679 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1051 | 1220648-2026-01501 | 24176679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24176679&pc=OZD | 02/02/2026 |  |  | 01/26/2026 |  |  | 01/16/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779529 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1052 | 1220648-2026-01502 | 24176761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24176761&pc=OZD | 01/17/2026 |  |  | 01/26/2026 |  |  | 01/17/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781149 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1053 | 1220648-2026-01504 | 24176803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24176803&pc=OZD | 01/15/2026 |  |  | 01/26/2026 |  |  | 01/16/2026 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1649323 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Overheating of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller experienced a battery overheat alarm while in use. no patient harm was reported. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D2a, d2b, h5, and h8 was erroneously entered on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1054 | 1220648-2026-01505 | 24176817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24176817&pc=OZD | 01/08/2026 |  |  | 01/26/2026 |  |  | 01/16/2026 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794986 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that implantation of an impella 5.5 was aborted due to the patient's anatomy. an impella cp was successfully implanted instead. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1055 | 1220648-2026-01506 | 24176868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24176868&pc=OZD | 01/15/2026 |  |  | 01/26/2026 |  |  | 01/15/2026 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1133147 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an automated impella controller generated a yellow controller error alarm upon startup. the controller was not used for patient support. | [5 empty] | The reported event did not occur during patient use. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1056 | 1220648-2026-01507 | 24176959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24176959&pc=OZD | 01/19/2026 |  |  | 01/26/2026 |  |  | 01/20/2026 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026780280 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 13 YR |  | Female | 60 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella 5.5 and experienced a separation in the anti-contamination sleeve during support. it was reported that the patient had been taken to the operating room for decannulation of extracorporeal membrane oxygenation (ecmo) and insertion of impella 5.5. the impella 5.5 was successfully implanted with no reported complications, and ecmo was successfully decannulated. following the decannulation of ecmo, the impella 5.5 flows dropped. attempts to troubleshoot with repositioning the pump were unsuccessful and the decision was made to replace the impella 5.5. it was noted that the initial pump flows were <1 l/min. a new impella 5.5 was successfully inserted, and support was initiated without reported complications. on support day 47, it was reported that the distal side of the anti-contamination sleeve was found to be separated from the distal locking hub. the medical team noted that the patient was very anxious and may have been pulling the catheter causing the separation of the sleeve. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the patient was successfully weaned from the impella 5.5 during the transplant procedure and underwent a successful heart transplant. the patient survived the procedure. this complaint involves two impella devices: pump 1: impella 5.5, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). this mdr specifically addresses pump 2: impella 5.5, with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1057 | 1220648-2026-01508 | 24176971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24176971&pc=OZD | 12/30/2025 |  |  | 01/26/2026 |  |  | 01/14/2026 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1446166 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an automated impella controller failed to recognize a purge cassette. despite replacing two cassettes the error persisted. the cassettes were functional on a different controller. the controller was replaced to continue patient support. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1058 | 1220648-2026-01510 | 24177082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24177082&pc=OZD | 01/15/2026 |  |  | 01/26/2026 |  |  | 01/15/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1289713 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Battery Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1059 | 1220648-2026-01511 | 24177102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24177102&pc=OZD | 01/09/2026 |  |  | 01/26/2026 |  |  | 01/15/2026 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1311372 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an automated impella controller entered battery mode despite being permanently connected to the electrical current, then spontaneously returned to receiving electrical power. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1060 | 1220648-2026-01509 | 24177253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24177253&pc=PYX | 01/15/2026 | 01/26/2026 |  | 01/26/2026 |  |  | 01/15/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1771596 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | R |  |  |  | Initial submission | Followup
1061 | 1220648-2026-01512 | 24177299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24177299&pc=PYX | 01/19/2026 | 01/30/2026 |  | 01/26/2026 |  |  | 01/19/2026 | 01/26/2026 |  |  | 02/28/2026 |  |  |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517696 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1062 | 1220648-2026-01513 | 24177518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24177518&pc=OZD | 01/19/2026 | 02/16/2026 |  | 01/26/2026 |  |  | 01/19/2026 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026793719 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 61 YR |  | Female | 75 KG | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 61-year-old female patient was placed on bipella impella support with the cp and rp flex. any underlying comorbidities that prompted the bipella support is unknown. the cp pump had suction alarms/low flows during the support, and any troubleshooting done to remedy was not shared. on the day after implant the patient expired. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's presenting native critical condition. | The complainant reported additional information upon request: "this impella is not available for return, as the patient expired on support and the devices were not removed prior to being taken to the morgue. as stated in now several previous emails, the patient expired (b)(6) around roughly 1500. this was a critically ill patient placed on bipella support who suffered extreme metabolic derangement and unfortunately had end-organ failure leading to expiration. " |  |  |  |  | The impella device has not been received from the customer; therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated. the investigation is still ongoing.
1063 | 1220648-2026-01514 | 24177633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24177633&pc=OZD | 01/19/2026 | 01/26/2026 |  | 01/26/2026 |  |  | 01/19/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766622 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic dissection | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1064 | 1220648-2026-01516 | 24177740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24177740&pc=OZD | 01/20/2026 | 01/26/2026 |  | 01/26/2026 |  |  | 01/20/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026773632 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1065 | 1220648-2026-01515 | 24177741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24177741&pc=OZD | 01/18/2026 |  |  | 01/26/2026 |  |  | 01/18/2026 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775996 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 78 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that the 5.5 cannula would not cross into calcified aorta. the pump was aborted and an impella cp was used. there was no reported patient harm. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. failure to advance: the cause of the delivery issue was determined to be patient condition due to calcification. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1066 | 1220648-2026-01517 | 24177772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24177772&pc=OZD | 01/19/2026 | 01/26/2026 |  | 01/26/2026 |  |  | 01/20/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633349 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1067 | 1220648-2026-01518 | 24177858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24177858&pc=OZD | 01/20/2026 | 01/26/2026 |  | 01/26/2026 |  |  | 01/20/2026 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026641066 |  | 1000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 59 YR |  | Female | 41 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella cp device was placed in a 59-year-old female patient for cardiogenic shock (scai stage d) in the setting of acute decompensated chronic non-ischemic cardiomyopathy and heart failure. during support, a pump stop was noted, and the device could not be restarted despite troubleshooting, with high motor current observed. a clinical decision was made to exchange the system and transition to a backup device. the patient experienced device revision or replacement. there were no associated signs, symptoms, or patient involvement related to the event. the patient remained clinically stable. comprehensive review did not identify evidence suggesting a causal relationship between the impella cp device and the reported patient event. the patient survived. | [5 empty] | Correction: e4 was inadvertently omitted on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1068 | 1220648-2026-01519 | 24177918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24177918&pc=OZD | 01/20/2026 | 01/26/2026 |  | 01/26/2026 |  |  | 01/20/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727199 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1069 | 1220648-2026-01522 | 24178597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24178597&pc=OZD | 11/12/2025 |  |  | 01/26/2026 |  |  | 11/12/2025 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026747002 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1070 | 1220648-2026-01523 | 24178838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24178838&pc=OZD | 01/21/2026 |  |  | 01/26/2026 |  |  | 01/21/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026763878 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1071 | 1220648-2026-01525 | 24178914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24178914&pc=OZD | 09/24/2025 |  |  | 01/26/2026 |  |  | 09/25/2025 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026633624 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella device was implanted in the patient for mechanical circulatory support. when attempting to place the cpsa, a 0.018-inch wire unintentionally slipped out of the left ventricle, obstructing the advancement of the impella. re-wiring was also considered, but the team determined that manipulating the wire would increase the risk of ventricular tachycardia or ventricular fibrillation. additionally, endothelial-like tissue was observed to be attached to the device's outlet portion. based on these factors, a decision was made to discontinue the placement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1072 | 1220648-2026-01526 | 24179086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24179086&pc=OZD | 09/24/2025 |  |  | 01/26/2026 |  |  | 09/25/2025 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026696201 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the purge flow rate had dropped by half since the start of support, to below 5ml/hr. the decrease in purge flow rate became noticeable on september 24th, so a pump replacement was performed, even though no alarm was sounded. there was no patient injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1073 | 1220648-2026-01528 | 24179198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24179198&pc=OZD | 09/24/2025 |  |  | 01/26/2026 |  |  | 09/24/2025 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1153140 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the physician did not feel comfortable proceeding with the ppci due to low placement signal and the catheter becoming kinked during insertion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1074 | 1220648-2026-01529 | 24179668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24179668&pc=OZD | 01/10/2026 |  |  | 01/26/2026 |  |  | 01/11/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026790053 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic dissection | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1075 | 1220648-2026-01530 | 24179848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24179848&pc=OZD | 01/20/2026 | 02/02/2026 |  | 01/26/2026 |  |  | 01/20/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026770381 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1076 | 1220648-2026-01531 | 24179933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24179933&pc=PYX | 01/20/2026 | 01/26/2026 |  | 01/26/2026 |  |  | 01/20/2026 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025536971 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Death | Yes | No | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1077 | 1220648-2026-01532 | 24183152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24183152&pc=OZD | 01/20/2026 |  |  | 01/26/2026 |  |  | 01/20/2026 | 01/26/2026 |  |  |  |  | 01/26/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026788405 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 71 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp and experienced aortic placement signal output that was not functioning as expected. it was reported that the impella cp was implanted and the automated impella controller was turned on with no reported issues. following priming of the pump, it was reported that the aorta placement output numbers indicated that it did not zero correctly. the abiomed representative advised that the placement signal would continue to be skewed and recommended replacing the impella cp. the decision was made to not replace the impella cp, and the patient¿s mean arterial pressure and arterial line pressures were monitored. the pump¿s position was confirmed with imaging. it was reported that there were no issues with impella flow or function throughout the case. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1078 | 1220648-2026-01539 | 24185872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24185872&pc=OZD | 12/18/2025 |  |  | 01/26/2026 |  |  | 12/29/2025 | 01/26/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0048-0003 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Literature | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1079 | 1220648-2026-01481 | 24175522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24175522&pc=OZD | 11/21/2025 |  |  | 01/25/2026 |  |  | 11/21/2025 | 01/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746063 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1080 | 1220648-2026-01479 | 24175528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24175528&pc=PYX | 11/20/2025 | 01/27/2026 |  | 01/25/2026 |  |  | 11/21/2025 | 01/25/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1669290 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Multiple Organ Dysfunction Syndrome | [34 empty] | R |  |  |  | Initial submission | Followup | Followup
1081 | 1220648-2026-01483 | 24175529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24175529&pc=PYX | 11/22/2025 |  |  | 01/25/2026 |  |  | 11/22/2025 | 01/25/2026 |  |  |  |  | 01/25/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1465870 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. during impella support, the automated impella controller (aic) experienced a brief controller error that self-resolved, occurring simultaneously with the placement signal (ps) screen blanking out. after the event, all waveforms were reported to be within normal limits. the rn monitored the device, located a backup aic in case the alarm returned, and notified the team. no additional corrective actions were required as the error self-resolved. no patient injury was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1082 | 1220648-2026-01487 | 24175533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24175533&pc=OZD | 12/01/2025 |  |  | 01/25/2026 |  |  | 12/01/2025 | 01/25/2026 |  |  |  |  | 01/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026666058 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. after 5 days of support prior to removal the pulses were present but lost post removal. the impella cp was explanted via surgical cutdown. the peel-away sheath remained in place. a clot was extracted. no pulse was detected following the initial repair, so the repair was repeated, after which pulses returned. the groin was then closed. the patient remained hemodynamically stable, with plans for bronchoscopy and anticipated extubation, followed by extracorporeal membrane oxygenation (ecmo) decannulation. the patient's outcome at the time of explant was survived. | [5 empty] | No product of logs were returned. in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details.
1083 | 1220648-2026-01488 | 24175700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24175700&pc=OZD | 12/07/2025 | 02/12/2026 |  | 01/25/2026 |  |  | 12/08/2025 | 01/25/2026 |  |  |  |  | 01/25/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026682017 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Priming Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support experienced an undetermined source of bleeding. a ct scan did not identify the origin of the bleeding. the patient received four units of packed red blood cells. hemolysis laboratory results were negative, and there was concern for possible disseminated intravascular coagulation (dic). the patient is reported to be stable. | [5 empty] | H6 updated. the investigation is ongoing. | Investigation summary: ppae (major bleed): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
1084 | 1220648-2026-01442 | 24173610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24173610&pc=PYX | 11/06/2025 | 02/18/2026 |  | 01/24/2026 |  |  | 11/06/2025 | 01/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1881868 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1085 | 1220648-2026-01443 | 24173612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24173612&pc=OZD | 11/09/2025 |  |  | 01/24/2026 |  |  | 11/09/2025 | 01/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026729491 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Increase in Pressure | [6 empty] | Ventricular Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1086 | 1220648-2026-01444 | 24173631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24173631&pc=OZD | 11/10/2025 |  |  | 01/24/2026 |  |  | 11/10/2025 | 01/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026749015 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Increase in Pressure | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1087 | 1220648-2026-01446 | 24173640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24173640&pc=OZD | 11/12/2025 |  |  | 01/24/2026 |  |  | 11/12/2025 | 01/24/2026 |  |  |  |  | 01/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1129883 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | During routine purge cassette and tubing replacement for the impella 5.5 system in the intensive care unit, the black component of the purge system broke for unknown reasons. the device remained functional, the patient was stable, and no harm occurred. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | D3 added manufacturer fax number as it was omitted from the initial report that was submitted. d6a should of been left blank on the initial report that was submitted. d8 revised as it was entered incorrectly on the initial report that was submitted. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated as they were entered incorrectly on the initial report that was submitted. the investigation was complete at the time the initial report was submitted. h11 the investigation was complete at the time the initial report was submitted. therefore, additional manufacturer narrative was updated. investigation summary: the investigation was completed. data review: the console logs were consistent with what was reported in the complaint. the logs revealed that the console issued the purge disc not detected alarm several times. device analysis: the console was sent back to field service for further investigation. the purge disc flag was confirmed to be broken. root cause: the purge disc flag on console was confirmed to be broken and resulted in the consoles inability to detect the purge disc. the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly. product history review summary: there were no issues related to the complaint discovered when reviewing the product history of the console.
1088 | 1220648-2026-01454 | 24173672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24173672&pc=OZD | 12/02/2025 |  |  | 01/24/2026 |  |  | 12/02/2025 | 01/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1142509 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1089 | 1220648-2026-01460 | 24174764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24174764&pc=OZD | 12/08/2025 |  |  | 01/24/2026 |  |  | 12/08/2025 | 01/24/2026 |  |  |  |  | 01/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773633 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 62 KG | Non Hispanic | White | Death | Yes | Yes | Deformation Due to Compressive Stress | Difficult to Advance | [5 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. during the procedure, wire access was lost in the groin, resulting in a femoral hemorrhage. attempts to sidebite the aorta were unsuccessful due to difficulty with cannula placement. the aorta was subsequently re-cannulated at a higher location, the heart was re-arrested and cross-clamped, and a graft was sewn in place. an impella 5.5 was then inserted using a wire and catheter technique. the abiomed 0.018 wire kinked during the initial attempt, and a v18 wire was used instead. after the device was positioned in the left ventricle, an attempt was made to wean the patient from cardiopulmonary bypass. the patient subsequently expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This follow-up report is being submitted to correct an inaccurately reported medical device problem code (h6) and to provide investigation results. the originally reported code a040601 (deformation due to compressive stress) was not applicable to the device involved in this mdr. the observed deformation was attributed to the guidewire and is reported within the guidewire mdr. the incorrect code has been removed from this report. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product returned. major bleed: patient had femoral hemorrhage along aassive amount of bleeding into leg, abdomen largely distended and concern for rp bleed. blood was noted in ett and tubing with the cause of the bleed was unable to be determined as insufficient clinical details were provided. cardiac arrest: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. difficult to advance: the cause of difficult to advance was unable to be determined as insufficient clinical details were provided. device history lot: device lot: 1991920 device history batch: subcomponent lot: n/a device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
1090 | 1220648-2026-01461 | 24174806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24174806&pc=OZD | 12/11/2025 |  |  | 01/24/2026 |  |  | 12/08/2025 | 01/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775825 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1091 | 1220648-2026-01462 | 24174807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24174807&pc=PYX | 12/08/2025 |  |  | 01/24/2026 |  |  | 12/08/2025 | 01/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1666954 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1092 | 1220648-2026-01463 | 24174810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24174810&pc=OZD | 12/03/2025 |  |  | 01/24/2026 |  |  | 12/10/2025 | 01/24/2026 |  |  |  |  | 01/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1814730 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 84 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during a purge cassette change, the automated impella controller (aic) was unable to detect the new purge disc. the nurse attempted multiple reinsertion attempts and performed de-airing of the aic, but the issue persisted. the patient continued on support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1093 | 1220648-2026-01469 | 24174838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24174838&pc=OZD | 12/18/2025 |  |  | 01/24/2026 |  |  | 12/18/2025 | 01/24/2026 |  |  |  |  | 01/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794864 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 114 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was being implanted with an impella 5.5 and the implant was aborted due to the device being unable to be delivered to the left ventricle. it was reported that despite the use of best practices, the physician was unable to turn down the aorta. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the patient's outcome at the time of explant was survived. | [5 empty] | Investigation summary ==> no product was returned. failure to advance: the cause of the failure to advance was determined to be patient condition as the patient had difficult vasculature and resistance at aorta preventing successful delivery of impella. device history lot ==> device lot: 1998858 device history batch ==> subcomponent lot: n/a device history review ==> device (sn: ) passed all post sterile inspection checks.
1094 | 1220648-2026-01471 | 24174840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24174840&pc=OZD | 12/18/2025 |  |  | 01/24/2026 |  |  | 12/18/2025 | 01/24/2026 |  |  |  |  | 01/24/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1261275 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 91 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient who was implanted with an impella 5.5 experienced hemolysis, a pump stop, and alarms for controller error and retrograde flow. it was reported on the first day of impella support, the patient experienced urine output that was reddish in color. it was reported that the impella was repositioned, and the hematuria subsequently resolved. additionally, it was reported that on the eleventh day of support the pump experienced an abrupt pump stop with alarms for controller error and retrograde flow. the impella was successfully restarted and the medical team reported there were no further issues. an onsite abiomed representative reviewed the device and noted that the pump was operating without issues and with appropriate pump flows for the performance level. it was additionally noted that the patient was also supported by veno-arterial extracorporeal membrane oxygenation. at the time of writing this report, the patient continues to be supported by impella 5.5. | [5 empty] | D6a should have been left blank on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1095 | 1220648-2026-01475 | 24174848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24174848&pc=OZD | 12/05/2025 |  |  | 01/24/2026 |  |  | 01/13/2026 | 01/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1484655 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1096 | 1220648-2026-01476 | 24174858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24174858&pc=PYX | 11/05/2025 |  |  | 01/24/2026 |  |  | 11/05/2025 | 01/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1717562 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1097 | 1220648-2026-01477 | 24174859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24174859&pc=PYX | 11/08/2025 |  |  | 01/24/2026 |  |  | 11/08/2025 | 01/24/2026 |  |  |  |  | 01/24/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1476670 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that when attempting to start the automated impella controller (aic) for a case, the aic generated alarms indicating a controller failure and a battery failure. a backup aic was brought in and used to proceed with the case. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1098 | 1220648-2026-01478 | 24174861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24174861&pc=OZD | 11/10/2025 |  |  | 01/24/2026 |  |  | 11/10/2025 | 01/24/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737094 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Poor Visibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1099 | 1220648-2026-01347 | 24158343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158343&pc=OZD | 11/28/2025 |  |  | 01/23/2026 |  |  | 11/28/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1420348 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a hospital perfusionist reported that when an automated impella controller (aic) was proactively powered on for a potential case, it displayed a ¿console error¿ alarm. due to this error, the team elected to use a backup console instead. the faulty primary console was swapped out and labeled for biomed inspection. a sign was placed on the primary console to indicate it needs evaluation. there was no patient involvement. | [5 empty] | E1. initial reporter first and last name is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1100 | 1220648-2026-01350 | 24158383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158383&pc=OZD | 12/01/2025 |  |  | 01/23/2026 |  |  | 12/01/2025 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1289312 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1101 | 1220648-2026-01349 | 24158384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158384&pc=PYX | 12/01/2025 |  |  | 01/23/2026 |  |  | 12/01/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1952690 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the console was displaying an alert indicating a fiber optical sensor issue. there was no patient involvement. | [5 empty] | E1. initial reporter first and last name is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1102 | 1220648-2026-01351 | 24158398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158398&pc=OZD | 12/01/2025 |  |  | 01/23/2026 |  |  | 12/01/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1606415 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the mechanical circulatory support (mcs) specialist reported that when beginning a new case, the purge disc was not detected. they attempted multiple times to insert the disc and proceed through the setup screens, but the error continued. a different console was then used, and the purge disc was successfully detected on the new unit. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | D2a and d2b revised as they were entered incorrectly on the initial report that was submitted. e1 added initial report phone number as it was omitted from the initial report that was submitted. g4 revised as it entered incorrectly on the initial report that was submitted. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. data review: the console logs were consistent with what was reported in the complaint. the logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. device analysis: the console was sent back to field service for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. root cause: the purge disc flag on the console was observed to be broken and was the root cause of the inability to detect the purge disc. the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly. product history review summary: there were no issues related to the complaint discovered when reviewing the product history of the console.
1103 | 1220648-2026-01357 | 24158432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158432&pc=OZD | 12/02/2025 |  |  | 01/23/2026 |  |  | 12/02/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1087350 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the device displayed an error message indicating a console failure. the console was exchanged without difficulty, and there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1104 | 1220648-2026-01359 | 24158448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158448&pc=PYX | 12/03/2025 |  |  | 01/23/2026 |  |  | 12/03/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1811668 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an alert indicating a controller failure appeared after the automated impella controller was powered on. the patient was not on support at the time, and the pump was not connected. the controller was replaced, and the removed controller was labeled and quarantined with the biomedical department. the complainant reported that the patient had been successfully weaned from impella support and was taken to the operating room for device removal. the device was explanted without complications, the graft was trimmed, and the pocket was closed. the patient tolerated the procedure well, and the attending physician expressed satisfaction with the case. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1105 | 1220648-2026-01360 | 24158454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158454&pc=OZD | 11/19/2025 |  |  | 01/23/2026 |  |  | 11/19/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1613084 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. during routine morning equipment checks, staff powered on the automated impella controller (aic) and the unit displayed a blue screen with the error message: ¿system self check failed. change console immediately.¿ the console was immediately marked ¿do not use¿ and removed from service. the team contacted the field representative. the malfunction occurred during equipment checkout only. the aic did not interact with any patients, and there was no patient involvement or harm. this event represents a reportable malfunction involving controller failure. | [5 empty] | D2a and d2b revised as they were entered incorrectly on the initial report that was submitted. d3 added manufacturer fax number as it was omitted from the initial report that was submitted. d9 device was returned and date received was added as it was omitted from the initial report that was submitted. e1 added initial reporter facility name as it was omitted from the initial report that was submitted. h6 added code b13 as it was submitted from the initial report that was submitted. code b17, c20 and d15 were reported incorrectly on the initial report that was submitted. codes should of reflected b01, c21 and d16 as the device was returned and the investigation was pending. type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. h11 the additional manufacturer narrative that was reported on the initial report was incorrect as the device was returned and was pending investigation completion. investigation summary: the investigation has been completed. data review: the console logs from the complaint date were consistent with what was reported in the complaint. the logs showed a communication issue with the power battery manager (pbm) during the initial bootup. due to the communication issue, the automated impella controller rebooted automatically (as designed) to attempt another power on. after rebooting, it displayed the "system self check failed" screen and was then forcefully shut down by the user. device analysis: the console was sent back for further investigation. the system self check failure screen was reproduced upon boot up during testing. visual inspection confirmed the pbm corrosion which had impacted a neighboring rs232 communication channel between the pbm and 4-in-1. root cause: the root cause of the ¿system self check failure¿ on the console was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. product history review summary: there were no issues related to the complaint discovered when reviewing the product history of the console. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1106 | 1220648-2026-01362 | 24158471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158471&pc=OZD | 11/27/2025 | 02/26/2026 |  | 01/23/2026 |  |  | 12/03/2025 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1387086 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1107 | 1220648-2026-01361 | 24158475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158475&pc=OZD | 12/03/2025 |  |  | 01/23/2026 |  |  | 12/03/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1095479 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 101 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was transferred from another facility on impella support, and the transport was uneventful. upon arrival, the patient was moved to the bed without difficulty, and all waveforms, flows, and alarms were appropriate. the transfer from one automated impella controller to another was completed, beginning with connection of the white cable followed by the purge cassette. the complainant reported that when the registered nurse attempted to disconnect the white cable, she attempted to twist it, and was instructed to pull it straight out instead. the pump was then connected and started as expected; however, suction occurred and the waveforms appeared distorted, resembling electrical static. the registered nurse was asked to confirm that the white cable was fully inserted, which it was. the clinical support center advised unplugging and reconnecting the cable, but this did not resolve the issue. device position was confirmed by echocardiography. the pump was adjusted to a lower setting to relieve suction. approximately one hour after the controller transfer, the waveforms returned to normal and there was no further suction, even at higher flow settings. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1108 | 1220648-2026-01363 | 24158477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158477&pc=OZD | 12/04/2025 |  |  | 01/23/2026 |  |  | 12/04/2025 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1071329 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1109 | 1220648-2026-01367 | 24158481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158481&pc=OZD | 12/05/2025 |  |  | 01/23/2026 |  |  | 12/05/2025 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026686529 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Restricted Flow rate | [6 empty] | Anemia | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1110 | 1220648-2026-01366 | 24158484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158484&pc=OZD | 12/03/2025 |  |  | 01/23/2026 |  |  | 12/05/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1144661 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that while preparing for an impella 5.5 insertion, the automated impella controller was powered on and immediately displayed a controller failure alarm. the device was removed from use, and an alternate controller was selected and operated without any issues. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1111 | 1220648-2026-01368 | 24158490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158490&pc=OZD | 12/05/2025 |  |  | 01/23/2026 |  |  | 12/05/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1177999 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a controller error alarm occurred, and the placement signal appeared flat. the alarm resolved after a few minutes but reoccurred within approximately thirty minutes. the clinical support center advised not to replace the automated impella controller at that time, as the alarm could potentially be related to the pump. throughout the event, flow and motor current remained stable and within expected ranges. a backup controller was positioned in the intensive care unit as a precaution. the following morning, the customer observed the alarm again and chose to replace the controller. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1112 | 1220648-2026-01369 | 24158492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158492&pc=OZD | 12/06/2025 |  |  | 01/23/2026 |  |  | 12/06/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026688405 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a male patient received a mechanical circulatory support device following valvular heart surgery for management of cardiogenic shock with low cardiac output. before placement of the device, the patient had been supported with an intra-aortic balloon pump, extracorporeal membrane oxygenation, and multiple inotropic and vasopressor medications. during support with the device, the patient developed gastrointestinal bleeding, which may have been associated with systemic anticoagulation therapy required for device use. the patient also experienced elevated purge pressures and decreased purge flow alarms, although no kinks in the purge line were observed. the purge solution was changed from a sodium bicarbonate¿based solution to a heparin-based solution without interruption of hemodynamic support. due to the severity of the underlying condition and in accordance with the family¿s wishes, mechanical support was withdrawn, and the patient died. the death will be conservatively reported for the mechanical circulatory support device. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product or data logs were returned for evaluation. major bleed: clinical details noted patient had gastrointestinal (gi bleed). the cause of the gi bleed was not determined due to insufficient clinical details. high purge pressure: clinical details noted elevated purge pressures and decreased purge flow alarms. no kinks in line observed. no additional details provided. the cause of the high purge pressure could not be determined as no product or logs were returned for evaluation and limited clinical details were provided.
1113 | 1220648-2026-01370 | 24158493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158493&pc=OZD | 12/06/2025 |  |  | 01/23/2026 |  |  | 12/06/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026653710 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the nursing staff stated during morning rounds that the patient experienced an episode of ventricular tachycardia during the night, which resulted in decreased extracorporeal membrane oxygenation flow and a reduction in the device performance level. the ventricular tachycardia resolved, and later in the shift the patient experienced another arrhythmic event consistent with torsades de pointes and ventricular tachycardia. amiodarone and magnesium were administered. during the second event, neither the mechanical circulatory support device nor the extracorporeal membrane oxygenation support was reduced. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1114 | 1220648-2026-01371 | 24158494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158494&pc=PYX | 12/06/2025 |  |  | 01/23/2026 |  |  | 12/06/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | The complainant reported that during patient transport to a computed tomography scan, a pump stop alarm and corresponding audible alarm occurred. clinical staff reported no loss of impella flow, with all displayed metrics remaining stable and unchanged. the patient¿s hemodynamic status also remained stable throughout the event. staff stated that they did not interact with the controller during the occurrence. the alarm appeared to be self-limiting and resolved without intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1115 | 1220648-2026-01372 | 24158498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158498&pc=OZD | 12/05/2025 |  |  | 01/23/2026 |  |  | 12/06/2025 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1302108 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
1116 | 1220648-2026-01374 | 24158499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158499&pc=PYX | 12/07/2025 |  |  | 01/23/2026 |  |  | 12/07/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1468353 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Open or Close | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller experienced an issue with the purge door, as the purge door would not open when the clinical team attempted to access it. no information was provided regarding whether a patient was being supported at the time of the event or whether any additional corrective actions were taken. the event was characterized as a reportable malfunction. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1117 | 1220648-2026-01373 | 24158501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158501&pc=OZD | 12/07/2025 |  |  | 01/23/2026 |  |  | 12/07/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1089014 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a registered nurse observed that while she was with the patient, the screen on the loaner automated impella controller suddenly went black and appeared to reboot. the staff confirmed that no one was interacting with the controller at the time and that the device was connected to an electrical outlet. the controller appeared to undergo a rebooting process that lasted approximately fifteen to thirty seconds. after rebooting, the impella automatically transitioned to performance level seven without receiving any user prompts. the registered nurse subsequently adjusted the device back to performance level eight, and no further issues were observed following this change. the registered nurse was unsure whether the impella pump stopped during the reboot sequence, as the patient remained stable and exhibited no change in hemodynamic status throughout the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1118 | 1220648-2026-01382 | 24160655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24160655&pc=OZD | 01/18/2026 | 02/02/2026 |  | 01/23/2026 |  |  | 01/18/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026741810 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1119 | 1220648-2026-01383 | 24160723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24160723&pc=OZD | 01/17/2026 | 02/02/2026 |  | 01/23/2026 |  |  | 01/18/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026726416 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1120 | 1220648-2026-01386 | 24160807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24160807&pc=OZD | 01/19/2026 | 01/23/2026 |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026807035 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 86 YR |  | Male | 122 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella support with a cp pump. the complainant reported bleeding from the patient¿s access site, requiring manual pressure and a dressing change. | [5 empty] | The impella device has not been received from the customer; therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1121 | 1220648-2026-01389 | 24160976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24160976&pc=OZD | 01/11/2026 |  |  | 01/23/2026 |  |  | 01/11/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766902 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Renal Failure | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1122 | 1220648-2026-01391 | 24161009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24161009&pc=OZD | 01/19/2026 |  |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1286795 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1123 | 1220648-2026-01393 | 24161068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24161068&pc=OZD | 01/18/2026 | 01/23/2026 |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026760962 |  | 000000000001000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1124 | 1220648-2026-01394 | 24161227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24161227&pc=PYX | 01/15/2026 |  |  | 01/23/2026 |  |  | 01/15/2026 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP PUMP SET, EU |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025594565 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male | 75 KG | Non Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp experienced excessively high placement signal values and low flows. an echocardiogram and an x-ray showed the pump was in the correct position. the pump was explanted. the patient was in critical condition. | [5 empty] | A6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1125 | 1220648-2026-01397 | 24161344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24161344&pc=OZD | 01/14/2026 |  |  | 01/23/2026 |  |  | 01/15/2026 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1085510 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 82 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an automated impella controller showed an error. the controller was swapped with another controller to continue patient support. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1126 | 1220648-2026-01398 | 24161435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24161435&pc=OZD | 01/19/2026 | 01/23/2026 |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026798839 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 72 YR |  | Female | 54 KG | Non Hispanic | White | Injury | Yes | No | Inaccurate Flow Rate | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 72-year-old female undergoing support for acute myocardial infarction and cardiogenic shock was supported with multiple impella systems during the reported hospitalization. during support with pump 1 (cp, sn (b)(6)), intermittent motor spikes and elevated left ventricular waveform pressures were observed, followed approximately ten minutes later by high pump motor current and a subsequent pump stoppage. the device was explanted on (b)(6) 2026 at 10:20 am due to suspected pump thrombosis; no product return is expected and no additional allegations were reported. a second device, pump 2 (impella 5.5, sn (b)(6)), was implanted via right axillary surgical access on (b)(6) 2026 at 9:00 am. during its support period, the system exhibited thrombosis, high or saturated pump flow, and an associated pump stop. the device was removed on (b)(6) 2026 at 3:10 pm, and product return is expected for evaluation. based on the information available, the events associated with pump 1 and pump 2 are consistent with suspected intrapump thrombosis leading to elevated motor current and pump stoppage. the root cause cannot be confirmed until product analysis is completed. pump 3 remains in use without reported malfunction. the impella cp and impella 5.5 will be conservatively reported for hemodynamic instability due to temporary hemodynamic support.interruption during the stoppage and exchange however the exchange was performed with no consequence to the patient and support was resumed without any known adverse events. | [5 empty] | The impella device has not been received from the customer; therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | E4 was inadvertently omitted on the initial manufacturer device report
1127 | 1220648-2026-01400 | 24161499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24161499&pc=OZD | 01/16/2026 |  |  | 01/23/2026 |  |  | 01/16/2026 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026726389 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male | 63 KG | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced suction alarms and a kink in the device. the pump was explanted and replaced with another impella cp. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1128 | 1220648-2026-01399 | 24161501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24161501&pc=OZD | 01/13/2026 |  |  | 01/23/2026 |  |  | 01/13/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026686656 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1129 | 1220648-2026-01401 | 24161583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24161583&pc=OZD | 01/18/2026 |  |  | 01/23/2026 |  |  | 01/18/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026788942 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1130 | 1220648-2026-01402 | 24161610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24161610&pc=OZD | 01/20/2026 |  |  | 01/23/2026 |  |  | 01/20/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2027817027 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1131 | 1220648-2026-01405 | 24161815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24161815&pc=OZD | 01/18/2026 | 02/03/2026 |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026725851 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced an alarm for a drop in purge pressure. the impella was explanted and a new impella cp was implanted. | [5 empty] | D9 updated; the product has been returned. the investigation is still ongoing. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1132 | 1220648-2026-01406 | 24162051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24162051&pc=OZD | 01/19/2026 |  |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773784 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male | 82 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced a blocked purge system. the purge was flushed without resolution. the pump was exchanged with another impella 5.5 to resolve the issue. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1133 | 1220648-2026-01407 | 24162152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24162152&pc=OZD | 01/18/2026 |  |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026729376 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male | 91 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced a purge system blocked alarm. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1134 | 1220648-2026-01412 | 24163880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24163880&pc=OZD | 01/20/2026 |  |  | 01/23/2026 |  |  | 01/20/2026 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026755748 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the device exhibited a placement signal not reliable alarm after the staffed turned the patient. the customer commented that they were going to check with an echocardiogram. flows 4.1 at p6. no changes in patient. additional information from the customer: the pump is not currently available to be sent back. the patient is still on support. we do plan and have notified the care team to save the pump for our team to send back for investigation once it is removed. | [5 empty] | D6b: added explant date. | D6b: the date of explant is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1135 | 1220648-2026-01416 | 24164646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24164646&pc=OZD | 10/06/2025 |  |  | 01/23/2026 |  |  | 10/07/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1429248 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that when attempting to start the arrhythmia induced cardiomyopathy (aic) during impella use, the self-diagnostic check failed, rendering it unusable. another aic was used to resolve the issue. additional information states that pump number one¿s control panel underwent a component swap involving an integrated circuit on the secondary controller. the work was performed over a period of several days. no symptoms or performance issues were observed during or after the activity. no revisions or replacements were required, and the system continued to operate normally without any startup concerns. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1136 | 1220648-2026-01417 | 24164654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24164654&pc=OZD | 10/05/2025 |  |  | 01/23/2026 |  |  | 10/07/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026641312 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the pump was advanced towards the heart. unfortunately, during removal of the placement wire, the impella was inadvertently withdrawn from the ventricle. additional information states that the initial impella device was implanted but required removal shortly thereafter due to a delivery-related issue. a second impella device was then implanted without reported complications and remained in use until it was subsequently removed later in the hospitalization. no additional concerns were reported for the second device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1137 | 1220648-2026-01419 | 24164759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24164759&pc=OZD | 10/07/2025 |  |  | 01/23/2026 |  |  | 10/07/2025 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026709461 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the physician reported that the patient exhibited mottling of both lower extremities, with the limbs noted to be cold and without detectable doppler signals. the patient was receiving high-dose catecholamine support. subsequently, the patient was described as hemodynamically stable but remained in critical condition, with distal lower-extremity perfusion assessed as adequate. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1138 | 1220648-2026-01425 | 24167392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24167392&pc=OZD | 01/17/2026 |  |  | 01/23/2026 |  |  | 01/17/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1583419 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1139 | 1220648-2026-01430 | 24168029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24168029&pc=OZD | 01/19/2026 | 01/23/2026 |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC.V |  |  | US |  | 2026788569 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1140 | 1220648-2026-01431 | 24168065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24168065&pc=OZD | 01/19/2026 | 01/23/2026 |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2027817200 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp was placed femorally for a 73 year old female patient admitted in with cardiogenic shock and acute myocardial infarction. the underlying and pre-existing knowledge of small vasculature prompted the team to place a fem-fem bypass prior to the cp being placed. any other cardiac and/or systemic comorbidities were not shared. on the day of implant the patient expired with hemodynamic instability. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's presenting native critical condition. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1141 | 1035166-2026-00004 | 24168070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24168070&pc=DYB |  | 01/23/2026 |  | 01/23/2026 |  |  | 01/05/2026 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | INTEGER |  |  | US |  | S9408529 |  | 0052-3006 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | We have been notified of a complaint involving an axillary insertion introducer (material#: 0052-3006, lot#: s9408529). bleeding noted during arteriotomy access. this is your notification of a damage/complaint of the axillary insertion introducer, lot: s9408529, part number: 0052-3006. 08/jan/2026 customer reported event occurred in the us. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. integer is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by integer which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, integer, or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, integer, or its employees, that the report constitutes an admission that the device, integer, or its employees caused or contributed to the event described in this report.
1142 | 1220648-2026-01432 | 24168249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24168249&pc=OZD | 01/22/2026 | 01/23/2026 |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026757636 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Restricted Flow rate | Connection Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1143 | 1220648-2026-01429 | 24168329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24168329&pc=OZD | 01/19/2026 | 01/30/2026 |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026753980 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Hematuria | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1144 | 1220648-2026-01434 | 24168638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24168638&pc=OZD | 01/19/2026 | 02/02/2026 |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026808683 |  | 1000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1145 | 1220648-2026-01433 | 24168809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24168809&pc=OZD | 01/19/2026 | 01/23/2026 |  | 01/23/2026 |  |  | 01/19/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026774078 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
1146 | 1220648-2026-01438 | 24171491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24171491&pc=OZD | 01/14/2026 |  |  | 01/23/2026 |  |  | 01/14/2026 | 01/23/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026780984 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1147 | 1220648-2026-01440 | 24171746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24171746&pc=OZD | 01/13/2026 |  |  | 01/23/2026 |  |  | 01/14/2026 | 01/23/2026 |  |  |  |  | 01/23/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026803135 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male | 75 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp was placed via the right femoral artery (/= 4 mm) to support the 53-year-old male patient who presented with acute myocardial infarction and cardiogenic shock. no other clinical details or pre-existing comorbidities were known. the patient was on systemic anticoagulation which allowed for acts between 160/180seconds and transferred to the icu for continued hemodynamic monitoring. after 2 days of support the limb accessed by the cp pump was noted to be ischemic. the leg had lost pulses. the team removed the pump and escalated to the axillary accessed impella 5.5 pump and had surgical repair performed to the cp pump access site. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1148 | 1220648-2026-01292 | 24144002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24144002&pc=OZD | 01/10/2026 |  |  | 01/22/2026 |  |  | 01/10/2026 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026634546 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1149 | 1220648-2026-01293 | 24144163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24144163&pc=OZD | 01/10/2026 |  |  | 01/22/2026 |  |  | 01/11/2026 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026791127 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1150 | 1220648-2026-01300 | 24145115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24145115&pc=OZD | 10/17/2025 |  |  | 01/22/2026 |  |  | 10/17/2025 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026758756 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Sepsis | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the impella pump had been in place for 8 days without issues. an echocardiogram performed the previous day showed a 3.4 cm position with no concerns. subsequently, the patient developed persistent suction alarms that could not be resolved. the pump appeared to have a motor issue, possibly related to thrombus formation and prolonged duration of support. a bicarbonate purge was initiated along with systemic heparin. the pump was originally scheduled for explant 4 days earlier; however, the patient became possibly septic, and the decision was made to keep the device in place through the weekend. the pump stopped functioning sometime late in the evening on (b)(6), and the team was notified the following morning. | [5 empty] | No product or logs were returned for evaluation. thrombosis, tachycardia, sepsis: clinical details noted that the patient was experiencing sepsis and tachycardia while on support. additionally, there was a concern for thrombus on the pump as well. in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pump stop: clinical details noted that the pump was experiencing low flows and spiking motor current before a pump stop occurred on the 8th day of support. the cause of the pump stop could not be determined as no product or logs were returned for evaluation. low pump flow: clinical details noted that low flows and suction alarms were occurring on the 8th day of support. suction alarms were not able to resolve and pump would not flow above 1.6 no matter the p-level. the cause of the low pump flows could not be determined as no product or logs were returned for evaluation. pump sn: (b)(6) passed all post-sterile inspection checks. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1151 | 1220648-2026-01308 | 24147712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24147712&pc=OZD | 01/06/2026 |  |  | 01/22/2026 |  |  | 01/15/2026 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026619232 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1152 | 1220648-2026-01310 | 24147899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24147899&pc=OZD | 01/19/2026 | 01/22/2026 |  | 01/22/2026 |  |  | 01/19/2026 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1285248 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 79 YR |  | Female | 72 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | This suction originated from a patient condition due to volume and not a ln automated impella controller (aic) malfunction. | The complainant reported a patient on impella cp support with an automated impella controller (aic). persistent suction occurred despite attempts at volume optimization. the suction occurred above p6. |  |  |  |  | B5 updated with additional information. d6a and d6b should have been left blank. e4 was inadvertently omitted on the initial manufacturer device report. h5 and h8 was erroneously entered on the initial manufacturer device report. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1153 | 1220648-2026-01309 | 24147908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24147908&pc=OZD | 01/09/2026 |  |  | 01/22/2026 |  |  | 01/09/2026 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782080 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1154 | 1220648-2026-01311 | 24148138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24148138&pc=OZD | 01/16/2026 |  |  | 01/22/2026 |  |  | 01/16/2026 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1696098 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Overheating of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1155 | 1220648-2026-01315 | 24149708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24149708&pc=OZD | 05/24/2025 |  |  | 01/22/2026 |  |  | 11/06/2025 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025578426 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 4 YR |  | Female | 10 KG | Non Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that information was received regarding three pediatric patients listed in the pediatric post-approval study report. one of these patients, identified as subject number twenty, was documented as having an internal sensor on the impella device that stopped functioning, which resulted in the inability to monitor the aortic and ventricular placement signal. this patient had previously been addressed in a prior complaint, which had already been closed. the complainant further reported that the patient ultimately passed away after a decision was made to withdraw care. | [5 empty] | The investigation into the reported issue has been completed. no product was received for evaluation. a paper stated that the impella' s internal sensor stopped working. looking at the case and patient updates, there was no mention of the "internal sensor not working." the only mentions of the placement signal was stated on (b)(6) where there was the ps low alarm but that was due to low diastolic blood pressure and then on (b)(6) it was stated that the placement of the device and the lv looked optimal. the data logs show a stable run throughout support with only a few instances of the placement signal low alarm at the time it was reported. there were no issues with the 5s parameters or any instances of the ps going out of range. the device passed all post sterile inspection checks. due to a lack of product returned and insufficient clinical details provided, the root cause of the optical sensor issue cannot be established.
1156 | 1220648-2026-01316 | 24150460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24150460&pc=OZD | 01/12/2026 |  |  | 01/22/2026 |  |  | 01/13/2026 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026783005 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 92 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Clinical rationale: the patient is a 63-year-old male with a past medical history significant for peripheral artery disease (pad) and peripheral vascular disease (pvd). the impella 5.5 (sn: (b)(6), (b)(4); gu case (b)(6) was surgically implanted via the right axillary/subclavian artery on (b)(6) 2026 at 1:15 pm for hemodynamic support in the setting of post-cardiotomy cardiogenic shock and low cardiac output syndrome. on (b)(6) 2026, the patient remained on v a ecmo and impella support until ecmo weaning and decannulation occurred at approximately 4:30 pm. shortly before midnight the same day, the patient lost pulses in the left leg¿previously the ecmo cannulation site. bivalirudin had been paused for decannulation. the patient was urgently taken to the operating room for a left lower-extremity thrombectomy. the patient is known to have significant underlying vascular disease prior to mechanical circulatory support. on (b)(6) 2026, the clinical team reported concern for infection due to an elevated wbc of 14; antibiotics were initiated and potentially broadened. no growth has been reported from cultures, and the impella insertion site appears clean with no signs of local infection. the device remains in situ, and the patient continues on impella support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1157 | 1220648-2026-01319 | 24151838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24151838&pc=OZD | 01/16/2026 |  |  | 01/22/2026 |  |  | 01/16/2026 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026754074 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1158 | 1220648-2026-01320 | 24152738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24152738&pc=OZD | 10/21/2025 |  |  | 01/22/2026 |  |  | 10/21/2025 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026710427 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 128 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | Renal Failure | Hematuria | [10 empty] | Death | Required Intervention |  |  |  |  | VA EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 device, escalation from an impella cp, for continued mechanical circulatory support (mcs) and experience complications. the patient would subsequently expire. an impella cp was inserted via the femoral artery in a patient who presented via emergency medical services (ems) with acute chest pain, hypotension, and flash pulmonary edema. upon ems arrival, the patient was diaphoretic, pale, and hypotensive with systolic blood pressures in the 80s. aspirin was administered, an electrocardiography test was completed, and the patient was transported to the emergency department. a st-segment elevation myocardial infarction (stemi) alert was activated, the patient was intubated, and the patient subsequently experienced cardiac arrest requiring approximately 20 minutes of cardiopulmonary resuscitation in the cardiac catheterization laboratory. due to cardiogenic shock, consistent with stage e, the decision was made to implant an impella cp for mechanical circulatory support (mcs). following impella cp implantation (day 0), the patient experienced runs of ventricular tachycardia requiring four defibrillation shocks and administration of lidocaine and amiodarone boluses. due to ongoing hemodynamic instability, the decision was made to escalate support to an impella 5.5 and initiate va ecmo. the impella 5.5 was implanted via a right axillary artery cutdown using a 10 mm × 15 cm vascutek gelweave graft. anticoagulation was achieved with heparin (activating clotting time [act] 312). the device was inserted and activated without complication; however, once on support, the impella 5.5 demonstrated restricted pump flow, suction events, and low pulsatility despite optimization attempts. hemodynamic assessment included swan-ganz catheter placement, echocardiographic evaluation, volume resuscitation, and escalation of inotropic support. despite these interventions, flows remained lower than expected. the patient was placed on veno-arterial extracorporeal membrane oxygenation (va ecmo), and the impella 5.5 was reduced to p2. in the days following implantation, the patient developed hematuria with reduced urine output (<30 cc/hr), progressed to renal failure requiring dialysis, and subsequently developed evidence of limb ischemia with absent flow in the legs and arms. the patient remained critically ill with a guarded prognosis. on day 2 of support care was withdrawn, and the patient expired. | [5 empty] | The exact date of death is unknown but has been captured as the date of explant. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | This report is being submitted to provide medical safety/clinical review. no new adverse event information is being reported. medical safety/clinical review: medical safety/clinical reviewed the event. an impella cp was inserted via the femoral artery in a [48-year-old male patient] who presented via ems with acute chest pain, hypotension, and flash pulmonary edema. the patient¿s only known comorbidity was obesity. upon ems arrival, the patient was diaphoretic, pale, and hypotensive with systolic blood pressures in the 80s. aspirin was administered, an ecg was completed, and the patient was transported to the emergency department. a stemi alert was activated, the patient was intubated, and the patient subsequently experienced cardiac arrest requiring approximately 20 minutes of cpr in the cardiac catheterization laboratory. due to cardiogenic shock, consistent with scai stage e, the decision was made to implant an impella cp for mechanical circulatory support. following impella cp implantation (day 0), the patient experienced runs of ventricular tachycardia requiring four defibrillation shocks and administration of lidocaine and amiodarone boluses. due to ongoing hemodynamic instability, the decision was made to escalate support to an impella 5.5 and initiate va ecmo. the impella 5.5 was implanted via a right axillary artery cutdown using a 10 mm × 15 cm vascutek gelweave graft. anticoagulation was achieved with heparin (activated clotting time 312 seconds). the device was inserted and activated without complication; however, once on support, the impella 5.5 demonstrated restricted pump flow, suction events despite optimization attempts. hemodynamic assessment included swan-ganz catheter placement, echocardiographic evaluation, volume resuscitation, and escalation of inotropic support. despite these interventions, flows remained lower than expected. the patient was placed on va ecmo, and the impella 5.5 was reduced to p-2. in the days following implantation, the patient developed hematuria with reduced urine output (<30 cc/hr.), progressed to renal failure requiring dialysis, and subsequently developed evidence of limb ischemia with absent flow in the legs and arms. the patient remained critically ill with a guarded prognosis. subsequently, care was withdrawn, and the patient expired. the event story involves two product complaints. impella cp - mdr 1220648-2026-01015: serious injury. impella 5.5 - mdr 1220648-2026-01320: death.
1159 | 1220648-2026-01321 | 24153056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24153056&pc=OZD | 12/22/2025 |  |  | 01/22/2026 |  |  | 12/22/2025 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026770248 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [28 empty] | Initial submission | Followup | [8 empty] | A 76-year-old female patient underwent a successful protected percutaneous coronary intervention (pci). subsequently, a retroperitoneal bleed was detected, requiring blood transfusions (complaint). the source was a ct-confirmed lesion of the uterine artery, which was successfully treated with coil embolization. bleeding events are listed in the impella ifu; however, this specific localization is unlikely a direct consequence of the impella pump. it represents a typical wire perforation injury, which can occur if the guidewire inadvertently deviates from the aorta into small side branches during catheter placement, leading to perforation. | [5 empty] | The investigation was completed and no product was returned. investigation summary ==> ppae ( retroperitoneal hemorrhage): the cause of the injury was not determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1160 | 1220648-2026-01270 | 24153424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24153424&pc=OZD | 01/16/2026 |  |  | 01/22/2026 |  |  | 01/16/2026 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026802398 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 101 KG |  |  | Injury | Yes | Yes | Fluid/Blood Leak | Decrease in Pressure | [5 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | An impella 5.5 was inserted via the right axillary subclavian artery in a 59-year-old male who presented with acute myocardial infarction complicated by cardiogenic shock. prior medical history was unknown. the patient was supported with extracorporeal membrane oxygenation, inotropes, vasopressors, and mechanical ventilation for respiratory support and was classified as scai stage d. following device placement and during closure of the access site, the patient experienced hemodynamic instability associated with a purge pressure low alarm. evaluation identified a fluid leak originating from the red or blue handle plug, with blood observed backing up into the purge side arm. purge flow was noted to be running at approximately 30 ml/hr. the clinical team was notified, and the purge solution requirements were discussed. due to the patient¿s unstable condition, the physician determined that device exchange was not feasible at that time and elected to leave the device in place while the patient was transferred to the intensive care unit for continued monitoring and management. additional information received states that bone wax and ioban were utilized during the procedure, and no blood products were administered. review of the event indicates that the reported hemodynamic instability and purge-related findings were managed clinically and are consistent with known challenges in critically ill patients receiving mechanical circulatory support in the setting of cardiogenic shock and concurrent ecmo support. comprehensive review did not identify evidence suggesting a causal relationship between the impella 5.5 device and the reported events. patient survived. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. after the impella 5.5 was placed and while the surgeon was closing the access site, purge pressure low alarm occurred. it appeared there was a leak from the red plug of the impella. physician and staff made aware and discussed the necessity of the purge solution. physician stated the patient was too unstable to exchange the device and decided to leave the pump in place and send patient to the unit. it was noted that blood was backing up into the purge side arm and purge was running at 30 ml/hr. |  |  |  |  | This case was further reviewed and was re-evaluated as serious injury the following fields were revised to reflect the change. b1, b5 revised to include clinical rationale. h1 and h6. | D6b: the explant date is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b: added explant date.
1161 | 1220648-2026-01322 | 24153482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24153482&pc=PYX | 12/16/2025 |  |  | 01/22/2026 |  |  | 12/16/2025 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1824431 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | When setting up for training, an automated impella controller continued to alarm ¿purge disc not recognized.¿ several attempts were made to utilize multiple purge cassettes. this was not found related to patient care or in a clinical setting. | [5 empty] | The device was returned. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. console (b)(6) was sent back to field service for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly.
1162 | 1220648-2026-01324 | 24157408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24157408&pc=OZD | 01/10/2026 |  |  | 01/22/2026 |  |  | 01/10/2026 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798616 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 70 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted via the right axillary artery in a 54-year-old male patient for cardiomyopathy. during preparation in the operating room, the patient underwent unintentional extracorporeal membrane oxygenation (ecmo) cannulation, resulting in prolonged operative time, escalating vasopressor requirements, and significant bleeding from the left groin arterial ecmo cutdown site. the patient developed coagulopathy during the case and required transfusion of packed red blood cells, platelets, cryoprecipitate, and fresh frozen plasma, with the most significant bleeding attributed to the ecmo access site. despite supportive measures, the patient¿s condition deteriorated and he expired. the reported events involve major bleeding, blood transfusion, and death in the setting of procedural complexity, unintended ecmo cannulation, and resultant coagulopathy. the bleeding is likely attributed to the ecmo cannulation, prolonged procedural course, and extensive transfusion requirements. the death is being conservatively reported, however, the device is unlikely to have contributed to the patient¿s death, which is attributed to ecmo-associated bleeding, coagulopathy and the patient¿s critical underlying condition. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1163 | 1220648-2026-01325 | 24157692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24157692&pc=OZD | 12/22/2025 |  |  | 01/22/2026 |  |  | 12/22/2025 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026756610 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Difficult or Delayed Positioning | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1164 | 1220648-2026-01326 | 24157766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24157766&pc=OZD | 12/22/2025 |  |  | 01/22/2026 |  |  | 12/22/2025 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026760012 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 41-year-old patient presented with stemi and underwent multivessel pci, including lad intervention. during the procedure, clinical deterioration occurred, and an impella cp was implanted according to best practices using micropuncture technique. despite mechanical circulatory support, no sustained improvement was achieved, leading to prolonged resuscitation and temporary pacemaker insertion. the patient died during the acute intervention despite all measures. this was reported as a complaint; however, mortality in cardiogenic shock remains approximately 50% even under optimal therapy and is very likely unrelated to impella. death was conservatively coded as most likely due to underlying disease and not device-related. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). (hemodynamic instability): the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1976092. device history batch: subcomponent lot: n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
1165 | 1220648-2026-01323 | 24157947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24157947&pc=OZD | 12/22/2025 |  |  | 01/22/2026 |  |  | 12/22/2025 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026734296 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 10 MO |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 72-year-old patient with acs complicated by cardiogenic shock received impella and extracorporeal membrane oxygenation (ecmo) support. limited case details suggest a critically ill patient, as indicated by high catecholamine requirements. the complaint reports discovery of a left ventricular thrombus on the first morning after impella placement, with ecmo also in place, and notes that no echocardiogram was performed prior to impella implantation. it is important to note that the ifu lists ventricular thrombus as a contraindication for impella therapy. it cannot be determined whether the thrombus was present before impella implantation (not uncommon in subacute mi), whether anticoagulation was temporarily interrupted or ineffective¿potentially promoting thrombus formation¿or whether ecmo or impella contributed to thrombus development. due to the poor prognosis, care was withdrawn and the patient expired. death was conservatively coded while most likely to be caused by the underlying disease and unrelated to impella. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. (thrombosis): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review the pump passed all the post sterile inspection checks.
1166 | 1220648-2026-01333 | 24157956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24157956&pc=OZD | 01/13/2026 |  |  | 01/22/2026 |  |  | 01/13/2026 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026765526 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Fever | Hemolysis | Unspecified Infection | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1167 | 1220648-2026-01332 | 24157981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24157981&pc=OZD | 01/17/2026 | 01/22/2026 |  | 01/22/2026 |  |  | 01/17/2026 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026654676 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1168 | 1220648-2026-01334 | 24157990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24157990&pc=OZD | 01/12/2026 | 01/22/2026 |  | 01/22/2026 |  |  | 01/13/2026 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026797960 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Perforation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1169 | 1220648-2026-01335 | 24158014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158014&pc=OZD | 01/10/2026 |  |  | 01/22/2026 |  |  | 01/11/2026 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026779244 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 device was inserted surgically into the right subclavian artery in a 71-year-old male patient presenting with acute myocardial infarction (ami) and cardiogenic shock (cgs). the patient presented in scai stage e shock, and was on extracorporeal membrane oxygenation (ecmo), multiple inotropes and vasopressors and was ventilated for respiratory support prior to initiation of support. five days post-implant, the patient was having neurological changes and found to have a catastrophic brain bleed. there was no intervention done due to the extent of the bleed. the patient was not a candidate for any escalation of care and was referred to organ donation. the patient expired on support. the impella 5.5 will be coded for stroke and conservatively reported for death; however, the device is unlikely to have contributed to the patient's death, which was most likely due to the patient's underlying critical clinical condition. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1170 | 1220648-2026-01329 | 24158063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158063&pc=OZD | 12/22/2025 |  |  | 01/22/2026 |  |  | 12/22/2025 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026701644 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [41 empty] | Initial submission
1171 | 1220648-2026-01330 | 24158084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158084&pc=OZD | 01/16/2026 | 01/22/2026 |  | 01/22/2026 |  |  | 01/16/2026 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2026798489 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male | 90 KG |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 74 year old male with an impella 5.5 due to cardiogenic shock. during support, a red alarm: ¿purge pressure high¿ (pf <1 ml/h) and pp 1070 mmhg had occurred. the line was checked for kinks and obstructions. a recommendation was made to replace the pump and tissue plasminogen activator lysis was also discussed. the patient was eventually successfully weaned. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1172 | 1220648-2026-01327 | 24158135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158135&pc=OZD | 01/16/2026 | 01/22/2026 |  | 01/22/2026 |  |  | 01/16/2026 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026798845 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 95 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old male patient on an impella 5.5 due to cardiogenic shock. it was reported that a decision was made to escalate to impella 5.5 from the cp and venoarterial extracorporeal membrane oxygenation (va ecmo). during the exchange, when the impella 5.5 was backloaded onto the wire, it was unable to pass the cannula across mid axillary/subclavian artery. the wire came out of the left venticle (lv). the lv wire access was regained with a 5f pigtail and 0.025 platinum plus wire placed across lv. when attempting again to advance the impella 5.5 through axillary/subclavian artery on 0.025 platinum plus, it would not advance past the mid axillary artery. subsequently, lv wire access was loss a total of 3 times and on the 4th attempt to backload the impella 5.5 onto the 0.025 platinum plus wire, a significant kink in the impella catheter near motor was noticed. with a kink in the catheter, there was concern that the 5.5 was compromised and a decision to abort impella 5.5 was made. a new impella 5.5 was used. the patient survived the procedure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | E4 was inadvertently omitted on the initial manufacturer device report.
1173 | 1220648-2026-01331 | 24158178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158178&pc=OZD | 01/16/2026 | 01/22/2026 |  | 01/22/2026 |  |  | 01/16/2026 | 01/22/2026 |  |  |  |  | 01/22/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1093078 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a male patient on a impella 5.5 due to acute myocardial infarction. bedside during support, there were multiple controller error alarms occurred. they did self resolve but continued overnight. staff performed an console (aic) to aic transfer today. the patient remained on support. | [5 empty] | D6a should have been left blank on the initial manufacturer device report. e4 was inadvertently omitted on the initial manufacturer device report. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1174 | 1220648-2026-01336 | 24158210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158210&pc=OZD | 01/16/2026 | 01/22/2026 |  | 01/22/2026 |  |  | 01/16/2026 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026798029 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | Hematuria | Vascular Dissection | Thrombocytopenia | [33 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1175 | 1220648-2026-01339 | 24158215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24158215&pc=OZD | 01/16/2026 | 01/21/2026 |  | 01/22/2026 |  |  | 01/16/2026 | 01/22/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026731229 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1176 | 1220648-2026-01222 | 24130384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24130384&pc=OZD | 01/06/2026 |  |  | 01/21/2026 |  |  | 01/06/2026 | 01/21/2026 |  |  |  |  | 01/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026747700 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male | 109 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: the impella 5.5 was placed as care escalation from a prior placed impella cp. the 5.5 was placed to continue the support of the 36 year old male with a history of cardiomyopathy, cardiogenic shock, and known cancer history. the pump was supporting when there was an observation made of self resolving ventricular tachycardia (vt). the team planned for a repeat of imaging for the pump positioning under echocardiography. it is unknown if pump malpositioning was the cause of any vt. | [5 empty] | The investigation was completed. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. the device history reviewed and passed all the post sterile inspection checks.
1177 | 1220648-2026-01225 | 24130666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24130666&pc=OZD | 01/14/2026 |  |  | 01/21/2026 |  |  | 01/14/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026771295 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1178 | 1220648-2026-01228 | 24130997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24130997&pc=OZD | 01/12/2026 |  |  | 01/21/2026 |  |  | 01/13/2026 | 01/21/2026 |  |  |  |  | 01/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026769164 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male | 75 KG |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella cp blood escaping from the red plug of the impella side arm. the side arm was pressed into the plug by the medical team, and the blood flow stopped. there were no further leaks, and it was noted that there were no issues with the pump flow, the purge flow, the placement signal, or pressures during the event. it was noted that there were no negative effects on therapy. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. | [5 empty] | Device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1179 | 1220648-2026-01230 | 24131310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24131310&pc=OZD | 01/10/2026 |  |  | 01/21/2026 |  |  | 01/10/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795527 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1180 | 1220648-2026-01231 | 24131360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24131360&pc=OZD | 01/14/2026 |  |  | 01/21/2026 |  |  | 01/14/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766851 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1181 | 1220648-2026-01232 | 24131565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24131565&pc=OZD | 01/10/2026 |  |  | 01/21/2026 |  |  | 01/11/2026 | 01/21/2026 |  |  |  |  | 01/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2027816961 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 88 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Cardiac Perforation | Pericardial Effusion | Heart Failure/Congestive Heart Failure | Hemodynamic instability | [8 empty] | Death | [5 empty] | DRAIN. | VENO ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Clinical review/rationale: a 58 year old male patient was admitted into the hospital in cardiogenic shock and acute myocardial infarction. the medical team first placed the patient on an intra-aortic balloon pump (iabp) for support. the iabp was placed due to multivessel coronary artery disease, but was explanted and escalated to the impella cp for a high risk percutaneous coronary intervention. after the cp was placed, the patient became hypotensive and a pericardial effusion was diagnosed. how the impella caused left ventricle perforation was not shared. the team performed pericardiocentesis, called for cardiothoracic surgery to evaluate, but the patient decompensated hours later, was placed on ecmo, however expired despite efforts. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. patient device interaction problem (pericardial effusion/cardiac perforation): the cause of this issue was not determined without returned product and sufficient clinical details. clinical symptoms (hypotension, hemodynamic instability, cardiac failure): the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
1182 | 1220648-2026-01237 | 24131935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24131935&pc=OZD | 01/12/2026 | 01/21/2026 |  | 01/21/2026 |  |  | 01/14/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026778580 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | [9 empty] | Injury | Yes | Yes | Break | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1183 | 1220648-2026-01236 | 24131957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24131957&pc=OZD | 01/08/2026 | 02/23/2026 |  | 01/21/2026 |  |  | 01/08/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773801 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Backflow | Malposition of Device | Device Sensing Problem |  |  |  |  | Stroke/CVA | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1184 | 1220648-2026-01239 | 24132261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24132261&pc=OZD | 01/14/2026 | 01/21/2026 |  | 01/21/2026 |  |  | 01/14/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026775609 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1185 | 1220648-2026-01240 | 24132292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24132292&pc=OZD | 01/14/2026 | 01/21/2026 |  | 01/21/2026 |  |  | 01/14/2026 | 01/21/2026 |  |  |  |  | 01/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026661769 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 57 YR |  | Female | 66 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella support with a 5.5 pump. the complainant reported that, due to patients anatomy and atherosclerotic vessel, the surgeon was unable to advance the device through the axillary artery and support was aborted. left axillary artery visualized under ultrasound and deemed too small to proceed. device was explanted and cut-down/closed without issue. | [5 empty] | Investigation summary: no product returned. failure to advance: the cause of the deliver issue was determined to be patient condition as the patient had small and difficult anatomy and atherosclerotic vessel preventing successful delivery of impella. device history review: device passed all post sterile inspection checks. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1186 | 1220648-2026-01243 | 24132579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24132579&pc=PYX | 01/14/2026 | 01/21/2026 |  | 01/21/2026 |  |  | 01/14/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794807 |  | 000000000001000323 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | [10 empty] | Injury | Yes | No | Device Sensing Problem | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1187 | 1220648-2026-01245 | 24132660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24132660&pc=OZD | 01/12/2026 |  |  | 01/21/2026 |  |  | 01/14/2026 | 01/21/2026 |  |  |  |  | 01/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1100888 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during the console check the console was switched on. the light turned green but the screen stayed black. the console gave an alarm. this is the same alarm sound as if the battery was switched off. no patient was involved. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1188 | 1220648-2026-01257 | 24134005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24134005&pc=OZD | 01/15/2026 |  |  | 01/21/2026 |  |  | 01/15/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026776007 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1189 | MW5182580 | 24134077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24134077&pc=OZD | 01/09/2026 | 01/14/2026 |  | 01/21/2026 |  |  |  | 01/21/2026 |  |  |  |  | 01/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | CATHETER HEART PUMP IMPELLA |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. | [6 empty] | 00813502011258 |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 51 YR |  | Male | 118 KG |  | White | Injury | No | Yes | Entrapment of Device | Difficult to Remove | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Vascular Problem | [11 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | Impella device used during cardiac intervention using the left femoral artery. at end of procedure device was being removed and became lodged in the right femoral artery requiring a cutdown with angioplasty and blood transfusion as a result.
1190 | 1220648-2026-01258 | 24134114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24134114&pc=PYX | 01/15/2026 |  |  | 01/21/2026 |  |  | 01/15/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1814736 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1191 | 1220648-2026-01259 | 24134452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24134452&pc=PYX | 01/12/2026 |  |  | 01/21/2026 |  |  | 01/14/2026 | 01/21/2026 |  |  |  |  | 01/21/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1502212 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the device was not reading the ao signal, and the customer requested a return. | [5 empty] | A: patient information is unavailable. e: initial reporter information is unavailable. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1192 | 1220648-2026-01262 | 24135002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24135002&pc=OZD | 01/14/2026 |  |  | 01/21/2026 |  |  | 01/14/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026774747 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1193 | 1220648-2026-01260 | 24135360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24135360&pc=OZD | 01/10/2026 | 01/21/2026 |  | 01/21/2026 |  |  | 01/11/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026814188 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Crack | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1194 | 1220648-2026-01263 | 24135665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24135665&pc=OZD | 01/16/2026 | 02/03/2026 |  | 01/21/2026 |  |  | 01/16/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775947 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1195 | 1220648-2026-01238 | 24136348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24136348&pc=OZD | 01/14/2026 | 01/21/2026 |  | 01/21/2026 |  |  | 01/14/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755416 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1196 | 1220648-2026-01241 | 24136363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24136363&pc=OZD | 01/14/2026 | 01/21/2026 |  | 01/21/2026 |  |  | 01/14/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026771412 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1197 | 1220648-2026-01267 | 24137464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24137464&pc=OZD | 01/12/2026 |  |  | 01/21/2026 |  |  | 01/12/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026657468 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [41 empty] | Initial submission | Followup
1198 | 1220648-2026-01268 | 24137791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24137791&pc=PYX | 01/14/2026 |  |  | 01/21/2026 |  |  | 01/16/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1991105 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | User facility | Company representation | [10 empty] | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1199 | 1220648-2026-01276 | 24140271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24140271&pc=OZD | 01/12/2026 |  |  | 01/21/2026 |  |  | 01/12/2026 | 01/21/2026 |  |  |  |  | 01/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026755550 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, the patient was found to be bleeding from the impella access site. it was confirmed that angle matching was done and pressure dressing applied. the patient was on anticoagulants and given 2 units of packed red blood cells (prbcs). the patient was successfully weaned and the device was explanted. it was reported that the patient was stable and on inotropes and pressers following the explant. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1200 | 1220648-2026-01278 | 24140785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24140785&pc=OZD | 01/11/2026 |  |  | 01/21/2026 |  |  | 01/12/2026 | 01/21/2026 |  |  | 01/31/2026 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025453897 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1201 | 1220648-2026-01281 | 24141162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24141162&pc=OZD | 01/09/2026 |  |  | 01/21/2026 |  |  | 01/12/2026 | 01/21/2026 |  |  |  |  | 01/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026769482 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp was inserted in a 83-year-old male with a known history of coronary artery disease and peripheral artery disease underwent impella support in the setting of a high-risk percutaneous coronary intervention. the patient developed ischemia. at the time of support, the patient was receiving inotropic support, vasopressors, and mechanical ventilation for respiratory support. an impella peel-away sheath was removed, and a bailout reperfusion sheath was used in a vessel with known peripheral artery disease. the impella cp will be conservatively reported for ischemia and death. the patient¿s death occurred in the setting of significant underlying comorbidities and critical illness. no information was identified to suggest a causal relationship between the impella device and the reported event. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the ischemia was unable to be determined due to insufficient clinical details. the patient¿s death occurred in the setting of significant underlying comorbidities and critical illness. no information was identified to suggest a causal relationship between the impella device and the reported event.
1202 | 1220648-2026-01282 | 24141176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24141176&pc=OZD | 01/12/2026 |  |  | 01/21/2026 |  |  | 01/12/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026692926 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Inaccurate Flow Rate | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1203 | 1220648-2026-01283 | 24141298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24141298&pc=OZD | 01/12/2026 |  |  | 01/21/2026 |  |  | 01/12/2026 | 01/21/2026 |  |  |  |  | 01/21/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026768629 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [28 empty] | Initial submission | Followup | [8 empty] | Clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 56-year-old male patient with a history of coronary artery disease who presented with cardiomyopathy. the primary clinical classification was acute decompensated chronic heart failure, with a secondary classification of non-ischemic etiology. the patient exhibited non-sustained ventricular tachycardia (nsvt). a transthoracic echocardiogram (tte) was ordered to evaluate device position but had not yet been completed at the time of reporting. the patient remains on impella 5.5 support. the reported event involves tachycardia requiring surgical intervention in the setting of advanced cardiomyopathy and ventricular arrhythmia. based on the available clinical information, the tachyarrhythmia is most plausibly related to the patient¿s underlying cardiac condition and hemodynamic instability. | Additional information received reporting the device was explanted. |  |  |  |  | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. based on the available clinical information, the tachyarrhythmia is most plausibly related to the patient¿s underlying cardiac condition and hemodynamic instability. | Additional information: b5 6b. explantation day, month and year.
1204 | 1220648-2026-01285 | 24141482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24141482&pc=OZD | 07/18/2024 | 01/27/2026 |  | 01/21/2026 |  |  | 01/13/2026 | 01/21/2026 |  |  | 01/31/2026 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025464198 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Myocardial Infarction | Thrombosis/Thrombus | Unspecified Tissue Injury | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1205 | 1220648-2026-01286 | 24141488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24141488&pc=OZD | 01/13/2026 |  |  | 01/21/2026 |  |  | 01/13/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026786891 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1206 | 1220648-2026-01287 | 24141490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24141490&pc=OZD | 01/13/2026 |  |  | 01/21/2026 |  |  | 01/13/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797216 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Increase in Pressure | [6 empty] | Thrombocytopenia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1207 | 1220648-2026-01288 | 24141498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24141498&pc=OZD | 01/12/2026 |  |  | 01/21/2026 |  |  | 01/13/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026688356 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1208 | 1220648-2026-01289 | 24141506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24141506&pc=OZD | 01/12/2026 |  |  | 01/21/2026 |  |  | 01/12/2026 | 01/21/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795352 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1209 | 1220648-2026-01105 | 24114364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24114364&pc=OZD | 01/09/2026 |  |  | 01/20/2026 |  |  | 01/09/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026730002 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 75 KG |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Death | [5 empty] | INTRA-AORTIC BALLOON PUMP (IABP) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review/rationale: the 75 year old male was admitted into the medical center for planned surgery. the patient was on a prior placed intra aortic balloon pump when the impella 5.5 was placed. the 5.5 was inserted and was supporting for 1 day when there was an observation made of hypotension and possible loss of placement signal. the team assessed pump position by echocardiographic imaging. the pump remained on for another 3 days when the patient expired. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's surgical intervention and presenting native critical condition, as presented in scai stage d. | [5 empty] | A5, a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's surgical intervention and presenting native critical condition, as presented in scai stage d.
1210 | 1220648-2026-01107 | 24114671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24114671&pc=OZD | 01/06/2026 |  |  | 01/20/2026 |  |  | 01/06/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1163557 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1211 | 1220648-2026-01108 | 24114681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24114681&pc=OZD | 01/10/2026 |  |  | 01/20/2026 |  |  | 01/10/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026800773 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | NA |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a hematoma. care was withdrawn and the patient expired. | [5 empty] | H6 e0506 added after re-evaluation: clinical rationale: an impella cp device was inserted into the left femoral artery in a male patient presenting with acute myocardial infarction (ami) and cardiogenic shock (cgs). during the procedure, access site bleeding was noted after insertion of the guidewire. about 2-3 hours post-procedure, it was noted that the patient had a hematoma at the access site while in the intensive care unit (icu). the family elected to withdraw care, and the patient expired. the impella cp will be coded for major bleed, hematoma, medical device removal, and conservatively reported for death; however, the device is unlikely to have contributed to the patient's death, which was most likely due to the underlying critical clinical condition. access site bleed is a a known risk per the device's instructions for use (ifu) due to anticoagulation and purge requirements. | A2, a4, a5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1212 | 1220648-2026-01109 | 24114688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24114688&pc=OZD | 01/10/2026 |  |  | 01/20/2026 |  |  | 01/10/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1768506 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 93 KG | Hispanic | White | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller knob requires a firm press and intermittently will not "select". no patient harm was reported. | [5 empty] | D2a and d2b were erroneously entered on the initial manufacturer device report. d6a and d6b should have been left blank on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1213 | 1220648-2026-01110 | 24114768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24114768&pc=OZD | 12/17/2025 |  |  | 01/20/2026 |  |  | 12/17/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1452168 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the device exhibited a battery failure alarm. console has been plugged up, battery switch checked. taken out of service to biomed to await arrival of service loaner from field services. | [5 empty] | A: patient information is not available. d6a and d7a: implant and explant dates are not available. e: information is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1214 | 1220648-2026-01114 | 24114879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24114879&pc=OZD | 01/12/2026 |  |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026772523 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1215 | 1220648-2026-01119 | 24115110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24115110&pc=OZD | 01/12/2026 |  |  | 01/20/2026 |  |  | 01/13/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1452267 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1216 | 1220648-2026-01120 | 24115124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24115124&pc=OZD | 12/18/2025 |  |  | 01/20/2026 |  |  | 12/18/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773623 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Unknown | 74 KG |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was being implanted with an impella 5.5 and the implant was aborted due to the patient¿s anatomy of the left axillary artery. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. failure to advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy of axillary artery along with tortuosity of vessels preventing successful delivery of impella. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | A3, a5, a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1217 | 1220648-2026-01121 | 24115341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24115341&pc=OZD | 01/13/2026 |  |  | 01/20/2026 |  |  | 01/13/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1610582 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller experienced a battery failure. the device would beep but would not turn on. the beeping resolved when the device was plugged in, and the device displayed a "battery failure" and zero percent charged. there was no patient use in the reported event as the device had a "do not use" sticker. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data review: console logs were consistent with what was reported in the complaint. device analysis: the reported battery failure alarm was unable to be reproduced. batttest telnet command was issued to monitor the battery voltages and status during charging and discharging cycles. no abnormalities were discovered that would indicate intermittent battery failure. root cause: the root cause of the reported battery failure alarm was likely because user didn¿t initially engage the transport connection battery switch prior to booting up the console. phr summary: there were no issues related to the complaint discovered when reviewing the product history of console (b)(4). | The reported event did not involve patient use. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1218 | 1220648-2026-01125 | 24115387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24115387&pc=OZD | 12/29/2025 |  |  | 01/20/2026 |  |  | 01/06/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026769787 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient had cardiogenic shock and was placed on extracorporeal membrane oxygenation (ecmo) with impella cp. after being removed from ecmo, the patient was managed with cp alone. it was reported that immediately after changing the purge fluid, a "purge pressure low" alarm sounded, and the purge pressure dropped below 300mmhg. following the instructions, the glucose concentration of the purge fluid was changed from 5% to 10%, and the alarm disappeared for approximately three hours, but then reappeared. changing the route temporarily improved the condition, but the purge pressure again decreased. the support level was then increased from p-2 to p-6, and although the purge pressure fluctuated, it improved to the point where the alarm no longer sounded, so the patient was managed in this manner overnight. the next day, the doctor was notified, and the patient was returned to p-2, and weaning was initiated. the issue was resolved by changing the support level. the removal was unrelated to this failure mode (removed because of recovery of the patient's cardiac function). | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1219 | 1220648-2026-01118 | 24115460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24115460&pc=OZD | 01/12/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026772507 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1220 | 1220648-2026-01127 | 24115500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24115500&pc=OZD | 01/13/2026 |  |  | 01/20/2026 |  |  | 01/13/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026800060 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Inaccurate Flow Rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1221 | 1220648-2026-01034 | 24115616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24115616&pc=OZD | 01/12/2026 | 01/19/2026 |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026795550 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Decrease in Pressure | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1222 | 1220648-2026-01128 | 24115734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24115734&pc=OZD | 01/12/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026784541 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1223 | 1220648-2026-01129 | 24115935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24115935&pc=OZD | 01/12/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026783574 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1224 | 1220648-2026-01130 | 24115957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24115957&pc=OZD | 01/10/2026 |  |  | 01/20/2026 |  |  | 01/10/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773631 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 98 KG | Non Hispanic | White | Death | No | Yes | Pumping Stopped | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during an attempted console exchange, the automated impella controller (aic) identified the pump as defective and would not allow the device to restart. as a result, the patient was placed on extracorporeal membrane oxygenation (ecmo). following discussion with the managing physician, the patient subsequently coded and expired, with brain death documented. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1225 | 1220648-2026-01132 | 24116228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24116228&pc=OZD | 01/13/2026 |  |  | 01/20/2026 |  |  | 01/13/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1327005 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Unknown | 80 KG |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) was being primed with an impella device when the aic had a system error for optical sensor immediately upon boot up. the aic unit was exchanged with another unit. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. | [5 empty] | A3, a5, a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | G4 was inadvertently omitted on the initial manufacturer device report. h5 and h8 was erroneously entered on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1226 | 1220648-2026-00985 | 24116271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24116271&pc=OZD | 10/18/2025 | 01/16/2026 |  | 01/20/2026 |  |  | 10/17/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026657744 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that medical doctor (md) called representative overnight for bleeding. best practices recommended and ultimately to let the bleeding stopped. patient was already planned for escalation today regardless of the bleeding. pump was removed but not for the reasons of bleeding. it was reported that the patient survived at explant. | [5 empty] | No product was returned for investigation. major bleed: bleeding occurred overnight in the groin. pressure held at site, site assessed for angle of insertion. bleeding stopped overnight. hemostasis was achieved with manual pressure. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device with sn: (b)(6) passed all post sterile inspection checks. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1227 | 1220648-2026-00982 | 24116281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24116281&pc=OZD | 10/16/2025 | 01/16/2026 |  | 01/20/2026 |  |  | 10/17/2025 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713123 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1228 | 1220648-2026-01133 | 24116293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24116293&pc=OZD | 01/08/2026 |  |  | 01/20/2026 |  |  | 01/08/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026771366 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 103 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical review rationale: an 80-year-old male underwent a cabg (coronary artery bypass grafting) procedure on (b)(6) 2026 and had direct impella 5.5 placed. during surgery of cabg times two, the patient required transfusion of three units of blood. with good faith efforts, the field representative responded in relation to events occurring during the bleeding "nothing out of the normal during cabgx2" and that heparin was given to the patient since it is required with cardiopulmonary bypass (amount unknown). bleeding is consistent with known device-related and patient-related factors documented in the instructions for use (ifu). bleeding is a foreseeable risk due to vascular access and anticoagulation requirements. | [5 empty] | Investigation summary: the investigation has been completed. the device was not returned for investigation. peri-procedural adverse event (major bleed): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
1229 | 1220648-2026-01135 | 24116307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24116307&pc=PYX | 11/06/2025 |  |  | 01/20/2026 |  |  | 11/06/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025528857 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Renal Failure | Multiple Organ Dysfunction Syndrome | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. significant hemolysis was noted while the patient was supported on bipella. the impella 5.5 position was confirmed to be appropriate, and the impella rp position was confirmed on x-ray. the patient expired. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1230 | 1220648-2026-00981 | 24116333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24116333&pc=OZD | 10/31/2025 | 01/16/2026 |  | 01/20/2026 |  |  | 11/01/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026756319 |  | 000000000001000482 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Ventricular Fibrillation | Thrombocytopenia | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the patient experienced episodes of ventricular fibrillation (vf) before, during, and after the implant procedure. a total of approximately 30 defibrillation attempts were required, along with medical therapy, to stabilize and control the situation. the clinical course was complicated by gastrointestinal ischemia that progressed to abdominal bleeding. due to the poor overall prognosis, therapy was ultimately withdrawn. the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1231 | 1220648-2026-01137 | 24116411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24116411&pc=OZD | 12/18/2025 |  |  | 01/20/2026 |  |  | 12/18/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026752204 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female | 49 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, there was bleeding at the impella insertion site. in addition, the patient was found to have hemolysis, with a new onset of red-tinged urine. no further information was provided as to the treatment; however, the patient remained on support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H 6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the root cause of the injuries could not determined due to insufficient clinical details.
1232 | 1220648-2026-01138 | 24116458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24116458&pc=OZD | 12/17/2025 |  |  | 01/20/2026 |  |  | 12/17/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026754466 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 60 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, there was a hematoma found at the insertion site. the patient was given 2 units of packed red blood cell (prbc) and pressor support was increased. no further information was provided. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product return: asae/hematoma: hematoma noted at insertion site. the cause of the access site adverse event was user relate as issue occurred after transfer to icu with act value 300 and the patient kicked leg up in bed. device history lot: device lot: 1973974. device history batch: subcomponent lot: n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
1233 | 1220648-2026-01140 | 24116505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24116505&pc=PYX | 01/13/2026 |  |  | 01/20/2026 |  |  | 01/13/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT,PACK'D, EU |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1647627 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that subsequent to performing preventative maintenance, the automated impella controller had a system self-check failure upon start up. there was no patient involvement. | [5 empty] | Section a has been left blank intentionally as there was no patient involvement. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
1234 | 1220648-2026-01141 | 24116861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24116861&pc=OZD | 01/06/2026 | 02/03/2026 |  | 01/20/2026 |  |  | 01/06/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795703 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that there was a placement signal not reliable. the patient was noted to be in stable condition. the patient remains on support. | [5 empty] | D6b updated. the investigation is ongoing. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1235 | 1220648-2026-01143 | 24117672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24117672&pc=OZD | 01/12/2026 | 01/22/2026 |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026792109 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1236 | 1220648-2026-01144 | 24117969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24117969&pc=OZD | 01/12/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026734898 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1237 | 1220648-2026-01146 | 24118045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24118045&pc=OZD | 01/12/2026 | 02/02/2026 |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610569 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Mitral Perforation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1238 | 1220648-2026-01147 | 24118363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24118363&pc=OZD | 01/12/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026807940 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 76 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient on impella support with a cp pump. the complainant reported that the patient had a traumatic intubation. a drop in hemoglobin required administration of blood products due to continuous mouth bleeding. heparin was also stopped. the patient remains intubated. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1239 | 1220648-2026-01148 | 24118437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24118437&pc=OZD | 09/17/2025 |  |  | 01/20/2026 |  |  | 01/09/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026678147 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Wound Dehiscence | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown gender, age, and race was admitted into the medical center for coronary artery bypass, and enrolled in the impact eu with an impella 5.5 for support. the pump had supported for a total of 4 days when explanted. there was adverse event reporting 4 months after the pump support. the reported adverse events were impella graft thrombosis treated by re-thoracotomy, thrombectomy and a sternal wound-healing disorder which was treated medically and healed before hospital discharge. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section a is unknown. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. wound dehiscence: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
1240 | 1220648-2026-01153 | 24118833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24118833&pc=PYX | 01/14/2026 |  |  | 01/20/2026 |  |  | 01/13/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT, PKGD, JP |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1622811 |  | 0042-0040-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller had a system self-check failure upon start up. the unit was rebooted several times, but the issue was not resolved. there was no patient involvement. the device is reportedly available for return; however, as of the time of this report the device has not yet been returned for evaluation. | [5 empty] | Section a has been left blank intentionally as there was no patient involvement. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1241 | 1220648-2026-01156 | 24119179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24119179&pc=OZD | 01/13/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/14/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026782814 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 47 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella support with a cp pump. the complainant reported that, after eight days of usage, the motor current was high and a pump stop occurred. it was not possible to start the pump again, so it was explanted. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1242 | 1220648-2026-01134 | 24119203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24119203&pc=OZD | 01/12/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026785040 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hemodynamic instability | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella support with a cp pump. the complainant reported that a patient began to decompensate after the explant. the family decided to withdraw care, and the patient eventually expired. the cause of the hemodynamic instability remains unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1243 | 1220648-2026-01150 | 24119206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24119206&pc=OZD | 01/12/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026761562 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1244 | 1220648-2026-01160 | 24119572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24119572&pc=OZD | 01/14/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/14/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739601 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1245 | 1220648-2026-01161 | 24119854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24119854&pc=OZD | 01/14/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/14/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026752376 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1246 | 1220648-2026-01163 | 24119931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24119931&pc=OZD | 01/08/2026 |  |  | 01/20/2026 |  |  | 01/08/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026730000 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1247 | 1220648-2026-01164 | 24119940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24119940&pc=OZD | 12/29/2025 |  |  | 01/20/2026 |  |  | 12/29/2025 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026780858 |  | 1000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1248 | 1220648-2026-01166 | 24120329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24120329&pc=OZD | 11/03/2025 |  |  | 01/20/2026 |  |  | 10/31/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746923 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Sepsis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced increasing wbc, they are concerned with sepsis. it was reported that the patient received volume and transthoracic echocardiogram (tte) was completed. then it was noted by nurse practitioner (np) and following up with internal medicine (ic) on-call, that right ventricle (rv) ¿not moving¿ and tricuspid annular plane excursion (tapse) 0.2. it was also reported that the right ventricle (rv) failure with impella cp having suction alarms and only able to stay at p-4 with advance practice provider (app). medical doctor (md) was being contacted by app and decided to take the patient to the operating room (or) to place on veno-arterial extracorporeal membrane oxygenation (va ecmo) due to decreasing spo2 to 90%. postponed gallbladder removal due to increasing wbc related to sepsis. it was reported that the patient expired at explant. | [5 empty] | Clinical symptoms (sepsis): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: the pump was not returned for investigation. data log shows the pump was not achieving adequate flow while running on p9 at the start of the case. pump p-level was reduced, and several suction alarms were noted with a drop in flow while running on steady p-level. no motor current spike or positioning issue was reported during the case. placement signal had a fluctuating trend throughout approximately 6 hours of the case run. clinical details indicate the patient experienced ongoing suction alarms while supported with the impella cp and was only able to maintain support at p-4. tte revealed severe right ventricular dysfunction (rv "not moving," tapse 0.2), and volume administration did not resolve the suction events. the suction alarms were likely related to right ventricular failure and associated low left-sided filling, resulting in inadequate preload for the pump. the cause of the suction issue was most likely patient condition¿related based on data log review and provided clinical details. the pump (b)(6) passed all post sterile inspection checks. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1249 | 1220648-2026-01145 | 24120422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24120422&pc=OZD | 01/12/2026 |  |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026751768 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1250 | 1220648-2026-01149 | 24120732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24120732&pc=OZD | 01/12/2026 |  |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026792185 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | Renal Failure | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1251 | 1220648-2026-01170 | 24120816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24120816&pc=OZD | 01/11/2026 |  |  | 01/20/2026 |  |  | 01/11/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795708 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1252 | 1220648-2026-01122 | 24120915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24120915&pc=OZD | 01/10/2026 |  |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026778354 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, the pump was running at p-4 when it was noticed that the purge pressure and the motor current had risen within a short period. shortly afterwards, the pump came to a stop. the pump was explanted. there was no reported harm or injury to the patient. the staff reported that clots could be observed in the cannula after the pump was explanted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1253 | 1220648-2026-01173 | 24121328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24121328&pc=OZD | 01/07/2026 |  |  | 01/20/2026 |  |  | 01/15/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1142497 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during bootup a controller error occurred. the device was troubleshooted by restarting. the device was removed due to the issue. no patient was involved. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1254 | 1220648-2026-01176 | 24122101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24122101&pc=OZD | 11/12/2025 | 12/11/2025 |  | 01/20/2026 |  |  | 11/12/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755660 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced low flow until the flow eventually stopped. the pump was explanted. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
1255 | 1220648-2026-01178 | 24122584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24122584&pc=OZD | 12/18/2025 |  |  | 01/20/2026 |  |  | 12/18/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026628993 |  | 1000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced thrombocytopenia and impella alarms for purge blocked. it was reported that the patient was not receiving anticoagulation due to thrombocytopenia. subsequently, the impella alarmed for purge blocked. the purge tubing and purge fluid were changed without resolution. the lines were checked for kinks and no kinks were observed. the medical team called abiomed customer support to request the protocol for adding tissue plasminogen activator (tpa) to the purge solution. it is unknown if tpa was added to the purge solution, and if the purge blocked alarm was resolved. the device was subsequently explanted three days later, successfully, after weaning. the patient's outcome at the time of explant was survived. | The tissue plasminogen activator protocol was initiated. it was successful. no explant was required. the patient was weaned and explanted based on clinical improvement. the pump was able to be retrieved and will be shipped for evaluation. |  |  |  |  | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | B5 additional information was received. d9 device was returned and date received was added. g1 manufacturing site name should of been left blank on the initial report that was submitted. | Product complaint # (b)(4). investigation summary: the cause of the injury could not be determined as insufficient clinical details were provided. the cause of the high purge pressure could not be determined as purge issue was not reproduced on returned product and no data logs were returned for evaluation. device history review: the pump sn (b)(6) passed all the post sterile inspection checks. | Section h6 health impact was previously incorrectly reported as f1204 permanent impairment. this has been corrected to f01 change in therapeutic response. based on currently available information, there was no evidence of irreversible deterioration or permanent damage to a body function or structure as defined by 21 cfr 803.3. this event was a temporary, reversible condition and does not meet fda criteria for permanent impairment.
1256 | 1220648-2026-01179 | 24122908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24122908&pc=OZD | 01/12/2026 |  |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026722235 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Cardiac Arrest | Hemolysis | Ischemia | Respiratory Failure | Thrombocytopenia | [33 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1257 | 1220648-2026-01181 | 24123194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24123194&pc=OZD | 01/11/2026 |  |  | 01/20/2026 |  |  | 01/13/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP. |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026735018 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1258 | 1220648-2026-01182 | 24123234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24123234&pc=OZD | 01/10/2026 |  |  | 01/20/2026 |  |  | 01/10/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1313338 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1259 | 1220648-2026-01183 | 24124323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24124323&pc=OZD | 12/19/2025 |  |  | 01/20/2026 |  |  | 12/22/2025 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026769777 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Unspecified Infection | Renal Failure | Hypovolemia | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1260 | 1220648-2026-01184 | 24124422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24124422&pc=OZD | 11/19/2025 |  |  | 01/20/2026 |  |  | 11/19/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026757849 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 89 KG | Non Hispanic | White | Death | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that multiple unsuccessful attempts were made to place an impella cardiac support device through a tortuous left axillary access pathway. during these attempts, the device became severely kinked. due to concern that the device might not function properly if advanced further, the clinical team elected to rewire the left ventricle and proceed with a new device. no additional procedural or clinical details were provided. | [5 empty] | No product was returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had calcification and tortuous left axillary preventing successful delivery of impella. product damage (catheter kink): the cause of the product was damage was most likely patient condition as the patient had calcification and tortuous left axillary preventing successful delivery of impella due to which the catheter kinked. device (b)(6) passed all post sterile inspection checks.
1261 | 1220648-2026-01186 | 24124469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24124469&pc=OZD | 11/19/2025 |  |  | 01/20/2026 |  |  | 11/20/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025610484 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 79 KG | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Anemia | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient developed a hematoma at the incision and insertion site overnight, accompanied by a decrease in hemoglobin to six grams per deciliter. during the initial insertion procedure, the clinician noted oozing at the graft site and placed additional sutures, believing the bleeding had been resolved. the surgical pocket was reported to have been closed without further issues in the operating room or upon arrival to the intensive care unit. overnight, the patient received four units of packed red blood cells in response to the drop in hemoglobin. the hemoglobin level subsequently increased to nine and one-half grams per deciliter. the insertion site remained bruised, but active bleeding was no longer present. the patient has expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1262 | 1220648-2026-01187 | 24124543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24124543&pc=OZD | 10/30/2025 | 11/28/2025 |  | 01/20/2026 |  |  | 10/30/2025 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller was used to provide support during a high-risk percutaneous coronary intervention. the impella device was inserted through the left femoral artery without difficulty, and vessel closure preparation was performed. percutaneous transluminal coronary angioplasty was conducted, followed by the use of shockwave balloons and cutting balloons in the left anterior descending artery. a decrease in pulsatility was noted during balloon inflations. three stents were implanted, and their placement was verified using intravascular ultrasound. following completion of the procedure, the impella device was gradually weaned and subsequently removed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1263 | 1220648-2026-01189 | 24124808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24124808&pc=OZD | 01/12/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/12/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026782883 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Difficult or Delayed Positioning | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1264 | 1220648-2026-01196 | 24125303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24125303&pc=OZD | 12/28/2025 |  |  | 01/20/2026 |  |  | 12/28/2025 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775753 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1265 | 1220648-2026-01202 | 24125517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24125517&pc=OZD | 11/10/2025 |  |  | 01/20/2026 |  |  | 11/10/2025 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746977 |  | 000000000001000100 |  | No |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Stroke/CVA | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | Renal Impairment | [32 empty] | Invalid/Incomplete |  |  |  | Initial submission
1266 | 1220648-2026-01207 | 24125874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24125874&pc=OZD | 02/29/2024 |  |  | 01/20/2026 |  |  | 12/01/2025 | 01/20/2026 |  |  | 06/30/2025 |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, EU |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2024356019 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the event occurred during the clinical management of the patient and was not associated with any malfunction of the device. the patient had remained on mechanical circulatory support for an extended period and developed various complications during the course of treatment. for this particular report, the clinical team identified the need to reposition the cannula because it was oriented toward the mitral valve. the team unlocked the catheter and advanced it into the intraventricular space to avoid the risk of valvular regurgitation. the user did not report any concerns to local technical support, as the event was attributed to the patient¿s condition and ongoing management rather than to a device issue. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details.
1267 | 1220648-2026-01210 | 24126043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24126043&pc=PYX | 01/13/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/13/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. |  |  | US |  | 1488654 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported by biomed at the user facility that when they attempted to insert purge cassette into the console, they received a "purge disc not detected" alarm. a second cassette was attempted with same result. the console was swapped out and the cassette performed as expected. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1268 | 1220648-2026-01211 | 24126202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24126202&pc=OZD | 01/13/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/13/2026 | 01/20/2026 |  |  |  |  | 01/20/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1148834 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 72 KG |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 60 year old male on a impella 5.5 for support during a high risk cardiac procedure. during support, a controller error occurred. the user facility considered exchanging the console but were advised not to. they have back up console on standby for now. all connections confirmed secure. no alarms except psnr at this time. patient remains on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6a should have been left blank. e4 was inadvertently omitted on the initial manufacturer device report.
1269 | 1220648-2026-01212 | 24126205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24126205&pc=OZD | 01/13/2026 | 01/20/2026 |  | 01/20/2026 |  |  | 01/13/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026727122 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1270 | 1220648-2026-01215 | 24126598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24126598&pc=OZD | 01/01/2026 |  |  | 01/20/2026 |  |  | 01/01/2026 | 01/20/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1081050 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1271 | 1220648-2026-01018 | 24106891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106891&pc=OZD | 10/23/2025 |  |  | 01/19/2026 |  |  | 10/23/2025 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026660144 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Fever | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Respiratory Failure | [33 empty] | Invalid/Incomplete |  |  |  | Initial submission
1272 | 1220648-2026-01020 | 24108636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24108636&pc=OZD | 10/24/2025 |  |  | 01/19/2026 |  |  | 10/24/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026748932 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella device. approximately two after start of the support, the patient removed the impella. the cardiologists do not want to insert a new impella for the time being because the patient appeared stable. there was no report or indication of adverse effects as a result of this event. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. device in wrong position (positioning issue) cause was determined to be an unintentional use error as the pump was pulled out of position by the patient.
1273 | 1220648-2026-01024 | 24109024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24109024&pc=OZD | 01/07/2026 |  |  | 01/19/2026 |  |  | 01/04/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026756912 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Difficult or Delayed Positioning | Difficult to Open or Close | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1274 | 1220648-2026-01025 | 24109095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24109095&pc=PYX | 01/06/2026 |  |  | 01/19/2026 |  |  | 01/06/2026 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1717562 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 84 KG | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male on an impella 5.5 for support during high-risk procedure. during support, placement signals were erratic with position alarms. there was no patient harm. the patient remains on support. | [5 empty] | Correction: d6a should have been left blank on the initial manufacturer device report. h5 and h8 was erroneously selected on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1275 | 1220648-2026-01027 | 24109381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24109381&pc=OZD | 01/08/2026 |  |  | 01/19/2026 |  |  | 01/08/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1276 | 1220648-2026-01029 | 24109491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24109491&pc=OZD | 01/06/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431034 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported their automated impella controller¿s (aic) purge disc was not recognized during setup. a new purge cassette was also not recognized. a new aic was used. | [5 empty] | Correction: d4 udi information and e4 was inadvertently omitted on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. console logs were consistent with what was reported in the complaint. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken root cause: the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly. phr summary :there were issues related to the complaint discovered when reviewing the product history of console.
1277 | 1220648-2026-01030 | 24109551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24109551&pc=OZD | 01/12/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026793766 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1278 | 1220648-2026-01031 | 24109730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24109730&pc=OZD | 01/12/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026790335 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella support with a cp pump. the complainant reported the patient had a recent history of non-st elevated myocardial infarction, percutaneous coronary intervention to the left anterior descending artery, and ventricular tachycardia/fibrillation (vt/vf). the patient endured another vt/vf arrest and was intubated while being taken to the intensive care unit (icu). multiple vt/vf episodes occurred thereafter, requiring 30-40 defibrillations in one night. the impella cp was placed, another vt episode occurred, and the patient¿s family decided to withdraw care. the patient eventually expired. | [5 empty] | Correction: d4 udi information and e4 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. ventricular fibrillation, tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer; therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1279 | 1220648-2026-01032 | 24110287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24110287&pc=OZD | 01/12/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026792954 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1280 | 1220648-2026-01033 | 24110796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24110796&pc=OZD | 01/12/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731968 |  | 1000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella support with a cp pump. the complainant reported that the cp¿s flows quickly dove and had complete suction at any p-level at or above p2. the pump was taken back across the valve and readvanced with no changes. this patient had st elevated myocardial infarction and multiple arrests prior to insertion. the patient eventually expired while on support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
1281 | 1220648-2026-01035 | 24111361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111361&pc=OZD | 01/11/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/11/2026 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026775913 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female | 47 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 56 year old female with an impella 5.5 due to acute myocardial infarction. during placement, the patient was found to have a left atria appendage clot and it was confirmed at that moment with tee. pt was given addition heparin bolus and impella was placed without difficulty using fluroscopy and run at p-7 flow 3.6 l/min. after placement, a tee was performed to confirm position and found out a clot is attached to impella inlet area and now left atrial appendage clots is no longer there. a decision was made to surgically remove the clot, the patient's chest was fully open, placed on bypass and clot was surgical removed. the patient remains on support. | [5 empty] | E4 was inadvertently omitted on the initial manufacturer device report. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1282 | 1220648-2026-01036 | 24111390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111390&pc=OZD | 01/10/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/10/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-CANADA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026791928 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1283 | 1220648-2026-01037 | 24111392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111392&pc=OZD | 01/10/2026 |  |  | 01/19/2026 |  |  | 01/10/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026774772 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Hematuria | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
1284 | 1220648-2026-01039 | 24111398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111398&pc=OZD | 01/11/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/11/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026795353 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1285 | 1220648-2026-01040 | 24111403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111403&pc=OZD | 01/11/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/11/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026775605 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1286 | 1220648-2026-01042 | 24111467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111467&pc=OZD | 01/11/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/11/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1299746 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
1287 | 1220648-2026-01043 | 24111485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111485&pc=OZD | 01/10/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026737066 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1288 | 1220648-2026-01044 | 24111507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111507&pc=PYX | 01/11/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER-EU |  |  | ABIOMED, INC. |  |  | US |  | 1653954 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during prep, both leds, the blue led and the flight-mode indicator, blink continuously. there was no patient involvement. | [5 empty] | Correction: d1 and e4 was inadvertently omitted on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1289 | 1220648-2026-01045 | 24111556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111556&pc=OZD | 01/08/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1112230 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female | 97 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 38 year old female with an impella cp due to acute myocardial infarction. during support, the impella console screen went completely black unable to see any waveforms. screen came back on its own without intervention from staff. there was a suction event just prior and some rhythm changes but the patient never lost consciousness. patient was shocked with the defibrillator. there were no other issues post pump placement after stated incident. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | D4 udi information and e4 was inadvertently omitted on the initial manufacturer device report. d6a and d6b should have been left blank on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1290 | 1220648-2026-01046 | 24111574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111574&pc=OZD | 01/12/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026774687 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1291 | 1220648-2026-01049 | 24111657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111657&pc=OZD | 01/11/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026763934 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | Yes | False Alarm | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Stroke/CVA | Ventricular Fibrillation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1292 | 1220648-2026-01047 | 24111672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111672&pc=OZD | 01/12/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736335 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombocytopenia | Thrombosis/Thrombus | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1293 | 1220648-2026-01051 | 24111685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111685&pc=OZD | 01/12/2026 | 01/19/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026792148 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Ischemia | Heart Failure/Congestive Heart Failure | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
1294 | 1220648-2026-01050 | 24111690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111690&pc=OZD | 01/12/2026 | 02/03/2026 |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026784578 |  | 1000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1295 | 1220648-2026-01054 | 24111841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111841&pc=OZD | 10/24/2025 |  |  | 01/19/2026 |  |  | 10/24/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630024 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 69 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support (mcs) and experienced ventricular tachycardia. the patient subsequently expired. the patient was noted to have multi-vessel disease, an intra-aortic balloon pump (iabp) was placed, and the patient was referred to cardiothoracic surgery. the patient underwent coronary artery bypass grafting (x5) with aortic valve replacement. post-op day 0, the patient decompensated in the intensive care unit. the patient was brought back to the operating room for emergent impella 5.5 placement and iabp removal. the impella was started on performance level p-3 and then ramped up to p-7. the patient went into ventricular tachycardia and was defibrillated one time. the physician elected for intra-aortic balloon pump augmentation. day 18 of impella mcs support noted the patient's prognosis was "very poor" and palliative team was in room with the family having end of life (eol) discussion. subsequently, care was withdrawn and the patient expired. | [5 empty] | The exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. investigation: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the ventricular fibrillation in order to make a root cause determination, all of the clinical details would have to be provided. the root cause was unable to be determined due to insufficient clinical details.
1296 | 1220648-2026-01055 | 24111961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24111961&pc=OZD | 10/23/2025 |  |  | 01/19/2026 |  |  | 10/24/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026745670 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 102 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the patient had a previous impella 5.5 device to assist with high-risk percutaneous coronary intervention (pci). shortly after the impella 5.5 device was explanted the patient decompensated. the patient was returned to the operating for a new impella 5.5. placement. through the night, the patient experienced episodes of ventricular tachycardia. the impella device was noted to have been repositioned. since early the next morning no ventricular tachycardia episodes were noted. the patient was noted to be on antiarrhythmic medication (amiodarone). the patient remained sedated and intubated. this same day it was noted family discussion was held, care was withdrawn and the patient expired. the patient was not a candidate for advanced therapies. | [5 empty] | The exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. investigation: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the tachycardia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause unable to be determined due to insufficient clinical details. related medical device: this is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the associated device.
1297 | 1220648-2026-01060 | 24112096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112096&pc=OZD | 01/11/2026 |  |  | 01/19/2026 |  |  | 01/11/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746263 |  | 1000080 |  | Device was returned to manufacturer |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission
1298 | 1220648-2026-01063 | 24112116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112116&pc=OZD | 10/31/2025 |  |  | 01/19/2026 |  |  | 10/31/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026760903 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 146 KG |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that patient decompensated and coded in the cath lab. impella was placed but unable to give adequate flows due to suction. patient was coded for 26 minutes, 8 epinephrine pushes and 11 rounds of compressions resulted in persistent pulseless electrical activity (pea) continually. medical doctors (md) spoke with family and code was stopped at around 15:39 pm. it was reported that the patient expired at explant. | [5 empty] | Clinical symptoms (cardiac arrest): the cause of the injury was not determined due to insufficient clinical details. low or blocked pump flow / pump suction: the pump was not returned for investigation, and the data log was not provided or available on the remote server. without the returned product and sufficient clinical information, the cause of the low pump flow and suction issue could not be determined. the pump (b)(6) passed all post sterile inspection checks.
1299 | 1220648-2026-01064 | 24112118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112118&pc=OZD | 11/03/2025 |  |  | 01/19/2026 |  |  | 10/31/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744160 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 93 KG |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Hypovolemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced lightheadedness due to hypovolemia, and was given 1 liter albumin. it was reported that impella weaned to p2 due to hypotensive, and impella was returned to p4. the patient was reported to have survived the procedure. | [5 empty] | No product was returned for investigation. hypovolemia/hypotension: the cause of the injury was not determined due to insufficient clinical details. device with sn: (b)(6) passed all post sterile inspection checks.
1300 | 1220648-2026-01066 | 24112123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112123&pc=OZD | 11/02/2025 |  |  | 01/19/2026 |  |  | 11/01/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026768296 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 53 KG | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced a pericardial effusion and would be going to the cardiac catheterization lab (ccl) with implanting interventional cardiologist (ic) to have it drained. it was reported that this event is not related to impella. it was reported that the patient expired at explant. | [5 empty] | Pericardial effusion: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. the pump s/n: 632538 passed all the post sterile inspection checks.
1301 | 1220648-2026-01067 | 24112129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112129&pc=OZD | 01/10/2026 |  |  | 01/19/2026 |  |  | 01/11/2026 | 01/19/2026 |  |  | 07/31/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2024376616 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1302 | 1220648-2026-01069 | 24112133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112133&pc=OZD | 11/19/2025 | 12/18/2025 |  | 01/19/2026 |  |  | 11/19/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704640 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 82 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a sixty-one-year-old male patient was receiving support with an impella cp device for cardiomyopathy. during support, the patient experienced multiple episodes of ventricular tachycardia during and after device insertion. the patient also demonstrated evidence of hemolysis based on laboratory results. epinephrine was initiated, and an echocardiogram was performed to facilitate device repositioning. the patient initially survived the procedure but later passed away while receiving palliative care. | [5 empty] | H6. investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position: clinical stated pump was shallow but unclear details provided why the pump became shallow. the root cause of the positioning issue cannot be determined due insufficient clinical data. hemolysis (miwb): the root cause of hemolysis issue most likely was improper positioning issue since hemolysis resolved after repositioning. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1303 | 1220648-2026-01071 | 24112134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112134&pc=OZD | 11/19/2025 |  |  | 01/19/2026 |  |  | 11/20/2025 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026714703 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1304 | 1220648-2026-01070 | 24112135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112135&pc=OZD | 11/19/2025 | 12/18/2025 |  | 01/19/2026 |  |  | 11/19/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768541 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a forty-four-year-old male patient was receiving support with an impella cp device for cardiomyopathy. during support, the patient developed red urine output. treatment considerations included lowering the performance level, reducing afterload, repositioning the device, and changing the purge solution to dextrose with bicarbonate. no additional treatment details were provided. the patient survived the procedure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1305 | 1220648-2026-01072 | 24112137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112137&pc=PYX | 11/20/2025 |  |  | 01/19/2026 |  |  | 11/20/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1692213 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the console unexpectedly shut off during use and requested that the unit be evaluated. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1306 | 1220648-2026-01076 | 24112153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112153&pc=OZD | 11/24/2025 |  |  | 01/19/2026 |  |  | 11/24/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765445 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 106 KG | Non Hispanic | White | Injury | Yes | No | False Alarm | Restricted Flow rate | Inaccurate Flow Rate | Malposition of Device | Device Sensing Problem | Patient Device Interaction Problem |  | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the physician entered the patient¿s room and repositioned the impella device. there was no change in the catheter motor marking, though some tension from torque buildup was relieved. the device had been positioned slightly deep. | [5 empty] | D9: added device return information. h6: added code a160105 based on information in the complaint file. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1307 | 1220648-2026-01077 | 24112161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112161&pc=PYX | 12/15/2025 |  |  | 01/19/2026 |  |  | 12/15/2025 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1906293 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1308 | 1220648-2026-01078 | 24112162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112162&pc=PYX | 12/15/2025 |  |  | 01/19/2026 |  |  | 12/15/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1701427 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that they were called to the intensive care unit (icu) due to red urine. after discussion with the intensive care unit (icu) advanced practice registered nurse (aprn), the decision was made to reposition the pump, although it had been noted to be in an excellent position. the afterload was reduced using medications, and the urine appearance showed slight improvement. the team continued to monitor the plasma free hemoglobin (pfhg) levels. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1309 | 1220648-2026-01084 | 24112171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112171&pc=OZD | 11/08/2025 |  |  | 01/19/2026 |  |  | 11/18/2025 | 01/19/2026 |  |  | 01/31/2026 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025544872 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Thrombocytopenia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1310 | 1220648-2026-01085 | 24112172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112172&pc=OZD | 11/18/2025 |  |  | 01/19/2026 |  |  | 11/18/2025 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748328 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Pain | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1311 | 1220648-2026-01086 | 24112175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112175&pc=OZD | 11/18/2025 |  |  | 01/19/2026 |  |  | 11/18/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713128 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. a ¿placement signal low¿ alarm was noted, as the placement signal did not correlate with the patient¿s blood pressure. the impella position was confirmed via bedside echocardiography. the alarm persisted despite the continued discrepancy between the placement signal and the patient¿s blood pressure. the patient survived the procedure. | [5 empty] | Investigation summary: no product was returned. the clinical details state that on the 18th of november, there were ps low alarms and it was not correlating to the patients blood pressure despite position being verified. additionally, on the 7th there were more ps low alarms with the ao diastolic in the 20-30's. the data logs show that the ps was stable throughout support with no issues seen in the 5s parameters. on 11/9 there are frequent ps low alarms but then ps monitoring is disabled. on 11/14 there are frequent ps low alarms that continuously continue throughout 11/18, but no further data logs were returned for the rest of the case. looking at an rt log during the first reported ps low alarm on 11/18, systolic values were around 50mmhg while diastolic was around 20mmhg - accurately triggering the ps low alarm. 5s parameters were stable during this time and motor current. there were no other issues noted with the pump run. due to the lack of data logs returned for the rest of the case as well as no product returned, the root cause cannot be established. device history lot: device lot: 1947668. device history batch: subcomponent: n/a. device history review: impella 611297 passed all post sterile inspection checks.
1312 | 1220648-2026-01089 | 24112184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112184&pc=OZD | 11/18/2025 |  |  | 01/19/2026 |  |  | 11/18/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765453 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the insertion site was oozing in the setting of a supratherapeutic partial thromboplastin time (ptt), and the heparin infusion was paused. the patient¿s care was subsequently withdrawn, and the patient expired. | [5 empty] | Investigation summary ==> no product return asae (major bleed) : site oozy with supratherapeutic pt, heparin paused. resolution was unknown. the cause of the access site adverse event was not determined due to insufficient clinical details. device history lot ==> device lot- 1988107. device history batch ==> subcomponent - n/a. device history review ==> the pump sn (b)(6) passed all the post sterile inspection checks.
1313 | 1220648-2026-01090 | 24112199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112199&pc=OZD | 01/08/2026 |  |  | 01/19/2026 |  |  | 01/08/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775979 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1314 | 1220648-2026-01091 | 24112200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112200&pc=OZD | 01/11/2026 |  |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026758326 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown | 120 KG |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, the pump indicated high purge pressure alarms. no other information was provided. the patient was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1315 | 1220648-2026-01092 | 24112202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112202&pc=OZD | 11/17/2025 | 02/18/2026 |  | 01/19/2026 |  |  | 11/16/2025 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1102528 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1316 | 1220648-2026-01093 | 24112203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112203&pc=OZD | 11/14/2025 |  |  | 01/19/2026 |  |  | 11/17/2025 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744195 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Respiratory Failure | Renal Impairment | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
1317 | 1220648-2026-01094 | 24112204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112204&pc=PYX | 11/11/2025 |  |  | 01/19/2026 |  |  | 11/17/2025 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1885741 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1318 | 1220648-2026-01095 | 24112205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112205&pc=OZD | 11/17/2025 |  |  | 01/19/2026 |  |  | 11/17/2025 | 01/19/2026 |  |  |  |  | 01/19/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026711568 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 88 KG | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal and lv waveforms were decoupled. pump position was confirmed to be appropriate. patient volume status and right-sided heart function were adequate. there was concern regarding the optical sensor¿s performance. customer requests the pump be returned post-explant for engineering evaluation. customer/provider also request an engineering explanation following completion of the evaluation. despite the sensor concerns, the pump continued to function within expected parameters, including the p-level, estimated flow and motor current. the patient subsequently expired. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1319 | 1220648-2026-01102 | 24112215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112215&pc=OZD | 11/17/2025 |  |  | 01/19/2026 |  |  | 11/17/2025 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1713329 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1320 | 1220648-2026-01104 | 24112259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24112259&pc=OZD | 01/11/2026 |  |  | 01/19/2026 |  |  | 01/12/2026 | 01/19/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795366 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hematuria | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1321 | 1220648-2026-01005 | 24106517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106517&pc=OZD | 10/20/2025 |  |  | 01/18/2026 |  |  | 10/21/2025 | 01/18/2026 |  |  |  |  | 01/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026729882 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation | Distributor |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, "placement signal not reliable" alarm occurred. the alarm occurred shortly after the impella device was activated. it was noted although there were no problems with the device's operation, after assessing the overall situation, the physician decided that the device could be removed. support was downgraded to an intra-aortic balloon pump (iabp), and treatment was successfully completed. there was no harm to the patient as a result of this event. | [5 empty] | Patient-specific information a2. age, a3a. sex, a3b. gender, and a4. weight is unknown at the time of this mdr writing. investigation: the product was received. no data logs were returned for analysis. device history record review was completed, and pump passed all post-sterile inspection checks. the clinical details state that shortly after starting the device, there was a placement signal not reliable (psnr) alarm. the returned product revealed that there was a kink in the catheter just distal to the anti-contamination sleeve and laser continuity test showed that the fiber was broken in that location as light was protruding from the kink. the root cause of the optical signal issue is due to a material fracture of the fiber line.
1322 | 1220648-2026-01007 | 24106647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106647&pc=OZD | 10/31/2025 |  |  | 01/18/2026 |  |  | 10/30/2025 | 01/18/2026 |  |  |  |  | 01/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026729485 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 55 KG | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for pre op placement support. patient with history of cad, previous stroke with left sided residuals, and carotid endarterectomy, had c/o shortness of breath for the last two weeks. she was initially admitted for sepsis and pneumonia. ekg found nstemi and brought to ccl where lhc revealed severe microvascular decompression patient was initially admitted to tele, but unfortunately, over night, patient's symptoms worsened and became hypotensive and requiring chemical support. patient was then intubated, started on 50 of levo and 10 of epi; brought to cvicu. pt was then scheduled today for impella placement. left axillary cut-down was performed followed by graft anastomosis. impella 5.5 was inserted through axillary sheath, but could not enter artery beyond anastomosis due to very small vessel size and calcification. after multiple attempts, surgeon removed impella 5.5 and replaced with impella cp sn # (b)(6), successfully. | [5 empty] | Product complaint # ==>(b)(4). investigation summary ==> no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had very small vessel size and calcification preventing successful delivery of impella. device history review ==> device (sn: (b)(6) passed all post sterile inspection checks. | This report is being submitted to provide the results of the medical safety/clinical review. there is no new complaint associated with the impella cp, and the device investigation was previously reported in the initial submission. no new information was identified that changes the previously submitted device investigation or its conclusions. medical safety/clinical review. medical safety/clinical reviewed the event. a 78-year-old female with a history of coronary artery disease, stroke and carotid endarterectomy presented to the hospital with shortness of breath which had been occurring over a course of two weeks. the patient was initially admitted for sepsis and pneumonia. an electrocardiogram (ekg) found non-st-segment elevation myocardial infarction. left heart catheterization revealed multi-vessel disease. the patient was accepted for surgery and was transferred to a higher-level care facility. the patient's symptoms worsened and the patient became hypotensive and requiring chemical support. the patient was intubated and impella placement took place. the initially selected 5.5 was unable to be implanted as, after multiple attempts, the device was unable to advance beyond anastomosis due to a very small vessel size and calcification. an impella cp was successfully placed without incident. the plan was for impella mechanical circulatory support (mcs) and coronary artery bypass graft (cabg) surgery the following monday, the next weekday, 2 days later. update the next day noted the patient remained hemodynamically stable. device position and function appeared stable with no alarms or evidence of complications at such time. the family ultimately decided to withdraw care and that monday planned for cabg surgery, the patient passed away. the patient's outcome occurred following the family¿s decision to withdraw life-sustaining treatment and to forgo planned therapeutic intervention.
1323 | 1220648-2026-00987 | 24106671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106671&pc=OZD | 09/09/2025 | 01/17/2026 |  | 01/18/2026 |  |  | 09/09/2025 | 01/18/2026 |  |  |  |  | 01/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026723861 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 60 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | Thrombocytopenia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 71-year-old female presented to the er with nausea and vomiting for four days. ekg revealed stemi, and the heart team was activated. the patient was intubated and transferred to the cardiac catheterization lab. past medical history included hyperlipidemia, hypertension, and type 2 diabetes mellitus. the patient experienced a complex clinical course following stemi and rca occlusion, requiring mechanical circulatory support. the first impella device failed due to kinked cannula, necessitating replacement. despite continued support with the second device, the patient developed ischemia, arrhythmia, and thrombocytopenia and ultimately expired on support. the complications of ischemia and thrombocytopenia are consistent with known device related and patient related factors documented in the instructions for use (ifu). device-related issues included low flow and cannula kink, while patient outcome was death due to underlying cardiac condition and complications. | [5 empty] | Medical safety/clinical reviewed the event. a 71-year-old female presented to the emergency department with nausea and vomiting for four days. an ecg demonstrated st-elevation myocardial infarction (stemi), and the heart team was activated. the patient was intubated and transferred to the cardiac catheterization laboratory. past medical history included hyperlipidemia, hypertension, and type 2 diabetes mellitus. coronary angiography identified a 100% occlusion of the right coronary artery, and a decision was made to intervene on the culprit vessel. during the procedure, the patient experienced cardiac arrest requiring four rounds of cpr and two defibrillation shocks. two high-dose vasopressors were initiated prior to mechanical circulatory support. an impella cp was placed for mechanical circulatory support; however, the device was noted to be kinked and was unable to achieve flows greater than 1.5 l/min. the impella cp was subsequently removed and successfully replaced with a second impella cp. on the day of replacement pump implantation, pulses were not detectable by doppler in either lower extremity. per the treating physician, this was attributed to peripheral vasoconstriction, and no immediate intervention was noted. two days later, pulses returned in the contralateral leg, while pulses remained absent in the impella access leg; however, the extremity was noted to be warm. the following day, plan of care documentation noted frequent ectopy and thrombocytopenia. as a result, the impella performance level was reduced to p-6. impella mechanical circulatory support (mcs) was continued with no plans for escalation of care due to the patient¿s prognosis. approximately two days later, life-sustaining care was withdrawn, and the patient expired. impella cp pump 1 is a malfunction and impella cp pump 2 is a death type of reportable event. h6 investigation type/finding/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | Product complaint # (b)(4) investigation summary ppae (arrhythmia, ischemia): the root cause of the injury was unable to be determined due to insufficient clinical details. (thrombocytopenia): the cause of the injury was not determined due to insufficient clinical details. device history review the pump sn (b)(6) passed all the post sterile inspection checks.
1324 | 1220648-2026-01011 | 24106673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106673&pc=OZD | 11/14/2025 |  |  | 01/18/2026 |  |  | 11/14/2025 | 01/18/2026 |  |  |  |  | 01/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1635221 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complaint at biomed reported that the automated impella controller(aic) shut down when in use with patient. the aic was replaced and no patient harm was reported. no other information is available. | [5 empty] | The patient information is not available. product complaint # (b)(4). investigation summary ==> product (b)(6) was returned to field service for investigation. data analysis: the complaint was not confirmed by log analysis. there was no evidence of the console shutting off while in use. however, ic9475 had a controller error alarm (opm comm crc error ¿ event set # 1045) triggered while running an impella. real-time (rt) logs confirmed that all signals received from the optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. given that this occurred on the console nine days before the complaint was submitted in salesforce and this was based off of a note on a console, most likely there was confusion of the placement signal disappearing with the console shutting down. device analysis: the console ran a known-good pump and purge without reproducing the momentary loss of placement signal and controller error alarm. root cause: the momentary loss of placement signal and controller error alarm was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. | E1 initial reporter phone was added as it was omitted from the initial report that was submitted. g1 manufacturing site name should of been left blank on the initial report that was submitted. h11 added product history review as it was omitted from the initial report that was submitted. product history review summary: there were no issues related to this complaint discovered when reviewing the product history of the console.
1325 | 1220648-2026-00993 | 24106674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106674&pc=OZD | 10/28/2025 | 01/17/2026 |  | 01/18/2026 |  |  | 10/28/2025 | 01/18/2026 |  |  |  |  | 01/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756401 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 64 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp in a 83 year old female patient for support of mechanical support. at the time of initial evaluation, patient in wide complex hemodynamically unstable tachycardia requiring cardioversion patient brought to cardiac catheterization laboratory (ccl) for emergent angiogram and possible mechanical circulatory support. patient arrived to ccl on levo at 8mcg, at one point required 20mcg for hypotension. dobutamine started at 5mcg in ccl. impella cp placed via radiofrequency ablation using best practices and placed in auto mode for duration of attempted percutaneous coronary intervention. unable to cross calcified left main coronary artery lesion despite multiple attempts. echocardiogram done in ccl which showed impella slightly deep, pulled back with improved waveforms on console. ao ps with minimal to no pulstaility, unchanged throughout procedure. doppler pulses on both extremities prior to case- post procedure unable to doppler pulses on both left and right lower extremities- likely needs to be warmed up per md. levo able to be weaned off completely during case. peel away exchange performed, repo sheath in place with x2 pre-close in. angle insertion maintained, patient transported out of ccl to icu. withdrawal of care, no opportunity for recovery or further intervention. | [5 empty] | No product was returned for investigation. hematoma: right groin with hematoma and pressure being held. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. ischemia: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the ischemia was unable to be determined due to insufficient clinical details. device with sn: (b)(6) passed all post sterile inspection check.
1326 | 1220648-2026-00997 | 24106675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106675&pc=OZD | 10/29/2025 | 01/17/2026 |  | 01/18/2026 |  |  | 10/29/2025 | 01/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026747132 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | Mitral Valve Insufficiency/ Regurgitation | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
1327 | 1220648-2026-01000 | 24106676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106676&pc=OZD | 10/29/2025 | 01/17/2026 |  | 01/18/2026 |  |  | 10/29/2025 | 01/18/2026 |  |  |  |  | 01/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026736443 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was planned to be implanted with an impella 5.5 for mechanical circulatory support. pump insertion attempted through right axillary graft, patient had bio prosthetic aortic valve that was stenotic with calcification that prevented all leaflets from fully opening, as well as odd angulation of the valve requiring navigation of diagnostic catheter and wire to engage and approach sideways. due to all complications of the valve, when placing the device into the body, the surgeon was unable to advance the pump across the valve. there were multiple attempts made, but no success in placing the device. due to the patient¿s acuity and unsuccessful pump placement, the decision was made to abort the procedure and cannula the patient for va ecmo. | [5 empty] | Product complaint # (b)(4). investigation summary: no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had bio prosthetic aortic valve that was stenotic with calcification that prevented all leaflets from fully opening, as well as odd angulation of the valve preventing successful delivery of impella. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
1328 | 1220648-2026-01006 | 24106677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106677&pc=OZD | 10/31/2025 | 01/18/2026 |  | 01/18/2026 |  |  | 10/30/2025 | 01/18/2026 |  |  |  |  | 01/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026764589 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 77 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient experienced elevated purge pressure , purge cassette was replaced to address the issue with no resolution. subsequently the device stopped and several attempts were unsuccessful. ct surgery decision to escalate patient to impella 5.5. patient is stable post replacement | [5 empty] | Product complaint # ==> 9b)(4). investigation summary ==> purge pressure high: the data logs confirm purge pressure high alarms. the logs show a gradual rise in purge pressure over about 20 minutes from 600mmhg to 850mmhg. values remained elevated before gradually increasing to over 1000mmhg and alarm was triggered. there were no motor current spikes prior and logs show after increase in purge pressure values, there was a 3 min bag change and a 3 min purge cassette change. pump was at p8 flowing at 3l/min and later became saturated. purge pressure values did not recover during support. the cause of the issue was determined to be biomaterial build up as evidenced by log analysis. pump stop: the purge pressure values were elevated and later there was an increase in motor current without a p-level change. pump was at p8 flowing at 3l/min before flows became saturated at about 4l/min. about 30 mins after initial flow saturation there were #13 impella stopped - motor current high alarms and unsuccessful restart attempts. the cause of the issue was determined to be biomaterial build up as evidenced by log analysis. device history review ==> the pump sn 631987 passed all post sterile inspection checks
1329 | 1220648-2026-01008 | 24106678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106678&pc=PYX | 10/31/2025 | 01/18/2026 |  | 01/18/2026 |  |  | 10/31/2025 | 01/18/2026 |  |  |  |  | 01/18/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765056 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 YR |  | Female | 37 KG |  |  | Injury | Yes | No | Restricted Flow rate | Device Sensing Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support a patient with hx of seizure disorder, autism, dcm, s/p hm3 12/2022 complicated by l middle cerebral artery stroke. patient had impella rp flex (sn (b)(6) that was removed on 10/28. since then, patient needed increase support, team decision made to place impella rp flex. physician concerns of elevated pa placement signal pressure 70, reduced flows 2.5-2.6, and purge flows of 3-3.8. reviewed rp with bedside rn. recommended placing swan for monitoring. impella reduced to p4. concerns for thrombus, tpa was ordered through purge system per dr. after discussion with medical affairs 3 units of red blood cells were administerd. patient was weaned as planned and is stable. | [5 empty] | Product complaint # (b)(4). investigation summary: rp flex & usb were returned. thrombosis: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. mechanical interaction with blood, low or blocked flow, placement signal issue: the clinical details state that there were concerns of elevated pa pressures and reduced pump flows. dark urine was noted and there was concern for thrombus so tpa was added to the purge system. the data logs show that after starting the pump there was a spike in motor current to about 1400ma following by a decrease in flows and increase in ps. motor current continued to trend up after that followed by the ps trending down with calculated flows trending up. during this time the purge pressure was slightly elevated with a decrease in purge flows. the purge pressure, motor current and flows remained variable for about 8 hours after until they stabilized. optical parameters were stable. the returned product did not have any evidence of biological material or any damages noted. the product was ran in a bucket and flows were within specification as well as the ps. the root cause of the mechanical interaction with blood is most likely due to biological material ingestion based on the data logs and clinical details provided. the root cause of the low or blocked flow is most likely due to biological material ingestion based on the data logs and clinical details provided. the root cause of the placement signal issue is most likely due to biological material ingestion based on the data logs and clinical details provided. device history review: impella 617990 passed all post sterile inspection checks.
1330 | 1220648-2026-01009 | 24106684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106684&pc=OZD | 10/31/2025 | 01/18/2026 |  | 01/18/2026 |  |  | 10/31/2025 | 01/18/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026718388 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1331 | 1220648-2026-01010 | 24106687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106687&pc=OZD | 10/31/2025 | 01/18/2026 |  | 01/18/2026 |  |  | 10/31/2025 | 01/18/2026 |  |  |  |  | 01/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026740974 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 119 KG | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient with vtach requiring defib 2 days ago brought to cardiac cath lab for left heart catheterization (lhc). lhc showed right coronary artery disease, during pci patient went into vtach requiring multiple defibs, md elected to place impella. impella prepped and inserted in rcfa. due to a patient anatomy pump was unable to advance. md elected to abort the procedure. no patient harm was reported due to the issue. | [5 empty] | Product complaint # (b)(4). investigation summary. no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. device history review. device (sn: (b)(6)) passed all post sterile inspection checks.
1332 | 1220648-2026-01012 | 24106691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106691&pc=OZD | 10/31/2025 |  |  | 01/18/2026 |  |  | 10/31/2025 | 01/18/2026 |  |  |  |  | 01/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026754106 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. md performed aortic arch replacement along w/head vessels, cabg x1 rca and decision made to place impella 5.5 as left ventricle was struggling. 5.5 prepped and inserted first without wire unsuccessfully with manual manipulation, then with a pig and .035 multiple times until successfully crossing and wiring the apex. 5.5 inserted and torqued into ideal position towards apex measuring 5.5cm. bypass weaned off and onto 5.5 successfully after a few attempts and efforts to control bleeding from the anastomosis. graft trimmed and ligated to the blue suture hub without issue and site closed. chest remained open and patient sent to the unit on support. shortly after patient arrived to the icu heart function reduced to almost nothing and bleeding was noted from the chest cavity. decision made to withdraw care and patient expired. | [5 empty] | Patient weight is unknown. product complaint#: (b)(4). investigation summary: no product was returned for investigation. major bleed: bypass weaned off and onto 5.5 successfully after a few attempts and efforts to control bleeding from the anastomosis. it was unknown whether the hemostasis was achieved. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: device with sn: 631856 passed all post sterile inspection checks.
1333 | 1220648-2026-01013 | 24106696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106696&pc=OZD | 10/20/2025 |  |  | 01/18/2026 |  |  | 10/21/2025 | 01/18/2026 |  |  |  |  | 01/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026679812 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | Death | Disability |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaint reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and during this support the patient experienced a cerebral hemorrhage and expired. the patient presented with ventricular fibrillation and was managed with extracorporeal membrane oxygenation (ecmo) and impella. the percutaneous coronary intervention (pci) was performed both the same day and the next day. heparin purge: 5% dextrose in water plus heparin 10u/1ml. the patient had anisocoria the day before the scheduled ecmo removal, and computed tomography scans revealed cerebral hemorrhage. difficulty in controlling coagulation was observed and activated partial thromboplastin time (aptt) exceeded 40 even without heparin. thrombosis was observed within the ecmo circuit; therefore, heparin was resumed. the specific date and time are unknown. arterial blood pressure was noted as 90-120/70-80 mean: 70-80. after confirming the cerebral hemorrhage, the physician decided to switch to sodium bicarbonate purging. while the cause is unknown, the physician commented that difficulty in controlling coagulation was considered a contributing factor. there was also trauma, but this was ruled out. additional information was received and noted the following: it is unknown whether the patient received a blood transfusion. fluid infusions were administered. there were no difficulties with delivery, but the colony was tortuous. although the circumflex artery was 100%, it was unable to be opened by pci after insertion. it was opened via retrograde the next day. the ruled out "trauma" was referring to head trauma that was noted upon arrival at the hospital, but the details are unknown. | [5 empty] | Patient-specific information a2. age and a3. weight is unknown at the time of this mdr writing. investigation: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the stroke, clinical details would have to be provided. the cause was unable to be determined due to insufficient clinical details.
1334 | 1220648-2026-01017 | 24106864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106864&pc=OZD | 10/21/2025 |  |  | 01/18/2026 |  |  | 10/21/2025 | 01/18/2026 |  |  |  |  | 01/18/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026738896 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 68 YR |  | Female | 78 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Death | Life Threatening | Required Intervention |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was unable to be weaned from bypass. the patient was cannulated for extracorporeal membrane oxygenation (ecmo) and impella 5.5 device was placed for left ventricular venting while on ecmo support. after successful implant of the impella 5.5, at performance level p-5, it was noted the patient was unable to be weaned/decannulated from ecmo onto sole impella support. volume was given for drop in central venous pressure. the patient was stable and sent to unit on ecpella (ecmo and impella) support. subsequently, a gastrointestinal bleed was noted and heparin was paused. discussion regarding transplant was paused during this time and it was noted continuous renal replacement therapy continued. additionally, it was noted there was no underlying cardiac rhythm currently; the patient was fully pacer dependent with intermittent ventricular tachycardia/ventricular fibrillation. eventually care was withdrawn, the patient expired and the pump was thereafter explanted. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. no product was returned for investigation. regarding the bleed, heparin was paused due to a gastrointestinal bleed. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. for the ventricular fibrillation, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details.
1335 | 1220648-2026-00988 | 24106393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106393&pc=OZD | 10/28/2025 |  |  | 01/17/2026 |  |  | 10/28/2025 | 01/17/2026 |  |  |  |  | 01/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703334 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female | 68 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. patient brought to cardiac cath lab (ccl), prepped and draped. already intubated fio2 100%, high peak pressures. pa-c placed by icu prior. currently requiring levo 80 mcg/min, dobut 5 mcg/kg/min, vaso 0.04 units/min, and amio 1 mg/min. bumex off. ketamine/fentanyl/precedex for sedation. after rcfa access gained, hr 130s quickly dropped to asystole. cpr/acls started. multiple rounds before rosc gained. md proceeded with perclose x2. abiomed 14f x 13 peel away inserted, act 214. lv accessed with 5f pigtail, lvedp 15. impella cp prepped and inserted in standard fashion without complications under fluoroscopy. guidewire removed and cp started in auto. after insertion, hr 150s to asystole x3 with inability to achieve rosc and decision made to place patient on va ecmo. . impella taken out of auto to p2 flowing 1.4l. maps 40-50s. st 110s. volume resuscitation in progress via ecmo circuit - total 500 albumin and 2 prbcs. after stabilization, dr khan proceeded with lhc - no cad noted. peel away removed and repositioning sheath advanced into place. impella position checked under fluoro and slack removed. patient experienced oozing from impella site. 1 unit packed red blood cells, and heprin was stopped. patient survived. | [5 empty] | The device or logs were not returned for investigation. the cause of the major blood was not determined since product was not returned and insufficient clinical details provided. the pump s/n: (b)(6) passed all the post sterile inspection checks.
1336 | 1220648-2026-00989 | 24106404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106404&pc=OZD | 10/28/2025 |  |  | 01/17/2026 |  |  | 10/28/2025 | 01/17/2026 |  |  |  |  | 01/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026739879 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. patient who presented on (b)(6) 2025 with stemi. provider placed 1 stent to the lad and patient was admitted to cvicu. patient condition worsened and decision was made to attempt impella cp at that time. physician was able to get the impella cp across the ao valve with great difficulty, but the cannula was kinked due to the angle of the patient¿s anatomy. staff initiated therapy and got suction immediately on auto with flows of 0.3 lpm. he was unable to position the pump without kinking from the femoral approach. decision was made to explant and try again from left axillary with a new impella cp. | [5 empty] | The pump was not returned for investigation. data log shows the pump was started twice, each time with low pump flow and active suction alarms. placement signal also confirms multiple pump insertion attempts during the case. as per the clinical details, advancing the device across the aortic valve was extremely difficult due to the patient¿s anatomical angle. the cannula became kinked during positioning attempts, leading to immediate and persistent suction alarms on auto mode with very low flows (0.3 lpm). optimal positioning could not be achieved from the femoral approach, so the device was replaced and planned for placement via the left axillary artery. difficult to deliver or advance / product damage (cannula kink): the delivery, and cannula kink issues were most likely patient condition related, attributed to challenging patient anatomy and vessel angulation, which resulted in difficulty positioning and cannula kinking. however, the exact damage to the cannula could not be verified without returned product investigation. pump suction: the data log review and clinical details were not sufficient to determine the cause of suction. therefore, the cause of the suction issue was not determined without the returned product investigation. the pump (b)(6) passed all post sterile inspection checks.
1337 | 1220648-2026-00990 | 24106405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106405&pc=OZD | 09/17/2025 |  |  | 01/17/2026 |  |  | 09/17/2025 | 01/17/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026681711 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1338 | 1220648-2026-00992 | 24106413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106413&pc=OZD | 10/28/2025 |  |  | 01/17/2026 |  |  | 10/28/2025 | 01/17/2026 |  |  |  |  | 01/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026750283 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was planned to be implanted with an impella cp device for mechanical circulatory support. patient was here with impella cp placed earlier from femoral approach and patient¿s anatomy caused the impella to kink due to angles. impella cp was unable to flow appropriately. dr. (b)(6) elected to explant and will now try percutaneous left axillary placement with new impella cp (reported under mwr: (b)(4). dr. (b)(6) access to the left axillary via percutaneous stick and placed impella 14fr short sheath. he made multiple attempts to insert the impella cp across the valve but was unable due to patient anatomy. he then removed the short sheath and inserted the long sheath. he again made multiple attempts to navigate the vascular angles and insert impella cp but was unsuccessful. impella cp aborted due to patient anatomy. | [5 empty] | Product complaint#: (b)(4). investigation summary: no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult vasculature anatomy near the aortic valve preventing successful delivery of impella. device history review: device (sn: (b)(6) passed all post sterile inspection checks.
1339 | 1220648-2026-00994 | 24106428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106428&pc=OZD | 10/28/2025 |  |  | 01/17/2026 |  |  | 10/28/2025 | 01/17/2026 |  |  |  |  | 01/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705271 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Inflammation | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. patient brought to cvor for escalation of therapy from impella cp to impella 5.5. successful axillary cutdown and implant of impella 5.5. wire first placed in lv with assistance from cardiology beside impella cp. cp was then pulled back into descending aorta and 5.5 advanced into lv with minimal time without support. 5.5 increased to p8. while on support, argatroban started due to low platelets. tongue was severely swollen as well as full body edema preventing extubating. patient became tachycardic so mirlinone being titrated down. patient was given packed red blood cells.. | [5 empty] | Tachycardia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. inflammation/ anemia: the cause of the injury was not determined due to insufficient clinical details. the pump s/n: (b)(6) passed all the post sterile inspection checks.
1340 | 1220648-2026-00995 | 24106438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106438&pc=OZD | 10/28/2025 | 11/25/2025 |  | 01/17/2026 |  |  | 10/28/2025 | 01/17/2026 |  |  |  |  | 01/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026745682 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient was presented with hx of tobacco presented to osh in cgs, w/hb. suspected myocarditis in setting positive influenza. successful implant of impella 5.5 while intra-aortic balloon pump (iabp) and extracorporeal membrane oxygenation in place. patient experienced ventricular fibrillation patient was shocked, paced, amino and lidocaine loaded and nothing has broke the patient¿s v-fib. patient expired. | [5 empty] | Ventricular fibrillation : in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details.
1341 | 1220648-2026-00996 | 24106445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106445&pc=OZD | 10/28/2025 |  |  | 01/17/2026 |  |  | 10/28/2025 | 01/17/2026 |  |  |  |  | 01/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763956 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during the removal of impella the failure of the perclose proglide suture-mediated closure devices (vcds) is strongly correlated with severe femoral artery calcification, which prevents proper knot seating or causes suture breakage. the inability to "dry close" necessitated prolonged, high-pressure manual compression to manage bleeding. patient coded on table. return of spontaneous circulation was not achieved and patient expired. | [5 empty] | Major bleed: the cause of the injury is most likely patient condition related since pt had bilateral perclose failed d/t calcification of femoral arteries. cardiac arrest: the cause of the injury was not determined due to insufficient clinical d.
1342 | 1220648-2026-00998 | 24106448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106448&pc=OZD | 10/17/2025 |  |  | 01/17/2026 |  |  | 10/17/2025 | 01/17/2026 |  |  |  |  | 01/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026737049 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Nerve Damage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with non-ischemic cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to decision to place impella 5.5 for bridge to advanced therapy. the following day the patient developed a fever that had resolved. approximately 18 days later it was noted the patient was having right arm weakness and right arm sensory issues over the last few days. the team believed this was likely a brachial plexus injury and less likely stroke since there are no other focal deficits. neurological consult initiated (outcome of the evaluation was unnoted). the patient continued on support and day 60, the patient was bridged to transplant with a survived outcome. | [5 empty] | Patient information a4. weight is unknown. investigation: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the fever and nerve compression the cause was not determined due to insufficient clinical details.
1343 | 1220648-2026-00999 | 24106451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106451&pc=OZD | 10/29/2025 |  |  | 01/17/2026 |  |  | 10/29/2025 | 01/17/2026 |  |  |  |  | 01/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025582859 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 79 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Low Blood Pressure/ Hypotension | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient had respiratory arrest with pea this morning around 7am. 24min of total cpr, ph of 7.01 and lactic of 10 at that time. lad had 100% occlusion and was ballooned. dissection noted in mid lad but stent unable to pass. rhc was preformed and impella cp placed for cardiogenic shock. patient t¿s urine output has decreased in the last few hours rn states care team holding off on crrt at this time. pulmonologist recommending not to proceed with ecmo at this time. patient developed a fever overnight of 102.6. cardiologist at bedside explaining to family the patient has a poor prognosis. patient struggling to maintain adequate bp on max impella flow and multiple inotropes. family does not wish to escalate, has realistic view on outcome and has changed patient to dnr. the physician did not believe the device was the cause or contributed to the patient¿s fever in any way. the patient was a post cardiac arrest the morning of the impella implant. it was a salvage situation when the impella was placed that evening and patient spiked fever a few hours afterwards. the family was aware of the grim prognosis and made the patient a dnr and then withdrew care completely. the device was pulled back into the aorta and turned off and remained in the aorta until the patient expired. the pump was not saved. it is unknown the source of infection could potentially be aspiration from when patient arrested and was intubated earlier in the day. | [5 empty] | Product complaint # (b)(4). investigation summary ==> fever/ hypotension: the cause of the injury was not determined due to insufficient clinical details. device history review ==> the pump sn:(b)(6) passed all post sterile inspection checks.
1344 | 1220648-2026-01003 | 24106462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106462&pc=OZD | 10/15/2025 |  |  | 01/17/2026 |  |  | 10/21/2025 | 01/17/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026735727 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Thrombosis/Thrombus | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
1345 | 1220648-2026-01004 | 24106483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24106483&pc=OZD | 10/21/2025 |  |  | 01/17/2026 |  |  | 10/21/2025 | 01/17/2026 |  |  |  |  | 01/17/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1375416 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp for mechanical circulatory support. during this support, the automated impella controller demonstrated a position waveform issue and returned to normal after monitoring. the aic was replaced when the patient was upgraded to an impella 5.5 device. there was no harm to the patient as a result of this event. | [5 empty] | Patient-information a4. weight is unknown at the time of this mdr writing. investigation: the product with event logs were received and the investigation was completed. product history review noted there were no issues related to the complaint discovered when reviewing the product history of console. data analysis: on the complaint date logs confirmed the reported failure mode given there was a controller error alarm triggered during the clinical procedure. rt logs confirmed that all signals received from optical bench (placement, signal not reliable, contrast, lamp, gain) are represented as -16384 values due to the error. device analysis: the console was tested, and the root cause of the controller error and the loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets. however, the aic software operated as designed.
1346 | 1220648-2026-00887 | 24092663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24092663&pc=OZD | 12/28/2025 |  |  | 01/16/2026 |  |  | 12/29/2025 | 01/16/2026 | [6 empty] | IMPELLA CPSA |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026763734 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1347 | 1220648-2026-00889 | 24092665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24092665&pc=OZD | 10/25/2024 |  |  | 01/16/2026 |  |  | 01/05/2026 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025547619 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with a history of severe aortic stenosis, coronary artery disease with anticipated high-risk pci during admission, chronic alcohol use, ejection fraction (ef) of 17%, chronic kidney disease stage 3, and severe vascular disease underwent high-risk aortic valve replacement. the impella 5.5 provided temporary mechanical circulatory support during high-risk surgical intervention in a patient with severe cardiac and vascular comorbidities. despite complications including atrial fibrillation with rapid ventricular response, oozing at the impella site which necessitate transfusion of blood products, and suspected stroke. at the time of pump explant, there was difficulty with removal and the team did a surgical cutdown to remove the pump. these reported events are most plausibly related to the patient's critical condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H4 was inadvertently omitted on the initial manufacturer device report number. the investigation was completed and no product returned. paroxysmal atrial fibrillation / stroke: patient had suspected l hemispheric stroke and afib while on support. the root cause of the injury was unable to be determined due to insufficient clinical details. asae/ major bleed: minor ooze noted at impella site and dressing was changed. oozed again 4 days later. the cause of the access site adverse event was not determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks.
1348 | 1220648-2026-00888 | 24092667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24092667&pc=OZD | 01/03/2026 |  |  | 01/16/2026 |  |  | 01/03/2026 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1161391 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller generated a purge system failure. the controller was exchanged for a new controller. the patient was in stable condition. | It was reported that the device also exhibited a priming problem. |  |  |  |  | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | A: added patient information . b5: added information per information in the complaint file. d4: updated udi. h6: added code per information in the complaint file.
1349 | 1220648-2026-00891 | 24092669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24092669&pc=OZD | 01/10/2026 | 02/27/2026 |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026798727 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Death | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1350 | 1220648-2026-00894 | 24092671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24092671&pc=OZD | 01/10/2026 | 01/15/2026 |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026719211 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1351 | 1220648-2026-00892 | 24092672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24092672&pc=OZD | 01/10/2026 | 01/15/2026 |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763097 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1352 | 1220648-2026-00893 | 24092673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24092673&pc=OZD | 01/10/2026 | 01/15/2026 |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1484655 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported their automated impella controller¿s (aic) placement signal was unreliable during a case. | [5 empty] | The impella device has not been received from the customer; therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1353 | 1220648-2026-00897 | 24092679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24092679&pc=OZD | 01/12/2026 |  |  | 01/16/2026 |  |  | 01/12/2026 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026709845 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced a sudden loss of placement signal. proper placement of the pump was confirmed. no patient harm was reported. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1354 | 1220648-2026-00898 | 24092688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24092688&pc=OZD | 01/07/2026 |  |  | 01/16/2026 |  |  | 01/08/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026762942 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1355 | 1220648-2026-00895 | 24092778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24092778&pc=OZD | 01/10/2026 | 01/16/2026 |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732225 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1356 | 1220648-2026-00899 | 24093630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093630&pc=OZD | 01/10/2026 |  |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026716125 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1357 | 1220648-2026-00902 | 24093637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093637&pc=PYX | 01/11/2026 |  |  | 01/16/2026 |  |  | 01/11/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1469546 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1358 | 1220648-2026-00901 | 24093638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093638&pc=OZD | 01/04/2026 |  |  | 01/16/2026 |  |  | 01/04/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1299746 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1359 | 1220648-2026-00903 | 24093640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093640&pc=OZD | 01/09/2026 |  |  | 01/16/2026 |  |  | 01/09/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795422 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1360 | 1220648-2026-00904 | 24093641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093641&pc=OZD | 01/10/2026 |  |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026761558 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Pseudoaneurysm | Vascular Dissection | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1361 | 1220648-2026-00905 | 24093643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093643&pc=OZD | 01/10/2026 |  |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775716 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1362 | 1220648-2026-00907 | 24093646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093646&pc=OZD | 01/10/2026 |  |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796231 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1363 | 1220648-2026-00909 | 24093647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093647&pc=OZD | 01/10/2026 |  |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026729996 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Cardiac Tamponade | Cardiac Perforation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1364 | 1220648-2026-00908 | 24093648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093648&pc=OZD | 12/17/2025 | 02/14/2026 |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026794569 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Tamponade | Cardiac Perforation | Pericardial Effusion | Thrombocytopenia | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that a patient was implanted with an impella cp for mechanical circulatory support and was escalated to an impella 5.5 a day later. pericardial effusion was observed on transesophageal echocardiogram after impella 5.5 placement. the patient arrived to the cardiovascular intensive care unit post impella 5.5 insertion and bleeding was noted where the previous coronary intervention was entered for tamponade. the patient continued to bleed. the patient was cold with core temperature of 33 degrees celsius. hemoglobin was 6.7, international normalized ration was 1.8 and platelets were 47. it was noted by the medical doctor that patient was on eliquis at home which was part of the hesitancy to upgrade the patient until the leg became ischemic. two units of red blood cells, two of fresh frozen plasma, and two of cryoprecipitate were ordered and given. the axillary site drainage was stable. the chest tube from the pericardial window has put out 800 milliliters in five hours and now was down to 20-30 milliliters per hour. the patient was responding well to volume and the team was able to increase p level and reduce vasopressors. ten days later, heparin-induced thrombocytopenia was noted. heparin was changed to bivalirudin. bleeding was reported to have continued. a computed tomography scan was negative for a retroperitoneal bleed. an esophagogastroduodenoscopy was planned to locate the source of bleeding. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1365 | 1220648-2026-00911 | 24093664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093664&pc=OZD | 12/17/2025 |  |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765170 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1366 | 1220648-2026-00913 | 24093688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093688&pc=OZD | 01/10/2026 |  |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026705477 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 74 YR |  | Male | 68 KG | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced gastrointestinal bleeding. | Impella 5.5 implanted in a 74-year-old male presenting with cardiogenic shock and a known history of cad. patient experienced gi bleeding. the reported gi bleeding is more plausibly related to the patient¿s critical condition but the procedural anticoagulation required for impella implant could be a contributing factor. the death is conservatively being reported to the impella 5.5 but is unlikely related and is most likely attributed to patients underlying critical condition as they presented in scai stage d shock n multiple inotropes and pressors and was ventilated for respiratory support. |  |  |  |  | A6 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: b5 amended. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
1367 | 1220648-2026-00915 | 24093724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093724&pc=OZD | 01/09/2026 |  |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737078 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1368 | 1220648-2026-00896 | 24093827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24093827&pc=OZD | 01/09/2026 | 01/16/2026 |  | 01/16/2026 |  |  | 01/10/2026 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572129 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella support with a cp pump. the complainant reported a low placement signal. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1369 | 1220648-2026-00918 | 24098150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24098150&pc=OZD | 01/09/2026 |  |  | 01/16/2026 |  |  | 01/09/2026 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026772526 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 67 KG | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient was to be implanted via the right surgical axillary/subclavian artery with an impella 5.5 for mechanical circulatory support. the cardiothoracic surgeon was unable to advance the impella 5.5 through the subclavian artery due to vessel stenosis, despite prior dilation of the vessel with a peripheral balloon. a hemashield platinum graft 10mm x 30cm was utilized. due to continued and significant resistance while trying to advance the pump, the decision was made to abort the impella 5.5 insertion and utilize an impella cp instead which was successfully implanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Correction: e4 the investigation of the reported failure mode has been completed. no product was received for evaluation. failure to advance: the cause of the failure to advance was determined to be patient condition as the patient had stenosis preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
1370 | 1220648-2026-00919 | 24098155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24098155&pc=OZD | 01/09/2026 |  |  | 01/16/2026 |  |  | 01/09/2026 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798705 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted via the right surgical axillary/subclavian artery with an impella 5.5 for mechanical circulatory support. the pump was dislodged due to patient movement and was unable to recross. the decision made to explant. the patient's outcome at time of explant was noted as survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. device in wrong position: the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position due to patient movement. device history review: the device passed all post sterile inspection checks.
1371 | 1220648-2026-00917 | 24098190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24098190&pc=OZD | 01/05/2026 |  |  | 01/16/2026 |  |  | 01/05/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026771364 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1372 | 1220648-2026-00922 | 24098281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24098281&pc=OZD | 12/18/2025 | 02/26/2026 |  | 01/16/2026 |  |  | 12/18/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1401384 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during prep a controller was powered on and then gave an ¿unexpected controller shutdown alarm¿. no pump was connected to it. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1373 | 1220648-2026-00928 | 24098492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24098492&pc=OZD | 12/19/2025 |  |  | 01/16/2026 |  |  | 12/19/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782060 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 73 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Congenital Anomaly | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. patient suffered distal limb ischemia from impella cp device prior to axillary impella 5.5 implant. during impella 5.5 support the patient's hemoglobin dropped overnight and received 2 units of packed blood cells. it was noted that the patient was on systematic heparin. this impella 5.5 device report is one of two devices associated with the event. the patient developed limb ischemia from another impella device (impella cp/serial number (b)(6), refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp. | [5 empty] | No product return ppae/major bleed: the cause of the injury was not determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1374 | 1220648-2026-00929 | 24098530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24098530&pc=OZD | 12/16/2025 |  |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026690367 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | Adverse Event Without Identified Device or Use Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup | Followup
1375 | 1220648-2026-00930 | 24098543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24098543&pc=OZD | 12/17/2025 |  |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026788419 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 79 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [28 empty] | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. after insertion of the impella cp, the patient went into ventricular tachycardia (v-tach) requiring cardioversion twice. sinus rhythm was restored and the impella was started as normal. there were no further episodes of v-tach and the patient was in stable condition. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1376 | 1220648-2026-00931 | 24101623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101623&pc=OZD | 12/17/2025 |  |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026754136 |  | 000000000001000080 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [41 empty] | Initial submission | Followup
1377 | 1220648-2026-00932 | 24101643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101643&pc=OZD | 12/17/2025 |  |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026784557 |  | 000000000001000080 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1378 | 1220648-2026-00933 | 24101646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101646&pc=OZD | 12/24/2025 |  |  | 01/16/2026 |  |  | 12/27/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026735669 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Pericardial Effusion | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1379 | 1220648-2026-00934 | 24101669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101669&pc=OZD | 12/27/2025 |  |  | 01/16/2026 |  |  | 12/27/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026760799 |  | 1000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Pumping Stopped | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1380 | 1220648-2026-00940 | 24101739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101739&pc=OZD | 12/17/2025 | 02/14/2026 |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026753131 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1381 | 1220648-2026-00943 | 24101773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101773&pc=OZD | 01/06/2026 |  |  | 01/16/2026 |  |  | 01/08/2026 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026709798 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1382 | 1220648-2026-00944 | 24101788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101788&pc=OZD | 11/21/2025 | 01/16/2026 |  | 01/16/2026 |  |  | 11/21/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722857 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella support with a 5.5 pump. the complainant reported that the patient was hypotensive, requiring increased pressor support. fevers with unknown etiology were also reported. the patient eventually expired. | [5 empty] | Investigation summary: ppae/ hypotension: the cause of the injury was not determined due to insufficient clinical details. ppae/fever: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1958866. device history batch: subcomponent no: n/a. device history review: this pump sn (b)(6) passed all post sterile inspection checks.
1383 | 1220648-2026-00945 | 24101802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101802&pc=OZD | 11/22/2025 | 01/16/2026 |  | 01/16/2026 |  |  | 11/22/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754516 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Distributor |  |  |  |  | 61 YR |  | Male | 94 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella support with a cp pump. the complainant reported the patient has continuous ventricular fibrillation runs overnight. care was eventually withdrawn, and the patient expired. | [5 empty] | Investigation summary: ppae (ventricular fibrillation): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1973937. device history batch: subcomponent lot: n/a. device history review the pump sn: (b)(6) passed all the post sterile inspection checks.
1384 | 1220648-2026-00947 | 24101841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101841&pc=OZD | 11/18/2025 |  |  | 01/16/2026 |  |  | 11/18/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026760944 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 58 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that, due to the patient¿s anatomy, the impella 5.5 could not be advanced or turned into the ascending aorta. the patient survived the procedure. | [5 empty] | Investigation summary: no product was returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had innominate artery and tortuosity preventing successful delivery of impella. device history lot: device lot: 1986731. device history batch: subcomponent lot: n/a. device history review: device (sn: 633142) passed all post sterile inspection checks.
1385 | 1220648-2026-00950 | 24101842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101842&pc=PYX | 11/11/2025 |  |  | 01/16/2026 |  |  | 11/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1885741 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | [10 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1386 | 1220648-2026-00949 | 24101843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101843&pc=OZD | 11/11/2025 | 02/03/2026 |  | 01/16/2026 |  |  | 11/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026747700 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1387 | 1220648-2026-00952 | 24101852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101852&pc=OZD | 11/18/2025 |  |  | 01/16/2026 |  |  | 11/18/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026769812 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. two covered stents were required in the right common femoral artery due to a small branch vessel perforation that occurred in the distal peripheral segment, away from the main femoral artery. there was no indication of interaction with the impella sheath, as the affected branch vessel originated near the femoral head at the initial access site and not near the sheath tip. the impella was explanted to allow placement of the covered stents after the sheath was withdrawn almost completely from the skin tract. the patient survived the procedure. | [5 empty] | Investigation summary: no product was returned. peripheral artery dissection: the cause of the injury was not determined due to insufficient clinical details. access site adverse event: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1979749. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
1388 | 1220648-2026-00951 | 24101854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101854&pc=OZD | 11/18/2025 |  |  | 01/16/2026 |  |  | 11/18/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026727551 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1389 | 1220648-2026-00954 | 24101859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101859&pc=OZD | 11/18/2025 |  |  | 01/16/2026 |  |  | 11/18/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766623 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. hemolysis was reported, and pump support was reduced from p8 to p6. hemolysis was confirmed by plasma-free hemoglobin laboratory testing and was managed by repositioning the pump. the patient subsequently expired. | [5 empty] | Investigation summary: no complaint product returned to evaluate. no logs to review. mechanical interaction with blood (hemolysis): the root cause of reported hemolysis issue cannot be determined since the complaint device not returned for analyze and insufficient clinical details provided. device history lot: lot#: 1979725. device history batch: sub-component lot#: n/a. device history review: the pump (b)(6) passed all post sterile inspection checks.
1390 | 1220648-2026-00946 | 24101888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101888&pc=OZD | 03/05/2025 |  |  | 01/16/2026 |  |  | 11/17/2025 | 01/16/2026 |  |  | 06/30/2025 |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2024373471 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female | 114 KG | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. loqi reported that the patient experienced bilateral lower limb ischemia, assessed as definitely related. the patient had an impella cp placed in the left groin, after which loss of pulses was noted in the left lower extremity by exam and doppler. arterial duplex demonstrated no flow to the limb. interventional cardiology subsequently placed a catheter from the left superficial femoral artery (sfa) to the right sfa in an attempt to restore perfusion. the impella was removed on (b)(6) 2025. due to worsening doppler findings, a two-incision, four-compartment fasciotomy of the left lower extremity was performed on (b)(6) 2025. posterior and superior compartments showed potentially viable musculature, while the lateral and anterior compartments were non-reactive and non-viable. progressive distributive shock resulted in the right lower extremity developing ischemia between 03-09 and (b)(6) 2025. vascular surgery recommended bilateral above-knee amputation (aka). the patient continued on medical management with vasoactive support, and no surgical interventions were performed. the patient subsequently expired. | [5 empty] | Investigation summary ppae (ischemia) : in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot device lot-1723784 device history batch subcomponent - n/a device history review the pump sn (b)(6) passed all the post sterile inspection checks.
1391 | 1220648-2026-00948 | 24101902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101902&pc=OZD | 11/17/2025 |  |  | 01/16/2026 |  |  | 11/17/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026753186 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 83 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the patient tested positive for heparin-induced thrombocytopenia (hit), with a platelet count of 37. the patient was not a candidate for transplant. the impella 5.5 was surgically removed, while the right ventricular assist device (rvad) remained in place to continue providing support at the time of explant. the patient survived the procedure and is currently stable. | [5 empty] | Investigation summary ppae ( thrombocytopenia): the cause of the injury was not determined due to insufficient clinical details. device history lot device lot- 1980887. device history batch subcomponent lot- n/a. device history review the pump sn (b)(6) passed all the post sterile inspection checks.
1392 | 1220648-2026-00956 | 24101974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101974&pc=OZD | 12/18/2025 |  |  | 01/16/2026 |  |  | 12/18/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025481865 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Backflow | Pumping Stopped | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient who was implanted with an impella 5.5 experienced hemolysis, a pump stop, and alarms for controller error and retrograde flow. it was reported on the first day of impella support, the patient experienced urine output that was reddish in color. it was reported that the impella was repositioned, and the hematuria subsequently resolved. additionally, it was reported that on the eleventh day of support the pump experienced an abrupt pump stop with alarms for controller error and retrograde flow. the impella was successfully restarted and the medical team reported there were no further issues. an onsite abiomed representative reviewed the device and noted that the pump was operating without issues and with appropriate pump flows for the performance level. it was additionally noted that the patient was also supported by veno-arterial extracorporeal membrane oxygenation. at the time of writing this report, the patient continues to be supported by impella 5.5. this complaint involves two impella devices: impella 5.5, with serial number (b)(6). automated impella controller with serial number (b)(6). this mdr specifically addresses impella 5.5, with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1393 | 1220648-2026-00958 | 24101984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24101984&pc=OZD | 12/17/2025 |  |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745077 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was being implanted with an impella 5.5 and the insertion was aborted. it was reported that the patient¿s anatomy was calcified and tortuous, and the impella 5.5 was unable to pass through the innominate artery. the patient was subsequently implanted with an impella cp. no information was provided on the impella cp. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product returned. difficult to advance: the cause of the delivery issue was determined to be patient condition related to the patient's anatomy, vessel tortuosity, and calcification. device history lot: device lot: 1978799. device history review: this pump sn (b)(6) passed all post sterile inspection checks.
1394 | 1220648-2026-00959 | 24102002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102002&pc=OZD | 12/18/2025 |  |  | 01/16/2026 |  |  | 12/18/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026754379 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | False Alarm | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1395 | 1220648-2026-00960 | 24102013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102013&pc=OZD | 11/29/2025 | 01/16/2026 |  | 01/16/2026 |  |  | 11/29/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026746310 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Distributor | [9 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1396 | 1220648-2026-00961 | 24102030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102030&pc=OZD | 12/19/2025 | 01/16/2026 |  | 01/16/2026 |  |  | 12/19/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026768316 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Distributor |  |  |  |  | 34 YR |  | Female | 98 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pain | Hematuria | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella support with a cp pump. the complainant reported that the patient experienced right arm pain, caused by a thrombus, and pink-tinged urine. | [5 empty] | Investigation summary: ppae: (thrombosis): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. (hematuria & pain): the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1983793. device history batch: subcomponent lot: n/a. device history review: this pump sn (b)(6) passed all post sterile inspection checks.
1397 | 1220648-2026-00963 | 24102040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102040&pc=OZD | 12/17/2025 | 01/16/2026 |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026774617 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | Yes | Difficult or Delayed Positioning | Increase in Pressure | Difficult to Open or Close |  |  |  |  | Stroke/CVA | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1398 | 1220648-2026-00964 | 24102063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102063&pc=OZD | 12/17/2025 | 01/16/2026 |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644330 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1399 | 1220648-2026-00966 | 24102079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102079&pc=OZD | 12/15/2025 | 01/16/2026 |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026737482 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella support with a 5.5 pump. the complainant reported that the impella was removed because the patient experienced occasional episodes of ventricular tachycardia (vtach) and ventricular fibrillation (vfib) while receiving support from the impella 5.5. an echocardiogram revealed that the impella was positioned at an appropriate depth within the free space of the left ventricular outflow tract. after its removal, the patient continued to have arrhythmias and sustained vtach. the patient was shocked multiple times and survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1400 | 1220648-2026-00968 | 24102096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102096&pc=OZD | 12/17/2025 | 01/16/2026 |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026754459 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1401 | 1220648-2026-00969 | 24102106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102106&pc=OZD | 12/17/2025 | 02/11/2026 |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779712 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1402 | 1220648-2026-00970 | 24102111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102111&pc=OZD | 11/12/2025 |  |  | 01/16/2026 |  |  | 11/12/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026758685 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 84 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient suffered a cardiac arrest. return of spontaneous circulation (rosc) could not be achieved. resuscitative efforts were terminated due to medical futility. the patient subsequently expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1403 | 1220648-2026-00971 | 24102124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102124&pc=OZD | 11/15/2025 |  |  | 01/16/2026 |  |  | 11/15/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026771481 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 73 KG | Non Hispanic | Black Or African American | Death | Yes | No | Restricted Flow rate | [6 empty] | Hemoptysis | Low Blood Pressure/ Hypotension | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was decompensated overnight with hypotension, decreased urine output, worsening laboratory values, and frequent suction alarms. impella flow was reduced to p2 due to suction events. echocardiogram confirmed appropriate impella positioning. the patient subsequently expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1404 | 1220648-2026-00972 | 24102372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102372&pc=PYX | 11/12/2025 |  |  | 01/16/2026 |  |  | 11/12/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1666954 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1405 | 1220648-2026-00973 | 24102451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102451&pc=OZD | 11/16/2025 |  |  | 01/16/2026 |  |  | 11/16/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026693049 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 67 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was worsened after placement of a left anterior descending (lad) artery stent. an impella device was inserted to support cardiac function. the patient required multiple rounds of cardiopulmonary resuscitation (cpr) and return of spontaneous circulation (rosc) could not be achieved, despite resuscitation efforts. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1406 | 1220648-2026-00974 | 24102508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102508&pc=OZD | 11/14/2025 |  |  | 01/16/2026 |  |  | 11/16/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026738922 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female | 82 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | Respiratory Failure | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient developed acute stroke symptoms, including increasing confusion and diaphoresis. a computed tomography (ct) scan was completed, revealing a hemorrhagic stroke. during the return to ct for a repeat scan and computed tomography angiography (cta), the patient experienced respiratory failure, requiring intubation and an epinephrine push. protamine was administered to reverse the heparin infusion, which has since been held. a repeat ct scan is scheduled at 1200. the patient's outcome at the time of explant was expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1407 | 1220648-2026-00976 | 24102529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102529&pc=OZD | 11/15/2025 |  |  | 01/16/2026 |  |  | 11/15/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026692930 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the bedside staff reported ectopy overnight and the intensivist performed a point-of-care ultrasound (pocus), finding the impella to be slightly deep. the impella support level was reduced from p3 to p2 in response. during morning rounds, cardiac surgery (cs) reviewed the case and decided not to adjust the impella depth at this time. ectopy improved by the next morning. the impella remained at p2, though with some episodes of preventing retrograde flow. the patient required cardioversion at that morning for rapid atrial fibrillation (afib). the patient's outcome at the time of explant was survived and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1408 | 1220648-2026-00977 | 24102559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102559&pc=OZD | 11/14/2025 |  |  | 01/16/2026 |  |  | 11/14/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025561197 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | Yes | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1409 | 1220648-2026-00978 | 24102612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102612&pc=OZD | 11/13/2025 |  |  | 01/16/2026 |  |  | 11/13/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1135097 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1410 | 1220648-2026-00979 | 24102674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102674&pc=OZD | 11/01/2025 |  |  | 01/16/2026 |  |  | 11/01/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026763829 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 104 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Tachycardia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the patient experienced cardiac arrest and was placed on veno-arterial extracorporeal membrane oxygenation (va ecmo). the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1411 | 1220648-2026-00980 | 24102690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24102690&pc=PYX | 11/01/2025 |  |  | 01/16/2026 |  |  | 11/01/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1570740 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient had been implanted with an impella device for mechanical circulatory support. during support, a controller failure alarm occurred, and none of the soft buttons were functional. no additional information was provided regarding how the error was resolved, and no patient data were available. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1412 | 1220648-2026-00983 | 24103321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24103321&pc=OZD | 10/17/2025 |  |  | 01/16/2026 |  |  | 10/17/2025 | 01/16/2026 |  |  |  |  | 01/16/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026711434 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that there was a failed delivery. the device was subsequently removed. the patient did not exhibit any signs or symptoms associated with the event. | [5 empty] | No product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy and aortic arch calcification preventing successful delivery of impella. device (sn: (b)(6)) passed all post sterile inspection checks. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1413 | 1220648-2026-00984 | 24103324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24103324&pc=OZD | 10/17/2025 |  |  | 01/16/2026 |  |  | 10/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1401651 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1414 | 1220648-2026-00962 | 24103334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24103334&pc=OZD | 12/17/2025 | 01/29/2026 |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720572 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation | [10 empty] | Death | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Arrest | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1415 | 1220648-2026-00965 | 24103339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24103339&pc=OZD | 12/17/2025 | 01/16/2026 |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026746387 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1416 | 1220648-2026-00967 | 24103344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24103344&pc=OZD | 12/17/2025 | 01/16/2026 |  | 01/16/2026 |  |  | 12/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026765807 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1417 | 1220648-2026-00955 | 24103346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24103346&pc=OZD | 08/04/2025 | 01/16/2026 |  | 01/16/2026 |  |  | 10/17/2025 | 01/16/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026699466 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Atrial Fibrillation | Low Blood Pressure/ Hypotension | Hematuria | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
1418 | 1220648-2026-00798 | 24079475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079475&pc=OZD | 09/26/2025 | 01/14/2026 |  | 01/15/2026 |  |  | 09/26/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026738911 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed ventricular tachycardia while on impella support at p9. the family was informed and subsequently elected to withdraw life-sustaining therapy. the impella device was retracted from across the aortic valve as part of the withdrawal process. the patient expired shortly thereafter, and the device was explanted. | [5 empty] | Medical safety/clinical reviewed the event. the patient presented to the emergency department on (b)(6) 2025 in an unresponsive state, requiring immediate intubation. no advance cardiopulmonary life support was required. electrocardiogram demonstrated st-elevation myocardial infarction (stemi). the patient was emergently taken to the cardiac catheterization laboratory. coronary angiography revealed calcified ostial left main (lm) disease, chronic total occlusion of the left anterior descending, and ostial right coronary artery disease. due to hemodynamic instability and critical lm disease, the physician elected to place an impella cp (pump 1) for mechanical circulatory support. after approximately three hours of intervention, one stent was placed in the left main artery. the patient remained on high-dose vasoactive medications and impella cp support at p-9. given persistent cardiogenic shock, the care team planned escalation to veno-arterial extracorporeal membrane oxygenation (va-ecmo) with impella support and transfer to a higher level of care. a prior echocardiogram performed approximately one month before the stemi showed an ejection fraction of 25%. later the same day, high purge pressure was noted and the purge system was found to be blocked. upon transfer to the receiving facility, an ¿air in purge¿ alarm was present. de-airing the purge system and exchanging the purge tubing and cassette did not resolve the alarm. tpa purge therapy was initiated per protocol, which is noted as off label use. the patient developed bilateral lower-extremity pulse deficits, and urine output became peach-pink to red in appearance. bedside echocardiography demonstrated the impella cp was positioned slightly deep at approximately 4.4 cm. the device was repositioned to a more optimal location. oozing was noted at the right femoral impella access site. the impella angle was corrected, and the device was initially weaned with plans for explant; however, echocardiography demonstrated worsening mitral regurgitation during weaning, and the decision was made to leave the pump in place despite recognized risks. the patient was extubated, vasopressors were discontinued, and afterload reduction was initiated. the access site was clean, dry, and intact, and left lower-extremity pulses were dopplerable. high purge pressure persisted, and tpa purge therapy was continued for two additional days. the patient subsequently developed frequent suction alarms with declining pressures, followed by pump stoppage. the impella cp (pump 1) was explanted due to pump stop in the setting of unresolved high purge pressure and replaced with a second impella cp device. the impella cp (pump 2) was implanted on (b)(6) 2025. following implantation, the patient developed ventricular tachycardia. after discussion with the care team, the family elected to withdraw care. the device was withdrawn across the aortic valve, care was withdrawn, and the patient expired shortly thereafter. impella cp pump 1 is a serious type of reportable event and impella cp pump 2 is a death. note: h6: investigation type/findings/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1419 | 1220648-2026-00794 | 24079478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079478&pc=OZD | 09/25/2025 | 01/14/2026 |  | 01/15/2026 |  |  | 09/25/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719081 |  | 000000000001000482 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. p-00736031 ((b)(6) 2025): the patient was awake and sedated in preparation for cardioversion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1420 | 1220648-2026-00805 | 24079706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079706&pc=OZD | 12/16/2025 | 01/15/2026 |  | 01/15/2026 |  |  | 12/16/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1651078 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Open or Close | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported their automated impella controller¿s (aic) purge door release button fails to open the door. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
1421 | 1220648-2026-00804 | 24079708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079708&pc=OZD | 01/09/2026 | 01/15/2026 |  | 01/15/2026 |  |  | 01/09/2026 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682803 |  | 000000000001000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1422 | 1220648-2026-00807 | 24079717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079717&pc=OZD | 01/07/2026 |  |  | 01/15/2026 |  |  | 01/07/2026 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798612 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1423 | 1220648-2026-00808 | 24079718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079718&pc=OZD | 12/19/2025 |  |  | 01/15/2026 |  |  | 12/19/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1084302 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) was turned on for the case, a ¿controller error¿ notification banner appeared. the aic was turned off and back on, but the same notification reappeared. the battery was then powered off and back on from underneath the console, which resolved the alarm. a new aic was obtained from the supply room and used for the case. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1424 | 1220648-2026-00809 | 24079727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079727&pc=OZD | 12/16/2025 |  |  | 01/15/2026 |  |  | 12/16/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1147694 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the purge disc reader on automated impella controller was not functioning. no further information was available at the time of reporting. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section a is unknown. section e initial reporter is unknown. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. console logs were consistent with what was reported in the complaint. device analysis: the console was sent back for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. root cause: the purge disc flag was observed to be broken and was the root cause of the console's inability to detect the purge disc. phr summary: there were no issues related to the complaint discovered when reviewing the product history of console.
1425 | 1220648-2026-00811 | 24079729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079729&pc=OZD | 12/22/2025 |  |  | 01/15/2026 |  |  | 12/22/2025 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026749022 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1426 | 1220648-2026-00812 | 24079744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079744&pc=OZD | 12/19/2025 |  |  | 01/15/2026 |  |  | 12/19/2025 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748305 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1427 | 1220648-2026-00816 | 24079790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079790&pc=OZD | 01/08/2026 |  |  | 01/15/2026 |  |  | 01/08/2026 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026726086 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1428 | 1220648-2026-00817 | 24079792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079792&pc=OZD | 01/09/2026 |  |  | 01/15/2026 |  |  | 01/09/2026 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797224 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 77 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 could not be completed due to patient anatomy. a buddy wire was used without success. the patient was implanted with an intra aortic balloon pump instead. no patient harm was reported. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. failure to advance: the cause of the delivery issue was determined to be patient condition related to patient anatomy, vessel tortuosity. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1429 | 1220648-2026-00819 | 24079796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079796&pc=OZD | 01/09/2026 |  |  | 01/15/2026 |  |  | 01/09/2026 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026726072 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male | 79 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp lost placement signal on the automated impella controller. the pump was replaced for a new pump to continue patient support. no patient harm was reported. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1430 | 1220648-2026-00820 | 24079797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079797&pc=OZD | 01/09/2026 |  |  | 01/15/2026 |  |  | 01/09/2026 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026754087 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1431 | 1220648-2026-00828 | 24079817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079817&pc=OZD | 01/08/2026 |  |  | 01/15/2026 |  |  | 01/08/2026 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797265 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1432 | 1220648-2026-00830 | 24079818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079818&pc=OZD | 01/09/2026 |  |  | 01/15/2026 |  |  | 01/09/2026 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026750033 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1433 | 1220648-2026-00827 | 24079819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079819&pc=OZD | 01/08/2026 |  |  | 01/15/2026 |  |  | 01/08/2026 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796350 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1434 | 1220648-2026-00832 | 24079835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24079835&pc=OZD | 01/08/2026 |  |  | 01/15/2026 |  |  | 01/09/2026 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026722284 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1435 | 1220648-2026-00837 | 24080145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24080145&pc=OZD | 01/05/2026 |  |  | 01/15/2026 |  |  | 01/05/2026 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026761549 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the surgeon placing axillary 5.5 pump and having difficulty placing device just after exiting the graft. then bend in the vasculature was difficult to push the pump past and the surgeon was forceful with the pump and 0.018¿ placement wire. while attempting placement, the placement wire had developed a bend mid-shaft, as well as the catheter kinking, with the surgeon stating the catheter had ¿ripped like someone had taken a knife to it.¿ the impella catheter was thrown off the field and was taken in order for the engineering department to identify causality of the possible catheter malfunction. the device was replaced and there was no patient harm. this report is for the pump and additional report will be sent for the guidewire (lot number 8854091). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1436 | 1220648-2026-00839 | 24080153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24080153&pc=PYX | 01/04/2026 |  |  | 01/15/2026 |  |  | 01/06/2026 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1431034 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp flex for mechanical circulatory support. it was reported that during impella rp flex ramp trial, the nurse went to p2 from p6. after assessment she returned to p6, but noticed the flows did not increase although the monitor said p6. she troubleshooted with the attending and when they went into flow control, the automated impella controller showed that the pump was on p2 not p6. the placement screen indicated p6. when they tried to go back to p6, the flows increased. there was no patient harm. this complaint is for the controller and additional report will be sent for the pump (serial number (b)(6)). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1437 | 1220648-2026-00846 | 24080160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24080160&pc=OZD | 12/17/2025 |  |  | 01/15/2026 |  |  | 12/17/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1370692 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that they were unable to perform self test of the automated impella controller. system self-check error occurred. a replacement was requested. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
1438 | 1220648-2026-00842 | 24080162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24080162&pc=OZD | 12/27/2025 | 02/27/2026 |  | 01/15/2026 |  |  | 12/27/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732045 |  | 1000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male | 94 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 35 year old male on impella cp due to acute myocardial infarction. the pump was implanted without complication and demonstrated appropriate placement signals at p-8 prior to transfer to cvicu. during support, intermittent unreliable placement signals occurred. repositioning of the impella was performed under bedside echocardiographic guidance with catheter adjustment and tuohy- lock secured. following repositioning, purge flow decreased to less than 2.0 ml/min with low purge flow alerts. patient remained hemodynamically stable with improved ejection fraction (ef) from ~30% to ~50% and mean arterial pressures (maps) in the 90s. the patient was taken to the cath lab and the impella removed successfully. patient was hemodynamically stable post removal. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
1439 | 1220648-2026-00844 | 24080163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24080163&pc=OZD | 12/19/2025 |  |  | 01/15/2026 |  |  | 12/19/2025 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1115975 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation | [10 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1440 | 1220648-2026-00845 | 24080164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24080164&pc=OZD | 12/19/2025 |  |  | 01/15/2026 |  |  | 12/19/2025 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1117232 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation | [10 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1441 | 1220648-2026-00736 | 24080209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24080209&pc=OZD | 01/10/2026 |  |  | 01/15/2026 |  |  | 01/10/2026 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET APAC |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026646596 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Fluid/Blood Leak | Decrease in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1442 | 1220648-2026-00833 | 24080260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24080260&pc=OZD | 01/09/2026 | 01/15/2026 |  | 01/15/2026 |  |  | 01/08/2026 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | OZD | [5 empty] | ABIOMED, INC |  |  | US |  | 1429248 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1443 | 1220648-2026-00633 | 24081006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081006&pc=OZD | 01/05/2026 |  |  | 01/15/2026 |  |  | 01/05/2026 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798006 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient was having premature ventricular contraction's (pvcs) and cannot rule out cause being the impella. imaging was used to check the pump position and the pump was in good position. additional information was received. the patient did not have any ventricular tachycardia. amiodarone and magnesium were given for the patient's pvcs. the medical doctor noted that the pvcs were frequent in the operating room and now much more infrequent. this is not thought to be impella related at this time since the pvcs are intermittent and random in nature. unfortunately, the pump was discarded by the hospital staff when explanted. | [5 empty] | Investigation summary: the investigation was completed. the device was not returned. peri-procedural adverse event (arrhythmia): in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
1444 | 1220648-2026-00850 | 24081080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081080&pc=OZD | 12/19/2025 |  |  | 01/15/2026 |  |  | 12/19/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026741851 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 76 KG | Non Hispanic | Black Or African American | Death | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. a false ¿impella position in aorta¿ alarm occurred due to low pulsatility. echocardiography confirmed appropriate impella position, and placement monitoring was disabled. the patient subsequently expired while on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. false alarm: clinical details report a false position in aorta alarm as position was confirmed and the patient had low pulsatility. since neither data logs nor product were returned for investigation, the cause of the false alarm could not be determined.
1445 | 1220648-2026-00853 | 24081125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081125&pc=OZD | 01/08/2026 |  |  | 01/15/2026 |  |  | 01/08/2026 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026780220 |  | 000000000001000100 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient was implanted with an impella 5.5 after removal of an impella cp. in addition, the patient was being supported by extracorporeal membrane oxygenation along with the 5.5. the following day the patient experienced a embolic stroke. no further information was provided. | [5 empty] | D6b explant date provided. the investigation of the reported failure mode has been completed. no product was received for evaluation. stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1446 | 1220648-2026-00855 | 24081131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081131&pc=OZD | 09/17/2025 |  |  | 01/15/2026 |  |  | 09/18/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026697079 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Fluid/Blood Leak | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that purge fluid was slowly leaking from the purge arm on the impella 5.5 catheter. the luer lock appeared to be intact, and the leak seemed to originate from underneath the protective shroud. impella purge pressures, flows, and motor currents were normal. the purge cassette was changed, but the leak persisted. | [5 empty] | The pump was not returned for investigation. data logs were only available until (b)(6). according to the clinical details, a leak was observed at the purge sidearm on (b)(6), and the issue was not resolved after changing the cassette. there were also placement signal discrepancies and suction alarms reported on (b)(6). purge leak ¿ pump: data log shows purge values were within range during the time of the reported leak issue. the cause of the purge leak was not determined without returned product and sufficient clinical information. pump suction: data log was not provided for the time of the reported issue. the cause of the suction issue was not determined without returned product investigation and sufficient clinical details. optical sensor issue: data log was not provided for the time of the reported issue. the cause of the optical sensor issue was not determined without returned product investigation and sufficient clinical details. the pump (b)(6) passed all post sterile inspection checks.
1447 | 1220648-2026-00854 | 24081136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081136&pc=OZD | 09/18/2025 |  |  | 01/15/2026 |  |  | 09/18/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026714664 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient has been experiencing short runs of ventricular tachycardia, all lasting less than five beats. | [5 empty] | No product was returned for investigation. arrhythmia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. device with sn: (b)(6) passed all post sterile inspection checks.
1448 | 1220648-2026-00856 | 24081140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081140&pc=OZD | 09/18/2025 |  |  | 01/15/2026 |  |  | 09/18/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026704337 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient became confused while on ventilation and veno-arterial extracorporeal membrane oxygenation and self-explanted. the blue suture hub sutures remained intact at the left groin. the repositioning sheath was pulled back on the catheter and was no longer in the left femoral artery. the impella was subsequently removed completely thereafter. | [5 empty] | No product was returned: device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the patient as the patient was confused. device (sn: (b)(6)) passed all post sterile inspection checks.
1449 | 1220648-2026-00859 | 24081146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081146&pc=OZD | 12/30/2025 |  |  | 01/15/2026 |  |  | 12/30/2025 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026658692 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1450 | 1220648-2026-00860 | 24081147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081147&pc=OZD | 12/31/2025 |  |  | 01/15/2026 |  |  | 12/31/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748244 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a physician called regarding a ¿pump pressure high¿ alarm. recommendations and best practices were shared during the discussion. the possibility of tissue plasminogen activator lysis was mentioned and requested by the physician. the complainant also advised storing the pump properly. | [5 empty] | B3 revised date as it was entered incorrectly on the initial report that was submitted. d6b explant date was added. d9 device was returned and the date received was added. h6 added code b13 as it was omitted from the initial report that was submitted. type of investigation, investigation findings, and investigation conclusions were updated accordingly based on the device being returned. h11 updated additional manufacturer narrative due to the device being returned. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1451 | 1220648-2026-00864 | 24081149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081149&pc=OZD | 12/30/2025 |  |  | 01/15/2026 |  |  | 12/31/2025 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026715005 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1452 | 1220648-2026-00862 | 24081151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081151&pc=OZD | 12/31/2025 |  |  | 01/15/2026 |  |  | 12/31/2025 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026778422 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1453 | 1220648-2026-00863 | 24081157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081157&pc=OZD | 12/31/2025 |  |  | 01/15/2026 |  |  | 12/31/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026792756 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 141 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was in multi-organ dysfunction as assessed by the provider team. the patient was producing less than 30 cc of pink to red urine. an echocardiogram was ordered, which revealed that the impella device was positioned slightly shallow in the left ventricle, measuring approximately 2.5 cm. the patient had experienced multiple cardiac arrests prior to the placement of the impella device. laboratory results showed elevated lactate dehydrogenase (ldh). plasma-free hemoglobin was not drawn at this time. the patient has expired. | [5 empty] | Correction: b1 product problem, and h6 health effect - impact code f02, f1904 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. hemolysis: the cause of the hemolysis was not determined without returned product investigation and sufficient clinical details. device in wrong position: the cause of the positioning issue was not determined without returned product investigation and sufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1454 | 1220648-2026-00861 | 24081158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081158&pc=OZD | 12/30/2025 |  |  | 01/15/2026 |  |  | 12/30/2025 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026678070 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Difficult to Open or Close | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1455 | 1220648-2026-00858 | 24081175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081175&pc=OZD | 12/30/2025 |  |  | 01/15/2026 |  |  | 12/30/2025 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026772532 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1456 | 1220648-2026-00857 | 24081178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081178&pc=OZD | 12/30/2025 | 02/25/2026 |  | 01/15/2026 |  |  | 12/30/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026722320 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Unspecified Infection | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an initial impella cp implant was performed for acute myocardial infarction and cardiogenic shock following an out-of-hospital cardiac arrest. the patient has a history of drug use and is on systemic heparin with partial thromboplastin times of 178.2, 30.5, and 75.1. the patient is also receiving aspirin and plavix daily. the device was later escalated to an impella 5.5, and the patient was placed on lovenox twice daily. an echocardiogram revealed an apical thrombus. there are no alarms or flow issues with the impella device. the patient was placed on eliquis and began experiencing a low-grade fever of 99 to 100 degrees fahrenheit. a bronchoscopy was performed, and cultures were sent; influenza and covid tests were negative. the patient is on antibiotics while awaiting culture results, and the physician suspects a respiratory infection. the impella insertion site is clean, dry, and intact with no issues noted, and the impella pump remains in place for ongoing support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b: added explant date h6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation investigation summary (infection/fever/thrombosis): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review the pump passed all the post sterile inspection checks.
1457 | 1220648-2026-00865 | 24081191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081191&pc=OZD | 01/07/2026 |  |  | 01/15/2026 |  |  | 01/07/2026 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796206 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | An impella 5.5 was used in a 69-year-old female presenting with cardiomyopathy for hemodynamic support. during support, nursing staff reported elevated purge pressure and reduced purge flow parameters, and tissue plasminogen activator (tpa) was administered for approximately 10 hours with minimal initial change in purge flow and pressure. at the time of consultation, purge flow was 1.9 ml/hr with a purge pressure of 1050 mmhg, and the motor current was within normal limits at 310 on p6. based on the information available at the time of reporting, the impella 5.5 continued to function as intended. | [5 empty] | B1 added selection as it was omitted from the initial report that was submitted. section 6 health effect - clinical code e2403 was incorrectly omitted in initial report and have been corrected to e0514. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. a4 is unknown.
1458 | 1220648-2026-00866 | 24081215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081215&pc=OZD | 09/13/2025 |  |  | 01/15/2026 |  |  | 09/13/2025 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654372 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1459 | 1220648-2026-00871 | 24081507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081507&pc=OZD | 01/09/2026 |  |  | 01/15/2026 |  |  | 01/09/2026 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1145430 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | User facility | Company representation | [10 empty] | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1460 | 1220648-2026-00868 | 24081615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24081615&pc=OZD | 10/16/2025 | 01/15/2026 |  | 01/15/2026 |  |  | 10/16/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026729374 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after running for about 50 minutes flows suddenly dropped. position was checked using transesophageal echocardiogram (tee), and additional volume was given neither intervention improved flow. catheter appeared well-positioned but was repositioned as a precaution. after consultation, it was suspected that the catheter may have ingested material and required replacement. catheter with sn (b)(6) was replaced with sn (b)(6). afterwards, impella was not able to run without suction above p5 despite giving volume, optimizing position and increasing inotropic support. patient was sent to icu in grave condition. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1461 | 1220648-2026-00876 | 24084035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24084035&pc=OZD | 12/17/2025 |  |  | 01/15/2026 |  |  | 12/17/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1143525 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 130 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, alarm ¿controller error¿ appeared and then resolved on its own with placement signal coming and going with alarm. it was reported that the patient survived explant. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1462 | 1220648-2026-00877 | 24084079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24084079&pc=OZD | 12/18/2025 |  |  | 01/15/2026 |  |  | 12/18/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026709916 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient that was implanted with an impella 5.5 experienced bleeding and a hematoma that required intervention during impella support. it was reported that during a procedure and on impella 5.5 support, the patient developed a hematoma at the axillary access site. a surgeon placed a jackson-pratt drain and a pressure dressing was applied. additionally, it was reported that the patient experienced bleeding at the impella access site. subsequent to the bleed, the patient received 5 units of blood products. it was noted that the patient¿s blood pressure and heart rate were stable, and the patient was successfully weaned and explanted. it was noted that the impella 5.5 was explanted due to the bleed. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H 6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the root cause of the injuries could not determined due to insufficient clinical details.
1463 | 1220648-2026-00878 | 24084093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24084093&pc=OZD | 12/16/2025 | 01/14/2026 |  | 01/15/2026 |  |  | 12/26/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704237 |  | 000000000001000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 66 YR |  | Male | 68 KG |  | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella support with a 5.5 pump. the complainant reported their automated impella controller (aic) could not recognize its purge disc. an alarm sounded. they ended up switching the purge cassette, disc, and impella to another console, and the alarm resolved. hematuria has also been reported. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary ==> ppae (hematuria): urinary output reported as concentrated/amber while on support. the cause of the injury was not determined due to insufficient clinical details. device history lot ==> device lot: 1939857. device history batch ==> subcomponent lot: n/a. device history review ==> the pump sn (b)(6) passed all post sterile inspection checks. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1464 | 1220648-2026-00880 | 24084205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24084205&pc=OZD | 01/07/2026 |  |  | 01/15/2026 |  |  | 01/07/2026 | 01/15/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1204376 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Read Input Signal | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
1465 | 1220648-2026-00881 | 24085765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24085765&pc=OZD | 12/27/2025 | 02/17/2026 |  | 01/15/2026 |  |  | 12/27/2025 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1558681 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 74 KG | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Insufficient Information | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella support with a left heart device, the automated impella controller (aic) alarmed and self-resolved. on the third day of impella support, the nurse called the customer support center to report that the placement signal and left ventricle signal disappeared from the aic screen and there was a controller alarm. it was reported that the error cleared on its own with no intervention. during the phone conversation, the alarm occurred twice more and self-resolved both times. the abiomed support team recommended to the nurse to check the white plug connection and verify the plug was seated securely, and to have a back-up aic ready if the error returned and could not be resolved. prior to the call ending, the waveforms were appropriate and there were no alarms. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. it was noted that the patient's outcome at explant was expired. additional clinical details, including perioperative course, contributing factors, and cause of death, were unavailable at the time of reporting. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 updated. the investigation is ongoing.
1466 | 1220648-2026-00884 | 24085792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24085792&pc=OZD | 01/06/2026 |  |  | 01/15/2026 |  |  | 01/07/2026 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST S2 SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746483 |  | 000000000001000459 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. high purge pressure was noted just after the impella 5.5 was inserted and running. the procedure was completed successfully with this device. no patient harm was noted. | [5 empty] | Additional information: the product was removed and replaced. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1467 | 1220648-2026-00882 | 24085815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24085815&pc=OZD | 01/04/2026 |  |  | 01/15/2026 |  |  | 01/04/2026 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026813688 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Pain | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with a history of cardiomyopathy, coronary artery disease, and diabetes mellitus underwent impella 5.5 implantation via the left axillary/subclavian artery. during device support, the patient experienced a sudden hemoglobin drop and back pain. imaging revealed a large spontaneous retroperitoneal bleed, likely related to anticoagulation therapy during impella support. the family declined intervention (coiling), and the patient was not a candidate for left ventricular assist device. comfort care was initiated, and the patient expired on (b)(6) 2026 after the pump was turned off. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. back pain/retroperitoneal hemorrhage: the cause of the injury was not determined due to insufficient clinical details.
1468 | 1220648-2026-00883 | 24085817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24085817&pc=OZD | 12/31/2025 |  |  | 01/15/2026 |  |  | 01/07/2026 | 01/15/2026 |  |  |  |  | 01/15/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026646697 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella cp was placed in a patient presenting with post cardiotomy cardiogenic shock (pccs) for hemodynamic support. the device experienced an emergency shutdown during support. two attempts to restart the pump and a replacement of the control device were unsuccessful. following the shutdown, the pump was surgically removed, and catecholamines were administered to maintain hemodynamic stability. the patient required a corrective invasive procedure (device removal), additional prescribed medication, and experienced hemodynamic instability due to loss of support, meeting criteria for serious injury. there was no evidence of device malfunction beyond the shutdown event, and the impella cp is considered an unlikely contributing factor due to the patient¿s underlying critical condition. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | D6a and d6b: correction to the explant and implant date based on new information.
1469 | 1220648-2026-00678 | 24066867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24066867&pc=OZD | 09/12/2025 | 01/13/2026 |  | 01/14/2026 |  |  | 09/26/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026685613 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation | Distributor |  |  | NA |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after coronary angiography (cag), coronary artery bypass grafting (cabg) was planned, and an attempt was made to insert an impella cp for stabilization. however, the patient¿s systemic vascular calcification was severe, and the iliac arteries were markedly tortuous. due to angulation of the common iliac artery (cia) and external iliac artery (eia), the pump could not be advanced. attempts to use a stiff wire and buddy wire were unsuccessful. the contralateral side was unusable due to prior major lower limb amputation; therefore, an intra-aortic balloon pump (iabp) was inserted from the same side. there was a lack of understanding regarding the creation of a clinical investigation (ci) for anatomical non-passage, which resulted in a delayed report. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1470 | 1220648-2026-00696 | 24067147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067147&pc=PYX | 01/07/2026 |  |  | 01/14/2026 |  |  | 01/07/2026 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026625517 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Heart Failure/Congestive Heart Failure | [40 empty] | Initial submission | Followup | Followup | Followup
1471 | 1220648-2026-00697 | 24067161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067161&pc=OZD | 12/14/2025 |  |  | 01/14/2026 |  |  | 12/14/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713515 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1472 | 1220648-2026-00703 | 24067207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067207&pc=OZD | 01/07/2026 |  |  | 01/14/2026 |  |  | 01/07/2026 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026771364 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1473 | 1220648-2026-00704 | 24067208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067208&pc=OZD | 01/07/2026 |  |  | 01/14/2026 |  |  | 01/07/2026 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1116350 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 150 KG | Non Hispanic | Black Or African American | Death | Yes | No | Application Program Problem | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | An impella cp was implanted in a 51-year-old male with cardiomyopathy for hemodynamic support. while on support, the aic 1931 controller suddenly displayed a ¿controller failure¿ alarm. despite the alarm, the device continued running with visible motor current; however, the placement signal and lv waveform disappeared. nursing staff and the attending physician were present at the bedside during the event. the patient remained hemodynamically stable. a replacement aic was obtained, and the patient was switched to the new controller without incident. the original aic was packed for return to headquarters for evaluation. the patient later expired. based on the available information, the controller failure alarm did not result in loss of pump function, and the patient remained stable during the event. the impella cp functioned as intended, and the controller was successfully replaced without complications. the patient¿s death is conservatively reported but is unlikely related to the device or the controller issue; it is more likely attributable to the patient¿s underlying cardiomyopathy and critical condition. therefore, no causal relationship between the device and the reported events has been identified. | Additional information was received indicating that the death date was incorrectly reported. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. | Additional information was received, and the h6 codes were updated accordingly. the b2 section was revised to correct the death date entered in the initial submitted report.
1474 | 1220648-2026-00707 | 24067215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067215&pc=OZD | 01/07/2026 |  |  | 01/14/2026 |  |  | 01/07/2026 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773833 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1475 | 1220648-2026-00705 | 24067216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067216&pc=OZD | 12/15/2025 |  |  | 01/14/2026 |  |  | 12/16/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766867 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 73 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Embolism | Hemodynamic instability | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. the nurse stated that the patient became unstable requiring more levophed. the medical doctor stated that the patient had a saddle pulmonary embolism (pe). the patient was on heparin and heparin assay was therapeutic. the patient was given options to treat the pe and refused them all and withdrew support on himself. support was withdrew and the next day the team was made aware that the patient had expired. | [5 empty] | Investigation summary: the investigation has been completed. peri-procedural adverse event (hemodynamic instability, pulmonary embolism): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1476 | 1220648-2026-00709 | 24067234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067234&pc=OZD | 01/01/2026 |  |  | 01/14/2026 |  |  | 01/01/2026 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026772530 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1477 | 1220648-2026-00692 | 24067242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067242&pc=PYX | 11/21/2025 |  |  | 01/14/2026 |  |  | 11/21/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1678954 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | Yes | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the device was received back from preventive maintenance (pm). during the safety inspection, the technician observed that the ground pin was very loose. due to this issue, the unit failed the safety check. the technician is returning the unit for repair. there was no patient involvement at the time of explant. | [5 empty] | E1. initial reporter first and last name are unknown.
1478 | 1220648-2026-00694 | 24067244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067244&pc=PYX | 11/21/2025 |  |  | 01/14/2026 |  |  | 11/21/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1678954 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | Yes | Yes | Degraded | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the console ic9995 was returned to the facility after being serviced. when biomed went to do a safety check they noticed the ground pin was very loose and it did not pass their safety test. they would like to send the console back. | [5 empty] | Data review: console logs did not contain any information relevant to the complaint. device analysis: console was tested after arriving in field service. console operated as intended when connected to a known good pump and purge cassette kit. inspection of the power cord confirmed that the ground pin on the power cord plug was slightly loose. the reported loose ground pin issue was confirmed. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | E1 is unknown.
1479 | 1220648-2026-00714 | 24067247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067247&pc=OZD | 01/08/2026 |  |  | 01/14/2026 |  |  | 01/08/2026 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1376304 |  | 0042-0000-US-RO |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) would not accept the purge disk and that a loaner aic was needed. the aic was removed due to the failure. | [5 empty] | The investigation was completed. console logs were consistent with what was reported in the complaint. imc logs revealed that the console issued the purge disc not detected alarm (event set #402) several times. device analysis: console (b)(4) was sent back to field service for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. root cause: the purge disc flag was observed to be broken and was the root cause of (b)(4) inability to detect the purge disc. es2023-0090 documents the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
1480 | 1220648-2026-00715 | 24067255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067255&pc=OZD | 01/04/2026 |  |  | 01/14/2026 |  |  | 01/04/2026 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026728734 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 69 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | Renal Failure | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | It was reported that a patient implanted with an impella 5.5 experienced a decrease in purge pressure. the cassette was replaced with no resolution. the patient also experienced hematuria and renal failure. the patient was put on dialysis. | Lysis solution was given through purge for the pp alarm. the pump is still implanted not sure of removal plan at this time. | An impella 5.5 device was inserted into a 56-year-old male patient via right axillary¿subclavian access for hemodynamic support. the patient presented with acute myocardial infarction and cardiogenic shock and was classified as society for cardiovascular angiography and interventions shock stage e. the patient was receiving inotropic therapy, vasopressors, and respiratory support via mechanical ventilation. during support, nursing staff reported that plasma free hemoglobin levels decreased from 10 to 3.5 compared to the previous night. the cassette was replaced without resolution. the registered nurse stated she would reach out to a second provider to discuss further management, including consideration of tissue plasminogen activator therapy. subsequent clinical findings included concentrated amber urine output, worsening renal function, and progression to renal failure requiring dialysis. the patient experienced hemolysis, renal failure, and required an additional device. comprehensive review suggests the hemolysis and renal failure were more likely attributable to the patient¿s underlying cardiogenic shock, acute myocardial infarction, severe hemodynamic instability, and critical illness. no evidence was identified to suggest a causal relationship between the impella 5.5 device and the reported events. |  |  |  | D6b explant date updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated with additional information. | Correction b1 product problem and h6 health effect - clinical code e0303 was inadvertently omitted on the initial manufacturer device report. b5 was amended from the initial manufacturer device report.
1481 | 1220648-2026-00706 | 24067807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067807&pc=OZD | 01/07/2026 | 01/14/2026 |  | 01/14/2026 |  |  | 01/07/2026 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026791754 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | White | Injury | Yes | No | Positioning Failure | Patient Device Interaction Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella cp was attempted to be placed in a 58-year-old male with cardiogenic shock (cgs) for hemodynamic support. the initial attempt was aborted due to difficulty advancing the pump through kinked vessels, raising concerns about potential vessel injury or pump damage. the device was removed and replaced with another impella cp, which was successfully inserted through a different puncture site. the patient remains on support with the second device. based on the available information, the reported inability to place the device and the need for replacement are most likely related to vascular anatomy and calcification rather than any device-related issue. vessel calcification can impede catheter passage and increase procedural complexity, providing a plausible explanation for the event independent of impella performance. case is being conservatively reported as a serious injury due to the patient requiring a corrective invasive procedure in which the impella cp was replaced but no causal relationship between the device and the reported events has been identified. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1482 | 1220648-2026-00717 | 24067812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067812&pc=OZD | 10/18/2025 |  |  | 01/14/2026 |  |  | 12/16/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026738715 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1483 | 1220648-2026-00720 | 24067823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067823&pc=OZD | 01/07/2026 |  |  | 01/14/2026 |  |  | 01/08/2026 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026734945 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Myocardial Infarction | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 was implanted in a 68-year-old male with cardiogenic shock (cgs) for hemodynamic support. the patient experienced a stemi the previous day, requiring escalation to ecmo support and coronary ballooning in the cath lab. additional contributing factors include ecmo support, a history of atrial fibrillation, a known left atrial appendage (laa) thrombus, and ongoing heparin therapy. the patient remains on impella support. based on the available information, the reported myocardial infarction and need for additional device support are most likely related to the patient¿s underlying cardiac condition, critical illness and the additional contributing factors rather than any device-related issue. the impella 5.5 is functioning as intended, and no evidence of causal relationship between the device and the reported events has been identified. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: myocardial infarction: the cause of the injury was not determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1484 | 1220648-2026-00721 | 24067834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24067834&pc=OZD | 12/19/2025 |  |  | 01/14/2026 |  |  | 12/19/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026732172 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1485 | 1220648-2026-00722 | 24068606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24068606&pc=OZD | 12/19/2025 |  |  | 01/14/2026 |  |  | 12/19/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1124467 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1486 | 1220648-2026-00724 | 24068623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24068623&pc=OZD | 01/07/2026 |  |  | 01/14/2026 |  |  | 01/07/2026 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026770631 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Hemolysis | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1487 | 1220648-2026-00726 | 24068653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24068653&pc=OZD | 12/16/2025 |  |  | 01/14/2026 |  |  | 12/16/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1117232 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 68 KG |  | White | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, there was a "purge disc not recognized" alarm on the console. the customer ended up switching the purge cassette, disc, and impella to another console and the alarm resolved. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. hematuria: urinary output reported as concentrated/amber while on support. the cause of the injury was not determined due to insufficient clinical details. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this aic controller device report is one of two devices associated with the event. refer to the related manufacturer report number 1220648-2026-00254 as noted in section h10 for the medical device report on the aic controller.
1488 | 1220648-2026-00727 | 24068675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24068675&pc=OZD | 12/30/2025 |  |  | 01/14/2026 |  |  | 12/31/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026774020 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1489 | 1220648-2026-00730 | 24069430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24069430&pc=OZD | 12/19/2025 |  |  | 01/14/2026 |  |  | 12/19/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1107434 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the purge cassette disk was not recognized by the automated impella controller (aic) despite multiple reinsertion attempts. after replacing the aic, the purge cassette and disk were successfully recognized. the patient outcome is still to be determined as the patient remains on support. | [5 empty] | There are two associated devices for the reported event. investigation summary: data review: console logs from the reported complaint date were consistent with what was reported in the complaint. during case start, the user was unable to advance past step 2 where it is instructed to insert the purge disc. a ¿purge disc not detected¿ alarm [event set #(b)(4)] was issued shortly after. (see ¿(b)(4)¿ in the attachments) device analysis: console (b)(6) was tested after arriving in field service. the issue was properly detecting the purge pressure disc was reproduced, and the purge flag was observed to be broken. (see "(b)(4)¿ in the attachments). root cause: the purge disc flag was observed to be broken and was the root cause of (b)(6)¿s inability to detect the purge disc. (b)(4) documents the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly. repair: before sending this console back to the customer, field service was instructed the following: - replace the purge disc detect flag [(b)(6)] and purge flag spring [(b)(6)]. - perform sections 10-12 of pm procedure [(b)(6)]. - perform a full functional check on the console and optical system [(b)(6)]. device history review: q1) service history review - are there any qns associated with the complaint device¿s most recent service report? no. q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode? n/a. q3) complaint history review - have there been any other complaints against this console? yes. there were 3 other complaints made against this console: (b)(4): issue recognizing pump insertion. likely result of pump plug not being fully inserted into the blue receptacle, caused by excessive contamination. (b)(4): placement signal ¿drift¿. likely due to pump (b)(6). (b)(4): rlm leds blinking simultaneously for over 1 minute. corruption of microsd card. phr summary: there were no issues related to the complaint discovered when reviewing the product history of console (b)(6).
1490 | 1220648-2026-00732 | 24069439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24069439&pc=OZD | 12/19/2025 |  |  | 01/14/2026 |  |  | 12/19/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026731915 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1491 | 1220648-2026-00733 | 24069441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24069441&pc=OZD | 09/18/2025 |  |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026617702 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that early in the morning, after extracorporeal membrane oxygenation insertion for ventricular tachycardia, the patient returned to the catheterization laboratory for impella cp insertion. due to tortuosity of the access vessel and aortic valve area being severely restricted, access to the left ventricle was not possible, and the insertion was aborted. | [5 empty] | No product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had tortuosity of the access vessel preventing successful delivery of impella. device (sn: (b)(6)) passed all post sterile inspection checks.
1492 | 1220648-2026-00734 | 24069447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24069447&pc=OZD | 09/18/2025 |  |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708950 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella was unable to cross the aortic valve. the valve area was severely restricted. attempts using a buddy wire and v-18 for delivery were all unsuccessful. | [5 empty] | No product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had known valve area less than 0.6. preventing successful delivery of impella. device (sn: (b)(6) passed all post sterile inspection checks.
1493 | 1220648-2026-00740 | 24072590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24072590&pc=PYX | 01/08/2026 |  |  | 01/14/2026 |  |  | 01/09/2026 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | PYX | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1465870 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the device exhibited a controller error alarm. the aic console was successfully switched to another aic console. it was reported that the patient survived explant. | [5 empty] | D4: the serial number provided by the customer cannot be found, the udi is not available. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1494 | 1220648-2026-00741 | 24072623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24072623&pc=OZD | 12/31/2025 | 02/25/2026 |  | 01/14/2026 |  |  | 12/31/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775998 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Hematuria | Pericardial Effusion | Heart Failure/Congestive Heart Failure | Hemodynamic instability | [32 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1495 | 1220648-2026-00634 | 24072629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24072629&pc=OZD | 01/07/2026 |  |  | 01/14/2026 |  |  | 01/07/2026 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026786755 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1496 | 1220648-2026-00743 | 24072679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24072679&pc=OZD | 12/27/2025 |  |  | 01/14/2026 |  |  | 12/27/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL, AIC, IMPELLA CONNECT, PKGD, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1815750 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 123 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during support with an impella device, the automated impella controller alarmed for a controller error and the aorta and left ventricle waveforms were no longer displayed. the issues self-resolved. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the impella device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1497 | 1220648-2026-00747 | 24072725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24072725&pc=OZD | 12/07/2025 | 01/14/2026 |  | 01/14/2026 |  |  | 12/31/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026687279 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Thrombocytopenia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1498 | 1220648-2026-00746 | 24072736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24072736&pc=OZD | 12/31/2025 |  |  | 01/14/2026 |  |  | 12/31/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026746991 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Pneumothorax | Tachycardia | Ventricular Fibrillation | [32 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup
1499 | 1220648-2026-00735 | 24072737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24072737&pc=OZD | 09/18/2025 | 01/14/2026 |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026721805 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that there was a sudden continuous yellow suction alarm and low pump flow despite the administration of volume. | [5 empty] | Data was not returned for review. product analysis: visual inspection found biomaterial inside the pump housing & wrapped around the impeller. no other issues were found on the rest of the pump. low or blocked pump flow/pump suction: the cause of the low flow was biomaterial ingestion. pump sn: (b)(6) passed all post sterile inspection checks.
1500 | 1220648-2026-00748 | 24076892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076892&pc=OZD | 12/20/2025 | 02/17/2026 |  | 01/14/2026 |  |  | 12/21/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2026790050 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1501 | 1220648-2026-00751 | 24076896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076896&pc=OZD | 12/30/2025 | 01/14/2026 |  | 01/14/2026 |  |  | 12/26/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026795716 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 95 KG | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71 year old male on an impella 5.5 for cardiomyopathy. during support, there was leading/high drain output noted from impella/ecmo bifurcated graft site. the patient received two units of packed red blood cells, prothrombin complex concentrate, and two units of fresh frozen plasma. additionally, a ¿purge system blocked¿ alarm occurred while attempting to verticalize patient. the tubing was free of kinks and had already been changed out purge tubing without resolution. the patient was started tissue plasminogen activation (tpa) purge infusion with improvement in purge flows. further details were provided which states that the patient expired. | [5 empty] | Corrected information b3. date of event. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1502 | 1220648-2026-00750 | 24076898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076898&pc=OZD | 12/26/2025 | 01/14/2026 |  | 01/14/2026 |  |  | 12/26/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026782199 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1503 | 1220648-2026-00752 | 24076899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076899&pc=OZD | 12/26/2025 | 01/14/2026 |  | 01/14/2026 |  |  | 12/26/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732157 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Atrial Fibrillation | Ischemia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1504 | 1220648-2026-00755 | 24076902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076902&pc=OZD | 12/27/2025 | 01/14/2026 |  | 01/14/2026 |  |  | 12/27/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025523427 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 68 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old female on a impella cp due to cardiomyopathy. during support, the patient's blood pressure dropped with a systolic in the 70's. volume was given (albumin) and the patient's numbers improved. the patient then had another hypotensive episode and after evaluation by the physician, the patient was taken to the operating room for a "washout" and a pluravac was placed for drainage. in addition, four units of blood and two units of fresh frozen plasma and 1 unit of platelets were given in the or. a cryo was just done in the unit and the patient received another one unit of blood. the physician stated that the impella was working well and he had no issues with the device. they are running a number of labs to rule out coagulopathy. the patient remains on support. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: (hypotension): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1505 | 1220648-2026-00759 | 24076905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076905&pc=OZD | 01/08/2026 |  |  | 01/14/2026 |  |  | 01/08/2026 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026798702 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1506 | 1220648-2026-00757 | 24076906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076906&pc=OZD | 01/07/2026 |  |  | 01/14/2026 |  |  | 01/07/2026 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025611380 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [41 empty] | Initial submission
1507 | 1220648-2026-00756 | 24076907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076907&pc=OZD | 12/27/2025 | 01/14/2026 |  | 01/14/2026 |  |  | 12/27/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026791190 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1508 | 1220648-2026-00758 | 24076909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076909&pc=OZD | 01/07/2026 |  |  | 01/14/2026 |  |  | 01/07/2026 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779614 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup
1509 | 1220648-2026-00760 | 24076918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076918&pc=OZD | 01/08/2026 |  |  | 01/14/2026 |  |  | 01/08/2026 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026797921 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1510 | 1220648-2026-00763 | 24076933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076933&pc=OZD | 09/14/2025 |  |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026679145 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient died of hemorrhagic shock caused by prolonged chest compression. the patient had a tendency to bleed upon arrival at the catheterization laboratory, and percutaneous coronary intervention was performed with veno-arterial extracorporeal membrane oxygenation and impella placement. due to bleeding, additional volume was administered through infusion, but this was not sufficient, and extracorporeal membrane oxygenation was insufficient. there were no issues with impella placement or the procedure, and no causal relationship exists. the following is a transcription of the doctor¿s comments: the patient died of hemorrhagic shock caused by prolonged chest compression. there were no complications with impella, and no causal relationship exists. percutaneous coronary intervention was performed without issue. | [5 empty] | Major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. arrhythmia : in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. the pump s/n: (b)(6) passed all the post sterile inspection checks.
1511 | 1220648-2026-00764 | 24076934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076934&pc=OZD | 09/18/2025 |  |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026668064 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that there was a placement signal unreliable alarm with no waveform or numerical readings. the intensive care unit team checked for cable kinks and proper securement in the automated impella console. the provider ordered the alarm audio to be disabled. they will verify the position using imaging as recommended. | [5 empty] | The data logs show that the placement signal and 5s parameters were stable throughout support until on (b)(6), the snr drops to 0, contrast oscillates between +3000 and -3000, lamp spiked, light decreased and gain was variable. this triggered psnr #1036 and controller error #1045. this lasted for 59 hours and then parameters returned to normal and remained within the expected range for the duration of the case. based on the data log review, it is apparent that the console is the potential cause for this optical signal issue. please refer to (b)(4) for an investigation into the console. (b)(4) was opened to investigate the legibility, completeness, and scannability of 2000193 batches. the batches were reworked and dispositioned as accept. impella (b)(6) passed all post sterile inspection checks.
1512 | 1220648-2026-00762 | 24076936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076936&pc=OZD | 09/17/2025 |  |  | 01/14/2026 |  |  | 09/17/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026705486 |  | 000000000001000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Unspecified Infection | Inflammation | Sepsis | Respiratory Failure | Hemodynamic instability |  |  |  |  | Required Intervention | Hospitalization | Death | Life Threatening | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that gross hemolysis was noted when the patient was transferred from a previous facility. hemolysis had started at the prior location, and the performance level was reduced in an attempt to alleviate it. plasma-free hemoglobin was elevated on arrival. the patient was escalated to a higher support device, and plasma-free hemoglobin decreased. in the afternoon, the patient became hypoxic and experienced spontaneous bronchial hemorrhage with significant blood loss. the patient underwent two bronchoscopies but became hypoxic again, and a pneumothorax was suspected. a right chest tube was placed. the patient then became hemodynamically unstable, with suspected profound septic shock of unknown etiology. lactic acid levels increased significantly, and white blood cell count was critically low. the patient became profoundly hypotensive and required multiple vasopressors. the patient achieved partial hemodynamic stability but later developed atrial fibrillation with rapid ventricular response, was shocked twice, and converted back to sinus rhythm. the patient has expired. | [5 empty] | No product was returned for investigation. data logs were returned for analysis. optical sensor issue: clinical details mention that placement signal unreliable (psnr) alarms were seen during the case on (b)(6) 2025. data log analysis confirmed psnr alarm on (b)(6) 2025, which did not resolve till the end of the case. prior to the psnr alarm, the placement signal showed no abnormal trends and was in-line with p-level changes. after the psnr alarm triggered, placement signal spiked to 3000, snr which was stable dropped to 0, contrast barcoded between values of +-3000, light decreased from approximately 2.7 to 0.5, gain decreased from approximately 1.5 to 1 but increased in amplitude, while lamp remained constant at 100. the root cause of the optical issue was most likely fiber break due to torquing of the optical fiber line based on the trends seen in the optical parameters from data log analysis. mechanical interaction with blood: clinical details mention that the patient was suffering from hemolysis on (b)(6) 2025. echo was done which verified that the pump was shallow at 2.3cm, which was moved to 3.8cm. patient suffered from ectopy the deeper the pump was moved inside the lv, so the pump was left in a shallow position. there were no impella in aorta alarms observed from the data logs, but impella position unknown and suction alarms were seen towards the beginning of the case. the root cause of the hemolysis was most likely positioning issues as the pump was shallow as confirmed by echocardiology. hypotension, major bleed, respiratory failure, inflammation, hemodynamic instability: clinical details mention that the patient suffered from ectopy during the case which increased the deeper the pump was positioned in the heart. it was also seen that the patient was hemodynamically unstable during the case, suffered from hypoxia with spontaneous bronchial hemorrhage, and became profoundly hypotensive after septic shock. the patient's lungs were also weak, with oxygenation issues. the patient also suffered from inflammation at near the pump's extremity where it was swollen and tight. gi bleeding and airway bleeding was also seen during the case. the root cause of the injury was not determined due to lack of sufficient clinical details. arrhythmia, sepsis, infection: clinical details mention that the patient suffered from arrhythmia on (b)(6) 2025. the patient had a septic shock on (b)(6) 2025. on (b)(6) 2025, the patient was administered antibiotics to treat infection caused by cpo. in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pump #(b)(6) passed all post-sterile inspection checks.
1513 | 1220648-2025-47627 | 24076940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076940&pc=OZD | 09/18/2025 |  |  | 01/14/2026 |  |  | 09/20/2025 | 01/14/2026 |  |  | 09/30/2025 |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2024401746 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Device Sensing Problem | Priming Problem | [5 empty] | Fever | Hematoma | Low Blood Pressure/ Hypotension | Unspecified Infection | Pain | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that there was some concern for a possible small hematoma around the incision site, but it was not confirmed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1514 | 1220648-2026-00769 | 24076943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076943&pc=OZD | 09/18/2025 |  |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026711249 |  | 1000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Bradycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | See case information #(b)(4). pump placed independently, pt. became bradycardic upon insertion, flow 0.7. tvp placed, heart rate increased, flows remained 0.7, i arrived onsite & had them replaced with a new cp. | [5 empty] | Impella cp sn (b)(6) + usb returned for evaluation. data review: data logs confirmed the low flow when pump started with fluctuated ps & suction alarms triggered. pump then was stopped manually & restarted with the same issue. the low flow remained to the rest of the case. there were no mc spikes or purge issues observed during support. product analysis: pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue during bench test. low or blocked pump flow: the cause of low flow was unable to be determined as limited clinical details were provided and no issue was found on the pump. bradycardia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pump sn: (b)(6) passed all post sterile inspection checks.
1515 | 1220648-2026-00770 | 24076944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076944&pc=OZD | 09/18/2025 |  |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026660136 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1516 | 1220648-2026-00766 | 24076945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076945&pc=OZD | 09/18/2025 |  |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026718743 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the catheterization laboratory technician noticed bleeding from the groin, so the physician placed a figure-eight stitch while the technician was holding pressure. the physician did not want to wait any longer due to bleeding concerns. one pre-close had already been placed prior to the impella introducer being inserted. the physician decided to wean the impella and remove it in the procedural area because the patient was on vasopressor support and the mean arterial pressure was stable. even after the impella was removed, hemostasis was obtained with one pre-close and forty minutes of manual pressure. lidocaine with epinephrine was also administered. | [5 empty] | No product was returned for investigation. major bleed: bleeding from groin site and was holding pressure, doctor did not want to wait any longer with bleeding concerns and decided to remove impella. impella weaned and hemostasis was obtained with one pre-perclose and 40 mins of manual pressure. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device with sn: (b)(6) passed all post sterile inspection checks. as per qn (b)(4), pump with serial number (b)(6), found to have particle inside and failed 100% visual inspection checks. rework of this unit has been successfully completed on rework workorder (b)(4), an "a" has been added to track sterilization. unit will be sent back to sterilizer.
1517 | 1220648-2026-00768 | 24076946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076946&pc=OZD | 09/18/2025 |  |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026740396 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient had bloody urine, and hemolysis was suspected. a lactate dehydrogenase test was ordered but has not yet been completed. | [5 empty] | No product was returned for investigation. hematuria: the cause of the clinical symptom was not determined due to insufficient clinical details. device with sn: (b)(6) passed all post sterile inspection checks.
1518 | 1220648-2026-00765 | 24076948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076948&pc=OZD | 09/18/2025 |  |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025521582 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the device operated for approximately three minutes upon insertion. while the repositioning sheath was being sutured, the pump abruptly changed to reduced flows. the performance level did not affect the flow rate. the pump was subsequently exchanged without issue. | [5 empty] | Data review: data logs confirmed the failure mode & clinical narrative. logs showed after pump started & run for 3 minutes, mc spiked followed low flow that triggered auto suction response & suction alarms. this event remained to the rest of the case. low or blocked pump flow: the cause of was low or blocked pump flow was most likely biomaterial ingestion since the clinical information & the data logs meet the pattern of the biomaterial ingestion. pump sn: (b)(6) passed all post sterile inspection checks.
1519 | 1220648-2026-00771 | 24076949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076949&pc=OZD | 09/18/2025 |  |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026739171 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1520 | 1220648-2026-00772 | 24076950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076950&pc=OZD | 09/18/2025 |  |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026720713 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during proactive calls, the registered nurse mentioned that the patient had a small hematoma which remained unchanged from the day shift, and some oozing was noted at the site, which was stable. | [5 empty] | Minor bleed/ hematoma: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. the pump s/n: (b)(6) passed all the post sterile inspection checks.
1521 | 1220648-2026-00773 | 24076954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076954&pc=OZD | 12/24/2025 | 01/14/2026 |  | 01/14/2026 |  |  | 12/24/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674432 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1522 | 1220648-2026-00775 | 24076958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076958&pc=OZD | 02/27/2024 |  |  | 01/14/2026 |  |  | 02/27/2024 | 01/14/2026 |  |  | 10/31/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2024415021 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1523 | 1220648-2026-00774 | 24076960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076960&pc=OZD | 12/27/2025 |  |  | 01/14/2026 |  |  | 12/27/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026775568 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 35-year-old female patient on an impella 5.5 due to acute myocardial infarction. during support, the ct scan revealed a catastrophic brain bleed post heart transplant. patient went into acute rejection after receiving her heart transplant and required ecmo and impella 5.5 and required explant of the 5.5. patient did not respond to multiple treatments and a discussion with staff and family to withdrawal from support and harvest remaining viable organs. the patient subsequently expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary : (stroke) : in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review the pump passed all the post sterile inspection checks.
1524 | 1220648-2026-00778 | 24076961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076961&pc=PYX | 12/21/2025 |  |  | 01/14/2026 |  |  | 12/21/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1678477 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female | 65 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that a purge cassette alarm occurred. the nurse stated she was unable to proceed with steps during a routine purge cassette change due to an active purge cassette/disc alarm. the exact alarm was not recalled but may have indicated that the purge cassette or disc was not detected. no damage or defect was observed on the blue purge disc holder. transferring the automated impella controller (aic) from (b)(6) resolved the alarm. the patient outcome is still to be determined as the patient remains on support. | [5 empty] | Correction: d6a and d6b should have been left blank. h3 and h8 was erroneously selected on the initial manufacturer device report the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data review: console logs were consistent with what was reported in the complaint. device analysis: the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. root cause: the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc and advance to the next step during purge cassette change. phr summary: there were no issues related to the complaint discovered when reviewing the product history of console. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1525 | 1220648-2026-00776 | 24076962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076962&pc=OZD | 09/18/2025 |  |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026707485 |  | 000000000001000100 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup | [8 empty] | The complainant reported that the aortic placement signal was significantly different from the patient¿s cuff blood pressure, although no alarms were present. the aortic signal was seventy-nine over forty-five with a mean of sixty-five, while the patient¿s blood pressure was ninety-seven over seventy-four with a mean of eighty-three. the left ventricular signal was seventy-seven over negative nineteen. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1526 | 1220648-2026-00777 | 24076964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076964&pc=OZD | 12/19/2025 |  |  | 01/14/2026 |  |  | 12/19/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773969 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that prior to device insertion, the patient was receiving multiple inotropes/pressors and had an intra-aortic balloon pump (iabp) in place. pink-red urine was noted upon arrival. after transfer to the icu with impella support, the family decided to withdraw care. the impella was turned off following best practices, and the patient subsequently expired. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the hematuria was not determined due to insufficient clinical details. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1527 | 1220648-2026-00779 | 24076965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076965&pc=OZD | 12/20/2025 |  |  | 01/14/2026 |  |  | 12/20/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713999 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 106 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the pump stopped and could not be restarted despite attempts at multiple performance levels. physician was at bedside and noted that tpa had been infused via the purge line. the plan was to retract the impella across the aortic valve. patient remained stable with increased pressor support. plan to exchange device was made, and the local team was notified. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1528 | 1220648-2026-00781 | 24076968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076968&pc=OZD | 12/03/2025 |  |  | 01/14/2026 |  |  | 12/20/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781123 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 108 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced ventricular tachycardia. an echocardiogram was performed, and the pump was repositioned, after which the ventricular tachycardia resolved. the patient was receiving intravenous amiodarone. the patient subsequently expired. additional information indicated a purge flow decrease alarm. purge flow dropped from 11 to 4, with pressure exceeding 1000. purge tubing was inspected and replaced without resolution. tpa was initiated through the pump. purge flow and pressure were within defined limits following intervention. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1529 | 1220648-2026-00780 | 24076970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076970&pc=OZD | 12/19/2025 |  |  | 01/14/2026 |  |  | 12/19/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025594741 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the patient experienced a new onset of atrial fibrillation with no prior history of the condition. the patient outcome is still to be determined as the patient remains on support. additional information indicates that atrial fibrillation (af) resolved following administration of amiodarone. | [5 empty] | Investigation summary: ppae (paroxysmal atrial fibrillation): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot- 1865703. device history batch: subcomponent lot- n/a. device history review: the pump sn 558752 passed all the post sterile inspection checks.
1530 | 1220648-2026-00749 | 24076976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076976&pc=OZD | 12/26/2025 | 02/27/2026 |  | 01/14/2026 |  |  | 12/26/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026782855 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 70 KG | Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male, who during insertion, the impella 5.5 was difficult to be pushed through the axillary artery. impella 5.5 kept getting coiled in the artery and was unable to be passed through the axillary into the aorta. there was no patient harm reported. a different device was successfully inserted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1531 | 1220648-2026-00782 | 24076987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076987&pc=OZD | 10/15/2025 |  |  | 01/14/2026 |  |  | 10/15/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026717600 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Death | No | Yes | Malposition of Device | [6 empty] | Stroke/CVA | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that heparin therapy was discontinued due to hemoptysis. the patient subsequently developed sepsis and elected to withdraw care. a computed tomography (ct) scan revealed multiple hemorrhagic strokes in the brain, which contributed to the overall poor prognosis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1532 | 1220648-2026-00783 | 24076991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076991&pc=OZD | 03/30/2025 |  |  | 01/14/2026 |  |  | 10/15/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025608721 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Impairment | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient developed acute kidney injury (aki) severe enough to placed on continuous renal replacement therapy (crrt), then transitioned to hemodialysis (hd), and eventually recovered kidney function enough to discontinue dialysis but with residual chronic kidney disease (ckd) that is worse than their baseline. additional information from caller that offered assistance for air in purge system alarms. patient had air in line alarms off the floor while in computed tomography (ct). air in line alarms resolved when new purge cassette was placed. patient was currently on d10w as they had no other bags available in the unit. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the acute kidney failure was not determined due to insufficient clinical details.
1533 | 1220648-2026-00784 | 24076995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076995&pc=OZD | 03/27/2025 |  |  | 01/14/2026 |  |  | 10/15/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025589948 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | Tachycardia | Ventricular Fibrillation | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
1534 | 1220648-2026-00785 | 24076997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076997&pc=OZD | 10/15/2025 | 11/14/2025 |  | 01/14/2026 |  |  | 10/15/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593733 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was bleeding at an impella insertion site with elevated partial thromboplastin time (ptt) (>200), managed with manual pressure and femstop until ptt normalized, and no further bleeding afterward. additional information reported that the impella placement done using pig/j-wire following best practices. no distal flow observed after stent placement and decision made to peel away the 14f sheath. distal perfusion catheters (dpcs) placed per hospital protocol. there was no patient complication. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1535 | 1220648-2026-00787 | 24076998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076998&pc=OZD | 10/15/2025 | 01/27/2026 |  | 01/14/2026 |  |  | 10/15/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH - 3007904765 |  |  | GM |  | 1383039 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported when the nurse started up the impella control device to set the time as part of a routine check, a blue screen displayed that the self-diagnosis had failed and this was reported to the clinical engineer (ce). the ce restarted the device and found that the issue was reproducible. an abiomed representative checked on-site and found that the issue was reproducible leading to a recall of the device. | [5 empty] | E1. initial reporter first and last name and h6. investigation conclusions is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1536 | 1220648-2026-00786 | 24076999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24076999&pc=OZD | 10/14/2025 |  |  | 01/14/2026 |  |  | 10/15/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025574121 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Rupture | Aortic Valve Insufficiency/ Regurgitation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1537 | 1220648-2026-00761 | 24077012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077012&pc=OZD | 09/18/2025 | 01/14/2026 |  | 01/14/2026 |  |  | 09/18/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1434576 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that after switching on the automated impella constroller (aic), the message ¿system check failed, replace console¿ was displayed. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1538 | 1220648-2026-00788 | 24077016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077016&pc=OZD | 12/19/2025 |  |  | 01/14/2026 |  |  | 12/19/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026756917 |  | 000000000001000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1539 | 1220648-2026-00792 | 24077025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077025&pc=OZD | 12/27/2025 | 01/14/2026 |  | 01/14/2026 |  |  | 12/27/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026767025 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1540 | 1220648-2026-00791 | 24077026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077026&pc=OZD | 12/27/2025 | 01/14/2026 |  | 01/14/2026 |  |  | 12/27/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766512 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1541 | 1220648-2026-00795 | 24077038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077038&pc=OZD | 09/24/2025 | 02/24/2026 |  | 01/14/2026 |  |  | 09/24/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026697912 |  | 000000000001000080 |  | No |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Unspecified Infection | Pulmonary Edema | Cardiogenic Shock | Respiratory Failure | Hematuria | Asystole | [31 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1542 | 1220648-2026-00796 | 24077039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077039&pc=OZD | 09/24/2025 |  |  | 01/14/2026 |  |  | 09/24/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026704368 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1543 | 1220648-2026-00793 | 24077041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077041&pc=OZD | 09/26/2025 |  |  | 01/14/2026 |  |  | 09/26/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719055 |  | 000000000001000482 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1544 | 1220648-2026-00797 | 24077042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077042&pc=OZD | 07/17/2025 |  |  | 01/14/2026 |  |  | 09/24/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026701045 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 56 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hypovolemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the patient¿s baseline platelet count on (b)(6) 2025, prior to the procedure, was 170. after the procedure on (b)(6) 2025, platelets decreased to 126. post-operatively, the patient was hypovolemic, which resolved with diuretics. during the hospital stay through (b)(6) 2025, the lowest platelet count recorded was 94. platelet levels began to recover after (b)(6) 2025, and by discharge on (b)(6) 2025, the count had increased to 307. platelets were monitored throughout, and no treatment was administered. the medical team was not concerned, and the impella 5.5 device was discarded. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1545 | 1220648-2026-00799 | 24077047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077047&pc=OZD | 09/29/2025 |  |  | 01/14/2026 |  |  | 09/29/2025 | 01/14/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026701797 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1546 | 1220648-2026-00800 | 24077048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077048&pc=PYX | 09/27/2025 |  |  | 01/14/2026 |  |  | 09/27/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1541587 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the nurse reported reduced impella flow following a motor current (mc) spike. review of the impella controller (ic) data indicated that the mc waveform flattened at 13:42, followed by a spike at 13:43, with mc exceeding 700 and flow decreasing to 0.7 l/min. metrics did not return to baseline, and mc continued to run higher in the 400s, while flow further declined to 1.4 l/min from a previous 2.5 l/min. additionally, the patient reported being unable to hear the pump operating during this period. although pump function resumed, the patient noted that it did not sound as it had previously. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1547 | 1220648-2026-00801 | 24077050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077050&pc=OZD | 09/30/2025 |  |  | 01/14/2026 |  |  | 10/01/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719352 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 67 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Impairment | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that there was a severe hemolysis contributing to progressive acute kidney injury (aki). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1548 | 1220648-2026-00802 | 24077051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077051&pc=OZD | 09/26/2025 |  |  | 01/14/2026 |  |  | 09/26/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026727111 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 67 KG | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Unspecified Infection | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient had a history of diabetes mellitus. an echocardiogram performed one month prior to the st-segment elevation myocardial infarction (stemi) demonstrated an ejection fraction (ef) of 25%. the patient presented to the emergency department unresponsive and required immediate intubation; advanced life support was not necessary. electrocardiogram (ekg) confirmed stemi, and the patient was taken emergently to the cardiac catheterization laboratory. coronary angiography revealed severe calcified ostial left main (lm) disease, chronic total occlusion (cto) of the left anterior descending (lad) artery, and ostial right coronary artery (rca) disease. due to hemodynamic instability and significant lm disease, the decision was made to proceed with impella support. following approximately three hours of intervention, a stent was successfully placed in the left main artery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1549 | 1220648-2026-00803 | 24077053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24077053&pc=OZD | 09/29/2025 |  |  | 01/14/2026 |  |  | 09/29/2025 | 01/14/2026 |  |  |  |  | 01/14/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026655189 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male | 83 KG | Non Hispanic | Asian | Injury | Yes | No | Difficult or Delayed Positioning | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during attendance at the impella 5.5 insertion, the device was introduced into the body without adequate priming, and air removal from the purge line was insufficient. as a result, air entrainment was confirmed on echocardiography. the situation was immediately communicated to the physician, and the procedure continued. fortunately, no serious complications occurred. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1550 | 1220648-2026-00587 | 24050245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050245&pc=OZD | 11/20/2025 |  |  | 01/13/2026 |  |  | 11/20/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025560826 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1551 | 1220648-2026-00588 | 24053233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053233&pc=OZD | 11/20/2025 | 01/12/2026 |  | 01/13/2026 |  |  | 11/20/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719774 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 104 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient failed spontaneous breathing trial (sbt) and was unable to extubate. patient undergoing supraventricular tachycardia (svt) again, but noted to be unlikely to pass due to excessive secretions and tachypnea. the physician assistant consulted regarding patient management and goals. possible device-related contribution to ventricular tachycardia (vt) during sbt. last echocardiogram showed device position remained at 5.3 cm. the pump was subsequently weaned and explanted. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1552 | 1220648-2026-00589 | 24053258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053258&pc=OZD | 11/21/2025 | 01/12/2026 |  | 01/13/2026 |  |  | 11/21/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026760943 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Flow Rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient left femoral artery access was obtained. the patient had a known left common iliac stent. the physician anticipated that the impella device would pass through the stent without issue; however, crossing the stent was unsuccessful. the impella was explanted, and the left femoral access site was closed. right femoral artery access was then obtained for percutaneous coronary intervention (pci). imaging of the left iliac and runoff demonstrated good flow. it was reported that the patient survived at explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1553 | 1220648-2026-00590 | 24053754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053754&pc=OZD | 01/06/2026 |  |  | 01/13/2026 |  |  | 01/06/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026741392 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1554 | 1220648-2026-00592 | 24053761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053761&pc=OZD | 01/07/2026 |  |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026807018 |  | 000000000001000080 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 133 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 68-year-old male underwent a percutaneous coronary intervention (pci) for acute myocardial infarction complicated by cardiogenic shock. an impella cp was placed post-pci for hemodynamic support. during cvicu assessment, a hematoma was observed at the right femoral access site used for impella insertion. the physician outlined the hematoma and applied light pressure; no active bleeding was noted. the patient remains on impella support with ongoing monitoring. this event is a known complication of large-bore femoral access required for impella placement. the device functioned as intended. hematoma formation may result from vessel trauma, anticoagulation, or post-procedural factors. | [5 empty] | Additional information: d6b added. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1555 | 1220648-2026-00591 | 24053762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053762&pc=OZD | 01/07/2026 |  |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026798092 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 51 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | An impella 5.5 was used in a 60-year-old patient with cardiomyopathy for hemodynamic support. during hospitalization, the patient had a quinton catheter placed and subsequently experienced bleeding at the catheter insertion site, which required medical intervention to control but did not require blood transfusion. the bleeding resolved after the critical care team placed an additional suture at the site. the patient was later reported to have three episodes of bloody bowel movements, prompting a gastroenterology consultation for evaluation of a potential gastrointestinal bleed. the bleeding at the quinton catheter site is attributable to a separate vascular access procedure, and the reported gastrointestinal bleeding is anatomically remote and physiologically unrelated to the impella 5.5. based on the information available at the time of reporting, the impella 5.5 is considered an unlikely contributing factor to the reported gastrointestinal bleed; however, the case was conservatively reported as a serious injury and the device continued to function as intended. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b explant date updated.
1556 | 1220648-2026-00595 | 24053763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053763&pc=OZD | 12/15/2025 |  |  | 01/13/2026 |  |  | 12/15/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1371823 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 84 KG | Non Hispanic | White | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient who was implanted with an impella cp experienced alarms that sounded different than expected on the automated impella controller. the unit was not replaced. it was reported that the unexpected sounding alarm occurred intermittently during use. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1557 | 1220648-2026-00593 | 24053764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053764&pc=OZD | 01/05/2026 |  |  | 01/13/2026 |  |  | 01/06/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714409 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1558 | 1220648-2026-00594 | 24053765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053765&pc=OZD | 01/07/2026 |  |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142497 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1559 | 1220648-2026-00597 | 24053767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053767&pc=OZD | 01/06/2026 | 02/02/2026 |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026749006 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1560 | 1220648-2026-00598 | 24053771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053771&pc=OZD | 01/04/2026 |  |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026734940 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1561 | 1220648-2026-00600 | 24053772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053772&pc=OZD | 01/06/2026 |  |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026775920 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1562 | 1220648-2026-00596 | 24053776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053776&pc=OZD | 01/07/2026 |  |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205049 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon start up of the automated impella controller a battery failure alert appeared. another controller was used to support the patient and no harm was reported. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D4. udi was added as it was omitted from the initial report that was submitted. d9. the device was returned and date was added. e1. added initial reporter phone number as it was omitted from the initial report that was submitted. h6. type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation and the device being returned. investigation summary: the investigation has been completed. data review: console logs were consistent with what was reported in the complaint. the logs confirmed the reported battery-related alarms. there were several battery failure alarms observed. device analysis: the console was sent back for further investigation. the battery failure alarm was reproduced upon boot up. results of running the batttest utility revealed that there was a cell imbalance observed in cell 2 of the battery in position 1. root cause: the battery failure was due to a locked-out battery because of internal cell imbalance. product history review summary: there were no issues related to the complaint discovered when reviewing the product history of console.
1563 | 1220648-2026-00599 | 24053779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053779&pc=OZD | 01/06/2026 |  |  | 01/13/2026 |  |  | 01/06/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739606 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1564 | 1220648-2026-00601 | 24053912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24053912&pc=OZD | 01/06/2026 |  |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026764701 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [41 empty] | Initial submission | Followup
1565 | 1220648-2026-00602 | 24054081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054081&pc=OZD | 01/06/2026 |  |  | 01/13/2026 |  |  | 01/06/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696927 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1566 | 1220648-2026-00603 | 24054084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054084&pc=OZD | 01/07/2026 |  |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026721350 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 37 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Poor Visibility |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | Additional information received reporting that the patient has gastrointestinal bleed. heparin was held. the nurse will discuss once physicians round for future plans on anticoagulation. it was known that the patient had low platelets pre-procedure and pre impella insertion. | A 77-year-old male underwent elective placement of an impella 5.5 for pchri cabg and valve surgery. due to the presence of a mechanical mitral valve, echocardiographic visualization was difficult, making optimal positioning challenging. impella flows were maintained at p4, but suction occurred above p5, likely related to right ventricular failure and/or positional factors. the patient was on five vasoactive drips, received blood products for volume optimization, and remained on impella support. urinary output was less than 30 cc/hr and appeared concentrated/amber. the reported event includes poor visualization, low or blocked pump flow, suspected device malposition, hematuria, and blood transfusion. these events are most consistent with patient-specific factors such as complex surgical status, hemodynamic instability, and rv dysfunction. per the instructions for use, impella 5.5 is intended for short-term ventricular support and requires careful monitoring of positioning and hemodynamics. visualization challenges, particularly in patients with prosthetic valves or altered anatomy, can complicate optimal placement. patient-specific factors such as severe cardiac disease, hypotension, and critical illness increase the risk of complications like suction events, hematuria, and bleeding. based on available information, there is no evidence of a causal relationship between the impella 5.5 and the reported events. |  |  |  |  | Additional information received; b5 and h6 updated based on information received from the clinical site. d4 udi information was inadvertently omitted on the initial manufacturer device report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b explant date provided.
1567 | 1220648-2026-00605 | 24054096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054096&pc=OZD | 01/05/2026 |  |  | 01/13/2026 |  |  | 01/05/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026781163 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1568 | 1220648-2026-00609 | 24054126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054126&pc=OZD | 01/02/2026 | 02/03/2026 |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026787484 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Decrease in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced a low pressure alarm. a leak was identified at the yellow luer connector. the physician attempted to rescrew the connector into place but it fell apart in his hands. the pump was replaced without major complications to the patient. | Clinical narrative an impella cp was implanted without issue in a patient with cardiogenic shock prior to pci. the device functioned as intended for over an hour during the procedure. after the procedure, while the nursing team was applying dressings and preparing to move the patient, a ¿low pressure in the irrigation system¿ alarm was triggered. inspection revealed a leak at the yellow luer connector. the physician attempted to tighten the connector, but it fell apart during handling. the pump was immediately replaced with a new one without any major complications for the patient. the reported event involves purge pressure low, fluid leak at the side arm, medical device removal, and device revision or replacement. there were no signs, symptoms, or adverse clinical impact on the patient. based on available information, the issue appears isolated to the connector component and does not indicate a systemic device malfunction. the impella cp performed as intended prior to the connector failure, and the replacement was completed promptly without patient harm. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 amended to include the clinical rationale. | D9 updated; the product has been returned. the investigation is still ongoing.
1569 | 1220648-2026-00607 | 24054137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054137&pc=OZD | 01/01/2026 |  |  | 01/13/2026 |  |  | 01/01/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026803703 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1570 | 1220648-2026-00608 | 24054154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054154&pc=OZD | 01/06/2026 |  |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720867 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 66 KG | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 57-year-old male with a history of acute myocardial infarction (ami) complicated by cardiogenic shock (cgs), known coronary artery disease (cad), diabetes mellitus, and renal insufficiency was transferred from an outside hospital to our facility with an impella cp in place. the patient was critically ill, intubated, and on multiple vasoactive drips. per report, pulses were not present prior to device insertion and remained absent bilaterally after placement. the peel-away sheath was left in, and no reperfusion sheath was placed. the patient also had peripheral vascular disease (pvd). despite ongoing care, the patient¿s condition deteriorated, and he was changed to dnr status. the patient expired. the reported event involves a peri-procedural adverse event, ischemia, and death. the impella cp remained in use without reported performance issues. these complications are most consistent with patient-specific factors, including severe underlying cardiac disease, critical illness, pre-existing absence of pulses, pvd, and hemodynamic instability, rather than device related. per the instructions for use, impella cp is intended for short-term ventricular support and requires careful monitoring of vascular access and perfusion. patient-specific factors such as advanced cad, pvd, and critical illness significantly increase the risk of ischemia and mortality independent of device function. based on available information, there is no evidence of a causal relationship between the impella cp and the reported events. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the ischemia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
1571 | 1220648-2026-00615 | 24054198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054198&pc=OZD | 09/19/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672045 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that impella 5.5 was implanted following best practice protocols. activated clotting time (act) maintained above 400. high motor current alarm occurred, followed by pump failure. attempts to restart at previous performance level and at p-2 were unsuccessful. for patient safety, pump removal and replacement were recommended. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1572 | 1220648-2026-00617 | 24054204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054204&pc=OZD | 09/19/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708154 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1573 | 1220648-2026-00616 | 24054205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054205&pc=OZD | 09/19/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713047 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Contamination | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that some blood stain was noted in the sterile sleeve. there was no active ooze or bleeding around the site. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1574 | 1220648-2026-00618 | 24054217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054217&pc=PYX | 09/19/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026712134 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the placement signal was unreliable. | [5 empty] | No product was returned. placement signal issue: clinical details noted a psnr alarm on the day of implant. data log review showed variation in ps/pap with increases, decreases, and going out of range throughout the case. psnr (alarm 1051) was triggered, which is associated with the dp sensor. during one increase in ps/pap on 19-sep-2025, there appeared to be periodic jumps in ps/pap with each increment being 10-20 minutes. it is unknown what was occurring with the patient at this time. pump flow was also affected by the variation in ps. the optical parameters and cvp were stable (besides a shift in 5s parameters when the console was exchanged on 19-sep-2025). the cause of the placement signal issue was not determined since insufficient clinical details were provided and no product was returned. the pump sn (b)(6) passed all post-sterile inspection checks.
1575 | 1220648-2026-00619 | 24054230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054230&pc=OZD | 09/19/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740369 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. it was reported that hemolysis was present with a plasma-free hemoglobin (pfhgb) level of 200. the previous pfhgb level was 100 while on extracorporeal membrane oxygenatio (ecmo) and prior to lv vent placement. an echocardiogram was ordered to verify impella cp positioning, and the performance level was reduced from p4 to p2. the ecmo team plans to discuss the potential addition of a second drainage cannula due to lower-than-expected circuit flows. hemoglobin remains stable (13.8 yesterday and 14.4 today), and no blood products have been administered. | [5 empty] | Pump and purge cassette were returned for investigation. no presence of biomaterial or damage was found on the impeller. the cannula was found kinked on the returned product; no other damage was noted. the pump was run per specification in a bucket of water. mechanical interaction with blood (hemolysis): as per the clinical details, the patient was on ecmo prior to impella placement. plasma free hemoglobin (pfhgb) increased after impella cp placement on the day of implant and was noted to decrease during the remaining 7 days of the case run. according to the data log, there was no suction, motor current spike, or major placement signal or flow trend during the time of hemolysis. no defect was found on the pump related to hemolysis. the cause of the hemolysis issue was not determined due to insufficient clinical information. device in wrong position: data log suggests several suction and position-related alarms during the end of the case run. clinical information suggests the impella was found deep in the lv. the returned product also confirms cannula kink damage, most likely due to the device being in the wrong position; however, the positioning issue was not determined since sufficient clinical information was not provided. clinical symptom (thrombosis/thrombus): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. the pump sn(b)(6) passed all post sterile inspection checks.
1576 | 1220648-2026-00620 | 24054240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054240&pc=OZD | 09/19/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710227 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1577 | 1220648-2026-00621 | 24054280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054280&pc=OZD | 09/26/2025 |  |  | 01/13/2026 |  |  | 09/26/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025541119 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. massive hemolysis occurred during the procedure, confirmed by laboratory indicators and visual inspection of the extracorporeal circuit. a significant reduction in flow was observed, decreasing from position 7 (p7) to position 3 (p3), which coincided with the onset of hemolysis. system parameters were adjusted to stabilize flow and minimize further hemolysis, and affected components were inspected. patient safety was prioritized throughout the intervention, and no additional adverse outcomes have been reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1578 | 1220648-2026-00622 | 24054324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24054324&pc=OZD | 10/16/2025 |  |  | 01/13/2026 |  |  | 10/16/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025536822 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Ventricular Fibrillation | Thrombocytopenia | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission
1579 | 1220648-2026-00629 | 24055303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24055303&pc=OZD | 01/08/2026 | 01/13/2026 |  | 01/13/2026 |  |  | 01/08/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1356082 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | No | Other | User facility | [11 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1580 | 1220648-2026-00630 | 24055321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24055321&pc=OZD | 01/08/2026 | 02/03/2026 |  | 01/13/2026 |  |  | 01/08/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026772535 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation | [10 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1581 | 1220648-2026-00631 | 24055457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24055457&pc=PYX | 01/05/2026 |  |  | 01/13/2026 |  |  | 01/05/2026 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1641550 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | Asian | Death | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal unreliable alarm occurred. consoles were switched and the alarm was resolved. the patient's outcome at explant was expired. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1582 | 1220648-2026-00626 | 24055463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24055463&pc=OZD | 11/25/2025 | 02/11/2026 |  | 01/13/2026 |  |  | 11/25/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026759966 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1583 | 1220648-2026-00625 | 24055471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24055471&pc=OZD | 01/07/2026 | 01/13/2026 |  | 01/13/2026 |  |  | 01/08/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026787539 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1584 | 1220648-2026-00636 | 24059227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059227&pc=OZD | 12/25/2025 |  |  | 01/13/2026 |  |  | 12/26/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795471 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1585 | 1220648-2026-00635 | 24059234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059234&pc=OZD | 12/01/2025 |  |  | 01/13/2026 |  |  | 12/01/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775558 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Nerve Damage | Renal Failure | Hematuria | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1586 | 1220648-2026-00642 | 24059309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059309&pc=OZD | 01/07/2026 | 01/13/2026 |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026741036 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1587 | 1220648-2026-00644 | 24059312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059312&pc=OZD | 12/18/2025 | 01/13/2026 |  | 01/13/2026 |  |  | 01/05/2026 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP , JP |  |  | OZD |  |  | IMPELLA CP , JP |  |  | ABIOMED, INC. |  |  | US |  | 2026786054 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Obstruction/Occlusion | Vascular Dissection | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 87 year old female on an impella cp for support due to acute myocardial infarction. during the procedure, chronic total occlusions were confirmed in both the right coronary artery and the left circumflex artery, with circumferential, severe calcified lesions. considering the condition of the lesion, prior to dilation using the cutting balloon, an intravascular lithotripsy (ivl) catheter was advanced under support of a guide extension catheter, and dilation was attempted; however, rupture was observed in a total of two ivl catheters. dilation was successfully achieved with a third ivl catheter, after which the subject cutting balloon device was advanced for dilation. during rapid pressurization and dilation of the lesion up to 20 atm, vascular perforation was confirmed. hemostasis was successfully achieved by overlapping implantation of covered stents. due to the patient¿s poor condition, initiation of ecmo was considered; however, the patient¿s family declined life-prolonging treatment with ecmo. subsequently, the patient passed away. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1588 | 1220648-2026-00643 | 24059315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059315&pc=OZD | 01/07/2026 | 01/13/2026 |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1419260 |  | 0042-0000-AU |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation | [9 empty] | Malfunction | No | Yes | Overheating of Device | Power Problem | No Apparent Adverse Event |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
1589 | 1220648-2026-00645 | 24059317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059317&pc=OZD | 01/27/2026 | 01/13/2026 |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - |  |  | US |  | 2026714713 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1590 | 1220648-2026-00647 | 24059319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059319&pc=OZD | 01/05/2026 | 02/03/2026 |  | 01/13/2026 |  |  | 01/05/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744379 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup
1591 | 1220648-2026-00648 | 24059331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059331&pc=OZD | 01/06/2026 | 01/13/2026 |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026762648 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male | 85 KG |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was successfully weaned from the pump. | The user facility reported an impella cp in a 57 year old male patient for support due to acute myocardial infarction. during support, the aortic placement signal had a big difference in numbers to hemodynamic monitoring from hospital. the aic showed aop and lv signal numbers after initial implant around 200, meanwhile blood pressure was only around 70. the pump was still providing flow and the patient was visibly stabilised after implant. the decision was made to leave the pump in-situ and the patient remains on full support in icu now. |  |  |  |  | Correction: e4. additional information received; b5 updated. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1592 | 1220648-2026-00649 | 24059576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059576&pc=OZD | 01/07/2026 | 01/13/2026 |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026751868 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1593 | 1220648-2026-00639 | 24059582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059582&pc=OZD | 01/04/2026 |  |  | 01/13/2026 |  |  | 01/05/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, JP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026759095 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1594 | 1220648-2026-00652 | 24059624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059624&pc=OZD | 01/07/2026 | 01/13/2026 |  | 01/13/2026 |  |  | 01/08/2026 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1439042 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 50 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported their automated impella controller (aic) could not recognize its purge disc. an alert sounded. the latch of the purge disc was broken, so the aic was swapped. no patient harm has been reported. | [5 empty] | Corrections: d3, e4, h5, h8. d6 should have been left blank. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1595 | 1220648-2026-00655 | 24059640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059640&pc=OZD | 09/19/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711533 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Patient-Device Incompatibility | Patient Device Interaction Problem | [5 empty] | Perforation of Vessels | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1596 | 1220648-2026-00656 | 24059642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059642&pc=OZD | 09/19/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026731729 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Sepsis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient had an elevated white blood cell count and was suspected to have developed septic shock, resulting in vasoplegia. the patient subsequently expired. additional information indicated that the pump was not considered to have contributed to the occurrence of sepsis or the patient¿s death. | [5 empty] | Investigation summary: no product was returned for investigation. sepsis: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the sepsis was unable to be determined due to insufficient clinical details. bradycardia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the bradycardia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1970145 device history batch: subcomponent lot: n/a. device history review (b)(4) device with sn: (b)(6) passed all post sterile inspection checks.
1597 | 1220648-2026-00654 | 24059644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059644&pc=OZD | 09/19/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695501 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1598 | 1220648-2026-00657 | 24059648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059648&pc=OZD | 09/19/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692918 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 138 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn |  |  |  |  | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient was transferred from an outside hospital on va ecmo and impella 5.5 support. neurological status was intact, and no issues were noted with impella support at performance level p-3. echocardiography measured 5 cm to the aortic valve annulus. the patient arrived with an existing axillary pocket hematoma and exhibited north-south syndrome. right ventricle was dilated. the plan was to consult surgery for possible adjustment of the drainage cannula, which was positioned too low, and for hematoma assessment. on (b)(6) 2025, the patient experienced a gastrointestinal bleed and was off anticoagulation until the morning of the report. it was also reported that there was a sleeve tear. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
1599 | 1220648-2026-00658 | 24059651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059651&pc=OZD | 09/19/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | OZDTEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686745 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 114 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | Difficult to Open or Close | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that while the patient was repositioned in bed, the impella device was not secured and was inadvertently displaced from the ventricle into the descending aorta. a decision was made to remove the impella and proceed with extracorporeal membrane oxygenation (ecmo) cannulation. | [5 empty] | Investigation summary: the pump was not returned for investigation. as per the ifu, the device should be in the locked position after initial placement unless repositioning is needed. according to the clinical details, the impella was not locked when the pump was pulled out. device in wrong position: the cause of the positioning issue was determined to be an unintentional use error, as the pump was pulled out of position during patient repositioning when the touhy was not locked. difficult to open/close catheter anchor or touhy: the cause of the touhy issue was unintentional use-related, based on the given clinical details. device history lot: device batch: 1931560. device history batch: subcomponent lot: na. device history review: the pump sn (b)(6) passed all post sterile inspection checks.
1600 | 1220648-2026-00659 | 24059660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059660&pc=OZD | 09/19/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713231 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1601 | 1220648-2026-00660 | 24059664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059664&pc=OZD | 08/12/2025 |  |  | 01/13/2026 |  |  | 09/19/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655251 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Fever | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported patient had a cp the complainant reported that a patient was implanted with an impella device for mechanical circulatory support and has been on impella 5.5 support for 35 days. on the day of the report, the aop signal displayed 67/43 mmhg with map 47, while the patient¿s actual arterial line blood pressure was 94/68 mmhg with map 77. the medical team was not concerned and determined that no adjustment to the placement signal was required. the patient was otherwise stable, ambulating in the hallway, and awaiting heart transplant, with plans to remain on support until then.. additional information indicated the patient was alert and oriented and on room air. an episode of ventricular tachycardia occurred overnight, during which the patient was shocked by their automatic implantable cardioverter defibrillator (aicd). oral amiodarone was initiated, and no further arrhythmias have occurred. the patient was febrile overnight; blood cultures were obtained and preliminary results are pending. the patient has been accepted to columbia for heart transplant evaluation but must complete an antibiotic course as per the care team. argatroban was temporarily held for line removal and is being restarted gradually; current partial thromboplastin time (ptt) is 53. no concerns related to the impella device were reported. the team was advised to contact clinical support center (csc) or the local team for any additional needs. | [5 empty] | Investigation summary ==> no product was returned. optical sensor issue - after 25 days on support, aop signal reading 67/43 map 47, while patient¿s actual blood pressure by a-line 94/68 map 77. medical team not concerned and does not require placement signal adjustment. data log review shows placement signal low alarms with no mc spikes, no ps adjustments made, no purge trends, and no flow trends. imc logs were not recovered. the cause of the optical signal issue was not determined as no product was returned, inconclusive data log review, and insufficient clinical details. patient codes - fever, infection: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the fever/infection was unable to be determined due to insufficient clinical details. device history lot ==> device lot: 1908573. device history batch ==> subcomponent lot: n/a device history review ==> pump s/n (b)(6) passed all post sterile inspection checks.
1602 | 1220648-2026-00661 | 24059665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059665&pc=OZD | 09/20/2025 |  |  | 01/13/2026 |  |  | 09/20/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727108 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1603 | 1220648-2026-00662 | 24059701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059701&pc=OZD | 01/06/2026 | 01/13/2026 |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2026781695 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intraoperative Cardiac Injury | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1604 | 1220648-2026-00663 | 24059752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059752&pc=PYX | 09/15/2025 |  |  | 01/13/2026 |  |  | 09/15/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025608403 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female | 101 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. aortic pressure immediately increased upon initiation. p level slowly increased to p9 indicated by echocardiogram on the heart. the patient was taken to icu for recovery after securing impella. the patient had a fever and was given antibiotics. it was reported that the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1605 | 1220648-2026-00664 | 24059861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059861&pc=OZD | 09/15/2025 |  |  | 01/13/2026 |  |  | 09/15/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026674743 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1606 | 1220648-2026-00665 | 24059919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059919&pc=OZD | 11/20/2025 |  |  | 01/13/2026 |  |  | 11/21/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026730400 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male | 109 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient underwent impella 5.5 axillary explant. it was reported the patient developed right arm mottling with non-dopplerable pulse. the patient was diagnosed with right subclavian thrombosis, requiring a vascular surgery consult. subsequently, the patient underwent right axillary artery stent placement. post-procedure, the patient was reported to have normal strength and sensation. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1607 | 1220648-2026-00666 | 24059953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24059953&pc=OZD | 11/21/2025 |  |  | 01/13/2026 |  |  | 11/21/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026686684 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 124 KG | Non Hispanic | White | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information received from the complainant reported the pump was weaned and explanted before fears of critical failure. the patient survived after the explant. | The complainant reported that a patient had a coughing spell, which caused the impella catheter to dive deep into the left ventricle (lv). catheter was repositioned under echo guidance. purge pressures increased to >1100 mmhg with flow <1 ml/hr. purge cassette was changed, but the issue persisted. physician decided to proceed with a procedure while maintaining impella support. intra-aortic balloon pump (iabp) was placed for support through the weekend. patient was on aggrastat and heparin drip; partial thromboplastin time (ptt) >180. |  |  |  |  | B5: additional information received. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. clinical details noted that shortly after pump was repositioned under echo, purge pressure suddenly increased to greater than 1100mmhg and purge flow less than 1ml/h. purge flow blocked alarms noted. decision was made to replace pump with iabp. the cause of the high purge pressure could not be determined as no product or logs were returned for evaluation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1608 | 1220648-2026-00670 | 24060333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060333&pc=OZD | 12/29/2025 |  |  | 01/13/2026 |  |  | 12/29/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1136926 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during setup of an impella cp and automated impella controller (aic), the pump was connected to the aic, primed, and subsequently had an alarm stating there was a fiber optic sensor malfunction. the medical team followed troubleshooting steps that appeared on the aic screen and the issue resolved. no issues were reported during the implantation of the impella cp. during the patient¿s procedure and while on impella support, there were alarms for placement signal not reliable, and the aorta placement signal was in and out of reading. the procedure continued without issue, and the impella pump was discarded after use. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves two impella devices: impella cp, with serial number (b)(6). automated impella controller with serial number ( b)(6). this mdr specifically addresses automated impella controller with serial number i(b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D6a and d6b implant and explant dates should have been left blank on the initial report that was submitted. d9 device was returned and date added. g1 manufacturing site name should have been left blank on the initial report that was submitted. h6 type of investigation, investigation findings, and investigation conclusions were updated accordingly based on the device being returned. h10 added related report number. h11 updated due to device being returned. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
1609 | 1220648-2026-00669 | 24060335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060335&pc=OZD | 01/06/2026 | 01/15/2026 |  | 01/13/2026 |  |  | 01/06/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734201 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1610 | 1220648-2026-00671 | 24060342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060342&pc=OZD | 12/28/2025 | 02/11/2026 |  | 01/13/2026 |  |  | 12/29/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026782171 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1611 | 1220648-2026-00673 | 24060352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060352&pc=PYX | 01/07/2026 |  |  | 01/13/2026 |  |  | 01/07/2026 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1701427 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1612 | 1220648-2026-00674 | 24060357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060357&pc=OZD | 01/02/2026 |  |  | 01/13/2026 |  |  | 01/03/2026 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026736989 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 79 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 61-year-old female on impella cp due to acute myocardial infarction. during support, the patient required a unit of blood overnight for two nights due to significant nosebleed. the patient was eventually weaned off the device. | [5 empty] | The investigation for the reported nosebleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the nosebleed could not be determined.
1613 | 1220648-2026-00676 | 24060578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060578&pc=OZD | 09/26/2025 |  |  | 01/13/2026 |  |  | 09/26/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026725123 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 86 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient¿s left ventricular assist device (lvad) and transplant evaluation has been placed on hold pending identification of the source of gastrointestinal bleeding. both colonoscopy and esophagogastroduodenoscopy (egd) have been performed, but results were inconclusive. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1614 | 1220648-2026-00675 | 24060579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060579&pc=OZD | 09/26/2025 |  |  | 01/13/2026 |  |  | 09/26/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026668977 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor | [9 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1615 | 1220648-2026-00677 | 24060586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060586&pc=OZD | 09/26/2025 |  |  | 01/13/2026 |  |  | 09/26/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713989 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor | [9 empty] | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1616 | 1220648-2026-00679 | 24060787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060787&pc=OZD | 10/11/2025 |  |  | 01/13/2026 |  |  | 10/11/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026761317 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1617 | 1220648-2026-00680 | 24060822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060822&pc=OZD | 10/07/2025 | 11/07/2025 |  | 01/13/2026 |  |  | 10/08/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025594266 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the device cannot move the impella in or out. pushing in the blue tab or inserted the yellow key and it still would not move. the impella will not go in or out. it was noted that the patient experienced ectopy and hemostasis no heparin and transfusion. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1618 | 1220648-2026-00683 | 24060858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060858&pc=OZD | 10/14/2025 |  |  | 01/13/2026 |  |  | 10/14/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026644324 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient had pink or red bloody urine output, despite confirming the impella cp position. heparin was restarted but awaiting bicarbonate purge solution from the pharmacy. the patient had been weaned from bilevel positive airway pressure (bipap) but had to be placed back on bipap due to intermittent low oxygen saturations. tachycardic with hypotension and rising creatinine. the patient was escalated to impella 5.5 for increased support and the impella cp was explanted with hemostasis achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1619 | 1220648-2026-00682 | 24060859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060859&pc=OZD | 10/14/2025 |  |  | 01/13/2026 |  |  | 10/14/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026758095 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1620 | 1220648-2026-00681 | 24060860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060860&pc=OZD | 10/09/2025 |  |  | 01/13/2026 |  |  | 10/10/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026758045 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the bedside nurse noticed loss of pulse in impella leg. lactic trended up and medical doctor elected to removed the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1621 | 1220648-2026-00684 | 24060864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060864&pc=OZD | 10/09/2025 |  |  | 01/13/2026 |  |  | 10/09/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026682470 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | No Apparent Adverse Event | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Sepsis | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was given antibiotics for positive blood cultures. additional information indicated in the data log shows multiple suction alarms on (b)(6) and (b)(6), with an overall decreasing trend in placement signal and pump flow compared to previous days of support. this was resolved after switching the pump to p8. later, on the 18th, the pump was reported on increased p-level (p9) with a fluctuating trend in placement signal and pump flow on the following day. on (b)(6), there was a drop in pump flow and motor current reported with suction alarms while running on a steady p-level for about one hour, with a steady decrease in flow that resolved later while running on p8. the pump continued to have fluctuating pump flow with intermittent suction alarms later during the case. no abnormalities were observed in motor current spikes, optical signal issue or positioning issues during the case. the cause of the suction issue was most likely related to patient condition, based on trends in the data log and provided clinical details. | [5 empty] | H6. medical device problem code is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1622 | 1220648-2026-00685 | 24060875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060875&pc=OZD | 10/14/2025 |  |  | 01/13/2026 |  |  | 10/14/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026633585 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Malposition of Device | [6 empty] | Pulmonary Edema | Renal Failure | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with renal failure who was extubated, now has pulmonary edema, and is being managed with possible escalation to p9 on impella support. the echo confirms no left ventricle (lv) distension and correct impella positioning. patient deteriorated despite optimal medical treatment, decision made to withdraw support. no fault or issue directed towards device, patient died. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary no product was returned for investigation. no data logs were returned for analysis. pump suction: clinical details mention that there were suction alarms on (b)(6) 2025 when the patient was mobilizing to side of bed. echocardiogram (echo) confirmed the pump to be slightly deep, and catheter was pulled back 2cm after which suction alarms were resolved. the root cause of the suction alarms was most likely user related attributed to the pump being in the wrong position. device in wrong position: clinical details mention that the pump was found to be slightly deep when the position was checked under echo after suction alarms were triggered when mobilizing to the side of the bed. the pump was repositioned and externally fixated at a better position. the root cause of the positioning issue was most likely user related as the pump was not externally fixated at the ideal position prior to triggering suction alarms based on provided clinical details. pulmonary edema, renal failure: clinical details mention that renal failure and pulmonary edema was recognized on (b)(6) 2025 while the patient was on impella support. the pump flow was increased to p9 and echo confirmed no left ventricle (lv) distension and correct impella placement. the cause of the injury was not determined due to insufficient clinical details. device history review pump passed all post-sterile inspection checks.
1623 | 1220648-2026-00686 | 24060884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060884&pc=OZD | 10/10/2025 |  |  | 01/13/2026 |  |  | 10/10/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026716668 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1624 | 1220648-2026-00687 | 24060894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060894&pc=OZD | 10/08/2025 |  |  | 01/13/2026 |  |  | 10/08/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713604 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the device impella cp was located in the papillary muscle and could not be optimally positioned. subsequently, the device could not be implanted due to the small diameter of the subclavian artery. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1625 | 1220648-2026-00688 | 24060905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24060905&pc=OZD | 10/14/2025 |  |  | 01/13/2026 |  |  | 10/14/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1155701 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male | 93 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while preparing for device implant, automated impella controller (aic) sn: (b)(6) would not recognize purge disc and would not prime or move past the ¿2b: click purge disc, then close door¿ screen. different aic (b)(6) obtained and team able to successfully prime and implant impella using the alternate aic. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1626 | 1220648-2026-00691 | 24061022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24061022&pc=OZD | 11/20/2025 |  |  | 01/13/2026 |  |  | 11/21/2025 | 01/13/2026 |  |  |  |  | 01/13/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026747633 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 81 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient experienced axillary pocket bleeding post-procedure, which required transfusion. the bleeding has since resolved. the attending physician attributed the cause to supratherapeutic anticoagulation levels in combination with an argatroban infusion administered during the case. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1627 | 1220648-2026-00693 | 24061085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24061085&pc=OZD | 11/20/2025 | 01/27/2026 |  | 01/13/2026 |  |  | 11/21/2025 | 01/13/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026705220 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1628 | 1220648-2026-00497 | 24031561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031561&pc=OZD | 09/11/2025 | 01/30/2026 |  | 01/12/2026 |  |  | 09/11/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703594 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the user facility reported patient had a 5.5 pump placed to support the surgical patient. the 5.5 was found to stop on day 3 and was able to be restarted. the restarted pump had higher "current consumption" prompted the team to wean and explant the 5.5 pump after the 3 days of support. | [5 empty] | Investigation summary: no product returned for evaluation. temporary pump stop: clinical details noted that overnight 3 days into support the pump had a sudden increase in motor current then stopped. pump was able to be restarted at p-3 with increased motor current. data logs were reviewed and lt logs showed a motor current spike at time of pump stop with alarms #115 and #13. pump was able to be restarted and lt log showed increased purge pressure and saturated flows for approximately 12 hours after alarm. the cause of the temporary pump stop was due to biomaterial ingestion based returned data log analysis. device history (lot). device lot: 1930314. device history (batch). subcomponent lot: n/a4. device history review. pump sn (b)(6) passed all post-sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1629 | 1220648-2026-00501 | 24031562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031562&pc=OZD | 12/29/2025 | 01/07/2026 |  | 01/12/2026 |  |  | 12/30/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026775521 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1630 | 1220648-2026-00503 | 24031563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031563&pc=OZD | 12/29/2025 | 01/09/2026 |  | 01/12/2026 |  |  | 12/29/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026726966 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1631 | 1220648-2026-00506 | 24031564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031564&pc=OZD | 12/29/2025 | 01/09/2026 |  | 01/12/2026 |  |  | 12/29/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026781152 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1632 | 1220648-2026-00507 | 24031565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031565&pc=OZD | 01/01/2026 |  |  | 01/12/2026 |  |  | 01/04/2026 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED EUROPE GMBH - 3007904765 |  |  | GM |  | 1383039 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Decrease in Pressure | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1633 | 1220648-2026-00508 | 24031567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031567&pc=OZD | 12/29/2025 | 01/09/2026 |  | 01/12/2026 |  |  | 12/29/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1399060 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation | [10 empty] | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1634 | 1220648-2026-00500 | 24031572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031572&pc=OZD | 11/20/2025 | 02/03/2026 |  | 01/12/2026 |  |  | 11/20/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026742179 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 65 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient had significant bleeding from impella insertion site, requiring the use of femstop, and bleeding was able to be controlled, with palpable pulse in distal limb. patient was given numerous units of prbcs, platelets, cryogenic, ffp overnight, and was able to achieve hemostasis at insertion site. plan moving froward will depend on how the patient progresses. it was also reported that the patient went for stat heat computed tomography (ct) this morning which revealed large intracerebral hemorrhage (ich) stroke with shift. bivalirudin infusion stopped due to intracranial hemorrhage (ich). it was reported that the patient expired at explant. | The complainant reported additional information upon request. the catheter was not saved. the cause of death reported was cardiogenic shock. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated. the investigation is ongoing.
1635 | 1220648-2026-00499 | 24031573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031573&pc=OZD | 09/11/2025 |  |  | 01/12/2026 |  |  | 09/11/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026714471 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the staff observed increased swelling and hematoma at the incisional site. the chest tube placed intraoperatively within the pocket drained approximately 500 ml of sanguineous output. hemoglobin and hematocrit decreased. the patient is being closely monitored. should hemoglobin/hematocrit continue to decline, the plan is to return to the operating room for washout. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1636 | 1220648-2026-00513 | 24031574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031574&pc=OZD | 01/06/2026 | 01/09/2026 |  | 01/12/2026 |  |  | 01/06/2026 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755399 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1637 | 1220648-2026-00498 | 24031575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031575&pc=OZD | 09/09/2025 |  |  | 01/12/2026 |  |  | 09/09/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026733042 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1638 | 1220648-2026-00505 | 24031579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031579&pc=OZD | 12/28/2025 | 02/24/2026 |  | 01/12/2026 |  |  | 12/28/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026786842 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1639 | 1220648-2026-00509 | 24031586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031586&pc=OZD | 01/06/2026 | 01/09/2026 |  | 01/12/2026 |  |  | 01/06/2026 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026736252 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1640 | 1220648-2026-00514 | 24031592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031592&pc=OZD | 01/06/2026 | 01/09/2026 |  | 01/12/2026 |  |  | 01/06/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026796345 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 65 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Material Split, Cut or Torn | [5 empty] | Hematuria | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | While patient rounding observation of hematuria. echo conducted at bedside. impella repositioned at bedside and pulled back 2 cm. | The user facility reported a 74-year-old female patient in whom the impella sheath and sterile cover were noted to have come apart. the site was covered, and the team elected not to reposition the device. the timing and cause of the issue were unclear, as it was noticed earlier that day. no additional information was provided. |  |  |  |  | Additional information received; b1, b2, b5, h1, and h6 updated based on the information. correction: e4. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1641 | 1220648-2026-00491 | 24031618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24031618&pc=OZD | 11/20/2025 | 01/09/2026 |  | 01/12/2026 |  |  | 11/20/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026726698 |  | 000000000001000482 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male | 150 KG |  | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the inlet was close to the mitral valve, which were not able to correct properly. it was noted that there was a positioning issue. it was reported that the patient experienced septic shock, an emergency impella was implanted due to acute glasgow coma scale (gcs). the patient was reported to have survived the procedure. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the positioning issue was unable to be determined due to insufficient clinical details. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1642 | 1220648-2026-00515 | 24035026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035026&pc=OZD | 12/29/2025 |  |  | 01/12/2026 |  |  | 12/29/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026777715 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1643 | 1220648-2026-00517 | 24035054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035054&pc=OZD | 11/21/2025 | 01/29/2026 |  | 01/12/2026 |  |  | 11/21/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766067 |  | 000000000001000080 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1644 | 1220648-2026-00518 | 24035187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035187&pc=OZD | 11/21/2025 | 02/10/2026 |  | 01/12/2026 |  |  | 11/21/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025573760 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 98 KG | Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hematoma | Pulmonary Edema | Respiratory Failure | Hematuria | Multiple Organ Dysfunction Syndrome | [7 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a hematoma was observed at the impella insertion site. the heparin infusion was paused, and manual pressure was applied for 20 minutes prior to patient transfer. the hematoma subsequently decreased in size and softened. on the same day, the patient was intubated and noted to be hypoxic with evidence of pulmonary edema. an echocardiogram revealed an ejection fraction of less than 20%, tapse of 1.1 cm, ivc measuring 1.3 cm, and impella depth between 1.5¿2 cm. hypoxia and impella positioning were discussed with the physician, who confirmed satisfaction with the device placement. plans were made to reassess hemodynamic data and arterial blood gases for potential ecmo evaluation. sedation was initiated with propofol and midazolam. the patient was febrile at 39.4°c and exhibited possible seizure-like activity; levetiracetam (keppra) was administered. later it was reported that the patient was in multi-organ failure with rising lactate levels and was deemed not stable enough for advanced therapies. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported hematoma, respiratory failure, organ failure, pulmonary edema, fever, hematuria has been completed. the device was not returned for evaluation. additionally, clinical details were insufficient to determine the root cause. therefore, the cause of reported events was not determined due to insufficient clinical details.
1645 | 1220648-2026-00516 | 24035247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035247&pc=OZD | 11/21/2025 |  |  | 01/12/2026 |  |  | 12/26/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026747049 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 74 KG | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal went to dashes momentarily about 10 seconds while in the room. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1646 | 1220648-2026-00520 | 24035302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035302&pc=OZD | 12/29/2025 |  |  | 01/12/2026 |  |  | 12/29/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026751799 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 55 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that an impella cp was unable to placed for mechanical circulatory support due to tortuosity in iliac and aorta. the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be recei as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | H6: updated codes based on the results of the investigation. investigation summary. no product was returned. failure to advance: the cause of the deliver issue was determined to be patient condition as the patient had tortuous vessels in iliac and aorta preventing successful delivery of impella. device history review: device passed all post sterile inspection checks.
1647 | 1220648-2026-00526 | 24035316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035316&pc=OZD | 11/21/2025 |  |  | 01/12/2026 |  |  | 11/21/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026770564 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1648 | 1220648-2026-00523 | 24035435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035435&pc=OZD | 08/18/2025 |  |  | 01/12/2026 |  |  | 08/18/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026698084 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella sheath was placed in the same access site as the previously explanted and bleeding site where the other impella device had been placed. the sheath, measuring approximately fourteen by thirteen centimeters, was kinked after delivery, and it was not possible to pass the impella device through. a decision was made to insert a fourteen french edward¿s sheath; however, the device was still unable to pass through. subsequently, a fourteen french by twenty-five centimeter sheath was inserted, but it also became kinked and the impella device could not be advanced. the decision was then made to abort the insertion. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1649 | 1220648-2026-00530 | 24035442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035442&pc=OZD | 11/22/2025 | 02/10/2026 |  | 01/12/2026 |  |  | 12/26/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026682014 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 75 YR |  | Male | 102 KG |  | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a hematoma approximately the size of a softball was noted around the impella sheath, and the automated impella controller (aic) alarm indicated that the impella was positioned in the aorta. the cicu fellow was applying manual pressure at the time of assessment. the impella site was covered with tegaderm extending down the sterile sleeve on the catheter, which had been looped into the knee immobilizer straps proximally and distally. after removal of the tegaderm and straps, proper angle alignment was achieved. a subsequent echocardiogram performed by another fellow revealed that only the pigtail segment of the catheter was across the aortic valve. the pump was repositioned so that the inlet crossed back into the appropriate position, and vascular surgery evaluated the patient, recommending continued manual pressure. the access site was stabilized, and the impella was repositioned under echocardiographic guidance across the valve, with flow increased to p2. later that night, recurrent bleeding from the groin site occurred, prompting vascular surgery to take the patient to the operating room for vascular repair. the impella cp device was removed, and no additional mechanical circulatory support device was placed. the patient subsequently expired overnight. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The following sections have been updated/corrected: b4, e1 (city and state), g3, g6, h6, h11. the investigation for the reported access site bleed and hematoma has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the cause of the bleed and hematoma was most likely use issue related since it occurred after the arrival of patient to facility.
1650 | 1220648-2026-00527 | 24035491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035491&pc=OZD | 11/25/2025 |  |  | 01/12/2026 |  |  | 11/25/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026753526 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 66 YR |  | Male | 65 KG | Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Stroke/CVA | Thrombocytopenia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported tha thrombocytopenia was identified. the patient tested positive for heparin-induced thrombocytopenia (hit), and surgical intervention was placed on hold pending serotonin release assay (sra) results. additionally, anemia was noted, leading to the suspension of bivalirudin therapy. there were no signs of active bleeding observed. subsequently, stroke was confirmed following overnight computed tomography, which revealed a left occipital lobe infarction with hemorrhagic transformation and bilateral cerebellar hemispheric infarctions. ultimately, the family elected to withdraw care, and the patient expired following device explanation. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1651 | 1220648-2026-00529 | 24035529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035529&pc=OZD | 11/21/2025 |  |  | 01/12/2026 |  |  | 12/26/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026740321 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1652 | 1220648-2026-00519 | 24035560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035560&pc=PYX | 12/29/2025 |  |  | 01/12/2026 |  |  | 12/29/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529006 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received. the pump is not available. it was removed and discarded. | An impella rp flex was implanted in a 75-year-old female for acute myocardial infarction complicated by cardiogenic shock. the device was successfully placed, and the groin access site was initially clean, dry, and intact in the catheterization lab. during patient transfer from the catheterization lab table to the bed and subsequently upstairs, the yellow sheath on the impella rp flex slid back slightly, resulting in significant bleeding. the bleeding was controlled by readvancing the yellow sheath and securing it with re-suturing under tension. no further complications with the device were noted. the event was classified as a serious injury/harm; however, based on the information available at the time of reporting, no causal relationship between the impella rp flex and the reported event was identified. |  |  |  |  | B5 additional information was received and added. h6 type of investigation code was revised to reflect the device being discarded. h11 revised additional manufacturer narrative to reflect that the device was discarded. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary: no product returned. access site adverse event/major bleed: the cause of the access site adverse event was user related as the repo unit was pulled out during patient transfer leading to bleeding which was resolved by readvancing it and securing with forward tension sutures. device history lot: device lot: 1813678. device history batch: subcomponent lot: n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
1653 | 1220648-2026-00521 | 24035564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035564&pc=OZD | 12/29/2025 |  |  | 01/12/2026 |  |  | 12/29/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025523427 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1654 | 1220648-2026-00524 | 24035573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035573&pc=OZD | 08/19/2025 |  |  | 01/12/2026 |  |  | 08/19/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | PUMP 371 14F IMPELLA CPLT CMR SET |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026722891 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the pump was unable to advance through the end of the peel-away sheath after multiple attempts were made by two different physicians. a kink was noted under fluoroscopy, and the motor housing was unable to advance. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1655 | 1220648-2026-00525 | 24035590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035590&pc=OZD | 11/21/2025 |  |  | 01/12/2026 |  |  | 11/21/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765486 |  | 000000000001000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1656 | 1220648-2026-00531 | 24035628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035628&pc=OZD | 11/22/2025 |  |  | 01/12/2026 |  |  | 12/26/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026759650 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor | [9 empty] | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Arrhythmia | Hemolysis | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1657 | 1220648-2026-00532 | 24035992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24035992&pc=OZD | 12/19/2025 |  |  | 01/12/2026 |  |  | 12/26/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026780957 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. vascular assessment revealed no flow to the right arm, prompting the vascular surgeon to perform an investigation and revascularization. during this intervention, the impella device continued to function as intended at p8 with stable flow. the surgeon successfully restored perfusion to the extremity; however, multiple blood products were administered during subclavian revascularization. while surgical closure was underway, the patient experienced pulseless electrical activity (pea). cardiopulmonary resuscitation (cpr) was performed for more than 20 minutes. the interventional cardiologist consulted with the family, and a decision was made to withdraw care. in accordance with team wishes, the impella device was left in place and was unavailable for return. the care team reported no concerns regarding the impella pump¿s function throughout the procedure. additional information noted that the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. ischemia & cardiac arrest: the root cause of the injury was unable to be determined due to insufficient clinical details.
1658 | 1220648-2026-00533 | 24036009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24036009&pc=PYX | 11/21/2025 |  |  | 01/12/2026 |  |  | 11/21/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026671771 |  | 000000000001000323 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 73 KG | Non Hispanic | Black Or African American | Death | Yes | No | Increase in Pressure | [6 empty] | Arrhythmia | Thrombocytopenia | Asystole | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. purge pressure was elevated. the plan was to discontinue (d/c) heparin in the purge solution once flow exceeded 2.0 liters per minute due to concern for heparin-induced thrombocytopenia (hit). the patient¿s vasopressor support continued to increase, and disseminated intravascular coagulation (dic) was suspected. the patient then developed cardiac ectopy and subsequently progressed to asystole. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1659 | 1220648-2026-00538 | 24036997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24036997&pc=OZD | 09/08/2025 |  |  | 01/12/2026 |  |  | 09/08/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713189 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1660 | 1220648-2026-00542 | 24037052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24037052&pc=OZD | 08/05/2025 | 02/11/2026 |  | 01/12/2026 |  |  | 08/06/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026712015 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the nurse stated the laboratory samples were hemolyzing and the urine appeared dark red. | [5 empty] | The cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1661 | 1220648-2026-00540 | 24037282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24037282&pc=OZD | 11/21/2025 |  |  | 01/12/2026 |  |  | 11/21/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026754384 |  | 000000000001000100 |  | No |  |  | No |  |  | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 65 KG | Hispanic | White | Malfunction | Yes | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. an impella 5.5 (a percutaneous left ventricular assist device) was unable to be placed due to the patient¿s vascular anatomy. the insertion procedure for the impella 5.5 was aborted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1662 | 1220648-2026-00536 | 24037303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24037303&pc=OZD | 09/11/2025 |  |  | 01/12/2026 |  |  | 09/11/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026659982 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that upon arrival to the cardiac catheterization laboratory, the impella was placed via the left groin. the impella was running in auto mode, and its position was verified under fluoroscopy. forward tension sutures were placed. the left groin was oozing, and the impella dressing was optimized and improved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1663 | 1220648-2026-00537 | 24037320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24037320&pc=OXY | 09/20/2025 | 11/05/2025 |  | 01/12/2026 |  |  | 09/20/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OXY |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026636140 |  | 000000000001000100 |  | No |  |  | No |  |  | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Hematuria | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. on (b)(6) 2025 (p-00733437), the patient was up in a chair and stated that he was glad the team chose to place the impella 5.5 (a percutaneous left ventricular assist device). the nurse reported that the patient had one episode of sustained ventricular tachycardia (vt) that morning but was not symptomatic. the patient converted back to normal sinus rhythm. on (b)(6) 2025 (p-00733845), the patient¿s urinalysis (ua) was positive for greater than 50 red blood cells (rbc) on the ua report. the urine was pink tinged but began clearing after heparin was paused. the milrinone dose was reduced by half. the patient continued to have runs of ventricular tachycardia (vt). the physician repositioned the impella 5.5 and stated that he would likely escalate the transfer to a ventricular assist device (vad) center. on (b)(6) 2025 (fm-53505), the physician requested insertion of a new impella 5.5 due to ongoing ventricular arrhythmias, which he believed were related to positioning along the septum with inability to achieve midventricular positioning. the doctor did not believe the ventricular ectopy was related to impella positioning; however, dr. devore continued to request replacement due to persistent ventricular arrhythmias, which he attributed to positioning near the septum. transesophageal echocardiogram (tee) showed the device appeared acceptable, but dr. devore felt it was causing ventricular ectopy. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1664 | 1220648-2026-00502 | 24037339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24037339&pc=OZD | 12/27/2025 | 01/09/2026 |  | 01/12/2026 |  |  | 12/29/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656878 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | False Alarm | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1665 | 1220648-2026-00541 | 24037386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24037386&pc=OZD | 09/04/2025 |  |  | 01/12/2026 |  |  | 11/21/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026706512 |  | 000000000001000100 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1666 | 1220648-2026-00543 | 24037425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24037425&pc=PYX | 01/03/2026 |  |  | 01/12/2026 |  |  | 01/04/2026 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025590577 |  | 1000323 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1667 | 1220648-2026-00544 | 24039872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24039872&pc=OZD | 12/25/2025 |  |  | 01/12/2026 |  |  | 12/26/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026754432 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 65 YR |  | Male | 74 KG | Non Hispanic | Asian | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the purge pressure was high. the patient outcome was reported as survived at explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary: no product was returned for evaluation. purge pressure high: clinical details noted that high purge pressure alarms were triggered on day 4 of support. insufficient data logs were returned for day of reported issue. returned data logs showed stable purge pressure and purge flow. no purge abnormalities in time frame of returned data logs. the cause of the high purge pressure could not be determined as no product and insufficient logs were returned for time of reported purge issue. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
1668 | 1220648-2026-00528 | 24048467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24048467&pc=OZD | 11/21/2025 | 01/12/2026 |  | 01/12/2026 |  |  | 12/26/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026763681 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 73 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that increased oozing was observed from all device and intravenous access sites, attributed to a supratherapeutic ptt. the medical team adjusted the heparin infusion accordingly, repeat laboratory tests are pending, and manual pressure was applied as an additional measure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1669 | 1220648-2026-00534 | 24048536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24048536&pc=PYX | 09/02/2025 | 01/12/2026 |  | 01/12/2026 |  |  | 09/03/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026676576 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Difficult or Delayed Positioning | Device Damaged by Another Device | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1670 | 1220648-2026-00547 | 24050050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050050&pc=OZD | 01/05/2026 |  |  | 01/12/2026 |  |  | 01/05/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1109484 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 77 KG | Hispanic |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that a piece broke off of automated impella controller near where the white cord goes into the controller. they switched to new controller and there was no patient harm. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data review: console logs did not contain any data relevant to the complaint. device analysis: console was tested after arriving and the optical connector dust cover was confirmed to be broken. root cause: the root cause of the broken pump connector dust cover was most likely related to excessive force during use. phr summary: there were no issues related to the complaint discovered when reviewing the product history of console. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1671 | 1220648-2026-00552 | 24050053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050053&pc=OZD | 01/04/2026 |  |  | 01/12/2026 |  |  | 01/05/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1228186 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Alarm System | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. it was reported that the placement signal has failed and the alarm "aic failur" is displayed. there is no backup automated impella controller on the ward. the "aic failur" warning disappears. the placement signal is still unavailable. the patient was noted in stable condition and was later successfully weaned from the pump. this report is for the aic and an additional report will be sent for the impella cp (serial number (b)(6)). | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1672 | 1220648-2026-00550 | 24050055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050055&pc=OZD | 01/04/2026 |  |  | 01/12/2026 |  |  | 01/04/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026793082 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 93 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. a nurse report tot the doctor that there were lost pedal pulses in both feet. unable to doppler. right foot (impella side) cold and white, no cap refill. able to doppler popliteal pulse. left foot cool to touch and has red blotching on toes, petechiae looking. the doctor would like to place external fem-fem. it was further reported that the doctor attempted for two hours to get left femoral access using ultrasound at bedside which was unsuccessful. doctor spoke with the intensivist and it was decided to place vascular consult. however, it was reported that the following day care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
1673 | 1220648-2026-00548 | 24050056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050056&pc=OZD | 01/05/2026 |  |  | 01/12/2026 |  |  | 01/05/2026 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2026754377 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1674 | 1220648-2026-00553 | 24050057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050057&pc=OZD | 01/05/2026 |  |  | 01/12/2026 |  |  | 01/05/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1185152 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they saw a message on the controller stating, "failure controller". it appears that the console was switched of within 5 minutes and no impella catheter was connected/ no case was started therefore no patient harm. | [5 empty] | D4 udi information was inadvertently omitted on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1675 | 1220648-2026-00551 | 24050063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050063&pc=OZD | 01/04/2026 |  |  | 01/12/2026 |  |  | 01/05/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026748827 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. it was reported that the placement signal has failed and the alarm "aic failure" is displayed. there is no backup automated impella controller on the ward. the "aic failure" warning disappears. the placement signal is still unavailable. the patient was noted in stable condition and was later successfully weaned from the pump. this report is for the impella cp and an additional report will be sent for the aic (serial number (b)(6)). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1676 | 1220648-2026-00549 | 24050064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050064&pc=OZD | 01/05/2026 |  |  | 01/12/2026 |  |  | 01/05/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683630 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella cp for mechanical circulatory support. it was reported that the patient has calcified iliacs visualized via angiogram (bilateral- left was more severe than right) prox lm 100% occluded. the medical doctor was aware of calcium but vessel was appropriate size. attempted insertion via right femoral artery - motor housing would not advance past prox iliac. intra-aortic balloon pump placed once impella removed. there was no patient harm. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1677 | 1220648-2026-00559 | 24050068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050068&pc=OZD | 01/05/2026 |  |  | 01/12/2026 |  |  | 01/06/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765165 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 69 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support the patient had an ischemic stroke. the patient was reported as stable. | Nothing was done to treat the stroke, they believe it is embolic to left middle cerebral artery (mca) and would like to get imaging to more thoroughly assess once the impella is explanted. the stroke is not resolved, the patient has expressive aphasia and was taken off the transplant list due to this. they are hoping for weaning the impella 5.5 and for him to go home on milrinone with no other advanced therapy. updated the impella was explanted on (b)(6) 2026 |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received: b5 and d6b updated. the investigation of the reported failure mode has been completed. no product was received for evaluation. stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
1678 | 1220648-2026-00568 | 24050070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050070&pc=OZD | 11/20/2025 |  |  | 01/12/2026 |  |  | 11/20/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026682031 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | Thrombosis/Thrombus | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1679 | 1220648-2026-00561 | 24050082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050082&pc=OZD | 12/24/2025 |  |  | 01/12/2026 |  |  | 12/29/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026701710 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 68 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Impella 5.5 was implanted in a 67-year-old male with a history of hfref (ef 10%), 15-year history of cardiomyopathy, hypertension, and lbbb. the patient was admitted with dyspnea and medically managed for adhf but progressed to worsening cardiogenic shock (pap 80s/30s, svo2 31, dilated lv, lactate 2.4). no device-related issues were reported. the plan was to extubate and proceed with lvad the following week. subsequently, the patient developed a large hemorrhagic brain bleed. neurologic prognosis remained poor without recovery, and the family elected to withdraw care. the reported events include a peri-procedural adverse event, stroke, and death. the impella 5.5 is intended for short-term ventricular support and requires careful monitoring of hemodynamics and neurologic status. the ifu notes that patient-specific factors such as severe underlying cardiac disease, hypotension, and critical illness can increase the risk of complications like stroke and bleeding. variations in device performance or optical sensor readings do not necessarily indicate malfunction and should be correlated with clinical parameters. based on available information, there is no evidence of a causal relationship between the impella device and the reported hemorrhagic stroke and death. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the stroke, in order to make a cause determination, all of the clinical details would have to be provided. the root cause was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
1680 | 1220648-2026-00558 | 24050084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050084&pc=OZD | 01/06/2026 |  |  | 01/12/2026 |  |  | 01/06/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026654138 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 76 KG |  | Black Or African American | Malfunction | No | Yes | Inaccurate Flow Rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 for mechanical circulatory support. purge flow was low and then a sudden purge system blocked alarm occurred. the care team started (after waiting a few hours for the pharmacy to mix) issue plasminogen activator in purge protocol. no resolution for multiple hours. the purge cassette was changed with no resolution. sudden flow saturation was noted with left ventricle placement signal well above aorta placement signal. the device was removed due to the failure mode. the introducer was flushed prior to the insertion. the patient's outcome is critical. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1681 | 1220648-2026-00566 | 24050085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050085&pc=OZD | 12/29/2025 |  |  | 01/12/2026 |  |  | 12/29/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026772379 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 70 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Impella 5.5 was implanted in a 64-year-old male with a history of coronary artery disease who was undergoing high-risk surgery. during the course of care, the patient developed a gastrointestinal bleed. there were no concerns regarding impella device performance. the patient was maintained on 5% dextrose in water with 25 meq/l sodium bicarbonate for purge solution and was not on heparin, which is acceptable per the instructions for use. despite clinical management, the patient expired. the reported events include a peri-procedural adverse event, serious injury, major bleeding, and death. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. based on available information, these complications are most consistent with the patient¿s comorbidity. per the ifu, impella 5.5 is intended for short-term ventricular support and allows for either heparinized purge solution or sodium bicarbonate-based purge solution to maintain purge flow and prevent thrombus formation. the ifu emphasizes careful monitoring of anticoagulation and hemostasis, as patient-specific factors such as severe cardiac disease, hypotension, and critical illness increase the risk of bleeding and multiorgan dysfunction. based on available information, there is no evidence of a causal relationship between the impella device and the reported gastrointestinal bleeding or mortality. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the bleed, the cause was not determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
1682 | 1220648-2026-00562 | 24050086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050086&pc=OZD | 12/15/2025 |  |  | 01/12/2026 |  |  | 12/15/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026781987 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1683 | 1220648-2026-00569 | 24050088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050088&pc=OZD | 01/03/2026 | 01/12/2026 |  | 01/12/2026 |  |  | 01/03/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026786616 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 104 KG | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Renal Failure | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 69 year old male on an impella cp for acute myocardial infarction. during support, the impella cp support started having suction alarms and as a result, it was suggested to remove the pump. however, the pump was repositioned but suction alarm remains. a swan was placed and the pump was removed. additionally, during support, it was reported that the patient experienced hematuria and renal failure. treatment details were not provided. the patient survived and remained on va -ecmo. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: low or blocked pump flow: the cause of the low or blocked pump flow issue was not determined without returned product investigation and sufficient clinical details, including data logs. hematuria: the cause of the injury was not determined due to insufficient clinical details. renal failure: the cause of the injury was not determined due to insufficient clinical details.
1684 | 1220648-2026-00565 | 24050089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050089&pc=OZD | 12/29/2025 |  |  | 01/12/2026 |  |  | 12/29/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026722160 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1685 | 1220648-2026-00574 | 24050090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050090&pc=OZD | 01/07/2026 | 01/12/2026 |  | 01/12/2026 |  |  | 01/07/2026 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026782035 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Anemia | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1686 | 1220648-2026-00572 | 24050091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050091&pc=OZD | 01/03/2026 | 01/12/2026 |  | 01/12/2026 |  |  | 01/03/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026796336 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 76 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | Additional information received that anticoagulation was stopped due to bleeding from the lungs. bedside nurse reported patient has history of metastatic lung cancer. the nurse also reported the patient has been getting inhaled tranexamic acid (txa), bleeding amount from endotracheal tube (ett) suctioning has improved team planning to bronch and trach patient today (b)(6) 2026. the patient remains on support. | The user facility reported a 73 year old male on impella 5.5 for cardiomyopathy. during support, there was concern for bleeding from surgical site prior to procedure and it was planned to place a jp drain post impella insertion. in the icu it was noted that the nurses required emptying the jp drain shortly after arriving to the icu. additionally, after placing the impella across the aortic valve and in the ventricle, the impella was attempted to start but was unable to because the wire had not been removed. once physician was made aware, the wire was immediately removed and the impella was started. the patient remains on support. |  |  |  |  | Additional information received b5 updated. correction: e4 | D6b: added explant date. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product/logs returned. temporary pump stop: no product/logs returned. the cause of the temporary pump stop was use-related as the wire was not removed prior to starting per clinical details. access site adverse event: the cause of the access site bleeding was not determined since insufficient clinical details were provided and product was not returned. major bleed: the cause of the injury was not determined due to insufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1687 | 1220648-2026-00570 | 24050093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050093&pc=OZD | 01/03/2026 | 01/12/2026 |  | 01/12/2026 |  |  | 01/03/2026 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754464 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1688 | 1220648-2026-00573 | 24050094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050094&pc=OZD | 01/06/2026 | 01/12/2026 |  | 01/12/2026 |  |  | 01/06/2026 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728753 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1689 | 1220648-2026-00571 | 24050095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050095&pc=OZD | 01/02/2026 | 01/12/2026 |  | 01/12/2026 |  |  | 01/03/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026780369 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported an impella cp being used on a patient for acute myocardial infarction. it was reported that during support, there was purge pressure high and purge flow low issues noticed. there were no purge system kinks or visible problems. the purge system was changed and the purge flow normalized but then the new alarm on aic "purge system failure - change aic". the purge system was exchanged and the patient was stabilized. after the exchange, no further placement-signal issues occurred on the new aic but the pump was working well with pulsatile motor current in range. troubleshooting was performed but the issue did not resolve and there was a recommendation to remove the pump. the pump remains on for support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional patient information received. a2 and a3a updated. h6 updated. | E4 was inadvertently omitted on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1690 | 1220648-2026-00554 | 24050136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050136&pc=PYX | 01/06/2026 | 01/12/2026 |  | 01/12/2026 |  |  | 01/06/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1765953 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported their automated impella controller (aic) was having a controller failure alarm. the aic was swapped to resolve the issue. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4 was inadvertently omitted on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1691 | 1220648-2026-00556 | 24050197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050197&pc=OZD | 01/06/2026 | 01/12/2026 |  | 01/12/2026 |  |  | 01/07/2026 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026773796 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Inaccurate Flow Rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1692 | 1220648-2026-00581 | 24050198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050198&pc=OZD | 09/17/2025 |  |  | 01/12/2026 |  |  | 09/18/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708852 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Anemia | Arrhythmia | Low Blood Pressure/ Hypotension | Respiratory Failure | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission
1693 | 1220648-2026-00579 | 24050199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050199&pc=OZD | 01/06/2026 | 01/12/2026 |  | 01/12/2026 |  |  | 01/06/2026 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026784429 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 82 YR |  | Male | 107 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella support with a cp pump. the complainant reported they were unable to get closure and hemostasis to the impella axillary site. a covered stent was placed to achieve hemostasis. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1694 | 1220648-2026-00580 | 24050200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050200&pc=OZD | 09/18/2025 |  |  | 01/12/2026 |  |  | 09/18/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720557 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient experiencing acute myocardial infarction and cardiogenic shock was implanted with an impella cp device to provide mechanical circulatory support. while under impella support, the patient experienced ischemia and hematuria. an angiogram of the right leg was performed during the placement of the pump, revealing an absence of pump flow. a femoral-femoral bypass was established, and the impella device was secured and sutured in place. furthermore, on the first day of support, the patient exhibited dark red urine. additional information indicated that a vascular surgeon was consulted for cold right foot post impella removal. vascular ultrasound of leg was completed and no clot or acute injury was found. the surgeon mentioned that there was horrible chronic arterial disease in lower leg. the patient subsequently required a right below-knee amputation (bka). the surgeon indicated that low cardiac output, underlying chronic peripheral vascular disease (pvd), and the presence of a femoral arterial sheath may have contributed to the need for bka. discussion with the implanting cardiologist revealed his opinion that chronic pvd was the primary contributing factor. the physician further noted that the femoral¿femoral bypass graft he placed demonstrated good flow distal to the antegrade sheath. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
1695 | 1220648-2026-00567 | 24050202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050202&pc=OZD | 12/29/2025 | 01/15/2026 |  | 01/12/2026 |  |  | 12/29/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755384 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1696 | 1220648-2026-00582 | 24050204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050204&pc=OZD | 09/11/2025 |  |  | 01/12/2026 |  |  | 09/19/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707320 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Backflow | Restricted Flow rate | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the system displayed a retrograde flow alarm along with high and erratic pump speed readings. additional information indicated that the caller observed low diastolic flows and a retrograde flow alarm accompanied by flattening of the motor current waveform, which self-resolved. no swan-ganz catheter or arterial line was in place. pump speed was being used for pressure monitoring due to difficulty obtaining cuff pressures. the patient had previously experienced low flows that improved with repositioning or mobilization. per report, the patient¿s condition remained unchanged. albumin was administered a few hours prior. review of waveforms showed consistent separation between aortic and left ventricular signals for several days. the caller was advised to perform echocardiography to confirm device position versus possible pressure signal noise. no suction events were noted despite flows of 2.9¿3.0 l/min at p9. the caller indicated that echocardiography could not be performed until morning and stated they would continue close monitoring. | [5 empty] | The pump was not returned for investigation. data log review was not required for this case run. the pump was utilized from (b)(6). data log shows that the placement signal and pump flow gradually became extremely fluctuating while running on a steady p-level. flow was noted to drop below 1 l/min while running on p6 and higher. there was a trend of snr dropping close to zero, with some spikes and loss of placement signal noted during periods when snr dropped to zero. other 5s parameters were not affected during the case run. the motor current was also closely monitored and showed some variation similar to the fluctuating trends in placement signal and flow. there were previous retrograde flow alarms noted while running on p5/p6 when flow dropped below 1 l/min. some motor speed and motor current ramp-up was observed during the time of the reported retrograde flow notification. low or blocked pump flow: the cause of the low pump flow issue was not determined without returned product investigation and sufficient clinical details. retrograde pump flow: the cause of the retrograde flow issue was not determined without returned product investigation and sufficient clinical details. optical sensor issue: the cause of the optical sensor issue was not determined without returned product investigation and sufficient clinical details.
1697 | 1220648-2026-00585 | 24050210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050210&pc=OZD | 09/18/2025 |  |  | 01/12/2026 |  |  | 09/18/2025 | 01/12/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026722247 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Anemia | Tachycardia | Thrombocytopenia | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
1698 | 1220648-2026-00584 | 24050211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24050211&pc=OZD | 09/19/2025 |  |  | 01/12/2026 |  |  | 09/19/2025 | 01/12/2026 |  |  |  |  | 01/12/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026653215 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient presented with elevated infection markers, likely secondary to pneumonia. sputum culture returned positive. crp is 298, temperature 38.3°c. no indication of impella-related complications. antibiotic therapy initiated. additional findings include thrombocytopenia and rectal bleeding. platelet count is low, activated partial thromboplastin time (aptt) at 47. platelet transfusion administered. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1699 | 1220648-2026-00426 | 24016817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016817&pc=OZD | 12/19/2025 |  |  | 01/09/2026 |  |  | 12/19/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATEDIMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1085518 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
1700 | 1220648-2026-00427 | 24016819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016819&pc=OZD | 12/19/2025 | 01/06/2026 |  | 01/09/2026 |  |  | 12/19/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026631253 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1701 | 1220648-2026-00429 | 24016824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016824&pc=OZD | 12/19/2025 |  |  | 01/09/2026 |  |  | 12/19/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026745684 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Thrombocytopenia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1702 | 1220648-2026-00430 | 24016828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016828&pc=OZD | 12/18/2025 |  |  | 01/09/2026 |  |  | 12/19/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026722127 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1703 | 1220648-2026-00431 | 24016830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016830&pc=OZD | 12/22/2025 |  |  | 01/09/2026 |  |  | 12/19/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026718798 |  | 0048-0024-JP |  | No |  |  | No |  |  | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1704 | 1220648-2026-00416 | 24016840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016840&pc=OZD | 12/22/2025 | 01/08/2026 |  | 01/09/2026 |  |  | 12/22/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026794989 |  | 000000000001000459 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male | 91 KG | Non Hispanic | Asian | Injury | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 50-year-old male with dilated cardiomyopathy and heart failure was supported with ecpella and underwent escalation to impella 5.5 under extracorporeal membrane oxygenation (ecmo) support. after removal of the impella cp, the first impella 5.5 pump was inserted via the right axillary graft without issue. approximately 10 minutes after initiation, a low purge pressure alarm occurred, with purge pressure around 200 mmhg and purge flow approximately 30 ml/h. troubleshooting was performed, including inspection for fluid leakage, reseating the purge cassette, and replacing the aic with 10% glucose solution; however, the condition did not resolve. the pump was subsequently replaced. following ecmo removal, blood pressure was supported pharmacologically, and a second impella 5.5 pump was implanted using a y-shaped graft extension and guidewire technique without issue. the second pump remains in use with the patient supported in the intensive care unit. the facility has requested pump analysis and replacement of the impella 5.5 pump and purge cassette. | [5 empty] | B3 revised date as it was entered incorrectly on the initial report that was submitted. e4 added selection as it was omitted from the initial report that was submitted. h6 added health effect - impact codes as they were omitted from the initial report that was submitted. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1705 | 1220648-2026-00428 | 24016842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016842&pc=OZD | 12/19/2025 |  |  | 01/09/2026 |  |  | 12/19/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773648 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 79 KG | Non Hispanic | White | Malfunction | Yes | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A second attempt was made to place an impella 5.5 device in the patient utilizing the left subclavian artery. access to the subclavian artery was obtained without incident, and the graft was sewn in the usual fashion. left ventricular access was achieved and confirmed using transesophageal echocardiography (tee). the physician applied propofol to lubricate the impella 5.5 prior to insertion. multiple attempts were made to advance the device without success. considering the prior unsuccessful attempt on the right side and repeated difficulty advancing the pump, the decision was made to abort the insertion to prevent potential vessel trauma to the patient. | [5 empty] | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. failure to advance: the cause of the delivery issue was determined to be patient condition as the patient had sub optimal vessel size preventing successful delivery of impella. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1706 | 1220648-2026-00418 | 24016844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016844&pc=OZD | 12/11/2025 | 01/08/2026 |  | 01/09/2026 |  |  | 12/11/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663219 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1707 | 1220648-2026-00432 | 24019458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24019458&pc=OZD | 01/03/2026 |  |  | 01/09/2026 |  |  | 01/04/2026 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026754445 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 71 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | Additional information was received stating that the nurse observed a change in aortic (ao) waveforms. an echocardiogram was performed and noted device positioning toward the mitral valve. the physician repositioned the device, and improvement in the aortic waveforms was subsequently reported. | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient was bleeding from all access sites including the impella insertion site requiring blood products. the patient had an mitral valve replacement today. imaging was used to check the pump position and the pump was in good position. two days later during rounds, the staff described the patient needing to go back to the operating room for surgical repair of the right axillary artery for bleeding. the repair was successful. imaging was used to check the pump's position and the pump was in good position. hematuria was also noted. urinary output was greater than 30 cubic centimeters and concentrated/amber. this is one of two related reports. this report represents the impella 5.5 and another report will be submitted to represent the impella rp flex. |  |  |  |  | D6b explant date updated. b1 product problem was inadvertently omitted on the initial manufacturer device report 1220648-2026-00432-1. | Additional information received and h6 codes were updated. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1708 | 1220648-2026-00434 | 24019461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24019461&pc=OZD | 01/03/2026 |  |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026747162 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1709 | 1220648-2026-00435 | 24019463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24019463&pc=OZD | 01/05/2026 |  |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026799053 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. the medical doctor reported a red alarm in the automated impella controller "pp high"; pp>1000ml" it was recommended to exchange the pump and for plasminogen activator - lyse therapy but these suggestions were rejected. possible risk for pump stop was pointed out. three days later, the perfusionist reported that therapy was withdrawn not due to impella related reasons. the patient's outcome is unknown. the pump was disposed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | H6. updated. the investigation is ongoing.
1710 | 1220648-2026-00438 | 24019476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24019476&pc=OZD | 01/03/2026 |  |  | 01/09/2026 |  |  | 01/03/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026793771 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1711 | 1220648-2026-00440 | 24019504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24019504&pc=OZD | 12/26/2025 | 02/17/2026 |  | 01/09/2026 |  |  | 12/26/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026791221 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1712 | 1220648-2026-00441 | 24019519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24019519&pc=OZD | 12/10/2025 |  |  | 01/09/2026 |  |  | 12/10/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026698915 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | Company representation | Distributor |  |  |  | 47 YR |  | Male | 82 KG | Non Hispanic | Asian | Death | Yes | No | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp lost distal pulses. the patient expired on impella support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1713 | 1220648-2026-00442 | 24019522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24019522&pc=OZD | 12/16/2025 |  |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759699 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced a suction alarm and low ao and lv position waveforms. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D4 udi information was inadvertently omitted on the initial manufacturer device report. | B3 date of event was entered on the initial manufacturer device report.
1714 | 1220648-2026-00443 | 24019635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24019635&pc=PYX | 01/05/2026 |  |  | 01/09/2026 |  |  | 01/04/2026 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502288 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Network Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon start up of an automated impella controller the device could not be switched to maintenance mode and could not connect to impella connect. | [5 empty] | No patient was involved in the reported event. section a was left blank. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data analysis: neither console logs nor journals were available for analysis. device analysis: device was tested and the reported issue had been confirmed. root cause: reported malfunction of the rlm on manifested by inability to enter maintenance mode and failure to connect to network, was caused by corruption of microsd card in the rlm. the exact cause of card corruption could not be determined. phr summary: there were no issues related to the complaint discovered when reviewing the product history of console.
1715 | 1220648-2026-00444 | 24021076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24021076&pc=OZD | 12/01/2025 |  |  | 01/09/2026 |  |  | 12/18/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026747134 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1716 | 1220648-2026-00445 | 24021228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24021228&pc=OZD | 12/29/2025 |  |  | 01/09/2026 |  |  | 12/29/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1192789 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1717 | 1220648-2026-00448 | 24021276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24021276&pc=OZD | 01/04/2026 |  |  | 01/09/2026 |  |  | 01/04/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732682 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1718 | 1220648-2026-00450 | 24021283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24021283&pc=OZD | 01/04/2026 |  |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026779284 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1719 | 1220648-2026-00449 | 24021285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24021285&pc=OZD | 01/04/2026 |  |  | 01/09/2026 |  |  | 01/04/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026779572 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1720 | 1220648-2026-00451 | 24021299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24021299&pc=OZD | 01/05/2026 |  |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768544 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1721 | 1220648-2026-00452 | 24021307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24021307&pc=OZD | 01/05/2026 |  |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682771 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Restricted Flow rate | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1722 | 1220648-2026-00453 | 24021308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24021308&pc=OZD | 01/06/2026 | 01/09/2026 |  | 01/09/2026 |  |  | 01/06/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026780143 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Stroke/CVA | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1723 | 1220648-2026-00455 | 24023463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24023463&pc=OZD | 12/03/2025 |  |  | 01/09/2026 |  |  | 12/10/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725128 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 84 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support the device exhibited an optical signal issue. no further information regarding the resolution of this issue was reported. the device was normally weaned and explanted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary no product returned. no logs available. optical signal issue: the cause of the optical signal issue is not determined due to no product or data logs returned. device history review device passed all post sterile inspection checks.
1724 | 1220648-2026-00456 | 24023548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24023548&pc=OZD | 01/04/2026 |  |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760023 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1725 | 1220648-2026-00031 | 24023642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24023642&pc=OZD | 08/29/2025 | 01/04/2026 |  | 01/09/2026 |  |  | 12/10/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1726 | 1220648-2026-00457 | 24023691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24023691&pc=PYX | 12/26/2025 |  |  | 01/09/2026 |  |  | 12/26/2025 | 01/09/2026 |  |  | 02/28/2026 |  |  |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478685 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1727 | 1220648-2026-00458 | 24023747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24023747&pc=OZD | 01/06/2026 | 01/09/2026 |  | 01/09/2026 |  |  | 01/06/2026 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1450701 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 76 KG | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 67-year-old male patient experienced a controller error alarm that self-resolved without additional alarms. the alarm occurred again, and the bedside nurse planned to discuss controller replacement with the managing physician. no further information was provided. | [5 empty] | Correction: d6a and d6b should have been left blank on the initial manufacturer device report. e4 was inadvertently omitted on the initial manufacturer device report. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1728 | 1220648-2026-00459 | 24023792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24023792&pc=PYX | 01/05/2026 |  |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529711 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 45 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced an unexpected pump stop while a catheter was being placed in the femoral vein. the pump could not be restarted so it was explanted and replaced with a new pump. the patient was in stable condition. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary pump stop: clinical reported team was placing a pa catheter through femoral vein then impella stopped. unable to restart. no product or logs were returned. the cause of the issue was not established as no product or logs were returned for analysis. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1729 | 1220648-2026-00460 | 24023816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24023816&pc=OZD | 01/05/2026 |  |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664659 |  | 1000080_ABM |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1730 | 1220648-2026-00461 | 24023827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24023827&pc=OZD | 01/05/2026 |  |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026771160 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1731 | 1220648-2026-00462 | 24023836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24023836&pc=OZD | 01/05/2026 | 01/14/2026 |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026781030 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Low Blood Pressure/ Hypotension | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1732 | 1220648-2026-00463 | 24023945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24023945&pc=OZD | 12/28/2025 |  |  | 01/09/2026 |  |  | 12/29/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1120454 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 109 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that critical care transport went to pick up a patient from an outside facility and when the transport team tried to switch over to a different automated impella controller, it gave them a "purge cassette not detected error". the procedure outcome was successful with another device. this is one of multiple related reports. this report addresses the automated impella controller used with an impella cp. other reports will be submitted to address reportable events associated with related devices. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1733 | 1220648-2026-00466 | 24024127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24024127&pc=OZD | 01/05/2026 |  |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026757630 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1734 | 1220648-2026-00467 | 24024212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24024212&pc=OZD | 09/12/2025 | 11/19/2025 |  | 01/09/2026 |  |  | 09/12/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659452 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 68 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support an apical thrombus was noted. in addition, there was a ¿placement signal not reliable¿ alarm. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1735 | 1220648-2026-00468 | 24024222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24024222&pc=OZD | 11/21/2025 |  |  | 01/09/2026 |  |  | 11/20/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026776771 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 74 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced hematoma formed at groin site, and noted hemolysis in urine. impella was repositioned via echocardiogram. it was also reported that the patient was cannulated with extracorporeal membrane oxygenation (emco). the patient was reported to have survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1736 | 1220648-2026-00469 | 24024247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24024247&pc=OZD | 11/21/2025 |  |  | 01/09/2026 |  |  | 11/20/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026770632 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that there were multiple failed attempts to cross into the left ventricle (lv). it was also reported that the device was unable to deliver or advance: the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy and resistance at innominate artery preventing successful delivery of impella. the patient was reported to have survived the procedure. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1737 | 1220648-2026-00470 | 24027149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027149&pc=OZD | 09/11/2025 |  |  | 01/09/2026 |  |  | 09/11/2025 | 01/09/2026 |  |  | 02/28/2026 |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025469460 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the patient developed ongoing ectopy that was unable to be controlled. a decision was made to remove the impella in the operating room and provide chemical support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1738 | 1220648-2026-00471 | 24027172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027172&pc=OZD | 09/11/2025 |  |  | 01/09/2026 |  |  | 09/11/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025544844 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Renal Failure | Sepsis | Ventricular Fibrillation | Thrombocytopenia | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that on (b)(6) 2025 , the patient experienced a run of ventricular tachycardia (vt) in the morning, requiring cardioversion. the following day, non-sustained ventricular tachycardia episodes were noted and were self-resolving. a change in mental status was observed the prior day; neurology was consulted. intracranial clots were identified, and a head ct was planned. the patient was reported to be adequately anticoagulated for the indicated duration. on (b)(6) 2025, the patient continues to exhibit renal and hepatic dysfunction requiring continuous venovenous hemofiltration (cvvh). possible sepsis noted. ejection fraction approximately 20%. possible heparin-induced thrombocytopenia (hit) under consideration. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1739 | 1220648-2026-00472 | 24027208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027208&pc=OZD | 09/12/2025 |  |  | 01/09/2026 |  |  | 09/12/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026641097 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1740 | 1220648-2026-00473 | 24027211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027211&pc=OZD | 09/15/2025 |  |  | 01/09/2026 |  |  | 09/15/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026674873 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Pump | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was changed to systemic argatroban due to concerns for heparin induced thrombocytopenia (hit). attempted to connect the pump to the automated impella console (aic) at approximately 1853, but an issue occurred because another aic was turned on at 1901, and the pump connected to the other aic instead. the second aic showed a number of errors, including two failures to connect. in total, 3 aics were turned on, but only aic (b)(6) and aic (b)(6) were used to connect the pump. aic (b)(6) was turned on to transfer the pump from another console. the pump was connected but failed to turn on. aic (b)(6) was on standby in the unit and was emergently turned on, and the pump was connected and failed again. aic (b)(6) was turned off and then turned back on for a second attempt. the pump successfully connected and was started. additional information stated that the onsite representative contacted the customer support center, who advised the aic issue was caused because the "start new case" button had not been pressed, causing the error sequence. there was no reported harm to the patient and care was transferred successfully to aic (b)(6). the patient experienced symptoms of thrombocytopenia and patient outcome required medication. | [5 empty] | There are four associated devices for the reported event. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1741 | 1220648-2026-00474 | 24027258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027258&pc=OZD | 09/14/2025 |  |  | 01/09/2026 |  |  | 09/14/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701060 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1742 | 1220648-2026-00475 | 24027273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027273&pc=PYX | 01/14/2026 |  |  | 01/09/2026 |  |  | 01/04/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1717178 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup
1743 | 1220648-2026-00476 | 24027274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027274&pc=OZD | 01/04/2026 |  |  | 01/09/2026 |  |  | 01/04/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026751835 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1744 | 1220648-2026-00477 | 24027283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027283&pc=OZD | 12/15/2025 |  |  | 01/09/2026 |  |  | 12/15/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1134175 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | No | Yes | Appropriate Term/Code Not Available | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | R |  |  |  | Initial submission | Followup
1745 | 1220648-2026-00479 | 24027284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027284&pc=OZD | 01/06/2026 | 01/12/2026 |  | 01/09/2026 |  |  | 01/06/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739120 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Fluid/Blood Leak | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1746 | 1220648-2026-00478 | 24027295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027295&pc=OZD | 01/06/2026 | 01/12/2026 |  | 01/09/2026 |  |  | 01/06/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026784263 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1747 | 1220648-2026-00480 | 24027302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027302&pc=OZD | 01/06/2026 | 01/09/2026 |  | 01/09/2026 |  |  | 01/06/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722136 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1748 | 1220648-2026-00481 | 24027324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027324&pc=OZD | 01/06/2026 | 01/09/2026 |  | 01/09/2026 |  |  | 01/06/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026752188 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation of Vessels | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1749 | 1220648-2026-00482 | 24027331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027331&pc=OZD | 01/05/2026 | 01/09/2026 |  | 01/09/2026 |  |  | 01/06/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026802742 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ischemia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1750 | 1220648-2026-00483 | 24027334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24027334&pc=OZD | 01/01/2026 | 01/14/2026 |  | 01/09/2026 |  |  | 01/05/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026795747 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | False Alarm | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup
1751 | 1220648-2026-00484 | 24028971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24028971&pc=OZD | 09/11/2025 |  |  | 01/09/2026 |  |  | 09/11/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713049 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1752 | 1220648-2026-00485 | 24028972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24028972&pc=OZD | 09/11/2025 |  |  | 01/09/2026 |  |  | 09/11/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026699721 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that there was an impella in aorta alarm overnight after the patient was turned. evaluated the pump position via ultrasound and repositioning will be attempted. in the lab, they were unable to readvance the pump back into the left ventricle, despite fluoroscopy guidance. additionally, the cannula was "kinking". the decision was made to place a new impella pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1753 | 1220648-2026-00486 | 24028976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24028976&pc=OZD | 09/11/2025 |  |  | 01/09/2026 |  |  | 09/11/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026645604 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Tachycardia | Asystole | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced an episode of ventricular tachycardia (vt). the patient coded overnight with multiple shocks given. return of spontaneous circulation (rosc) was achieved; however, the patient remained hypotensive. a do-not-resuscitate (dnr) order was implemented. the patient subsequently developed asystole later that day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1754 | 1220648-2026-00487 | 24028978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24028978&pc=OZD | 09/11/2025 |  |  | 01/09/2026 |  |  | 09/11/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739566 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. after the procedure, impella cp was weaned due to leg ischemia. upon explant, left ventricle function improved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1755 | 1220648-2026-00488 | 24028979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24028979&pc=OZD | 09/15/2025 |  |  | 01/09/2026 |  |  | 09/16/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1381389 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Use of Device Problem | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was changed to systemic argatroban due to concerns for heparin induced thrombocytopenia (hit). attempted to connect the pump to the automated impella console (aic) at approximately 1853, but an issue occurred because another aic was turned on at 1901, and the pump connected to the other aic instead. the second aic showed a number of errors, including two failures to connect. in total, 3 aics were turned on, but only aic (sn: (b)(6)) and aic (sn:(b)(6)) were used to connect the pump. aic (sn:(b)(6)) was turned on to transfer the pump from another console. the pump was connected but failed to turn on. aic (sn: (b)(6)) was on standby in the unit and was emergently turned on, and the pump was connected and failed again. aic (sn:(b)(6)) was turned off and then turned back on for a second attempt. the pump successfully connected and was started. additional information stated that the onsite representative contacted the customer support center, who advised the aic issue was caused because the "start new case" button had not been pressed, causing the error sequence. there was no reported harm to the patient and care was transferred successfully to aic. | [5 empty] | There are four associated devices for the reported event. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1756 | 1220648-2026-00490 | 24028980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24028980&pc=OZD | 09/15/2025 |  |  | 01/09/2026 |  |  | 09/15/2025 | 01/09/2026 |  |  |  |  | 01/09/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1082388 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Failure to Read Input Signal | Use of Device Problem | Application Program Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was changed to systemic argatroban due to concerns for heparin induced thrombocytopenia (hit). attempted to connect the pump to the automated impella console (aic) at approximately 1853, but an issue occurred because another aic was turned on at 1901, and the pump connected to the other aic instead. the second aic showed a number of errors, including two failures to connect. in total, 3 aics were turned on, but only aic (sn: (b)(6) and aic (sn: (b)(6) were used to connect the pump. aic (sn: (b)(6) was turned on to transfer the pump from another console. the pump was connected but failed to turn on. aic (sn: (b)(6) was on standby in the unit and was emergently turned on, and the pump was connected and failed again. aic (sn: (b)(6) was turned off and then turned back on for a second attempt. the pump successfully connected and was started. additional information stated that the onsite representative contacted the customer support center, who advised the aic issue was caused because the "start new case" button had not been pressed, causing the error sequence. there was no reported harm to the patient and care was transferred successfully to aic (sn: (b)(6). | [5 empty] | There are four associated devices for the reported event. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1757 | 1220648-2026-00489 | 24028981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24028981&pc=PYX | 09/15/2025 |  |  | 01/09/2026 |  |  | 09/16/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1668177 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | Yes | Yes | Use of Device Problem | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1758 | 1220648-2026-00492 | 24028989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24028989&pc=OZD | 11/20/2025 |  |  | 01/09/2026 |  |  | 11/20/2025 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026659277 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1759 | 1220648-2026-00494 | 24029013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24029013&pc=OZD | 01/04/2026 |  |  | 01/09/2026 |  |  | 01/04/2026 | 01/09/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645961 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1760 | 1220648-2026-00305 | 24000241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000241&pc=OZD | 01/02/2026 | 01/07/2026 |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | UNK_IMPELLA CP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Difficult or Delayed Positioning | Use of Device Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1761 | 1220648-2026-00304 | 24000243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000243&pc=OZD | 01/02/2026 | 01/07/2026 |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768784 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 40-year-old female whose impella cp device was removed. a subtotal occlusion of the femoral artery was noted. the patient had adequate distal flow and was supported with an intra-aortic balloon pump. no additional information was provided. | [5 empty] | Correction: e4 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1762 | 1220648-2026-00307 | 24000247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000247&pc=OZD | 01/02/2026 | 01/07/2026 |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671752 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1763 | 1220648-2026-00309 | 24000250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000250&pc=OZD | 01/02/2026 | 01/07/2026 |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026793970 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Flow Rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 65-year-old male patient whose pump saturation was noted 30 minutes after implant and initial ramp-up. no additional information was provided. | [5 empty] | Correction: e4. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. high or saturated pump flow: the cause of saturated pump flow was likely due to biomaterial observed beneath the impeller/purge gap based on product analysis.
1764 | 1220648-2026-00225 | 24000251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000251&pc=OZD | 01/02/2026 |  |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026762282 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1765 | 1220648-2026-00310 | 24000266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000266&pc=OZD | 12/24/2025 | 01/07/2026 |  | 01/08/2026 |  |  | 12/25/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1132457 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1766 | 1220648-2026-00104 | 24000274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000274&pc=OZD | 10/28/2025 |  |  | 01/08/2026 |  |  | 10/28/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744981 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1767 | 1220648-2026-00103 | 24000275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000275&pc=OZD | 10/28/2025 |  |  | 01/08/2026 |  |  | 10/28/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765706 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female | 65 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support during a percutaneous coronary intervention (pci). the impella cp was implanted without complications. after a successful pci, the impella was weaned to performance (p) level 2, however the medical team was concerned for possible decompensation from myocardial stunning, and the decision was made to leave the impella in overnight to unload the heart overnight. after arriving to the intensive care unit, it was noted that the impella access site was bleeding. manual pressure was held, and the physician attempted to tighten the perclose sutures. it was noted that the bleeding slowed after tightening the sutures. the patient was transfused with one unit of red blood cells as a result of the blood loss. an echocardiogram was performed and revealed that the impella location was deeper than expected. the physician declined to reposition the impella as there were no other issues. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | Investigation summary. no product was returned. device in wrong position (positioning issue): echocardiogram done. impella position slightly deep. the cause of the positioning issue was unable to be determined due to insufficient clinical details. access site adverse event: bleeding was after transfer to intensive care unit and act was still high. unsure if bleeding was movement to the intensive care unit or activated clotting time issue, or if the act was just still high because there was not enough time for it to come down. the cause of the bleed was unable to be determined due to insufficient clinical details. device history review. device (b)(6) passed all post sterile inspection checks.
1768 | 1220648-2026-00251 | 24000293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000293&pc=OZD | 12/30/2025 |  |  | 01/08/2026 |  |  | 12/30/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025596362 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1769 | 1220648-2026-00263 | 24000296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000296&pc=OZD | 12/26/2025 |  |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026795444 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male | 82 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that when priming, placement signal drifted up to 9/9. tried to unplug and plug back in but wouldn¿t fix. the pump was explanted and replaced. there was no patient harm. | [5 empty] | Product complaint # (b)(4). investigation summary impella 5.5 pump sn643813 returned. placement signal issue: . data logs were not available for investigation. the pump was returned. when connected to a console, all 5s parameters were within the expected range. calibrated and factory g0 were 24 apart, suggesting no abnormalities with the sensor calibration. the pump also passed laser continuity testing. finally, the pump was connected to the console again, and the ps was reading ~5 mmhg in air. it was then run in a bucket of water at p-8 for several hours and the ps was reading ~10 mmhg, which is not abnormal. since limited clinical details were provided and data logs were not available for review, the cause of the placement signal issue could not be determined. device history lot device lot: 2003696 device history batch subcomponent: na device history review pump sn (b)(6) passed all post-sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1770 | 1220648-2026-00286 | 24000297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000297&pc=OZD | 11/01/2025 | 01/06/2026 |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026741795 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | False Alarm | [6 empty] | Atrial Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1771 | 1220648-2026-00098 | 24000305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000305&pc=OZD | 12/24/2025 |  |  | 01/08/2026 |  |  | 12/24/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026782029 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 82 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 could not be completed due to patient anatomy. the pump could not pass the heavily calcified area of innominate artery. the patient was implanted with an impella cp instead. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1772 | 1220648-2026-00311 | 24000325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000325&pc=OZD | 12/18/2025 |  |  | 01/08/2026 |  |  | 12/19/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748484 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | Distributor | [9 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1773 | 1220648-2026-00312 | 24000332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000332&pc=OZD | 12/17/2025 |  |  | 01/08/2026 |  |  | 12/19/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026771271 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 71 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on ecpella support (veno-arterial extracorporeal membrane oxygenation and impella 5.5). the patient had a hemorrhagic stroke and subsequently care was withdrawn. the patient expired. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the stroke was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
1774 | 1220648-2026-00313 | 24000346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000346&pc=OZD | 12/19/2025 |  |  | 01/08/2026 |  |  | 12/20/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026777463 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Foreign | Health Professional | [10 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [41 empty] | Initial submission | Followup
1775 | 1220648-2026-00316 | 24002714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002714&pc=OZD | 12/26/2025 |  |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687213 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male | 70 KG | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced an unexpected pump stop and alarms of high motor current. attempts to restart the pump were unsuccessful. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1776 | 1220648-2026-00317 | 24002717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002717&pc=OZD | 01/02/2026 |  |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026790880 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemolysis | Hematuria | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1777 | 1220648-2026-00318 | 24002730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002730&pc=OZD | 01/08/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026789362 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup
1778 | 1220648-2026-00314 | 24002733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002733&pc=OZD | 12/28/2025 | 01/08/2026 |  | 01/08/2026 |  |  | 12/29/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026794857 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1779 | 1220648-2026-00320 | 24002741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002741&pc=OZD | 01/04/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026776900 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1780 | 1220648-2026-00321 | 24002751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002751&pc=OZD | 01/03/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026733348 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1781 | 1220648-2026-00322 | 24002779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002779&pc=OZD | 01/03/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026770343 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1782 | 1220648-2026-00323 | 24002793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002793&pc=OZD | 01/04/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1452168 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Unintended Collision | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1783 | 1220648-2026-00324 | 24002847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002847&pc=OZD | 01/03/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026712859 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1784 | 1220648-2026-00326 | 24002861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002861&pc=OZD | 12/30/2025 |  |  | 01/08/2026 |  |  | 12/30/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1119117 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 109 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. the automated impella controller (aic) had a controller failure alarm. a new aic was used and support continued. the patient's outcome is stable. this report represents the automated impella controller used with an impella 5.5. separate report(s) will be submitted for associated devices with reportable events. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D6a and d6b: added implant and explant date.
1785 | 1220648-2026-00328 | 24002880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002880&pc=OZD | 12/30/2025 |  |  | 01/08/2026 |  |  | 12/30/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026789401 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Ischemia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1786 | 1220648-2026-00329 | 24002883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002883&pc=OZD | 01/02/2026 |  |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664659 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Ischemia | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup
1787 | 1220648-2026-00319 | 24002893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002893&pc=OZD | 01/04/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1454040 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Overheating of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1788 | 1220648-2026-00330 | 24002906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24002906&pc=OZD | 01/03/2026 |  |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026796264 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | False Alarm | [6 empty] | Atrial Fibrillation | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1789 | 1220648-2026-00332 | 24003359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003359&pc=OZD | 01/02/2026 |  |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026752212 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1790 | 1220648-2026-00331 | 24003363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003363&pc=OZD | 01/02/2026 |  |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026782150 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1791 | 1220648-2026-00333 | 24003365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003365&pc=OZD | 01/02/2026 |  |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026793619 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1792 | 1220648-2026-00335 | 24003368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003368&pc=OZD | 01/03/2026 | 01/13/2026 |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026774600 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1793 | 1220648-2026-00334 | 24003370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003370&pc=OZD | 01/03/2026 | 02/02/2026 |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026760978 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1794 | 1220648-2026-00336 | 24003377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003377&pc=PYX | 01/02/2026 |  |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529006 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [41 empty] | Initial submission | Followup
1795 | 1220648-2026-00337 | 24003378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003378&pc=OZD | 01/02/2026 |  |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026751771 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Positioning Failure | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1796 | 1220648-2026-00338 | 24003901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003901&pc=OZD | 01/02/2026 | 01/13/2026 |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608736 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1797 | 1220648-2026-00342 | 24003902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003902&pc=OZD | 01/02/2026 |  |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026767837 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1798 | 1220648-2026-00341 | 24003904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003904&pc=OZD | 01/04/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639685 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1799 | 1220648-2026-00344 | 24003905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003905&pc=OZD | 01/05/2026 |  |  | 01/08/2026 |  |  | 01/05/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026774620 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1800 | 1220648-2026-00343 | 24003907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003907&pc=OZD | 01/04/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755831 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 93 KG | Non Hispanic | Asian | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. hematuria noted during intensive care unit check in. doctor elected to leave device in place overnight and assess for removal of the device. the device was later explanted the same day. patient was in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1801 | 1220648-2026-00340 | 24003909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003909&pc=OZD | 01/02/2026 |  |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704415 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 98 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a loss of placement signal. the patient's procedure was completed successfully and the patient was weaned from the pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. since neither product nor logs were available for investigation, the cause of the placement signal issue could not be determined. device history review: the device passed all post sterile inspection checks.
1802 | 1220648-2026-00345 | 24003940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24003940&pc=OZD | 12/24/2025 |  |  | 01/08/2026 |  |  | 12/24/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025471700 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient had a large hematoma extending into the abdomen with visible bruising. external dressing was intact, indicating no external bleeding. hemoglobin levels had dropped, suggesting internal bleeding or ongoing blood loss. a decision was being made to withdraw care, so there was no intent to treat the dropping hemoglobin. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1803 | 1220648-2026-00347 | 24004033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24004033&pc=OZD | 12/24/2025 | 02/16/2026 |  | 01/08/2026 |  |  | 12/24/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026787389 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 93 KG | Non Hispanic | Black Or African American | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient¿s urine appeared black. all laboratory hemolyzed. the echo performed showing left ventricle small and toward the mitral valve. pump repositioned under echo. subsequently, it was reported that the patient expired. | [5 empty] | Investigation summary: the pump was not returned for investigation. pump run time was around 50 hours. data log shows a few short instances of pump position in aorta/ventricle during the case run, without affecting motor current. apart from that, there was drop/fluctuation in motor current and pump flow noted while the pump was running on a steady p-level. placement signal was trending downward mid-case and showed a fluctuating trend while running on a steady p-level. no issue was noted on motor current spike or optical signal issue correlated with the placement signal trend. mechanical interaction with blood (hemolysis): clinical details indicate that the patient developed black urine and all labs were hemolyzed after pump implantation via left femoral artery. echo showed the left ventricle was small and toward the mitral valve; the pump was repositioned under echo guidance, and imaging confirmed good pump position. the cause of the issue was not determined since the data log was inconclusive without sufficient clinical details and returned product. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1995162. device history batch: subcomponent lot: n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1804 | 1220648-2026-00348 | 24004036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24004036&pc=OZD | 12/24/2025 |  |  | 01/08/2026 |  |  | 12/24/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026778397 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1805 | 1220648-2026-00346 | 24004055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24004055&pc=OZD | 12/24/2025 |  |  | 01/08/2026 |  |  | 12/24/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026695330 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 94 KG | Non Hispanic | White | Death | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that the device generated a purge system blocked alarm. purge flow dropped to < 1 ml/hr. purge pressure elevated to 1141 mmhg. tpa (tissue plasminogen activator) is running in the purge line to address suspected thrombus or blockage. the patient was reported to have survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. purge pressure high: clinical details noted that "purge system blocked" alarm was triggered. tpa added to purge but no clinical details were provided as to if alarms and increased purge pressure resolved. the cause of the high purge pressure could not be determined as no product or logs were returned and limited clinical details were provided. | B2: added death and death date. d6b: added explant date. h1: changed reportability to death. h6: added code f02. additional information: care was withdrawn and the pump was left in along with other lines, the customer is unable to send anything back for investigation. clinical rationale: an impella 5.5 device was inserted via the right axillary artery in a 55-year-old female patient presenting with acute myocardial infarction complicated by cardiogenic shock (ami/cgs), scai shock stage e, with a history of known coronary artery disease. during support, a purge system blocked alarm was noted, with purge flow decreasing to less than 1 ml/hr and purge pressure elevated to 1141 mmhg. in response to these findings, tissue plasminogen activator (tpa) was administered in the purge, consistent with standard troubleshooting for suspected purge obstruction. at the time of reporting, tpa infusion in the purge was ongoing. subsequently, care was withdrawn, and the patient expired. the reported events are a purge system issue requiring medication (tpa) and death. high purge pressures and reduced purge flow may be evaluated as device-related; however, these findings can also occur in the setting of patient-specific and clinical factors, including critical illness, prothrombotic state, and severe cardiogenic shock. the patient¿s death is conservatively being reported on the impella 5.5, as the device was in use at the time of death. however, based on the available information, the death is unlikely to be related to device function and is most plausibly attributed to the patient¿s underlying critical condition, including ami complicated by scai stage e shock.
1806 | 1220648-2026-00349 | 24004148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24004148&pc=OZD | 01/02/2026 | 01/14/2026 |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026813785 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | Material Split, Cut or Torn |  |  |  |  | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1807 | 1220648-2026-00351 | 24010829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24010829&pc=OZD | 12/19/2025 |  |  | 01/08/2026 |  |  | 12/19/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026778855 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | False Alarm | Malposition of Device | [5 empty] | Cardiac Arrest | Ischemia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1808 | 1220648-2026-00352 | 24011059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24011059&pc=OZD | 12/19/2025 |  |  | 01/08/2026 |  |  | 12/19/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1159475 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Computer Operating System Problem | Device Fell | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1809 | 1220648-2026-00353 | 24011197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24011197&pc=OZD | 12/19/2025 |  |  | 01/08/2026 |  |  | 12/19/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026778407 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that upon removal of impella cp from right femoral artery (rfa), a thrombus was identified in rfa, and thrombus evacuated by accessing left femoral artery (lfa) and removing the clot. it was noted that for perclose attempted twice in rfa but failed. good distal flow was observed after clot evacuation. the patient was reported to have survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: thrombosis: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review the pump passed all the post sterile inspection checks.
1810 | 1220648-2026-00354 | 24011665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24011665&pc=OZD | 12/19/2025 |  |  | 01/08/2026 |  |  | 12/19/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1238447 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 100 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical, circulatory support. during support, it was reported that controller error¿ warning presented on the, automated impella controller (aic). screen recommendation was to switch to new aic, which staff was able to do. the patient was reported to have survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1811 | 1220648-2026-00355 | 24016576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016576&pc=OZD | 12/09/2025 |  |  | 01/08/2026 |  |  | 12/09/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026721184 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 66 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient¿s axillary artery was exposed and observed to be undersized. the physician still wished to proceed with an attempted insertion of the 5.5 device. the 5.5 device was inserted, although it was unable to make the turn down the arch. the physician then decided to remove the device due to the patient¿s anatomy and place a centrifugal pump instead. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product was returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had small axillary artery and aortic arch preventing successful delivery of impella. device history review: device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1812 | 1220648-2026-00358 | 24016577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016577&pc=OZD | 12/09/2025 | 02/24/2026 |  | 01/08/2026 |  |  | 12/09/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026712365 |  | 1000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1813 | 1220648-2026-00357 | 24016578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016578&pc=OZD | 12/08/2025 |  |  | 01/08/2026 |  |  | 12/09/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1122151 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1814 | 1220648-2026-00356 | 24016579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016579&pc=OZD | 12/09/2025 |  |  | 01/08/2026 |  |  | 12/09/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026774799 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 54 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after the impella was placed, distal limb pulses could not be obtained. the physician was informed and elected not to intervene despite recommendations from the complainant and the critical care physician. the pump is not available for return as it was explanted and disposed of. the family elected to withdraw care for the patient, and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1815 | 1220648-2026-00359 | 24016581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016581&pc=OZD | 12/09/2025 |  |  | 01/08/2026 |  |  | 12/09/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026741019 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Cardiac Perforation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1816 | 1220648-2026-00360 | 24016587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016587&pc=PYX | 11/21/2025 |  |  | 01/08/2026 |  |  | 11/21/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026726711 |  | 000000000001000323 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1817 | 1220648-2026-00363 | 24016596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016596&pc=OZD | 12/26/2025 | 02/05/2026 |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026763868 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1818 | 1220648-2026-00364 | 24016597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016597&pc=OZD | 12/26/2025 |  |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779622 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that decreased blood flow was noted down impella leg. a distal perfusion sheath was placed to provide flow. the patient outcome is still to be determined as the patient remains on support. additional information indicates that a hematoma was noted at groin site after transition to a repositioning sheath. manual pressure was applied, the perclose was cinched further, and angle matching was completed. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the ischemia and hematoma were unable to be determined due to insufficient clinical details. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1819 | 1220648-2026-00361 | 24016598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016598&pc=OZD | 12/30/2025 |  |  | 01/08/2026 |  |  | 12/30/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026776460 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Difficult or Delayed Positioning | Restricted Flow rate | [5 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1820 | 1220648-2026-00362 | 24016599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016599&pc=OZD | 12/30/2025 |  |  | 01/08/2026 |  |  | 12/30/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1420348 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 93 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that both aortic (ao) and left ventricular (lv) placement signals were no longer displayed on automated impella controller (aic) serial number (b)(6) (not (b)(6)-which was a forced selection). motor current values remained displayed. within less than five minutes of the loss of placement signals, an alert prompted switching to a different console. the system was transitioned to a loaner console to complete the procedure, and the pump continued to operate without further issues. the patient care was withdrawn and the patient has expired. additional information indicated that the case was complex in nature. the procedure initially began as a primary percutaneous coronary intervention (ppci); however, a dissection of the circumflex artery occurred, and the patient was emergently transferred to the operating room for coronary artery bypass grafting (cabg) with impella support. per the surgeon, the aortic arch was noted to be heavily calcified, requiring multiple repositioning attempts of the aortic cross-clamp. postoperatively, the patient¿s abdomen was observed to be significantly enlarged and distended compared to baseline at the start of the case. there was concern that embolized calcific material may have resulted in vascular obstruction and subsequent ischemia. despite mechanical support and attempted repositioning, the patient¿s lactic acid levels continued to rise. during the cabg, the impella pump was noted to migrate; although it remained positioned across the aortic valve, it was angled toward the mitral valve. pump flows were limited and could not be increased beyond p5. attempts to reposition the device, including withdrawing it approximately 1¿2 cm, were unsuccessful. the cause of death was not related to impella as per physician. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1821 | 1220648-2026-00365 | 24016600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016600&pc=OZD | 12/26/2025 |  |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026789343 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1822 | 1220648-2026-00366 | 24016602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016602&pc=OZD | 12/26/2025 | 02/05/2026 |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708087 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1823 | 1220648-2026-00369 | 24016603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016603&pc=OZD | 12/26/2025 | 02/17/2026 |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1431841 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 73 KG | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | Computer Operating System Problem | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Clinical assessment: 83-year-old female with known history of coronary artery disease presents with acute myocardial infarction and in cardiogenic shock, stage c. labs prior included act 290, with no lactate or arterial blood gas draw. her echo demonstrated left ventricular ejection fraction of 25% and left ventricular end diastolic pressue of 25. no pulmonary artery catherter placed. it is noted she was on 1 inotrope and had her respiratory system supported. a cp was placed but she unfortuanltey died 8 hours later while on support. the automated impella controller (aic) alarm was noted prior to connection to impella and a new aic was used during the case. the automated impella controller (aic) will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange however the exchange was performed with no consequence to the patient and support was resumed without any known adverse events. engineering assessment: during startup, automated impella controller (aic) experienced startup failure. team stated that the automated impella controller (aic) was unable to perform the system self-test at startup and message read ¿system self check failed. change console immediately¿. console turned off and back on, with same result. console was removed from service. this is considered a reportable malfunction. | The complainant reported that when the team member turned on the automated impella controller (aic), an error message displayed after several minutes stating, ¿system check failed, change console immediately.¿ the console was turned off and back on with the same result. biomedical engineering was notified, a complaint was filed, and the aic was taken out of use. the patient has expired. |  |  |  |  | B5 updated/corrected. a1102 and f1905 removed from h6; h6 updated with new codes. the investigation is ongoing. | D4. expiration date is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1824 | 1220648-2026-00371 | 24016604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016604&pc=OZD | 10/27/2025 |  |  | 01/08/2026 |  |  | 12/16/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1420045 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that pressure sensor not responding (psnr) was observed and pressure sensor pump number two (ps pump #2) was absent, escalated from pump number one (control panel [cp]). at the time of observation, the system was still supporting suction despite pressure sensor loss and sensing issues, with no signs or symptoms. a subsequent issue for the console was later noted, showing the same failure modes. there were no signs, symptoms, patient involvement, or consequences. the issue appears related to optical sensor performance. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1825 | 1220648-2026-00372 | 24016605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016605&pc=OZD | 10/27/2025 |  |  | 01/08/2026 |  |  | 12/17/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1200772 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella console displayed a placement signal not reliable alarm. prior to this alarm, an impella in aorta alarm appeared, and the device was repositioned at the bedside using point-of-care ultrasound by advancing approximately four centimeters. the placement changed from two centimeters to approximately four centimeters in depth across the aortic valve to the mid-inlet. just before the alarm, the purge bag was switched from dextrose 5 percent in water to dextrose 5 percent in water with sodium bicarbonate. the device continued to function after the alarm, and the care team monitored the pump overnight and gradually weaned the device from performance level six to performance level two. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1826 | 1220648-2026-00375 | 24016607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016607&pc=OZD | 12/26/2025 |  |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026773616 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a bedside nurse called regarding a placement signal alarm. the yellow alarm indicating that the placement signal was unreliable appeared on the screen. the complainant was able to guide the nurse to ensure that the plug was fully inserted into the automated impella controller (aic). it was confirmed to be properly connected; however, when the nurse adjusted the plug slightly, the placement signal began functioning again. it appears to be working intermittently and remains unreliable. additional information stated that the patient is still receiving support, and the team will make every effort to recover the device after explant. a return kit is already available at home, and notification will be provided if and when the device is shipped. the patient experienced a hemorrhagic stroke while the 5.5 device was in place. the team does not believe the 5.5 device contributed to the stroke, as it was associated with another medical condition. the patient underwent an emergent craniotomy and evacuation, and this remains an ongoing concern for the patient. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary no product return stroke/peri-procedural adverse event: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review the pump passed all the post sterile inspection checks. | Additional information received; d6b. explantation day, month and year. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1827 | 1220648-2026-00374 | 24016608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016608&pc=OZD | 12/26/2025 |  |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1124472 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a bedside nurse called regarding a placement signal alarm. the yellow alarm indicating that the placement signal was unreliable appeared on the screen. the complainant was able to guide the nurse to ensure that the plug was fully inserted into the automated impella controller (aic). it was confirmed to be properly connected; however, when the nurse adjusted the plug slightly, the placement signal began functioning again. it appears to be working intermittently and remains unreliable. additional information stated that the patient is still receiving support, and the team will make every effort to recover the device after explant. a return kit is already available at home, and notification will be provided if and when the device is shipped. the patient experienced a hemorrhagic stroke while the 5.5 device was in place. the team does not believe the 5.5 device contributed to the stroke, as it was associated with another medical condition. the patient underwent an emergent craniotomy and evacuation, and this remains an ongoing concern for the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received: 6b. explantation day, month and year d9 device available for evaluation? d9 is device returned to manufacturer? d9 date device returned to manufacturer. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
1828 | 1220648-2026-00373 | 24016609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016609&pc=OZD | 12/26/2025 |  |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026768530 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Decrease in Pressure | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1829 | 1220648-2026-00368 | 24016611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016611&pc=OZD | 12/26/2025 |  |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026731835 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation of Vessels | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1830 | 1220648-2026-00376 | 24016612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016612&pc=OZD | 12/26/2025 |  |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026771218 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 98 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that following device insertion, a low placement signal alarm was observed. the placement signal measured 5/2 (3), with a systemic arterial line map of 60 mmhg. at p4, flows were approximately 3.4 l/min; however, no placement signal was detected, and the left ventricle placement signal was negative. device position was confirmed via x-ray and transesophageal echocardiography (tee). a new device was opened and inserted without issue. no patient harm was reported, and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1831 | 1220648-2026-00377 | 24016617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016617&pc=OZD | 12/26/2025 |  |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026796226 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Cardiac Perforation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1832 | 1220648-2026-00339 | 24016622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016622&pc=OZD | 01/02/2026 | 01/08/2026 |  | 01/08/2026 |  |  | 01/02/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727439 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1833 | 1220648-2026-00380 | 24016623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016623&pc=OZD | 01/04/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760949 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1834 | 1220648-2026-00382 | 24016626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016626&pc=OZD | 12/26/2025 | 01/08/2026 |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026738787 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1835 | 1220648-2026-00383 | 24016629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016629&pc=OZD | 12/27/2025 | 01/08/2026 |  | 01/08/2026 |  |  | 12/28/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026775377 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 73 YR |  | Male | 72 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella support with a cp. the complainant reported the patient bled out of their nasogastric tube. care was eventually withdrawn and the patient expired. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the epistaxis was unable to be determined due to insufficient clinical details. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
1836 | 1220648-2026-00385 | 24016630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016630&pc=OZD | 12/26/2025 | 01/08/2026 |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728848 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1837 | 1220648-2026-00386 | 24016634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016634&pc=OZD | 01/04/2026 |  |  | 01/08/2026 |  |  | 01/05/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026767795 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1838 | 1220648-2026-00388 | 24016635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016635&pc=OZD | 01/05/2026 |  |  | 01/08/2026 |  |  | 01/05/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026797220 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1839 | 1220648-2026-00387 | 24016636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016636&pc=OZD | 12/16/2025 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1316944 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
1840 | 1220648-2026-00389 | 24016648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016648&pc=OZD | 12/30/2025 | 01/08/2026 |  | 01/08/2026 |  |  | 12/30/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026785093 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 59 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an attempt was made to implant an impella 5.5 for mechanical circulatory support; however, the device could not be advanced due to inadequate vessel size. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. failure to advance: the cause of the delivery issue was determined to be patient condition due to small vessel size. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1841 | 1220648-2026-00391 | 24016652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016652&pc=OZD | 12/30/2025 |  |  | 01/08/2026 |  |  | 12/30/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026783689 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 109 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that on insertion of an impella 5.5 for mechanical circulatory support, the pump was directed towards the apex and was deep. the pump was pulled back to five centimeters from the aortic valve but was directed to the mitral. attempts were made to reposition but were unsuccessful. the pump function was within normal limits so the decision was made to leave as is. the patient's outcome is stable. this report represents the impella 5.5. separate report(s) will be submitted for associated devices with reportable events. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary. no product was returned. device in wrong position: the cause of the positioning issue was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D6b: added explant date.
1842 | 1220648-2026-00390 | 24016654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016654&pc=OZD | 12/30/2025 |  |  | 01/08/2026 |  |  | 12/30/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026779551 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was admitted as having st-elevation myocardial infarction. multivessel disease was noted. ejection fraction was 25% and end-diastolic pressure was 24. percutaneous transluminal coronary angioplasty (pcta) to mid/distal left circumflex artery was noted. ptca/percutaneous coronary intervention to the proximal left anterior descending artery was done. the impella cp was placed. angioplasty revealed no flow to the right femoral artery below the impella access. the doctor noted an occlusion and wanted to consult with the cardiovascular surgeon to escalate the patient to an impella 5.5 axillary due to occlusion. no pulse was noted. both limbs were ischemic. the impella remained in with no issues at this time. the cardiovascular surgeon was in surgery so the patient was awaiting consult with them. the patient was taken to the operating room. the impella cp was removed successfully with right femoral artery cutdown and repair completed re-establishing blood flow to right lower extremity. no thrombectomy was needed. due to the complexity of the patient's condition and grim outcome, the doctor spoke with the family and decided not to place an additional device at this time. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: (ischemia): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
1843 | 1220648-2026-00392 | 24016655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016655&pc=OZD | 01/04/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1301514 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male | 88 KG | Non Hispanic | White | Death | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | Unfortunately the pump was discarded, so no return kit is needed. the doctors replaced the controller at the end when an increasing in motor current was observed. however, it was clearly stated that this increase was not caused by a pump defect, but that it was most likely due to a thrombogenic event. the doctors clearly stated that the patient did not die because of a pump failure, but that his clinical condition deteriorated very rapidly, leading to this state. | The user facility reported that a patient had been implanted with an impella cp for mechanical circulatory support. the medical doctor reported that the pump had stopped. an alarm "aic failure" was displayed. it is recommended to change the automated impella controller (aic). there was no backup aic on the ward. one must be retrieved from the catheterization lab. the patient was impella-dependent and needed to be stabilized with rising catecholamine levels. it was discussed with the medical doctor that it might be necessary to replace the pump. abiomed best practices were shared. the procedure outcome was that the patient expired on support. this is one of two related reports. this report represent the automated impella controller and another report will be submitted to represent the impella cp. |  |  |  |  | Additional information received b5 updated. | This report is being submitted to provide medical safety/clinical review and to update reportability based on medical assessment of previously reported information. no new clinical information was received; the update reflects assessment of reportability criteria based on the existing event details. patient coding was updated to align with change in reportability. add: hemodynamic instability (e2343) and death (f02) remove: problem identified before clinical use/exposure (f2601) and no clinical signs, symptoms or conditions (e2403) medical safety/clinical review. medical safety/clinical reviewed the event. the patient was being treated for acute myocardial infarction complicated by cardiogenic shock. therapies administered within one hour prior to impella support included inotropes and vasopressors. there was no information indicating the patient was receiving extracorporeal membrane oxygenation (ecmo) support. an impella cp device was implanted for mechanical circulatory support. on day 7 of impella support, the impella cp pump stopped functioning. an ¿aic failure¿ alarm was displayed, and the clinical team was advised to exchange the automated impella controller (aic). at the time of the event, there was no backup aic available on the ward, and one had to be retrieved from the cardiac catheterization laboratory (it is unknown if the aic was available. additional information received noted the aic was replaced at the end when an increasing in motor current was observed). the patient was reported to be impella-dependent and required stabilization with escalating catecholamine support following the loss of pump function. the patient ultimately expired the same day as the pump stop event. it is unknown whether the patient expired immediately following the device event. based on the event details, there was an immediate loss of hemodynamic support in a patient with cardiogenic shock who was dependent on impella support. this likely resulted in acute hemodynamic deterioration. the event story involves two product complaints: impella cp - mdr 1220648-2026-00393: death. automated impella controller - mdr 1220648-2026-00392: death. related medical device reports: this automated impella controller device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1844 | 1220648-2026-00393 | 24016656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016656&pc=OZD | 01/04/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749178 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation | [9 empty] | Death | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1845 | 1220648-2026-00394 | 24016657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016657&pc=PYX | 01/03/2026 |  |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026772363 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1846 | 1220648-2026-00395 | 24016659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016659&pc=OZD | 11/26/2023 |  |  | 01/08/2026 |  |  | 12/11/2025 | 01/08/2026 |  |  | 09/30/2025 |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397102 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 102 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Renal Failure | Hemodynamic instability | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported a patient presenting in cardiogenic shock secondary to an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and subsequently experienced limb ischemia requiring medical intervention. the impella cp device was inserted prior to percutaneous coronary intervention (pci) of the left anterior descending artery, during which a stent was placed. at the time of implantation, the patient had a reported left ventricular ejection fraction of approximately 15%, 3+ mitral regurgitation, and significant volume overload. following the procedure, the patient was transported to the intensive care unit on impella cp mcs. on the following day, the patient developed worsening renal function with rising creatinine and rising potassium and was initiated on continuous renal replacement therapy. overnight, the patient developed right lower extremity limb ischemia. the patient also experienced worsening hemodynamic instability and was emergently cannulated for venoarterial extracorporeal membrane oxygenation (va-ecmo) the next morning. the impella cp device was ultimately removed. the patient was taken emergently to the operating room, where vascular surgery performed a mechanical thrombectomy and emergent fasciotomy. the patient remained hemodynamically unstable on va-ecmo support. vasopressors were being weaned, and lactic acid levels were reported to be decreasing from 6 mmol/l to 3 mmol/l. right foot pulses were unable to be detected by doppler, and the extremity remained "significantly" mottled. of note, the patient was on pressers which can potentially make the limb ischemia worse. a consultation was initiated for transfer to a higher level of care, and at the conclusion of impella cp support, the patient was reported as surviving and remained on va-ecmo support. | [5 empty] | H6: added codes to health effect-impact code, health effect-clinical code. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: hemodynamic instability/renal failure: the cause of the injury was not determined due to insufficient clinical details. ischemia: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1847 | 1220648-2026-00399 | 24016663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016663&pc=OZD | 01/04/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026779302 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1848 | 1220648-2026-00398 | 24016664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016664&pc=OZD | 01/04/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768750 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1849 | 1220648-2026-00400 | 24016665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016665&pc=OZD | 01/04/2026 |  |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750822 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1850 | 1220648-2026-00397 | 24016666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016666&pc=OZD | 01/02/2026 | 01/14/2026 |  | 01/08/2026 |  |  | 01/04/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026786129 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Arrest | Cardiac Perforation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1851 | 1220648-2026-00381 | 24016668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016668&pc=OZD | 12/31/2025 |  |  | 01/08/2026 |  |  | 01/05/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026746507 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1852 | 1220648-2026-00402 | 24016670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016670&pc=OZD | 12/26/2025 | 01/08/2026 |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1597910 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | [9 empty] | Malfunction | No | Yes | Pumping Stopped | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission
1853 | 1220648-2026-00401 | 24016671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016671&pc=OZD | 12/26/2025 | 02/05/2026 |  | 01/08/2026 |  |  | 12/26/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766546 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1854 | 1220648-2026-00403 | 24016676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016676&pc=OZD | 12/24/2025 |  |  | 01/08/2026 |  |  | 12/24/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026749030 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1855 | 1220648-2026-00404 | 24016713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016713&pc=OZD | 01/03/2026 | 01/08/2026 |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763074 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1856 | 1220648-2026-00406 | 24016715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016715&pc=OZD | 01/03/2026 | 01/08/2026 |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026796341 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Difficult or Delayed Positioning | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1857 | 1220648-2026-00405 | 24016716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016716&pc=OZD | 01/03/2026 | 01/08/2026 |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1100033 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | Operating System Becomes Nonfunctional | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1858 | 1220648-2026-00408 | 24016718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016718&pc=OZD | 01/03/2026 | 01/08/2026 |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026770275 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | Valvular Insufficiency/ Regurgitation | Mitral Valve Insufficiency/ Regurgitation | [33 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1859 | 1220648-2026-00411 | 24016719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016719&pc=OZD | 01/03/2026 | 01/08/2026 |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732047 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1860 | 1220648-2026-00410 | 24016721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016721&pc=OZD | 01/03/2026 | 01/08/2026 |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732602 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1861 | 1220648-2026-00412 | 24016728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016728&pc=OZD | 01/03/2026 | 01/08/2026 |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564779 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1862 | 1220648-2026-00413 | 24016730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016730&pc=OZD | 01/06/2026 | 01/12/2026 |  | 01/08/2026 |  |  | 01/06/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026737673 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | No Apparent Adverse Event | [5 empty] | Anemia | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1863 | 1220648-2026-00414 | 24016731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016731&pc=PYX | 12/12/2025 | 01/08/2026 |  | 01/08/2026 |  |  | 12/12/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1678954 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 121 KG | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that a 72 year old male experienced an unexpected battery power loss while being transferred from the cath lab to the intensive care unit (icu). after only several minutes on battery power, the system displayed a loss of battery power notification despite indicating 50% battery life remaining. battery status later returned to 100%. the icu team switched the patient to a backup aic. the facility noted this has been a frequent occurrence and expressed frustration with the lack of resolution. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
1864 | 1220648-2026-00415 | 24016736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016736&pc=OZD | 01/05/2026 | 01/08/2026 |  | 01/08/2026 |  |  | 01/05/2026 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026797959 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 85 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported notification of a stroke in a 68-year-old male patient. the event was reported by nursing staff, and the care team became aware of the event during rounds. no additional information was provided. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the stroke, the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. d4 added primary udi number as it was omitted from the initial report that was submitted. h6 health effect - impact code. was corrected from f24 to f12. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1865 | 1220648-2026-00417 | 24016738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016738&pc=OZD | 01/06/2026 | 01/08/2026 |  | 01/08/2026 |  |  | 01/06/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1399832 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1866 | 1220648-2026-00409 | 24016748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016748&pc=OZD | 01/03/2026 | 02/02/2026 |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026765841 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1867 | 1220648-2026-00407 | 24016755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016755&pc=OZD | 01/02/2026 | 01/08/2026 |  | 01/08/2026 |  |  | 01/03/2026 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026774651 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1868 | 1220648-2026-00419 | 24016774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016774&pc=OZD | 11/21/2025 |  |  | 01/08/2026 |  |  | 11/21/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026743348 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Inaccurate Flow Rate | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1869 | 1220648-2026-00420 | 24016778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016778&pc=OZD | 11/21/2025 |  |  | 01/08/2026 |  |  | 11/21/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026727733 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1870 | 1220648-2026-00421 | 24016788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016788&pc=OZD | 11/21/2025 |  |  | 01/08/2026 |  |  | 11/21/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026726971 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. the impella device was explanted in the procedural area following a successful hrpci. two perclose sutures were placed at the start of the procedure for pre-closure. after removal of the 14 fr x 25 cm peel-away sheath, the team attempted to secure the perclose sutures. at this time, act was 283. both sutures broke during the process. initially, dr. chaturvedula suspected the breakage was due to pressure applied at the groin site by an rn. three additional perclose sutures were attempted, for a total of five, all of which broke in a similar manner even after the rn adjusted technique. an 8 fr angio-seal was then attempted and also broke. contralateral access was obtained, and temporary balloon tamponade was used for hemostasis for approximately 20¿30 minutes before deflation. during this time, 20 mg of protamine was administered, and act decreased to 109. hemostasis was nearly achieved, so additional manual pressure was applied, and the team prepared for femstop per dr. request. dr. did not suspect the impella or peel-away sheath as the cause, but suggested that something in the subcutaneous tissue may have contributed. flow was visualized down the leg with pulses present and no thrombus noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. investigation summary ==> major bleed/ access site adverse event: slight trickle of blood at groin noted after 5 perclose sutures and an 8fr angio seal broke when trying to secure and achieve hemostasis. patient vessel condition noted to be tortuous. the cause of the issue was determined to be patient condition related as evidenced by the numerous perclose sutures breaking even after technique changes as well as 8fr angioseal breaking. device history lot ==> device lot number: 1960386 device history batch ==> subcomponent lot: n/a device history review ==> the pump sn (b)(6) passed all post sterile inspection checks.
1871 | 1220648-2026-00422 | 24016793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016793&pc=OZD | 12/19/2025 | 01/16/2026 |  | 01/08/2026 |  |  | 12/19/2025 | 01/08/2026 |  |  | 02/28/2026 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025471257 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1872 | 1220648-2026-00423 | 24016799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016799&pc=OZD | 12/19/2025 |  |  | 01/08/2026 |  |  | 12/19/2025 | 01/08/2026 |  |  |  |  | 01/08/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026787423 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | After the patient was transferred to the icu following impella cp implantation in the cardiac catheterization lab (ccl), bleeding and oozing were observed at the femoral access site. the dressing was removed, and manual pressure was applied by ccu staff at the bedside. after approximately 30 minutes of manual pressure, the bleeding and oozing resolved. the femoral site was cleaned and redressed. no hematoma was noted at the access site. | [5 empty] | No product return: asae/ minor bleed: bleeding/oozing noted at fem site. dressing removed and manual pressure held which resolved bleeding. the cause of the access site adverse event was not determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1873 | 1220648-2026-00424 | 24016804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016804&pc=OZD | 12/19/2025 |  |  | 01/08/2026 |  |  | 12/19/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026771163 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1874 | 1220648-2026-00425 | 24016812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24016812&pc=PYX | 12/19/2025 | 01/29/2026 |  | 01/08/2026 |  |  | 12/19/2025 | 01/08/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026772936 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1875 | 1220648-2026-00068 | 23988459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23988459&pc=OZD | 12/11/2025 | 01/04/2026 |  | 01/07/2026 |  |  | 12/11/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026763987 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 86 KG | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient needing an impella device for mechanical circulatory support. during implantation, the device was unable to be placed due to poor vessel size. no further information was provided. there was no reported harm or injury to the patient. | [5 empty] | Investigation summary: no product returned. difficult to advance: the cause of the delivery issue was determined to be patient condition related to the patient's poor vessel size. device history lot: device lot: 1986747. device history batch: subcomponent lot: n/a. device history review: qn (b)(4) was opened for s/n (b)(6) of l/n 1986747 for failing ppm. mrb confirmed s/n (b)(6) to have passing ppm results. this is unrelated to the failure mode. the complaint pump s/n (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1876 | 1220648-2026-00136 | 23988723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23988723&pc=OZD | 12/15/2025 | 01/02/2026 |  | 01/07/2026 |  |  | 12/16/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026788915 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male | 78 KG | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 68-year-old male was implanted with an impella cp device for support during a high-risk cardiac procedure. it was reported that the device did not pass the brachiocephalic trunk due to incorrect angulation. attempts to correct the angulation resulted in the device repeatedly slipping out of the left ventricle. after several hours, further attempts were discontinued and support was switched to extracorporeal membrane oxygenation. no additional information was provided. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1877 | 1220648-2026-00210 | 23988773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23988773&pc=OZD | 12/19/2025 | 01/27/2026 |  | 01/07/2026 |  |  | 12/18/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781986 |  | 000000000001000080 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1878 | 1220648-2026-00211 | 23988799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23988799&pc=OZD | 12/21/2025 |  |  | 01/07/2026 |  |  | 12/18/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026776563 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female | 64 KG | Hispanic | White | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the surgeon dissected aorta when placing the companion sheath, and no issues with impella insertion prior to that. it was reported that patient experienced hemolysis from 3 point drop of hemoglobin. the medical doctor (md) planned to wean and remove impella, lactate dehydrogenase (ldh) lab ordered for confirmation and no other signs symptoms of hemolysis noted or reported. it was reported also that upon removal of the device a tissue was noted at the end of the impella pigtail which impacts the papillary muscle. patient reportedly had vascular and tissue abnormalities. multiple medical doctors (md) even unable to start an a line due to fragile arteries. the impella was removed smoothly nothing out of the ordinary noted with removal. the patient was reported to have survived the procedure. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be file.
1879 | 1220648-2026-00212 | 23988849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23988849&pc=OZD | 12/22/2025 |  |  | 01/07/2026 |  |  | 12/19/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1085763 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor | [9 empty] | Malfunction | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1880 | 1220648-2026-00213 | 23988875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23988875&pc=OZD | 12/19/2025 |  |  | 01/07/2026 |  |  | 12/19/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779188 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Distributor | [10 empty] | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1881 | 1220648-2026-00217 | 23991793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991793&pc=OZD | 01/01/2026 |  |  | 01/07/2026 |  |  | 01/01/2026 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026745676 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 12 KG |  | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella 5.5 was aborted as the pump could not pass through the patient's anatomy. no patient harm was reported. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. failure to advance: the cause of the deliver issue was determined to be patient condition as the patient had stenosis and resistance at aortic valve preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
1882 | 1220648-2026-00216 | 23991795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991795&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026792013 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup
1883 | 1220648-2026-00218 | 23991798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991798&pc=OZD | 01/02/2026 | 01/13/2026 |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1135095 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
1884 | 1220648-2026-00215 | 23991814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991814&pc=OZD | 01/01/2026 |  |  | 01/07/2026 |  |  | 01/01/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026774875 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1885 | 1220648-2026-00214 | 23991827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991827&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755557 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1886 | 1220648-2026-00223 | 23991840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991840&pc=OZD | 01/02/2026 | 01/13/2026 |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599109 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1887 | 1220648-2026-00221 | 23991842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991842&pc=OZD | 01/02/2026 | 01/14/2026 |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760224 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup
1888 | 1220648-2026-00224 | 23991862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991862&pc=OZD | 01/02/2026 | 01/12/2026 |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592603 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Rupture | No Clinical Signs, Symptoms or Conditions | Intraoperative Cardiac Injury | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1889 | 1220648-2026-00228 | 23991864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991864&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026796357 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 45 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced an ischemic cerebrovascular accident and expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
1890 | 1220648-2026-00219 | 23991894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991894&pc=OZD | 12/31/2025 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026774673 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Difficult to Insert | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1891 | 1220648-2026-00227 | 23991899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991899&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732764 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1892 | 1220648-2026-00222 | 23991907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991907&pc=OZD | 12/25/2025 |  |  | 01/07/2026 |  |  | 12/25/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755423 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 45 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding from the axillary site requiring a transfusion of one unit of packed red blood cells. the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1893 | 1220648-2026-00226 | 23991912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991912&pc=OZD | 01/02/2026 | 01/13/2026 |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026769452 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1894 | 1220648-2026-00231 | 23991944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23991944&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666184 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1895 | 23992615 | 23992615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23992615&pc=OZD | 11/01/2025 | 11/10/2025 | 11/10/2025 | 01/07/2026 | 11/10/2025 | 01/07/2026 |  | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 65 YR |  | Female |  |  |  | Malfunction | No | Yes | Inappropriate Waveform | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella 5.5 lost waveforms. they came back when impella was increased to p6. impella remains in place.
1896 | 1220648-2026-00232 | 23993710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993710&pc=OZD | 12/24/2025 |  |  | 01/07/2026 |  |  | 12/24/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026741026 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1897 | 1220648-2026-00233 | 23993744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993744&pc=OZD | 12/24/2025 |  |  | 01/07/2026 |  |  | 12/24/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026762886 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. it was reported that before leaving the lab, the medical doctor performed a lower extremity ultrasound that showed flow at that time. the patient had doppler pulses. the patient's outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
1898 | 1220648-2026-00234 | 23993746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993746&pc=OZD | 12/30/2025 | 02/13/2026 |  | 01/07/2026 |  |  | 12/30/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026779567 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Difficult to Open or Close | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1899 | 1220648-2026-00235 | 23993750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993750&pc=OZD | 12/30/2025 |  |  | 01/07/2026 |  |  | 12/30/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026774623 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 98 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient experienced placement signal (ps) low alarm. the ps appeared to have lots of artifact and extremely elevated ps numbers. the patient was bearing down but was unchanged when stopped bearing. an echocardiogram was ordered to confirm positioning. the pump was in good position. this is one of three pumps used on the patient. this report addresses the third pump used. additional reports will be submitted to represent the other associated devices with reportable events. | [5 empty] | Additional information received: 6b. explantation day, month and year. corrected information: 3. date of event. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1900 | 1220648-2026-00237 | 23993761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993761&pc=OZD | 12/23/2025 |  |  | 01/07/2026 |  |  | 12/23/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026795488 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup
1901 | 1220648-2026-00236 | 23993762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993762&pc=OZD | 12/23/2025 |  |  | 01/07/2026 |  |  | 12/23/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026752074 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 64 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp. the patient had severe peripheral artery disease but was implanted with an impella cp due to severe cardiogenic shock. after several hours, the leg was cold and pulses could not be found so the decision was made to remove the impella cp and place an intra-aortic balloon pump. the patient's outcome was stable. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the ischemia was unable to be determined due to insufficient clinical details. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral, vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
1902 | 1220648-2026-00239 | 23993780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993780&pc=OZD | 12/31/2025 | 01/29/2026 |  | 01/07/2026 |  |  | 01/01/2026 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709622 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male | 90 KG |  | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hematuria | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | [8 empty] | The complainant report that a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the purge fluid dropped below two milliliters per hour (ml/h) and "pp high" alarm occurred. motor current was to be monitored. the pump runs well without any problems and the decreased purge fluid was addressed. local issue plasminogen activator (tpa) lysis was mentioned and requested by the doctor. fluid was changed to 50 milliliters of aqua added and tpa was performed together. the following day, the medical doctor reported that the lysis was successful and is wanting to know about purge fluid and purge pressure trend.. purge fluid was slightly decreased again to 3.5 ml/h. four days later, the medical doctor reported that purge fluid trend has again shown a negative development in recent days causing the purge fluid to drop below three ml/h during the past night. the doctor wants to apply tpa again. hematuria was noted. urinary output was greater than 30 cubic centimeters per hour and was pink-red. this is one of two related reports. this report represent the impella 5.5 and another report was submitted to represent the automated impella controller. refer to section h10 for the related report number. | The complainant also reported that the optical sensor was defective; reliable placement and left ventricle signals were no longer showing on the accompanying automated impella controller. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation is still ongoing.
1903 | 1220648-2026-00238 | 23993782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993782&pc=OZD | 04/02/2025 | 02/16/2026 |  | 01/07/2026 |  |  | 12/22/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628637 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Unspecified Infection | Inflammation | Cardiogenic Shock | Pericardial Effusion | [32 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1904 | 1220648-2026-00240 | 23993788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993788&pc=OZD | 12/25/2025 | 02/24/2026 |  | 01/07/2026 |  |  | 12/26/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026785151 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1905 | 1220648-2026-00241 | 23993814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993814&pc=OZD | 12/26/2025 |  |  | 01/07/2026 |  |  | 12/26/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026778728 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. it was reported intermittent ventricular tachycardia with pump positioning, resolved when repositioned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1906 | 1220648-2026-00243 | 23993857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993857&pc=OZD | 12/26/2025 |  |  | 01/07/2026 |  |  | 12/26/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026769380 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1907 | 1220648-2026-00245 | 23993894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993894&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026774882 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the cp went in to aorta and when repositioning to get back across the av device went too deep and patient went into ventricular tachycardia (vt). decreased from p4 to p2, and continue to reposition under transthoracic echocardiogram. eventually vt improved and did not need to shock patient and impella looked to be in a better position. it was further reported that the patient was having frequent premature ventricular contractions (pvcs), team attempted repositioning impella several times however the pigtail was thought to be causing pvcs, decision was to explant device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position: the cause of the device in wrong position was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
1908 | 1220648-2026-00246 | 23993902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993902&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026721540 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1909 | 1220648-2026-00242 | 23993907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993907&pc=OZD | 12/25/2025 |  |  | 01/07/2026 |  |  | 12/25/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026790147 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 84 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. it was reported urinary output - anuric and the patient was put on dialysis. it was also noted that the patient expired. no additional information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1910 | 1220648-2026-00249 | 23993946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993946&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026752067 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1911 | 1220648-2026-00244 | 23993961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993961&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1081050 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Biomed reported from field service to make aware that controller is showing an error and needs to be sent out for repair. no other information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1912 | 1220648-2026-00248 | 23993992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993992&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026730406 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Positioning Failure | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1913 | 1220648-2026-00247 | 23993998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23993998&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026778819 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1914 | 1220648-2026-00250 | 23994033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23994033&pc=OZD | 12/24/2025 |  |  | 01/07/2026 |  |  | 12/24/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026782967 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1915 | 1220648-2026-00253 | 23994081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23994081&pc=OZD | 12/30/2025 |  |  | 01/07/2026 |  |  | 12/30/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765581 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1916 | 1220648-2026-00252 | 23994100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23994100&pc=PYX | 12/26/2025 |  |  | 01/07/2026 |  |  | 12/30/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1805843 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A user facility reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the silver door on the automated impella controller broke off when opening to plug in impella catheter. the pump functioned optimally and was explanted. the catheterization lab notified abiomed of the broken piece. no patient harm was noted. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data review: console logs did not contain any data relevant to the complaint. device analysis: console was tested after and the optical connector dust cover was confirmed to be broken during visual inspection of the console. aside from the broken optical connector dust cover, the console operated as intended while testing with a known good pump and purge cassette kit. root cause: the root cause of the broken optical connector dust cover was most likely related to excessive force during use. phr summary: there were no issues related to the complaint discovered when reviewing the product history of console | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
1917 | 1220648-2026-00255 | 23994136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23994136&pc=OZD | 12/24/2025 |  |  | 01/07/2026 |  |  | 12/24/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699700 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1918 | 1220648-2026-00257 | 23994169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23994169&pc=OZD | 12/25/2025 |  |  | 01/07/2026 |  |  | 12/26/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026779528 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 134 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support the patient coded and was shocked that resolved the problem. the patient remains on support. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the arrhythmia was unable to be determined due to insufficient clinical details.
1919 | 1220648-2026-00254 | 23994212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23994212&pc=OZD | 12/15/2025 |  |  | 01/07/2026 |  |  | 12/15/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026721513 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1920 | 1220648-2026-00048 | 23999851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999851&pc=OZD | 12/01/2025 | 02/04/2026 |  | 01/07/2026 |  |  | 12/01/2025 | 01/07/2026 |  |  | 11/30/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | 1000100 |  |  | ABIOMED, INC. |  |  | US |  | 2024427624 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Difficult to Remove | Malposition of Device | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1921 | 1220648-2026-00260 | 23999868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999868&pc=OZD | 12/26/2025 |  |  | 01/07/2026 |  |  | 12/30/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716937 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 97 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp for mechanical circulatory support. the surgeon notified the team that there was concern for distal limb ischemia post impella cp explant. the foot was cold. the patient's outcome is stable. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
1922 | 1220648-2026-00262 | 23999870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999870&pc=OZD | 12/24/2025 |  |  | 01/07/2026 |  |  | 01/01/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026697083 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1923 | 1220648-2026-00265 | 23999872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999872&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026795391 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1924 | 1220648-2026-00264 | 23999873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999873&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026781821 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1925 | 1220648-2026-00266 | 23999875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999875&pc=OZD | 01/01/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1587053 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1926 | 1220648-2026-00267 | 23999916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999916&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1347548 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
1927 | 1220648-2026-00268 | 23999922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999922&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 |  |  | 02/28/2026 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481172 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Failure to Advance | Deformation Due to Compressive Stress | Difficult to Advance |  |  |  | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1928 | 1220648-2026-00271 | 23999926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999926&pc=OZD | 12/26/2025 | 02/14/2026 |  | 01/07/2026 |  |  | 12/26/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026754371 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Ventricular Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1929 | 1220648-2026-00270 | 23999927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999927&pc=OZD | 12/23/2025 | 01/06/2026 |  | 01/07/2026 |  |  | 12/25/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2026781715 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1930 | 1220648-2026-00272 | 23999928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999928&pc=OZD | 12/24/2025 | 01/06/2026 |  | 01/07/2026 |  |  | 12/26/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026782171 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1931 | 1220648-2026-00273 | 23999929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999929&pc=OZD | 12/29/2025 | 01/07/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026776837 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 109 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70 year old male patient in whom a second impella cp device was implanted after removal of an initial device due to clinical instability. the replacement procedure was performed without incident. the device was initially positioned deep in the left ventricle, and support was limited to p4 during sedation and paralysis for sheath removal. after peel-away sheath removal, the insertion site was clean, dry, and intact. the device was repositioned by approximately 3 cm, resulting in improved flows; however, support could not be increased beyond p6. the device position measured 4.1 cm within the left ventricle, and no further repositioning was performed. this complaint pertains to the second catheter. no additional information was provided. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. low or blocked pump flow: clinical reported impella was initially deep, but impella was pulled back 3cm and flows increased but still unable to flow past p6. impella measured 4.1 in the left ventricle (lv). unsure if flows are due to anatomy or position. the cause of the issue was not established as no logs or product were returned for analysis. device history review: the pump passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1932 | 1220648-2026-00274 | 23999930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999930&pc=OZD | 12/29/2025 | 01/07/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026753139 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1933 | 1220648-2026-00275 | 23999931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999931&pc=OZD | 12/24/2025 | 01/07/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026786721 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Decrease in Pressure | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 69-year-old male who was implanted with an impella cp device for support during a high-risk cardiac procedure. it was reported that during an aar the perfusionist noted a leaking purge sidearm. the patient experienced several episodes of delirium and restlessness, and it was reported that the sidearm may have been damaged during these events. a purge filter bypass was performed per instructions, after which purge values were stable and no further leaking was observed. the pump later stopped and was explanted. no patient harm or additional issues were reported. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1934 | 1220648-2026-00276 | 23999932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999932&pc=OZD | 12/29/2025 | 01/07/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528951 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 70 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old female patient in whom an impella 5.5 device was implanted via a 10 mm graft into the subclavian artery. it was reported that during insertion, the cannula was noted to be bent or kinked. no additional information was provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1935 | 1220648-2026-00277 | 23999935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999935&pc=OZD | 12/29/2025 | 02/13/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754132 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1936 | 1220648-2026-00278 | 23999937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999937&pc=OZD | 12/29/2025 | 02/11/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026789339 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1937 | 1220648-2026-00280 | 23999938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999938&pc=OZD | 12/29/2025 | 01/06/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026775637 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1938 | 1220648-2026-00282 | 23999940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999940&pc=OZD | 12/28/2025 | 01/12/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026799133 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the aic displayed several impella in aorta alarms. despite the alarms, motor current was noted to be pulsatile, and transthoracic echocardiography confirmed that the impella position was correct. no additional information was provided. | [5 empty] | The following sections have been updated: b1-added adverse event, b2-selected death and added date of death, b4, g3, g6, h2, h6 and h11. additional information was provided on the patient outcome stating that care was withdrawn and the patient subsequently expired. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1939 | 1220648-2026-00281 | 23999941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999941&pc=OZD | 12/29/2025 | 01/06/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026796290 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1940 | 1220648-2026-00283 | 23999942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999942&pc=OZD | 12/29/2025 | 01/06/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026715923 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 67 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Heart Block | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | After impella cp was positioned in the lv pt converted to intermittent third degree heart block. dr. (b)(6) elected to place a temporary trans venous pacemaker for the duration of the pci. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
1941 | 1220648-2026-00284 | 23999943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999943&pc=OZD | 12/29/2025 | 02/13/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026767060 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1942 | 1220648-2026-00285 | 23999946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999946&pc=OZD | 12/29/2025 | 01/06/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026792129 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | [10 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1943 | 1220648-2026-00287 | 23999949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999949&pc=OZD | 12/29/2025 | 01/06/2026 |  | 01/07/2026 |  |  | 12/29/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026782073 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1944 | 1220648-2026-00288 | 23999952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999952&pc=OZD | 12/23/2025 | 01/06/2026 |  | 01/07/2026 |  |  | 12/26/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768720 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1945 | 1220648-2026-00289 | 23999954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999954&pc=OZD | 12/26/2025 | 01/06/2026 |  | 01/07/2026 |  |  | 12/26/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026774047 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1946 | 1220648-2026-00290 | 23999975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999975&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736494 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1947 | 1220648-2026-00291 | 23999978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999978&pc=OZD | 01/04/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754047 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1948 | 1220648-2026-00292 | 23999981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999981&pc=OZD | 12/30/2025 | 01/07/2026 |  | 01/07/2026 |  |  | 12/30/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733775 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1949 | 1220648-2026-00294 | 23999986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999986&pc=OZD | 01/02/2026 |  |  | 01/07/2026 |  |  | 01/02/2026 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026794594 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1950 | 1220648-2026-00295 | 23999990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999990&pc=OZD | 12/30/2025 |  |  | 01/07/2026 |  |  | 12/30/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1081051 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 82 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the waveforms were no longer displayed. the patient survived the procedure and remains on extracorporeal membrane oxygenation (ecmo) support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
1951 | 1220648-2026-00297 | 23999992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999992&pc=OZD | 12/30/2025 |  |  | 01/07/2026 |  |  | 12/30/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026765841 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Tachycardia | Hematuria | Hemodynamic instability | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1952 | 1220648-2026-00296 | 23999994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999994&pc=OZD | 12/30/2025 |  |  | 01/07/2026 |  |  | 12/30/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026738810 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 76 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it reported that the patient readmitted with purulent drainage from the chest tube site. ct chest was concerning for infection. per app, blood cultures are currently negative. the patient was taken to the operating room for chest washout. the patient survived the procedure and was successfully weaned at postoperative day 5.5 status post aortic, mitral, and tricuspid valve replacements (avr/mvr/tvr). transferred to the unit on a low-dose epinephrine infusion for right ventricular support. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. (infection): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1953 | 1220648-2026-00298 | 23999997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23999997&pc=OZD | 12/30/2025 |  |  | 01/07/2026 |  |  | 12/30/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026790285 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Anemia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1954 | 1220648-2026-00300 | 24000000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000000&pc=OZD | 12/30/2025 | 01/12/2026 |  | 01/07/2026 |  |  | 12/30/2025 | 01/07/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026791186 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1955 | 1220648-2026-00303 | 24000080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=24000080&pc=OZD | 12/19/2025 |  |  | 01/07/2026 |  |  | 12/18/2025 | 01/07/2026 |  |  |  |  | 01/07/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026764257 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 88 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported during high risk percutaneous coronary intervention (hrpci), short 14 french scale (fr) peel away was inserted, left ventricular (lv) access obtained, 0.18 was exchanged for 0.35 and impella was backloaded. impella cp would not advance past left illiac bifurcation so decision was made to remove impella and intra-aortic ballon pump (iabp). iabp wire went up and it is thought that impella cp cannula displaced plaque and iabp wire went into false lumen. catheter and wire would then not advance up into descending aorta and patient was found to have dissection in left common fem illiac. vascular doctor was called to fix vessel. dissection was mitigated with ballooning and stenting of illiac. impella then used through stent to support patient through percutaneous coronary intervention (pci) for acute myocardial infarction associated cadiogenic shock (amics) patient. the patient was reported to have survived the procedure, and patient outcome was stable. | [5 empty] | Product complaint # (b)(4), investigation summary: no product was returned. difficult to advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy and calcification preventing successful delivery of impella. device history lot: device lot: 1979718. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1956 | 1220648-2025-49686 | 23975119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23975119&pc=PYX | 12/11/2025 |  |  | 01/06/2026 |  |  | 12/11/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026709767 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1957 | 1220648-2025-49647 | 23975131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23975131&pc=OZD | 12/15/2025 |  |  | 01/06/2026 |  |  | 12/15/2025 | 01/06/2026 |  |  |  |  | 01/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026773663 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia while leaning forward and coughing. the arrhythmia resolved when the patient changed position. the pump was confirmed to be in a good position and no intervention was required. the patient was in stable condition and impella support continues. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1958 | 1220648-2025-49644 | 23975142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23975142&pc=OZD | 12/15/2025 |  |  | 01/06/2026 |  |  | 12/15/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601125 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1959 | 1220648-2025-49305 | 23976118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23976118&pc=DQX | 12/11/2025 | 02/10/2026 |  | 01/06/2026 |  |  | 12/11/2025 | 01/06/2026 |  |  |  |  | 01/06/2026 |  | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | GUIDE WIRE 0.018"PTFE COATED |  |  | ABIOMED, INC. |  |  | US |  | 9151126 |  | 0052-3005 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 84 KG | Non Hispanic | Black Or African American | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was admitted for acute myocardial infarction/cardiogenic shock. the patient was placed on extracorporeal membrane oxygenation (ecmo) with an impella cp. the patient was being upgraded to an impella 5.5. the placement went well without complications. post placement, the patient started to become unstable. impella flows dropped and ecmo flows dropped as well. transesophageal echocardiogram showed an acute effusion. the surgeon elected to attempt a pericardial window. a clot was evacuated and the patient was stabilized. after some time, the decision was made to open the patient's chest to check for any source of bleeding. once opened, a small puncture was found on the posterior left ventricle. although it cannot be for sure, it appeared that a wire went through the ventricle in the area of his infarction. sutures were placed and the bleeding was controlled. the chest tube was aced and the chest was closed. at the end of the case, the patient was stable and was sent to the intensive care unit. the surgeon was happy with the outcome. additional information was received. the patient had native heart recovery and the impella 5.5 was explanted. the pump was not saved. this is one of two related reports. this report represents the guidewire. another report will be submitted for the impella 5.5. refer to section h10 for the related report number. | [5 empty] | The guidewire was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The following sections have been updated: h6 clinical code the investigation for the cardiac perforation has been completed. no product was returned. without additional clinical details, the root cause of the perforation issue cannot be determined due to complaint device not returned and lack of clinical details.
1960 | 1220648-2026-00183 | 23977630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977630&pc=OZD | 12/16/2025 |  |  | 01/06/2026 |  |  | 12/16/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026752258 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Renal Failure | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1961 | 1220648-2026-00184 | 23977636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977636&pc=OZD | 12/23/2025 |  |  | 01/06/2026 |  |  | 12/23/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026747630 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1962 | 1220648-2026-00185 | 23977642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977642&pc=OZD | 01/01/2026 |  |  | 01/06/2026 |  |  | 01/01/2026 | 01/06/2026 |  |  |  |  | 01/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754130 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced experienced malpositioning of the pump when it was pulled into the aorta. the pump could not be repositioned so it was replaced with a new impella cp. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product return. device in wrong position: the cause of the device in wrong position was unable to be determined due to insufficient clinical details.
1963 | 1220648-2026-00187 | 23977651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977651&pc=OZD | 01/01/2026 | 01/12/2026 |  | 01/06/2026 |  |  | 01/01/2026 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026792073 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1964 | 1220648-2026-00188 | 23977653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977653&pc=OZD | 01/01/2026 |  |  | 01/06/2026 |  |  | 01/01/2026 | 01/06/2026 |  |  |  |  | 01/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750963 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced significant oozing and a hematoma at the access site. the patient's hemoglobin dropped. a suture and pressure dressing were placed and the patient was transfused multiple units of blood. the patient was in stable condition. | [5 empty] | D6b: added explant date | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1965 | 1220648-2026-00189 | 23977662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977662&pc=OZD | 12/12/2025 |  |  | 01/06/2026 |  |  | 12/12/2025 | 01/06/2026 |  |  |  |  | 01/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756523 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 76 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. vessel noted to be suboptimal prior to placement. patient suffered pulseless electrical activity arrest so decision proceed with placement of impella understanding it would most likely be occlusive. at end of case an angiogram was performed through sheath verifying no flow distal to the sheath. various options discussed, including external fem-fem bypass. given patient was not on any pressers decision made to remove pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1966 | 1220648-2026-00192 | 23977671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977671&pc=OZD | 01/01/2026 |  |  | 01/06/2026 |  |  | 01/01/2026 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703386 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1967 | 1220648-2026-00190 | 23977672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977672&pc=OZD | 01/05/2026 |  |  | 01/06/2026 |  |  | 01/01/2026 | 01/06/2026 |  |  |  |  | 01/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765496 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced a placement signal not reliable alarm. the placement signal appeared intermittently. proper positioning of the pump was confirmed using an echocardiogram and no tubing kinks were found. no patient harm was reported. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. tachycardia: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. placement signal issue: the cause of the placement signal issue was not established as no definitive pump failure pattern was observed and no product was returned for analysis. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B3, date of event, has been updated to jan 5 2026. additional patient information received from the reporter states the patient experienced a short run of ventricular tachycardia that resolved without intervention. h6 codes have been updated to reflect this information.
1968 | 1220648-2026-00191 | 23977675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977675&pc=OZD | 01/01/2026 |  |  | 01/06/2026 |  |  | 01/02/2026 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708688 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1969 | 1220648-2026-00193 | 23977679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977679&pc=OZD | 01/02/2026 | 01/12/2026 |  | 01/06/2026 |  |  | 01/02/2026 | 01/06/2026 |  |  |  |  | 01/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026793853 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 70 KG |  | White | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | Additional information was received indicating that the patient had bleeding at the access site around the sheath. the pump was eventually weaned and explanted, and the patient survived the event. | It was reported that a patient implanted with an impella cp experienced hematuria. the pump was positioned deep and was pulled back into position. impella support continues. |  |  |  |  | Additional information received and h6 codes were updated. | D6a: added implant date. h6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product return. hematuria: the cause of the injury was not determined due to insufficient clinical details. access site adverse event/minor bleed: the cause of the access site adverse event was not determined due to insufficient clinical details. device in wrong position: the cause of the positioning issue was unable to be determined due to insufficient clinical details | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The following sections have been updated: b4, g3, g6, h2, h11 additional information was provided which states that impella cp was explanted and bridged to the impella 5.5. the patient survived the procedure.
1970 | 1220648-2026-00194 | 23977683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977683&pc=OZD | 01/01/2026 |  |  | 01/06/2026 |  |  | 01/01/2026 | 01/06/2026 |  |  |  |  | 01/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766639 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 112 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received. the bleeding was resolved. unfortunately, the customer disposed of the pump so it is not available for evaluation. | It was reported that a patient implanted with an impella cp experienced bleeding at the access site. surgicel and a redressing were used to obtain hemostasis. the patient was in stable condition. |  |  |  |  | B5 additional information was received that was not reported on the initial report that was submitted. d6b explant was added. h6 revised investigation conclusions code as it was previously entered incorrectly on the initial report that was submitted. updated type of investigation code to reflect information that was received. h11 updated additional manufacturer narrative as the device was disposed of. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product return. access site adverse event/major bleed: the cause of the access site adverse event was determined to use related to patient movement as the bleeding occurred after the patient coughed hard. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1971 | 1220648-2026-00195 | 23977690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977690&pc=OZD | 12/31/2025 | 01/12/2026 |  | 01/06/2026 |  |  | 01/01/2026 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026794821 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Restricted Flow rate | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1972 | 1220648-2026-00196 | 23977707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977707&pc=OZD | 01/04/2026 |  |  | 01/06/2026 |  |  | 01/04/2026 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026733475 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1973 | 1220648-2026-00197 | 23977714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977714&pc=OZD | 01/01/2026 |  |  | 01/06/2026 |  |  | 01/01/2026 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026783751 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1974 | 1220648-2026-00198 | 23977719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977719&pc=OZD | 12/23/2025 |  |  | 01/06/2026 |  |  | 12/23/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691888 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1975 | 1220648-2026-00199 | 23977720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977720&pc=OZD | 01/01/2026 |  |  | 01/06/2026 |  |  | 01/01/2026 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719386 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1976 | 1220648-2026-00200 | 23977879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23977879&pc=OZD | 01/01/2026 |  |  | 01/06/2026 |  |  | 01/01/2026 | 01/06/2026 |  |  |  |  | 01/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701038 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a hematoma at the access site. a femstop was placed and the patient received a blood transfusion. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. asae/hematoma/major bleed: the cause of the access site adverse event was user related as bleeding was linked to a procedural factor involving a steep angle of entry. device history review: the device passed all post sterile inspection checks.
1977 | 1220648-2026-00201 | 23978233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23978233&pc=OZD | 12/24/2025 |  |  | 01/06/2026 |  |  | 12/24/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026733364 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1978 | 1220648-2026-00202 | 23978256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23978256&pc=OZD | 12/23/2025 | 02/14/2026 |  | 01/06/2026 |  |  | 12/23/2025 | 01/06/2026 |  |  |  |  | 01/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026769738 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported further information that the hemolysis was treated by repositioning the catheter and running the catheter at a lower p level. | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. bedside rn suspicious of hemolysis due to increased ldh. |  |  |  |  | B5 updated. the investigation is still ongoing. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1979 | 1220648-2026-00203 | 23978284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23978284&pc=OZD | 12/30/2025 |  |  | 01/06/2026 |  |  | 12/30/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026746975 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | Patient Device Interaction Problem | [5 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1980 | 1220648-2026-00204 | 23978294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23978294&pc=OZD | 12/30/2025 |  |  | 01/06/2026 |  |  | 12/30/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026774624 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | Patient Device Interaction Problem |  |  |  |  | Vascular Dissection | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1981 | 1220648-2026-00205 | 23978305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23978305&pc=OZD | 12/24/2025 |  |  | 01/06/2026 |  |  | 12/24/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636118 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Edema | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1982 | 1220648-2026-00207 | 23978308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23978308&pc=OZD | 12/30/2025 |  |  | 01/06/2026 |  |  | 12/30/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732157 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1983 | 1220648-2026-00206 | 23978309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23978309&pc=DQX | 12/29/2025 | 02/25/2026 |  | 01/06/2026 |  |  | 12/30/2025 | 01/06/2026 | [6 empty] | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | GUIDE WIRE 0.018" PTFE COATED |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0052-3005 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1984 | 1220648-2026-00049 | 23978324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23978324&pc=OZD | 12/09/2025 | 01/04/2026 |  | 01/06/2026 |  |  | 12/10/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677583 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Use of Device Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1985 | 1220648-2026-00044 | 23978328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23978328&pc=OZD | 12/09/2025 |  |  | 01/06/2026 |  |  | 12/09/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026780637 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | Connection Problem | Device Sensing Problem | [5 empty] | Cardiogenic Shock | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1986 | 1220648-2026-00208 | 23981504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23981504&pc=OZD | 12/23/2025 | 01/09/2026 |  | 01/06/2026 |  |  | 12/24/2025 | 01/06/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026727346 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1987 | 1220648-2026-00209 | 23981906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23981906&pc=OZD | 12/24/2025 |  |  | 01/06/2026 |  |  | 12/24/2025 | 01/06/2026 |  |  |  |  | 01/06/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026784462 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | Yes | Decrease in Pressure | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after the patient was placed with impella cp and running for approximately 20 mins, doctor trimmed 10mm graft like an impella 5.5 (leaving only some graft outside skin). they did not want impella cp repo sheath in axillary artery, made decision to cut/trim portion of impella cp repo sheath off in order to fit inside graft and not have repositioning sheath in right axillary artery. no issues during this time. graft secured to blue suture collar with suture. shortly after (minutes), received low pp (process pressure) alarms on aic (automated impella controller), purge flow to 30. purge tubing was assessed by a biomedical vad (ventricular assist device) engineer for leaks and found none. blood was then visualized inside sterile sleeve and all the way back at red handle. the doctor suspected to have knicked 9 french impella catheter shaft, purge line, a local team agreed. decision was made by doctor to remove impella cp and replace with another impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1988 | 1220648-2026-00076 | 23956127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956127&pc=OZD | 12/15/2025 |  |  | 01/05/2026 |  |  | 12/16/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719957 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the device had a liquid leak from the glucose side arm. water droplets were observed near the pressure reservoir. there were traces of dripping found inside the case, along with signs of leakage on the exterior. additionally, a droplet of glucose solution was noted at the base of the control unit stand. the patient outcome is still to be determined as the patient remains on support. additional information indicated that the leak stopped after observation. no specific measures were taken and there were no other abnormalities found with the device other than the leak. additional information indicated that the patient expired on (b)(6) 2025. the cause of death was the inability of cardiac function to recover, and the patient was placed to end-of-life care. according to the physician, the death was not related to impella¿s flow measurement (fm). no leaks or malfunctions from the purge side arm have been observed since the flow measurement (fm) event (occurred on (b)(6) and resolved). | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1989 | 1220648-2026-00077 | 23956131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956131&pc=OZD | 12/15/2025 |  |  | 01/05/2026 |  |  | 12/16/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026744012 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1990 | 1220648-2026-00078 | 23956158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956158&pc=OZD | 12/16/2025 |  |  | 01/05/2026 |  |  | 12/16/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026763321 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
1991 | 1220648-2026-00079 | 23956520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956520&pc=OZD | 09/16/2025 |  |  | 01/05/2026 |  |  | 09/16/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740273 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Inaccurate Flow Rate | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
1992 | 1220648-2026-00080 | 23956545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956545&pc=OZD | 11/07/2025 |  |  | 01/05/2026 |  |  | 11/07/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026763601 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
1993 | 1220648-2025-45962 | 23956635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956635&pc=OZD | 07/31/2025 |  |  | 01/05/2026 |  |  | 07/31/2025 | 01/05/2026 |  |  | 01/31/2026 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528689 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1994 | 1220648-2026-00081 | 23956643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956643&pc=OZD | 09/09/2025 |  |  | 01/05/2026 |  |  | 09/09/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683611 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that after the patient was transferred to a different hospital, a kink in the cannula was visualized on transesophageal echocardiogram, restricting flows. the pump was explanted. no patient harm was reported. | [5 empty] | A4 is unknown. the pump was returned for investigation and a kink was observed on the cannula near the motor housing. the cause of the cannula kink could not be determined due to insufficient clinical details. the device passed all post sterile inspection criteria.
1995 | 1220648-2026-00082 | 23956667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956667&pc=OZD | 09/20/2025 |  |  | 01/05/2026 |  |  | 09/20/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699034 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Break | Difficult to Advance | [5 empty] | Foreign Body In Patient | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
1996 | 1220648-2026-00083 | 23956685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956685&pc=OZD | 09/21/2025 |  |  | 01/05/2026 |  |  | 09/24/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1700366 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp encountered purge system failures. the purge cassette was replaced but this did not resolve the issue. the supporting automated impella controller was exchanged to resolve the failures. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1997 | 1220648-2026-00085 | 23956865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956865&pc=OZD | 12/28/2025 |  |  | 01/05/2026 |  |  | 12/28/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026786861 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
1998 | 1220648-2026-00086 | 23956874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956874&pc=OZD | 12/28/2025 |  |  | 01/05/2026 |  |  | 12/28/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026790769 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 114 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Patient presented with present bilateral pedal pulses. impella cp placed in the cath lab successfully. pedal pulses rechecked in the cath lab before departure to intensive care unit (icu). intermittent doppler signals received. extremities had the same temperature at this time. in the icu patient complaining of right leg numbness and right foot felt cold to touch, but the remainder of the leg was warm. vascular surgery and the cath lab called to bedside. both specialties concurred that the cause of change is likely due to patient¿s prior medical history, rather than the impella placement. it was further reported that the patient expired on support however no additional information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
1999 | 1220648-2026-00087 | 23956882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956882&pc=OZD | 12/25/2025 |  |  | 01/05/2026 |  |  | 12/25/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026742699 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
2000 | 1220648-2026-00088 | 23956886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956886&pc=OZD | 12/25/2025 | 02/05/2026 |  | 01/05/2026 |  |  | 12/25/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026772324 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 60 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support it was reported that the patient received one unit of packed red blood cells. no other information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary no product was returned. asae/ major bleed: the cause of the bleed was unable to be determined due to insufficient clinical details. device history lot device lot: 1988163 device history batch subcomponent lot: n/a device history review device sn: (b)(6) passed all post sterile inspection checks.
2001 | 1220648-2026-00089 | 23956888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956888&pc=OZD | 12/25/2025 |  |  | 01/05/2026 |  |  | 12/25/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732697 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2002 | 1220648-2026-00090 | 23956891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956891&pc=OZD | 12/25/2025 |  |  | 01/05/2026 |  |  | 12/25/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732582 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 117 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis and hematuria after repositioning of the pump occurred. the patient was in stable condition. | The complainant reported that the impella was repositioned again, which resolved the pink tinged urine. |  |  |  |  | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary analysis summary: no product return. no logs available hematuria: the cause of the injury was not determined due to insufficient clinical details. mechanical interaction with blood/hemolysis: patient continues to have hemolysis urine after repositioning. the cause of hemolysis was not determined due no product returned and insufficient clinical details provided. device history review the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated. the investigation is still ongoing.
2003 | 1220648-2026-00091 | 23956920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956920&pc=PYX | 12/25/2025 |  |  | 01/05/2026 |  |  | 12/25/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025584362 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Restricted Flow rate | Decrease in Pressure | Device Sensing Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2004 | 1220648-2026-00092 | 23956929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956929&pc=OZD | 12/25/2025 |  |  | 01/05/2026 |  |  | 12/25/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026776832 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
2005 | 1220648-2026-00093 | 23956936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956936&pc=OZD | 12/25/2025 | 02/05/2026 |  | 01/05/2026 |  |  | 12/25/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026775458 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 84 KG | Non Hispanic | Black Or African American | Injury | No | Yes | Use of Device Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanting an impella cp the team forgot to connect to the purge line and prime the pump. a new impella cp was prepared and implanted correctly. no patient harm was reported. | [5 empty] | Investigation summary: priming problem/use against labeling: clinical reported prior to device start, team forgot to connect pump to purge line and pump was not primed prior to entering body. the cause of this issue was determined to be use related as evidenced by clinical details mentioning pump insertion without priming. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot : 1994175. device history batch: subcomponent lot : n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2006 | 1220648-2026-00094 | 23956940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956940&pc=OZD | 12/25/2025 |  |  | 01/05/2026 |  |  | 12/25/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744116 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male | 72 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 was suspected of having hemolysis. the pump was explanted and clots were noted on the inflow and outflow. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Device status. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. d9 has been updated to reflect product was returned for investigation. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect product was returned and investigation is underway. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. mechanical interaction with blood/hemolysis: the clinical details state that the "physician suspected hemolysis. beside echocardiogram obtained. inflow positioned mid-cavity in the left ventricle (lv). outflow well about aortic valve. decision made to bring patient back to the or. device removed, clot noted on both inflow and outflow." there were no data logs returned. the product was returned and there was clot in the cannula and wrapped around the impeller. there were no damages noted of the pump. since the clinical details confirmed good positioning, a clot on explant, and that the patient was stable after the pump was removed along with clot noted in the returned product, the root cause of the mechanical interaction with blood is most likely due to biological material, but the cause of the biological material cannot be established. device history review: impella passed all post sterile inspection checks.
2007 | 1220648-2026-00095 | 23956941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956941&pc=PYX | 12/24/2025 |  |  | 01/05/2026 |  |  | 12/24/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026729264 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2008 | 1220648-2026-00096 | 23956945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956945&pc=OZD | 12/28/2025 |  |  | 01/05/2026 |  |  | 12/28/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation | [7 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Thrombosis/Thrombus | Aortic Valve Insufficiency/ Regurgitation | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission
2009 | 1220648-2026-00097 | 23956947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956947&pc=OZD | 12/24/2025 |  |  | 01/05/2026 |  |  | 12/24/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731887 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female | 48 KG |  | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted due to patient anatomy. no patient harm was reported. | [5 empty] | Product complaint # (b)(4). investigation summary: no product return. failure to advance: the cause of the deliver issue was determined to be patient condition as the patient had anatomical issues and calcification preventing successful delivery of impella. device history lot: device lot: 1961734. device history batch: subcomponent lot : n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks. | A5 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2010 | 1220648-2026-00099 | 23956965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956965&pc=OZD | 12/19/2025 |  |  | 01/05/2026 |  |  | 12/24/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026786174 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Pumping Stopped | Device Damaged by Another Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2011 | 1220648-2026-00100 | 23956972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23956972&pc=OZD | 09/05/2024 | 02/25/2026 |  | 01/05/2026 |  |  | 12/28/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025489185 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 70 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that 24 days after an impella 5.5 was explanted a mobile thrombus-like structure was found in the patient's ascending aorta. the thrombus was removed. the patient was in stable condition. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction b3. date of event updated.
2012 | 1220648-2026-00101 | 23957187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957187&pc=OZD | 10/28/2025 |  |  | 01/05/2026 |  |  | 10/28/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754827 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Failure to Advance | Device Damaged by Another Device | [5 empty] | Low Blood Pressure/ Hypotension | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
2013 | 1220648-2026-00102 | 23957196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957196&pc=OZD | 10/28/2025 |  |  | 01/05/2026 |  |  | 10/28/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764203 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
2014 | 1220648-2026-00108 | 23957229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957229&pc=OZD | 10/27/2025 |  |  | 01/05/2026 |  |  | 10/27/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026747644 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was being implanted with an impella 5.5 for mechanical circulatory support after decompensating in the hospital. the impella 5.5 was attempted to be implanted via the right axillary artery, but the 0.035 guidewire was unable to pass through the iliac crest. another attempt was made with the 0.018 glidewire and judkins right 4 diagnostic catheter but was unable to pass into the innominate artery, and the axillary artery became occluded. the right axillary insertion was aborted, and the left axillary was attempted. the 0.035 guidewire was able to pass the innominate artery; however, it was unable to cross the aortic valve. the team switched equipment to a multi-purpose catheter and a lundquist extra-stiff wire and was able to cross the aortic valve under fluoroscopy. the lundquist wire was backloaded onto impella 5.5, however the impella 5.5 was unable to pass through the innominate artery due to calcification. the patient had a poor prognosis and was noted to not be candidate for a durable left ventricle assist device or heart transplant. the medical team decided that they would abort the impella 5.5 implant and attempt to implant an impella cp. the impella cp was prepped and backloaded on to the wire and was unable to pass innominate artery and the implant was aborted. patient outcome is unknown as outcome information was not available in the complaint record accessible for mdr assessment at time of reporting. this complaint involves two automated impella devices: pump 1: impella 5.5, with serial number (b)(6) pump 2: impella cp, with serial number (b)(6) this mdr specifically addresses pump 1: impella 5.5, with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | A4 is unknown. investigation summary no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy near aortic valve and resistance at axillary artery preventing successful delivery of impella. device history review device (sn: (b)(6)) passed all post sterile inspection checks.
2015 | 1220648-2026-00105 | 23957236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957236&pc=OZD | 10/27/2025 |  |  | 01/05/2026 |  |  | 10/27/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760611 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2016 | 1220648-2026-00107 | 23957237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957237&pc=PYX | 10/27/2025 |  |  | 01/05/2026 |  |  | 10/27/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026741998 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2017 | 1220648-2026-00109 | 23957240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957240&pc=OZD | 10/27/2025 |  |  | 01/05/2026 |  |  | 10/27/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026745758 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male | 96 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Inflammation | Hematuria | Multiple Organ Dysfunction Syndrome | [9 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was escalated to an impella 5.5 from an impella cp for mechanical circulatory support. the impella was implanted with no complications. on support day 6, it was reported that the patient¿s haptoglobin was <20 mg/dl and the patient experienced hematuria. the physician noted that there was early onset gross hemolysis. the impella device position was reported to be adequate, however it was suggested to confirm position to minimize shear-related hemolysis. the patient was noted to have transaminitis and hyperbilirubinemia. additionally, plasma-free hemoglobin is elevated. the medical team continued to monitor of liver function tests, hemolysis markers, and clinical symptoms to assess progression. on support day 8, it was reported that the patient had chest swelling at the site of the impella insertion. it appeared to be stable and was monitored for changes. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. additional information regarding the impella cp was not available in the complaint record accessible for mdr assessment at time of reporting. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: corrected the medical device problem code to a01. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned for investigation. no data logs available. organ failure/hematuria/inflammation: the cause of the injury was not determined due to insufficient clinical details. hemolysis: hemolysis reported while on support with elevated plasma free hemoglobin (pfhg). pump position adequate. the cause of the hemolysis was not determined due to lack of returned product and data logs. device history review: device passed all post sterile inspection checks.
2018 | 1220648-2026-00110 | 23957246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957246&pc=OZD | 10/25/2025 |  |  | 01/05/2026 |  |  | 10/25/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | +1(978)739-0590 |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700169 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was escalated to an impella 5.5 from an impella cp. during support of impella 5.5 on (b)(6) 2025, critical care team was concerned patient may have gastrointestinal bleed (gi bleed) and possibly stop systemic heparin this evening. systemic heparin was turned off as patient continues to have melena, bloody gastric content and bloody sputum. multiple blood products were administered. care was withdrawn, patient expired. additional information regarding the impella cp is unknown | [5 empty] | A4 is unknown: b2. date of death information was not available in the complaint record accessible for mdr assessment at time of reporting. investigation summary: major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
2019 | 1220648-2026-00112 | 23957252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957252&pc=OZD | 10/24/2025 |  |  | 01/05/2026 |  |  | 10/24/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026740908 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in cardiogenic shock was escalated from an impella cp to an impella 5.5 for care escalation. the impella 5.5 was successfully via the right axillary artery, and the impella cp was explanted without complications. on day 6 of impella 5.5 support, it was reported that the patient experienced hematuria and melena. systemic anticoagulation was withheld. a colonoscopy was performed and revealed the patient had ischemic colitis from myocardial infarction. additional information regarding the impella cp is unknown, and the final patient outcome is unknown. | [5 empty] | A4 is unknown. investigation summary: no product was returned for investigation. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review device with sn: (b)(6) passed all post sterile inspection checks.
2020 | 1220648-2026-00111 | 23957254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957254&pc=OZD | 10/24/2025 |  |  | 01/05/2026 |  |  | 10/24/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501619 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
2021 | 1220648-2026-00114 | 23957258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957258&pc=OZD | 10/23/2025 |  |  | 01/05/2026 |  |  | 10/23/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763828 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. the patient was in the cardiac catheterization laboratory for a percutaneous coronary insertion, during the procedure the patient decompensated, and the decision was made to place an impella cp for bailout. the short peel away sheath was successfully placed, however while attempting to insert the pump via the guidewire, the guidewire access was lost. both guidewire and pump were removed, and the peel away sheath was replaced with a long peel away sheath. the second attempt was also unsuccessful, and it was noted that the patient had severe aortic stenosis. the medical team decided to abort the impella cp implant and attempt to place an intra-aortic balloon pump. no patient harm was reported in relation to this event. no additional information was provided, and the final patient outcome is unknown. | [5 empty] | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2022 | 1220648-2026-00116 | 23957268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957268&pc=OZD | 10/15/2025 |  |  | 01/05/2026 |  |  | 10/15/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026695301 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 62 KG | Hispanic | White | Injury | Yes | Yes | False Alarm | Inaccurate Flow Rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 after decompensating following removal of impella cp for mechanical circulatory support. the patient presented to the emergency department on (b)(6) 2025 with complaints of abdominal pain with nausea and vomiting for the past 3 days. an electrocardiogram was not immediately done and delayed by 3 hours, and revealed the patient was in st elevation myocardial infarction. the patient was then brought to the cardiac catheterization laboratory where an impella cp was placed via the right femoral artery without complications and supported a successful percutaneous coronary intervention. after the procedure was completed, the impella position was verified, the impella was secured, and the patient was transferred to the medical intensive care unit (micu). upon arrival at micu, a hematoma was noted in the right groin and pressure was held. on day 2 of impella cp support, the right groin was noted to have bruising but no more oozing. epinephrine and lidocaine were injected into the right groin impella insertion site the previous day and pressure dressing was applied to resolve the hematoma and bleeding. the patient¿s ejection fraction was noted to be 10% and the medical team planned on escalating the patient to an impella 5.5 in the next day or two. on day 3 of impella cp support, the patient experienced hematuria that began the previous night. of note, no laboratory samples were noted to have been hemolyzed, and a plasma free hemoglobin was drawn and sent out for analysis. echocardiogram revealed the impella was at an adequate depth of 3.7cm, and clear of the anterior mitral valve leaflet, but positioned more towards the septum. additionally, it was noted that the patient had a fever. on day 4 of impella cp support, it was noted that the patient¿s hematuria was improving and had downward trending lactate dehydrogenase. the patient was transfused with a unit of prbcs for a downward trending hemoglobin and hematocrit with no signs or symptoms of bleeding. systemic anticoagulation therapy was withheld for hematuria and transfusion requirements. the patient was weaned from the impella cp and explanted. overnight, the patient developed a fever and worsening shortness of breath, with questionable aspiration. the patient required increased dobutamine, and the decision was made to place an impella 5.5 for further support. the impella 5.5 was successfully implanted via the right axillary artery with no complications. the patient was transfused with two units of packed red blood cells (prbcs) during the procedure. on day 1 of impella 5.5 support, there were noted alarms with purge flow increasing and decreasing. the purge cassette was replaced as a result. the patient was noted to be moderately hypertensive, which resulted in decreased pump flows. on day 6 of impella 5.5 support, it was noted that the patient was positive for heparin induced thrombocytopenia. on day 8 of impella 5.5 support, the medical team attempted to wean the patient from impella 5.5, however the patient did not tolerate performance (p) level 2, and it was noted that the patient experienced the same issues during impella cp wean. on day 28 of impella 5.5 support, the nurse called the abiomed customer support center concerned with the impella stopping and requested information on what signs to look for when the pump may stop pumping. it was noted that the medical team disabled placement signal monitoring due to false impella in aorta alarms. due to pump saturation and concerns for pump stop, the medical team made the decision to replace the impella 5.5 which was placed via right axillary artery with new impella 5.5 placed via left axillary artery. the new impella 5.5 was placed without complications, and the previous impella 5.5 was explanted successfully. product has been returned; therefore, the impella 5.5 has been explanted. a request for additional information regarding explant date and patient outcome will be made. this complaint involves three impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). purge cassette used with impella 5.5, with unknown serial number. this mdr specifically pump 2: impella 5.5, with serial number 603001, which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Investigation summary: impella 5.5 #603001 and purge cassette were returned for investigation. clinical details report false "position in aorta" alarms on (b)(6) 2025. at the same time, saturated flows were noted. pump position was confirmed with echocardiogram. data logs were reviewed and confirmed the reports. on the reported date, there was a sudden motor current (mc) spike to 1200 ma with a ms deviation followed by a drop in mc modulation, saturated pump flows, and "position in aorta" alarms. there was another subsequent mc spike, and at that time, flows returned to specification and the positioning alarms resolved. this pattern is indicative of biomaterial ingestion. returned product was investigated and significant biomaterial was found around the outflow cage and impeller. based on the location of the biomaterial and the fact that pump flows returned to normal in the field, there is potential that this was picked up on explant. that being said, borescope images show some material in the rotation path of the impeller, so some of it could have been present during the case. the pump passed laser continuity testing and all 5s parameters were within the expected range when connected to a console. no other abnormalities were noted. the pump was run in a bucket of water and pump flows were initially low/variable. the pump was stopped and as much biomaterial was removed as possible, and pump flows got progressively closer to the p-9 specification. false alarm: returned product confirmed no issues with the optical sensor system via 5s and laser continuity testing. based on the patterns in data logs supported by the confirmation in clinical details that the pump's position was confirmed with imaging, the cause of the false alarm was determined to be biomaterial ingestion. saturated flow: based on the patterns noted in data logs, the cause of the saturated flows was determined to be biomaterial ingestion. thrombocytopenia/major bleed: clinical details report that packed red blood cells were given during the insertion procedure. additionally, the patient was heparin induced thrombocytopenia positive on (b)(6) 2025. the cause of the injuries could not be determined due to insufficient clinical details. device history review: per qn (b)(4), pump sn (b)(6) had a particle inside the tray. mrb dispositioned the pump as rework. pump sn (b)(6) passed all post-sterile inspection checks.
2023 | 1220648-2026-00115 | 23957269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957269&pc=OZD | 10/19/2025 |  |  | 01/05/2026 |  |  | 10/19/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026712877 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Malposition of Device | [5 empty] | Hematuria | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in st elevation myocardial infarction /cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient presented to an outside hospital with complaints of chest pains for the previous two days. the patient was diagnosed with non-st elevation myocardial infarction that evolved into st elevation myocardial infarction. prior to being transported to the complainant hospital the patient became unresponsive after being disconnected from the monitor. a precordial thump was performed and the patient regained consciousness. the patient was successfully transferred and brought to the cardiac catheterization laboratory for treatment. following procedures to address the left anterior descending coronary artery, the patient was transferred from the table to a bed and after the transfer the patient decompensated. the patient was transferred back to the table, and the coronary arteries were observed again with no blockages found. an emergent valvuloplasty was performed, and the patient was sent to the unit for recovery. the patient did well overnight but decompensated the following day where it was reported that the patient became diaphoretic, oxygenation saturation dropped, and the patient developed pulmonary edema. the patient was then taken to the ccl for impella placement. the impella cp was placed via the right femoral artery without complications, and the patient was transferred back to the unit with no issues. on support day 2, the physician performed an echocardiogram and found the impella to be deep. the physician pulled the impella back to reposition, but the impella migrated too far back and crossed the aortic valve. the physician was able to advance the pump back wirelessly into the left ventricle. after repositioning, there were noted suction alarms. the performance (p) level was decreased to p6. the physician pulled the device back another 2cm and a 250ml bolus of normal saline was ordered for the patient. it was additionally noted that the patient¿s urine output was red but clearing. the physician believed the cause of the red urine was from a traumatic foley insertion. on support day 3, the patient experienced more shortness of breath that began overnight and worsening conditions. the pump continued to have frequent suction alarms overnight and it was noted that the patient¿s fluid status was more than adequate. the medical staff confirmed that the pump did not have suction alarms until it was repositioned the previous day. it was recommended to the medical staff to perform additional imaging to confirm the impella placement. on support day 4, the patient was intubated due to worsening patient condition. an echocardiogram was performed and the impella was observed to be in optical position despite ongoing suction alarms. the p level was decreased to p5. on support day 5, it was noted that the patient was no longer making urine. lasix was ordered. the patient then declined overnight, and the decision was made not to escalate the patient¿s care due to lack of advanced therapy options. care was withdrawn, and the patient expired while on impella cp support. | [5 empty] | Although the patient's presenting condition may be more likely have impacted the event the impella pump cannot be excluded as a possible contributing factor in the patient's demise. a4 is unknown. investigation summary: no product was returned for investigation. data logs were returned for analysis. data log analysis confirmed suction alarms occurring at multiple instances throughout the case, with the majority of the alarms occurring towards the beginning of the case, from 8:45 am to 10:20 pm on (b)(6) 2025. just a few minutes before the onset of this instance of suction alarms, impella position in aorta alarm was triggered around 8:40 am and resolved within a minute. a purge pressure spike outside of purge bag changes was seen around 7:40 am on (b)(6) 2025, and a gradual rise in motor current was seen. a motor current spike was seen around 6:50 pm on (b)(6) 2025, which coincided with a sharp rise in purge pressure, drop in purge flow, and transition from low pump flow to saturated pump flow for the same p-level. suction alarms occurred towards the end of the case around 7pm on (b)(6) 2025 and coincided with a drop in motor current and pump flow for the same p-level. placement signal low alarms were seen triggering at the same time as this instance of suction alarms. impella position unknown alarm was also seen around 6:12pm on the same day but resolved with the same minute. no other abnormalities were seen. pump suction: clinical details mention that the patient suffered from suction alarms after repositioning was done on (b)(6) 2025 and therefore the pump was unable to be run above p-6. per clinical information on (b)(6) 2025, there were no suction alarms prior to repositioning. on (b)(6) 2025, it was noted from clinical details that the suction alarms kept triggering despite positioning was deemed optimal under echocardiography observation. data log analysis confirmed the suction alarms, along with the observed trends in motor current, purge and pump flow, and purge pressure. the root cause of the pump suction was not determined due to inconclusive evidence from the data logs and clinical details for suction alarms occurring on (b)(6) 2025. device in wrong position: clinical details reveal that the patient suffered from severe aortic stenosis and was diagnosed with non-st elevation myocardial infarction which evolved into a st elevation myocardial infarction prior to impella insertion. during transfer of the patient from the table to the bed, the patient decompensated. aortic stenosis in the patient worsened. after impella insertion, per echocardiology on (b)(6) 2025, the pump appeared to be deep and was pulled back. impella popped back across the valve, but the doctor was able to advance it back in the left ventricle (lv). the patient was also administered fluid boluses to resolve suction alarms and suffered from worsening respiratory status and low ejection fraction (ej) values. the root cause of the positioning issues was not determined based on the provided clinical details, as it was not clear whether the patient's anatomical issues contributed to the positioning issues. hematuria: clinical details mention the patient producing dark red urine output on the day of implant, which the doctor attributed to a previous traumatic foley insertion. the cause of the injury was not determined due to insufficient clinical details. device history review: pump #(b)(6) passed all post-sterile inspection checks.
2024 | 1220648-2026-00118 | 23957990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23957990&pc=OZD | 12/10/2025 |  |  | 01/05/2026 |  |  | 12/22/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685874 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2025 | 1220648-2026-00119 | 23958177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958177&pc=OZD | 12/20/2025 |  |  | 01/05/2026 |  |  | 12/22/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026770691 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Female | 51 KG | Non Hispanic |  | Malfunction | No | Yes | Crack | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon insertion of an impella cp the catheter cracked. the device was removed and a new impella was successfully implanted. no patient harm was reported. | [5 empty] | D4: added device information: h6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product returned. pd-any device-(crack): the cause of the product damage was unable to be determined as insufficient clinical details were provided. failure to advance: the cause of the failure to advance was unable to be determined as insufficient clinical details were provided. device history review: the pump passed all the post sterile inspection checks. | A6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2026 | 1220648-2026-00120 | 23958194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958194&pc=PYX | 12/22/2025 |  |  | 01/05/2026 |  |  | 12/23/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1693068 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
2027 | 1220648-2026-00122 | 23958199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958199&pc=OZD | 12/20/2025 | 01/09/2026 |  | 01/05/2026 |  |  | 12/20/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482055 |  | 000000000001000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
2028 | 1220648-2026-00121 | 23958204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958204&pc=OZD | 12/23/2025 |  |  | 01/05/2026 |  |  | 12/23/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768578 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2029 | 1220648-2026-00123 | 23958208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958208&pc=OZD | 12/20/2025 | 01/02/2026 |  | 01/05/2026 |  |  | 12/20/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591192 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 49 YR |  | Female | 101 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Increase in Pressure | Malposition of Device | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support. the complainant reported pump in aorta alarms and placement signal with high pressures in the 200s. the motor current and flow were unchanged. the doctor stated they were reluctant to reposition the pump. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the pump was not returned for investigation. data log was not provided or retrieved from the remote server. placement signal issue: clinical details indicate intermittent high placement signal pressures, suspected to be a sensor issue. the cause was not determined without returned product investigation and sufficient clinical details, including data logs. device in wrong position: the cause was not determined without returned product investigation and sufficient clinical details, including data logs.
2030 | 1220648-2026-00124 | 23958222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958222&pc=OZD | 12/20/2025 | 02/11/2026 |  | 01/05/2026 |  |  | 12/20/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750172 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2031 | 1220648-2026-00117 | 23958225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958225&pc=OZD | 12/28/2025 |  |  | 01/05/2026 |  |  | 12/28/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026775473 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2032 | 1220648-2026-00125 | 23958260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958260&pc=OZD | 12/19/2025 | 01/02/2026 |  | 01/05/2026 |  |  | 12/20/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026784406 |  | 000000000001000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2033 | 1220648-2026-00126 | 23958266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958266&pc=OZD | 12/19/2025 | 01/02/2026 |  | 01/05/2026 |  |  | 12/20/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649019 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | [10 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2034 | 1220648-2026-00128 | 23958350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958350&pc=OZD | 12/24/2025 |  |  | 01/05/2026 |  |  | 12/24/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026772315 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 89 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia. the patient was defibrillated and treated with anti-arrhythmia drip. the patient was in stable condition. | The pump is not available for return. patient has passed. the patient was originally admitted with this issue so the pump was not thought be the culprit according to care team. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received. b2 updated to death, however date of death and date of explant were not reported. b5 updated with patient outcome, h1, and h6 updated based on patient outcome.
2035 | 1220648-2026-00129 | 23958354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958354&pc=PYX | 12/28/2025 |  |  | 01/05/2026 |  |  | 12/28/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1635044 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 58 KG | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Bedside nurse reported controller error. it was recommended controller exchange. upon arrival, automated impella controller (aic) had a grey controller error alarm and no placement signal on display. perfusion and bedside nurse proceeded to aic exchange. purge bag & cassette were transferred first and the aic connector. pump went to prior p level (6) & flows to 3.8 (yellow), placement signal not reliable alarm started showing. 5 minutes after, they called saying impella stopped. upon second arrival, a new aic was in the room and everything had been transferred. p level had to be manually input. placement signal was present and no alarms showed. patient remained stable during all aic exchanges. this report is for the first aic. additional report will be sent for the 2nd aic and another for the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E1 added initial reporter phone number as they were omitted from the initial report that was submitted.
2036 | 1220648-2026-00130 | 23958413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958413&pc=OZD | 12/25/2025 | 01/09/2026 |  | 01/05/2026 |  |  | 12/26/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755275 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Failure to Pump | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced a placement signal low alarm and a suction alarm. the impella would not start. another impella cp was used to continue patient support. | [5 empty] | D9 updated; the product has been returned. the investigation is still ongoing. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2037 | 1220648-2026-00050 | 23958506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958506&pc=OZD | 12/10/2025 | 01/05/2026 |  | 01/05/2026 |  |  | 12/10/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026776725 |  | 000000000001000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | [10 empty] | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2038 | 1220648-2026-00067 | 23958603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23958603&pc=PYX | 12/11/2025 | 01/05/2026 |  | 01/05/2026 |  |  | 12/11/2025 | 01/05/2026 |  |  | 09/03/2021 |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1606415 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support there was a ¿communication error¿ with the automated impella controller. no further information was provided. there was no reported harm or injury to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2039 | 1220648-2026-00135 | 23960348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23960348&pc=PYX | 12/15/2025 | 01/02/2026 |  | 01/05/2026 |  |  | 12/16/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1940063 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male | 130 KG | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that aic operated without issue for approximately one hour. the control unit screen then went black for approximately two minutes before returning to the pump position screen. a brief ¿control unit malfunction¿ alarm appeared and immediately cleared. as a precaution, the control unit was replaced. no additional information was provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
2040 | 1220648-2026-00140 | 23960367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23960367&pc=OZD | 12/26/2025 |  |  | 01/05/2026 |  |  | 12/26/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026734319 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Access site of the impella and mac has a notable hematoma that was seen after the peelaway sheath was removed. manual pressure held, femstop was placed and nitro drip was added for blood pressure management. blood pressure was high >170/100/126. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E1 added initial reporter phone number as they were omitted from the initial report that was submitted.
2041 | 1220648-2026-00139 | 23960369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23960369&pc=OZD | 12/16/2025 | 01/02/2026 |  | 01/05/2026 |  |  | 12/17/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026780047 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2042 | 1220648-2026-00141 | 23960381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23960381&pc=OZD | 12/28/2025 |  |  | 01/05/2026 |  |  | 12/28/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1654451 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup
2043 | 1220648-2026-00142 | 23960560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23960560&pc=OZD | 12/28/2025 |  |  | 01/05/2026 |  |  | 12/28/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026779575 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Pump | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
2044 | 1220648-2026-00127 | 23960597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23960597&pc=OZD | 12/20/2025 | 01/05/2026 |  | 01/05/2026 |  |  | 12/20/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026773879 |  | 000000000001000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella cp support. the complainant reported decreasing the p level from p9 to p7 due to hemolysis. labs were hemolyzed, and hematuria was present. the pump had been running at p9 with the patient¿s mean arterial pressure (map) in the mid-80s with pressor support. abiomed recommended to run the pump at the lowest p level with minimal-to-zero pressor support to maintain a map of 60-80. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
2045 | 1220648-2026-00143 | 23961037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23961037&pc=OZD | 12/28/2025 | 01/26/2026 |  | 01/05/2026 |  |  | 12/28/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1187999 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 138 KG |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | During set up process for impella 5.5 error message ¿defective impella¿ occurred. replaced the impella catheter with a new pump and restarted the set up process. during the connection of the second impella catheter the same ¿defective impella¿ message occurred. customer then switched out the console and no further alarms appeared. there was no patient harm. | [5 empty] | Correction: d4 and h4 updated with the correct product information. the investigation is still ongoing. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2046 | 1220648-2026-00144 | 23961046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23961046&pc=OZD | 12/13/2025 |  |  | 01/05/2026 |  |  | 12/13/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026785089 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2047 | 1220648-2026-00145 | 23961150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23961150&pc=OZD | 12/23/2025 | 02/27/2026 |  | 01/05/2026 |  |  | 12/23/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1319205 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the battery had an issue on alarm, and no information about error code. subsequently they switched the automated impella controller (aic), as follows there is a need of replacement of the device. no patient injury was reported. the patient outcome is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9 updated; the product has been returned. the investigation is still ongoing.
2048 | 1220648-2026-00146 | 23962250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23962250&pc=OZD | 12/14/2025 | 02/27/2026 |  | 01/05/2026 |  |  | 12/14/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026792648 |  | IMPELLA 5.5 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Renal Failure | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2049 | 1220648-2026-00147 | 23962304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23962304&pc=OZD | 12/23/2025 | 02/11/2026 |  | 01/05/2026 |  |  | 12/24/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026760425 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Alarm System | Decrease in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2050 | 1220648-2026-00131 | 23963997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23963997&pc=OZD | 12/22/2025 |  |  | 01/05/2026 |  |  | 12/22/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | OZD |  |  | IMPELLA CPSA C9+ SET, OUS |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026757386 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male | 110 KG |  |  | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support the purge sidearm and luer lock was leaking, the resolution for this issue is unknown. it was reported that the patient survived explant, however the patient outcome was expired. the cause of death was hypoxic brain damage. | [5 empty] | Correction: b2 date of death. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2051 | 1220648-2026-00149 | 23963999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23963999&pc=OZD | 12/29/2025 |  |  | 01/05/2026 |  |  | 12/29/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760933 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2052 | 1220648-2026-00148 | 23964001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23964001&pc=OZD | 12/28/2025 |  |  | 01/05/2026 |  |  | 12/28/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026790185 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 77 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support the patient developed a gastrointestinal bleed. family decided to withdraw care and move patient to comfort care. the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the gi bleed.was unable to be determined due to insufficient clinical details.
2053 | 1220648-2026-00137 | 23964195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23964195&pc=OZD | 12/16/2025 | 01/05/2026 |  | 01/05/2026 |  |  | 12/17/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026743972 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
2054 | 1220648-2026-00132 | 23964349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23964349&pc=OZD | 11/21/2025 | 01/05/2026 |  | 01/05/2026 |  |  | 11/21/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763671 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
2055 | 1220648-2026-00138 | 23964350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23964350&pc=OZD | 12/15/2025 | 01/08/2026 |  | 01/05/2026 |  |  | 12/17/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026758652 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2056 | 1220648-2026-00133 | 23964359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23964359&pc=OZD | 11/21/2025 | 01/05/2026 |  | 01/05/2026 |  |  | 11/21/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026777305 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 80 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The user facility reported a 63-year-old male patient whose impella device was weaned and removed at the end of the procedure. a thrombus was noted on the pump, and echocardiography revealed a large left ventricular thrombus. no additional information was provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2057 | 1220648-2026-00150 | 23964866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23964866&pc=OZD | 12/28/2025 |  |  | 01/05/2026 |  |  | 12/28/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512628 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2058 | 1220648-2026-00151 | 23968007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23968007&pc=OZD | 12/20/2025 | 01/02/2026 |  | 01/05/2026 |  |  | 12/20/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732030 |  | 000000000001000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was on impella cp support. it was reported that the patient developed a large hematoma and had their blood pressure dropped. the issue resolved with manual pressure being held. | Information was reported that the patient lost about one unit of blood. blood products were given. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | B2 revised selection as intervention was required. e1 added zip code extension as it was omitted from the initial report that was submitted. e4 added selection as it was omitted from the initial report that was submitted. g1 manufacturer site addr. street line 1 and manufacturer site fax should of been left blank on the initial report that was submitted. g2 added report source company rep as it was omitted from the initial report that was submitted. h6 revised type of investigation code as the device was discarded. added health effect, clinical codes that were omitted from the initial report that was submitted. revised health effect, impact code due to information added in b5. h11 revised as the device was discarded. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root causes for the hematoma, bleed and blood pressure drop were unable to be determined due to insufficient clinical details.
2059 | 1220648-2026-00153 | 23968059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23968059&pc=OZD | 12/19/2025 |  |  | 01/05/2026 |  |  | 12/19/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026743531 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 94 KG | Non Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced hematuria and expired on support. | [5 empty] | Product complaint #(b)(4). investigation summary: product & data were not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the pump sn (b)(6) passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B2 added date of death as it was omitted from the initial report that was submitted. e1 added initial reporter phone as it was omitted from the initial report that was submitted.
2060 | 1220648-2026-00152 | 23968079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23968079&pc=OZD | 12/15/2025 | 02/25/2026 |  | 01/05/2026 |  |  | 12/16/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026753799 |  | 1000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported they are planning on returning the pump. the complainant also reported that no harm or complications to the patient occurred; the pump could be changed without any problems. | The user facility reported a 65 year old female was implanted with a impella cp for support during a high risk cardiac procedure. the pump stopped following high purge pressure and low pump flow. no additional information was provided. |  |  |  |  | B5 updated. the investigation is still ongoing. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Updated information: d4: catalog number. h6: component code changed from g04105 to g07003. the investigation for the high purge pressure has been completed. the cause of the high purge pressure could not be determined as limited clinical details were provided and no product was returned for analysis in attempt to reproduce issue. the cause of the pump stop could not be determined as no product was returned for evaluation and data logs were not conclusive to distinguish between biomaterial ingestion or bearing wear.
2061 | 1220648-2026-00154 | 23970839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970839&pc=OZD | 12/19/2025 |  |  | 01/05/2026 |  |  | 12/19/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026795433 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 118 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [27 empty] | Initial submission | Followup | [8 empty] | The complainant reported that the physician stated there was a potential for left ventricular perforation during insertion, as the patient required a return to the operating room for a washout after being transferred back to the intensive care unit. the impella 5.5 pump was implanted through direct surgical access via the right side of the aorta. the patient later expired following explant. product return status is unknown and pending further evaluation. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. patient device interaction problem/cardiac perforation: the cause of the injury was not determined due to insufficient clinical details.
2062 | 1220648-2026-00155 | 23970840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970840&pc=OZD | 07/10/2025 |  |  | 01/05/2026 |  |  | 12/19/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026691765 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Renal Failure | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
2063 | 1220648-2026-00157 | 23970844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970844&pc=PYX | 12/19/2025 |  |  | 01/05/2026 |  |  | 12/19/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1454040 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm was triggered, and the placement signal waveforms became absent. during this alarm, a secondary alarm state occurred, and the console began alarming for controller error followed by controller failure. both alarms resolved independently. the console maintained flow throughout, the patient remained hemodynamically stable, and the activated partial thromboplastin time was 62. a second automated impella controller was present in the room as backup. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2064 | 1220648-2026-00156 | 23970845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970845&pc=OZD | 12/19/2025 |  |  | 01/05/2026 |  |  | 12/19/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026741042 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2065 | 1220648-2026-00158 | 23970847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970847&pc=OZD | 12/19/2025 |  |  | 01/05/2026 |  |  | 12/19/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026733303 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 66 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was admitted yesterday with a non-st elevation myocardial infarction and an impella cp device was implanted. later that afternoon, the patient underwent an off-pump coronary artery bypass grafting procedure with the impella cp in place. the patient remained stable overnight; however, her hemoglobin and hematocrit levels decreased during the night. this morning, the patient required multiple units of blood. the team does not believe that this drop in hemoglobin and hematocrit is related to the impella device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed and no product was returned. investigation summary ==> ppae (anemia): the cause of the injury was not determined due to insufficient clinical details. ppae (ischemia): the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks.
2066 | 1220648-2026-00159 | 23970848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970848&pc=OZD | 12/16/2025 | 01/09/2026 |  | 01/05/2026 |  |  | 12/19/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026735819 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while speaking with the patient, the patient mentioned experiencing a stroke earlier in the week, which was likely the reason for delaying the ventricular assist device procedure. this was immediately confirmed with the care team. at the time of the suspected stroke, the patient reported right upper extremity weakness. a neurology consult was performed, and the patient underwent a computed tomography scan, which showed no large vessel occlusion in the neck or head. the patient has a history of left ventricular thrombus, but no thrombus was observed at the time of implant or during care. the patient has been anticoagulated throughout impella support, with activated partial thromboplastin time ranging from 80 to 130. the patient has regained function in the right upper extremity but continues to experience difficulty with gripping using the right hand. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: ppae: (stroke): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot device lot: 1969442. device history batch subcomponent lot: 1969442. device history review this pump sn (b)(6) passed all post sterile inspection checks.
2067 | 1220648-2026-00160 | 23970850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970850&pc=OZD | 12/19/2025 |  |  | 01/05/2026 |  |  | 12/19/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719995 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2068 | 1220648-2026-00162 | 23970853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970853&pc=OZD | 12/19/2025 | 02/17/2026 |  | 01/05/2026 |  |  | 12/19/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026787442 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the care team spontaneously received an ¿impella in aorta¿ alarm, accompanied by a flat motor current and a left ventricular waveform completely underlying the aortic placement signal. the team noted that the patient had been turned slightly about thirty minutes earlier but had not moved since. the touhy needle remained secure, and the catheter marking appeared unchanged. echocardiography confirmed that the impella inlet cage was no longer in the left ventricle, although the pigtail appeared to still be across the valve. | [5 empty] | Investigation summary: no product return. device in wrong positioning: impella in aorta alarmed with flat motor current. the patient was turned 30 minutes prior but had not moved since. data log review confirmed the alarm triggered correctly with a drop in motor current and maintained placement signal pulsatility. the cause of device in wrong positioning was most likely due to patient movement based on the report of patient movement prior to the alarm and data log review. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2069 | 1220648-2026-00161 | 23970854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970854&pc=OZD | 12/19/2025 |  |  | 01/05/2026 |  |  | 12/19/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026781030 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Hematoma | Tachycardia | Hematuria | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2070 | 1220648-2026-00163 | 23970855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970855&pc=PYX | 12/19/2025 |  |  | 01/05/2026 |  |  | 12/19/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026639329 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | Device Sensing Problem | Patient Device Interaction Problem |  |  |  | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | [8 empty] | The complainant reported that the patient was placed on rp flex for arrhythmic right ventricular cardiomyopathy as a bridge to transplant and for right ventricular failure. the patient was initially covid-positive and on isolation during implant but was later cleared and became ambulatory. the patient remained stable on room air and dobutamine infusion. on the morning of the event, while raising the head of the bed, the console began alarming ¿impella stopped, motor current high.¿ the alarms were associated with sharp spikes in motor current, intermittent inability to flow above performance level p2, and at one point, the console displayed zero flow at p0 before self-resolving. the pump was gradually ramped back to p5, and tissue plasminogen activator was administered through the purge line as a precaution. throughout the event, the patient remained hemodynamically stable, without pain or shortness of breath, and imaging and transesophageal echocardiography were unchanged. later, the pump was functioning normally at p7 with stable waveforms and no alarms. subsequently, while preparing for a planned heart transplant, the console displayed pulmonary artery placement signals of 130s/100s with calculated mean flows of 0.0 liters per minute. due to the absence of impella connect, trend data was unavailable. the rp flex was removed shortly after for the transplant procedure. | [5 empty] | H6 codes have been updated the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2071 | 1220648-2026-00166 | 23970865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970865&pc=OZD | 12/18/2025 |  |  | 01/05/2026 |  |  | 12/18/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713963 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 94 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient was taken back to the operating room to evaluate the 5.5 access site for a suspected hematoma. interrogation was performed, and the patient returned with adequate hemostasis. the patient continued on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2072 | 1220648-2026-00164 | 23970866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970866&pc=OZD | 12/15/2025 |  |  | 01/05/2026 |  |  | 12/15/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026768413 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that during pump preparation for insertion, the console remained in continuous prime mode. fluid was observed at the outflow, and no leaks were identified. the console was switched to ic2425 with the issue reported resolved with case progress advancing. additional information indicated that hemolysis was attributed to improper positioning, which was corrected, resulting in clearance of urine and hemolysis. the patient experienced significant myocardial stunning following wire perforation during primary percutaneous coronary intervention (ppci), ultimately requiring surgical intervention. the family was consulted regarding transfer and escalation to impella 5.5 for recovery support; however, they declined and opted for comfort care. the patient later expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product returned. data log analysis: insufficient data logs returned, only 7 minutes available on (b)(6) and the pump is not turned on during this time. hemolysis/mechanical interaction with blood: the cause of the hemolysis was use related to improper positioning as the hemolysis resolved after the pump was repositioned per clinical. hematuria: the cause of the injury was not determined due to insufficient clinical details. the complaint pump s/n (b)(6) passed all post sterile inspection checks.
2073 | 1220648-2026-00165 | 23970867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970867&pc=OZD | 12/18/2025 |  |  | 01/05/2026 |  |  | 12/18/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026787424 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | False Alarm | Inaccurate Flow Rate | [5 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
2074 | 1220648-2026-00167 | 23970868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970868&pc=OZD | 12/18/2025 |  |  | 01/05/2026 |  |  | 12/18/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026789156 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | High Blood Pressure/ Hypertension | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
2075 | 1220648-2026-00168 | 23970869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970869&pc=OZD | 12/17/2025 | 02/14/2026 |  | 01/05/2026 |  |  | 12/17/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026759667 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Inaccurate Flow Rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2076 | 1220648-2026-00169 | 23970872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970872&pc=OZD | 12/18/2025 |  |  | 01/05/2026 |  |  | 12/18/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026751040 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient was noted to have bloody urine shortly after foley catheter placement. the nurse indicated this may be due to traumatic insertion. echocardiography showed the impella positioned at 3.6 cm, centered in the left ventricle (lv) and away from surrounding structures. the physician was consulted and decided to monitor the situation for improvement over the next few hours. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2077 | 1220648-2026-00170 | 23970890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970890&pc=OZD | 12/18/2025 | 01/09/2026 |  | 01/05/2026 |  |  | 12/18/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026704597 |  | 0048-0045 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
2078 | 1220648-2026-00171 | 23970891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970891&pc=OZD | 12/18/2025 |  |  | 01/05/2026 |  |  | 12/18/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026750951 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2079 | 1220648-2026-00172 | 23970892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970892&pc=OZD | 12/18/2025 |  |  | 01/05/2026 |  |  | 12/18/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775772 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2080 | 1220648-2026-00174 | 23970901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970901&pc=OZD | 11/18/2025 | 01/05/2026 |  | 01/05/2026 |  |  | 11/18/2025 | 01/05/2026 |  |  | 11/30/2025 |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430133 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 62 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Atrial Fibrillation | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Sepsis | Thrombocytopenia | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 78 year old male that was implanted with a impella cp for support during high risk cardiac procedure. the patient was started on a sepsis protocol due to increased lactate levels, and no imaging was performed overnight. plasma-free hemoglobin was planned to be sent to assess for hemolysis; a prior value was reported as 30, and there were no plans to reposition the device unless signs of hemolysis developed. the patient developed a fever, and antibiotics were initiated. heparin was held, and a hit panel was sent. heparin remained on hold due to bleeding from chest tube sites and a decrease in platelet count. the plan of care later included initiation of amiodarone for atrial fibrillation. no additional information was provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. major bleed: the cause of major bleed was unable to be determined as limited clinical details were provided and no product was returned for review. paroxysmal atrial fibrillation/sepsis: the root cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details.
2081 | 1220648-2026-00175 | 23970903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970903&pc=OZD | 11/28/2025 | 01/05/2026 |  | 01/05/2026 |  |  | 11/28/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026772344 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission
2082 | 1220648-2026-00178 | 23970909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970909&pc=OZD | 12/24/2025 | 01/05/2026 |  | 01/05/2026 |  |  | 12/25/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026795682 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2083 | 1220648-2026-00179 | 23970912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970912&pc=OZD | 12/25/2025 | 01/05/2026 |  | 01/05/2026 |  |  | 12/25/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732196 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 62 KG | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 71-year-old female that was implanted with an impella cp. it was reported that the patient had continuous oozing from the impella insertion site. standard troubleshooting was performed, and the oozing was reported to be resolving per nursing staff. distal pulses were present by doppler. the patient received four units of packed red blood cells overnight. laboratory values and device parameters were reported, and the impella device remained in use at the time of reporting. the patient subsequently expired. no additional information regarding the circumstances of death was provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2084 | 1220648-2026-00180 | 23970913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970913&pc=OZD | 12/25/2025 | 01/05/2026 |  | 01/05/2026 |  |  | 12/25/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026786584 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 78 year old male that was implanted with a impella cp for support during high risk cardiac procedure. it was reported that there was concerns for hemolysis. the care team planned to evaluate the position of both devices and optimize positioning as needed. no additional information was provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2085 | 1220648-2026-00181 | 23970914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970914&pc=OZD | 12/25/2025 | 01/05/2026 |  | 01/05/2026 |  |  | 12/25/2025 | 01/05/2026 |  |  |  |  | 01/05/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755193 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 116 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 58 year old male was implanted with a impella cp for support during a high risk cardiac procedure. it was reported that the patient had bleeding at the groin site requiring transfusion of two units of packed red blood cells. the perclose device was tightened, and the swan catheter was repositioned. the bleeding was reported to be slowing. no additional information was provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2086 | 1220648-2026-00182 | 23970916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23970916&pc=OZD | 12/14/2025 | 02/27/2026 |  | 01/05/2026 |  |  | 12/17/2025 | 01/05/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711662 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
2087 | 1220648-2026-00035 | 23954140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954140&pc=OZD | 12/09/2025 |  |  | 01/04/2026 |  |  | 12/09/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026765495 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that there was a bleeding from the axillary site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the reported event, bleeding, could not be determined due to insufficient clinical details.
2088 | 1220648-2026-00036 | 23954141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954141&pc=OZD | 12/09/2025 | 01/29/2026 |  | 01/04/2026 |  |  | 12/09/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750790 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2089 | 1220648-2026-00028 | 23954152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954152&pc=OZD | 12/08/2025 |  |  | 01/04/2026 |  |  | 12/08/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026775783 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2090 | 1220648-2026-00030 | 23954163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954163&pc=PYX | 12/09/2025 |  |  | 01/04/2026 |  |  | 12/09/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025489625 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the right percutaneous device was delivered via the left femoral vein and advanced through a very large right atrium. upon removal of the peel-away sheath, the device migrated back across and could not be re-crossed. no patient harm was reported. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product was returned. device in wrong position: the cause of the device in wrong position was determined to be an unintentional use error since the pump came out of position upon peel away removal. device history review: device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2091 | 1220648-2026-00038 | 23954180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954180&pc=OZD | 12/09/2025 |  |  | 01/04/2026 |  |  | 12/09/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603874 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  | White | Malfunction | No | Yes | Shipping Damage or Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that upon opening the impella catheter, it was observed that the easy-guide lumen was not properly positioned within the pigtail. when attempting to reposition the easy-guide lumen and thread the guidewire, the physician damaged the wire beyond functionality. the team elected to open a new device. no patient harm was reported as the issue occurred prior to use. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2092 | 1220648-2026-00039 | 23954186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954186&pc=OZD | 12/09/2025 |  |  | 01/04/2026 |  |  | 12/09/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1329220 |  | 0042-0000-AU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that upon insertion of the impella plug into the automated impella controller, a ¿controller failure¿ alarm occurred with instruction to switch to the backup console. the team switched to the backup console, after which the device functioned appropriately. no additional information was provided. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
2093 | 1220648-2026-00040 | 23954196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954196&pc=OZD | 12/09/2025 | 01/03/2026 |  | 01/04/2026 |  |  | 12/09/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026742636 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Perforation of Vessels | Cardiac Perforation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
2094 | 1220648-2026-00041 | 23954244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954244&pc=OZD | 12/09/2025 |  |  | 01/04/2026 |  |  | 12/09/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026779623 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient had a ventricular tachycardia episode and was being medically managed. | [5 empty] | Investigation summary: the investigation has been completed. peri-procedural adverse event (tachycardia): in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
2095 | 1220648-2026-00043 | 23954266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954266&pc=OZD | 11/21/2022 | 01/03/2026 |  | 01/04/2026 |  |  | 12/09/2025 | 01/04/2026 |  |  | 08/31/2024 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250761 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Foreign | Study | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Inflammation | Pain | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
2096 | 1220648-2026-00045 | 23954267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954267&pc=OZD | 12/09/2025 |  |  | 01/04/2026 |  |  | 12/09/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026769340 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2097 | 1220648-2026-00047 | 23954268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954268&pc=OZD | 12/10/2025 | 01/03/2026 |  | 01/04/2026 |  |  | 12/10/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026773860 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female | 105 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was implanted with a impella cp for support during high risk cardiac procedure. hemolysis, with a haptoglobin level less than 10 was reported. no additional information was provided. | Information was received. a chest tube was placed and the patient was going for a computed tomography. blood products were being given. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. hemolysis/mechanical interaction with blood: the cause of the hemolysis was not determined due to lack of product and data logs. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | B5 added information that was omitted from the initial report that was submitted. e1 added zip code extension as it was omitted from the initial report that was submitted. e4 added selection as it was omitted from the initial report that was submitted. h6 updated type of investigation code from not returned to discarded. updated health effect - impact code due to information added to b5. h11 revised as the device was discarded. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
2098 | 1220648-2026-00051 | 23954270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954270&pc=OZD | 12/10/2025 | 01/08/2026 |  | 01/04/2026 |  |  | 12/10/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026784435 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. on the second day of support, it was reported that the patient experienced oozing at the impella access site at the repositioning sheath around the introducer hub valve. it was noted that the patient received a blood transfusion. it was reported that hemostasis was achieved by advancing the repositioning sheath. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: no product returned. asae/major bleed: oozing observed at impella insertion site. the cause of the access site adverse event was not determined due to insufficient clinical details. device history lot: device lot: 1995728. device history batch: subcomponent lot: n/a. device history review: device sn (b)(6) passed all post sterile inspection checks.
2099 | 1220648-2026-00052 | 23954271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954271&pc=OZD | 12/11/2025 |  |  | 01/04/2026 |  |  | 12/11/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026747097 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2100 | 1220648-2026-00053 | 23954273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954273&pc=PYX | 12/11/2025 |  |  | 01/04/2026 |  |  | 12/11/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | +1(978)739-0590 |  |  | ABIOMED, INC. |  |  | US |  | 1739460 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella rp flex for mechanical circulatory support. it was reported there were no complications during the insertion of the impella rp flex. it was reported that during the new case start up on the automated impella controller (aic), there was an optical sensor alert. the medical team followed guidance from the onsite abiomed representative to unplug the power cord of the aic and then replug the power cord. the aic performed a self-check on startup and the alert cleared. prior to removing the impella rp flex guidewire and starting support, the pulmonary artery (pa) placement signal was negative on the aic. the wire was not removed, and the abiomed representative called customer support for further guidance. the decision was made to remove the wire and start support. after support was initiated, the pa placement signal numbers improved and appeared to be as expected. the abiomed representative confirmed the aic readings and waveforms and the pump operated as expected. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. the device is reportedly available for return; however, as of the time of this report the device has not yet been returned for evaluation. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2101 | 1220648-2026-00056 | 23954275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954275&pc=OZD | 12/11/2025 |  |  | 01/04/2026 |  |  | 12/11/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617661 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2102 | 1220648-2026-00054 | 23954278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954278&pc=OZD | 12/11/2025 | 01/14/2026 |  | 01/04/2026 |  |  | 12/11/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626238 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Pain | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. on the second day of support, the patient complained of abdominal and back pain. it was reported that the patient had occasional bleeding overnight from the impella access site and the patient¿s hemoglobin was noted to be dropping. the medical team believed the patient to have a retroperitoneal bleed, and a computed tomography (ct) scan was ordered. it is unknown what the results of the ct scan were and if vascular repair was required. it was noted that the patient was stable. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. the device is reportedly available for return; however, as of the time of this report the device has not yet been returned for evaluation. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary no product returned. pain: patient complained of stomach and back pain. the cause of the injury was not determined due to insufficient clinical details. asae/ retroperitoneal hemorrhage: patient had occasional bleeding overnight from her groin. the cause of the access site adverse event was not determined due to insufficient clinical details. device history lot device lot-1905373 device history batch subcomponent lot-n/a device history review the pump sn 563271a passed all the post sterile inspection checks.
2103 | 1220648-2026-00055 | 23954279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954279&pc=OZD | 12/18/2025 | 02/14/2026 |  | 01/04/2026 |  |  | 12/18/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026782246 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
2104 | 1220648-2026-00057 | 23954284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954284&pc=OZD | 12/11/2025 | 02/12/2026 |  | 01/04/2026 |  |  | 12/11/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026779714 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Sepsis | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2105 | 1220648-2026-00058 | 23954285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954285&pc=OZD | 12/15/2025 |  |  | 01/04/2026 |  |  | 12/15/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768606 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female | 70 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Obstruction/Occlusion | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a hematoma at the access site and hypotension requiring levo and a blood transfusion. an arterial occlusion was identified. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. asae/hematoma: a large hematoma developed at lfa access site. the cause of the access site adverse event was not determined due to insufficient clinical information. arterial occlusion/hypotension/major bleed: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2106 | 1220648-2026-00059 | 23954286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954286&pc=OZD | 12/15/2025 | 01/27/2026 |  | 01/04/2026 |  |  | 12/15/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026781168 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
2107 | 1220648-2026-00061 | 23954290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954290&pc=OZD | 12/15/2025 |  |  | 01/04/2026 |  |  | 12/15/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026767273 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Backflow | Restricted Flow rate | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
2108 | 1220648-2026-00060 | 23954291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954291&pc=OZD | 12/15/2025 |  |  | 01/04/2026 |  |  | 12/15/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707491 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2109 | 1220648-2026-00062 | 23954293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954293&pc=OZD | 12/15/2025 | 02/25/2026 |  | 01/04/2026 |  |  | 12/15/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026749457 |  | 1000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Inaccurate Flow Rate | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2110 | 1220648-2026-00063 | 23954294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954294&pc=OZD | 12/15/2025 |  |  | 01/04/2026 |  |  | 12/16/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026781783 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 75 KG | Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced ectopy including non sustained ventricular tachycardia and aberrant sustained ventricular tachycardia with right bundle branch block. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1: added product problem. h6: added code f1904. additional information the customer experienced positional issues where it was repositioned multiple times by at least 2 practitioners over the course of treatment. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. arrhythmia/tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position: the cause of the device in wrong position was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2111 | 1220648-2026-00065 | 23954295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954295&pc=OZD | 12/15/2025 |  |  | 01/04/2026 |  |  | 12/17/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026782143 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced low flow. repositioning of the pump did not resolve the issue. the pump was explanted and a clot was observed in the suction port. the patient is in stable condition. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the root cause of the issues was not determined without returned product investigation and sufficient clinical details, including data logs. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2112 | 1220648-2026-00064 | 23954296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954296&pc=OZD | 12/16/2025 | 01/02/2026 |  | 01/04/2026 |  |  | 12/16/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600808 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported fluctuations on the patient¿s ecg. patient information was unknown. no patient consequences were reported, and the event was noted as a peri-procedural adverse event. no additional information was available. | [5 empty] | B3: corrected the date of the event. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. (arrhythmia): in order to make a cause determination, the clinical detailswould have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: this pump passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. patient information unknown.
2113 | 1220648-2026-00066 | 23954300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954300&pc=PYX | 12/17/2025 |  |  | 01/04/2026 |  |  | 12/18/2025 | 01/04/2026 |  |  | 01/31/2026 |  |  |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025495321 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
2114 | 1220648-2026-00069 | 23954309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954309&pc=OZD | 12/16/2025 |  |  | 01/04/2026 |  |  | 12/16/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026781114 |  | 000000000001000100 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 114 KG |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced impella in aorta alarms. proper placement of the pump was confirmed via echocardiogram. the alarms began when the patient sat upright. the pump was explanted and no patient harm was reported. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2115 | 1220648-2026-00046 | 23954313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954313&pc=OZD | 11/27/2025 | 01/04/2026 |  | 01/04/2026 |  |  | 12/10/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025536784 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was implanted with a impella cp for support during a high-risk cardiac procedure. the care team decided to wean and explant the impella 5.5 due to the patient¿s inability to tolerate systemic intravenous anticoagulation. while on a heparin infusion, the patient experienced oropharyngeal bleeding requiring a blood transfusion and episodes of tachycardia. the device was successfully explanted and the patient was transitioned to an intra-aortic balloon pump. two days later, a small hematoma was noted at the insertions site that looked good. approximately, a week later the patient received one unit of blood. | Information was received. the patient was having oral bleeding and was currently packed. bumex was decreased. right leg arterial duplex was done and was negative. a head and chest computed tomography (ct) was done. the head ct was negative. a soft hematoma was present on the axillary site (nine centimeters (cm) by four and a half cm)). a right chest hematoma around the subclavian with no obvious source of bleeding was noted. the next day the patient went to the operating room for an axillary wash out. the bleeding resolved after the pump was removed. the patient no longer required systemic heparin. the pump is not available for return. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. no data logs returned for analysis. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed & hematoma: the cause of the injury was not determined due to insufficient clinical details. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | B31: revised date of event as it was entered incorrectly on the initial report that was submitted. b5: added information that was omitted from the initial report that was submitted. e1: added initial reporter phone number and zip code extension as they were omitted from the initial report that was submitted. e4: added selection as it was omitted from the initial report that was submitted. g3: date should of reflected 27-nov-2025 on the initial report that was submitted. h6: added health effect - impact codes due to information in b5.
2116 | 1220648-2026-00032 | 23954316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954316&pc=PYX | 09/20/2025 | 01/04/2026 |  | 01/04/2026 |  |  | 12/11/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1603745 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission
2117 | 1220648-2026-00070 | 23954333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954333&pc=OZD | 12/24/2025 |  |  | 01/04/2026 |  |  | 12/24/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604724 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
2118 | 1220648-2026-00071 | 23954342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954342&pc=OZD | 12/23/2025 |  |  | 01/04/2026 |  |  | 12/23/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501284 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 79 KG |  | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that after the impella cp was placed, there was pressure inadvertently placed on repo sheath by the staff causing the repo sheath to be pulled out of vessel approximately one inch and bleeding to the site began. subsequently, the cardiologist returned to advance the repo sheath into the vessel and the impella was locked in place and the device appeared stable on fluoroscopy as well. in the intensive care unit, it was reported that the site continued to bleed and urine became more pink-tinged in color, and no identified alarms were noted on the device. it was observed that the mean arterial pressure (map) was in the 100¿s on placement signal and it was recommended by the cardiologist/intensivist to perform an echocardiogram to confirm placement. the inlet appeared to be towards the papillary muscle on the echocardiogram. the initial measurement was 4.84 centimeters and was retracted to 3.3 centimeters per the intensivist. the intensivist was okay with leaving the device at this measurement after repositioning. it was reported that the patient experienced increase lactate dehydrogenase which resulted to 730 and protein was found with hemoglobin at 30. urine began to gradually clear but was still pink-tinged. ultimately, the decision was made by the cardiologist to pull the device. the patient was reported to have survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
2119 | 1220648-2026-00072 | 23954353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954353&pc=OXD | 12/22/2025 | 02/19/2026 |  | 01/04/2026 |  |  | 12/23/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OXD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1748105 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
2120 | 1220648-2026-00073 | 23954367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954367&pc=OZD | 12/22/2025 | 01/02/2026 |  | 01/04/2026 |  |  | 12/22/2025 | 01/04/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026753361 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2121 | 1220648-2026-00075 | 23954398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23954398&pc=OZD | 12/20/2025 | 01/02/2026 |  | 01/04/2026 |  |  | 12/20/2025 | 01/04/2026 |  |  |  |  | 01/04/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712263 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 93 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Tachycardia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was transferred to another facility and expired on an unknown date. | The complainant reported a patient on impella 5.5 support had a hematoma and oozing at the insertion site with crepitus noted. heparin was not initiated due to the issue, but no intervention was performed. the hematoma and oozing eventually resolved. the patient also experienced small episodes of ventricular tachycardia when coughing. the patient also had right-sided weakness and unresponsiveness so a stroke code was called and the patient was sent for a computed tomography scan. |  |  |  |  | Additional information received that the patient expired on an unknown date. b2, b5, h1 and h6 updated accordingly. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2122 | 1220648-2026-00021 | 23953637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23953637&pc=OZD | 12/24/2025 |  |  | 01/03/2026 |  |  | 12/24/2025 | 01/03/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026785131 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
2123 | 1220648-2026-00023 | 23953638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23953638&pc=PYX | 12/24/2025 |  |  | 01/03/2026 |  |  | 12/24/2025 | 01/03/2026 |  |  |  |  | 01/03/2026 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025562137 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 87 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that staff was alerted to the console display showing ¿0.0¿ on impella flow readings, accompanied by a significant increase in pulmonary artery pressures to 203/187 mmhg, while the pulsatile motor current remained. a stat chest x-ray and transesophageal echocardiogram were performed to confirm inlet and outlet positioning, with color doppler flow used as secondary confirmation of appropriate flow from the outlet above the pulmonic valve. laboratory tests were drawn for trending. the cardiac support center team was consulted, and findings indicated a failed sensor, as the console displayed ¿impella flow unreliable¿ and showed a hashed pulmonary artery signal. the patient remained hemodynamically stable without suction alarms, and laboratory values and urine output confirmed the absence of hemolysis. the patient was well-supported, and the physician was updated in person. a request was made to send the pump to engineering and schedule a meeting with the engineering team to review the device sensor failure. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2124 | 1220648-2026-00025 | 23953639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23953639&pc=OZD | 12/24/2025 |  |  | 01/03/2026 |  |  | 12/24/2025 | 01/03/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | [10 empty] | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2125 | 1220648-2026-00024 | 23953640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23953640&pc=OZD | 12/24/2025 | 02/25/2026 |  | 01/03/2026 |  |  | 12/24/2025 | 01/03/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026755553 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2126 | 1220648-2026-00026 | 23953641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23953641&pc=OZD | 12/24/2025 |  |  | 01/03/2026 |  |  | 12/24/2025 | 01/03/2026 |  |  |  |  | 01/03/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026777891 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 78 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp device was removed from the right femoral artery due to concern for ischemia in the leg. doppler assessment showed pulses present but with monophasic flow. left ventricular function had improved since initial placement, and vasoactive medications were weaned, leading the team to decide on device removal. the pump had already been discarded upon arrival to the unit. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2127 | 1220648-2026-00027 | 23953642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23953642&pc=OZD | 12/24/2025 |  |  | 01/03/2026 |  |  | 12/24/2025 | 01/03/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1401308 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup
2128 | 1220648-2026-00029 | 23953644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23953644&pc=OZD | 12/08/2025 | 01/02/2026 |  | 01/03/2026 |  |  | 12/08/2025 | 01/03/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732666 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
2129 | 1220648-2026-00002 | 23942783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23942783&pc=OZD | 12/14/2025 |  |  | 01/02/2026 |  |  | 12/14/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755504 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 63 KG | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient came into cath lab in cardiac arrest. right common femoral artery accessed via micropuncture technique. 14f peel-away sheath inserted. pigtail catheter and wire advanced into left ventricle and abiomed .018 wire placed into left ventricle. impella cp was advanced into sheath and advanced to the right common iliac artery where it became stuck. the right common iliac artery was then accessed and a 6f sheath placed. iliac angiography performed which showed left common iliac artery calcification. left common iliac artery was treated with plain old balloon angioplasty. at this point the patient coded again and return of spontaneous circulation was unable to be achieved therefore the patient expired. | [5 empty] | The investigation was completed and no product was returned. failure to advance: the cause of the failure to advance was not determined due to insufficient clinical details. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2130 | 1220648-2026-00003 | 23943008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943008&pc=OZD | 12/25/2025 |  |  | 01/02/2026 |  |  | 12/25/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700813 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male | 78 KG | Non Hispanic | White | Injury | No | Yes | Increase in Pressure | Failure to Sense | Device Sensing Problem |  |  |  |  | Low Blood Pressure/ Hypotension | Renal Failure | No Clinical Signs, Symptoms or Conditions | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received. the nurse reported a sudden ¿placement signal low¿ alarm of unclear origin, accompanied by decreasing blood pressure and cardiopulmonary resuscitation with return of spontaneous circulation after 10 minutes. transesophageal echocardiography confirmed correct pump position. the patient is currently stable, and the pump is running without alarms. | Additional information was received indicating that the patient experienced renal failure and required an additional device. | It was reported that a patient implanted with an impella 5.5 experienced high purge pressure. thrombolysis protocol was performed. additionally, a false positioning alarm occurred and there was a discrepancy between the ao and blood pressure reading. the patient is in stable condition. |  |  |  | Additional information was received and h6 codes updated. | Additional information was received and h6 codes updated/corrected. | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure, which triggers the ¿purge pressure high¿ alarm message, could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿.
2131 | 1220648-2026-00004 | 23943031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943031&pc=OZD | 11/24/2025 |  |  | 01/02/2026 |  |  | 12/25/2025 | 01/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026769748 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Myocardial Infarction | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup
2132 | 1220648-2026-00005 | 23943068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943068&pc=OZD | 12/25/2025 |  |  | 01/02/2026 |  |  | 12/26/2025 | 01/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597545 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Cardiac Perforation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission
2133 | 1220648-2026-00006 | 23943103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943103&pc=OZD | 12/19/2025 |  |  | 01/02/2026 |  |  | 12/19/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758922 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 73 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. an echocardiogram showed the pump was behind the papillary muscle so it was repositioned at the bedside and the p level was reduced. additionally, the patient experienced ventricular fibrillation requiring reintubation and vasopressors. the pump was explanted and the patient was supported with venoarterial extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position: the cause of the device in wrong position was unable to be determined due to insufficient clinical details.
2134 | 1220648-2026-00007 | 23943159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943159&pc=OZD | 12/20/2025 |  |  | 01/02/2026 |  |  | 12/20/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527662 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male | 130 KG | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding from the repositioning sheath. the bleed comes out of the sheath between the white sheath and the connection of the blue suture hub. no bleeding from the puncture was noted. the bleeding resolved with out intervention. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2135 | 1220648-2026-00008 | 23943186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943186&pc=OZD | 12/19/2025 |  |  | 01/02/2026 |  |  | 12/19/2025 | 01/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026770418 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup
2136 | 1220648-2026-00009 | 23943250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943250&pc=OZD | 12/14/2025 | 01/27/2026 |  | 01/02/2026 |  |  | 12/14/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1423376 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon being turned on, the automated impella controller (aic) said aic failure. there was no patient harm. | [5 empty] | Investigation summary: data review: the device logs confirmed a controller error (event set #1039) alarm during startup. this alarm is triggered when the light which reaches the ccd sensor array is less than 0.75v, which indicate a malfunction of the optical lamp. the alarm was intermittently triggering. device analysis: the reported controller error was not able to be reproduced upon console start up and testing. perform 5s parameter check, ccd distortion check, and confirmed lamp power output were all within range. conclusion: the root cause of the controller error is likely due to a defective optical bench lamp assembly. device history lot n/a. device history batch subcomponent name: optical bench lamp assembly. material number: 2450-0038. lot number: 1845269. serial number: (b)(6). device history review: q1) service history review - are there any qns associated with the complaint device¿s most recent service report? console ic7317 underwent rework due to a nonconformance according to the fsr 300072386 which was related to the failure mode. q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode? no. q3) complaint history review - have there been any other complaints against this console? yes (b)(4): optical signal issue. phr summary: there were no issues related to the complaint discovered when reviewing the product history of console ic7317. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
2137 | 1220648-2026-00010 | 23943309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943309&pc=OZD | 12/19/2025 |  |  | 01/02/2026 |  |  | 12/19/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026771009 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced an impella in aorta alarm after the patient moved. the impella was repositioned but the alarm did not resolve. an echocardiogram with doppler was performed to confirm correct placement. the patient was in stable condition. | [5 empty] | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2138 | 1220648-2026-00011 | 23943354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943354&pc=OZD | 10/02/2025 |  |  | 01/02/2026 |  |  | 10/02/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026761523 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 64 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient decompensated requiring cardiopulmonary resuscitation while in the cardiac catheterization laboratory for a percutaneous coronary intervention. the medical team decided to implant the impella cp, and after the impella cp was placed, support was initiated however the pump¿s flows were lower than expected at one liter per minute. fluoroscopy was performed to reveal a kink in the impella cannula. multiple attempts were made to reposition the impella, but all were unsuccessful. return of spontaneous circulation was achieved, and the physician decided to replace the impella with a new device. the second impella cp was placed successfully. it was noted that the patient had severe peripheral artery disease and there was no flow to the right lower extremity distal to the impella. additionally, there were suction alarms, and the p-level was reduced to p4. due to ischemia complications, the decision was made to bring the patient to the operating room for an escalation to an impella 5.5. the patient was successfully escalated to the impella 5.5, and the second impella cp was explanted. it was reported that during impella 5.5 support, the patient was given blood products for hematoma at the impella access site. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves three impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella cp with serial number (b)(6). pump 3: impella 5.5, with serial number (b)(6). this report specifically addresses pump 1: impella cp, with serial number (b)(6), which is the subject of this report. separate reports will be submitted for the other devices associated with this complaint. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data review: data logs confirmed low pump flow when pump started and remained to the rest of the case. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. visual inspection found a kink on can about 15 mm from of cage and can bonding site. no other issues were found on the rest of the pump. low or blocked pump flow: the cause of low or blocked pump flow was most likely kinked cannula as per clinical information and observed on the returned product. product damage (cannula kink): the cause of the product damage (cannula kink) was most likely due to patient condition since patient had tavr. device history review: the device passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
2139 | 1220648-2026-00012 | 23943380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943380&pc=OZD | 10/07/2025 |  |  | 01/02/2026 |  |  | 10/07/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711504 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 86 KG |  |  | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Fever | Hemoptysis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to heart transplant. on support day four, the patient experienced episodes of ventricular tachycardia (vt) resulting in a shock from his implanted implantable cardioverter defibrillator (icd). a formal echocardiogram was performed to confirm impella placement, but images were not yet available. the following day it was noted that the patient continued to have vt and the impella pump was found to be flipped back towards the septal wall. the patient was placed on mandatory bedrest. an echocardiogram (echo) was ordered to facilitate impella 5.5 repositioning. despite the clinical events, the patient was reported to be doing well, feeling good, and anxious to mobilize. the impella was repositioned and the following day it was noted that after repositioning the impella, the patient did not experience any further episodes of vt or icd shocks. on support day nine, the patient was noted to have hemoptysis and as a result heparin was withheld. on support day 10, the impella experienced a spike in motor current and a significant drop in purge flow rates (from 10-11 milliliters per hour (ml/hr) down to 4 ml/hr), indicating a potential pump issue, possibly related to thrombus formation. the clinical team elected to initiate treatment with tissue plasminogen activator (tpa) delivered via the purge line rather than systemic heparin due to the patient¿s bleeding risk. additionally, the patient spiked a fever and started on antibiotics. the following day, the patient completed tpa administration and was switched back to bicarbonate purge solution. the hemoptysis was noted to be resolving, and low-dose systemic heparin was initiated. the patient was feeling well, eating, and ambulating two laps. on support day 17, the impella experienced several suction alarms over the last few preceding days and had the patient experienced runs of vt the previous night and during the morning. an echo revealed the pump was positioned deep at just over 5.5cm from the aortic annulus, and close to the septal wall. in consultation with the physician, the pump was repositioned under echo guidance and was pulled back approximately 1 cm and tilted away from the septum. the pump was re-secured with proper three-point fixation. following the repositioning, the pump performance level was successfully increased from p-6 to p-7, and purge flows stabilized at 11-12 ml/hr. the patient's partial thromboplastin time was therapeutic at 57 seconds, and the patient was instructed to rest while awaiting heart transplant evaluation. an echo technician was scheduled to confirm final measurements and images. the patient was successfully weaned from the impella 5.5 during the transplant procedure and underwent a successful heart transplant. the patient survived the procedure. | [5 empty] | A5 and a6 are unknown. investigation summary: the investigation has been completed. fever/hemoptysis: the cause of the injury was not determined due to insufficient clinical details. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. positioning issues: on two occasions during the case, the pump was seen close to the septal wall and required repositioning. the product was not returned for investigation. the cause of the positioning issues could not be determined due to insufficient clinical details and no product returned. pump suction: clinical details reported a number of intermittent suction events around the times where the pump was positioned deep or near the septal wall. the product was not returned. data log review confirms suction alarms triggering from 10/17 - 10/20. prior to the event, a motor current spike did occur, followed by a drop in purge flow, as well as elevated purge pressure and pump flow levels. this was most likely due biomaterial ingested at the impeller purge gap, but these conditions resolved after tissue plasminogen activator (tpa) was put through the purge line, and would not relate to the suction events later on in the case. during the reported suction events, pressure waveforms briefly became aortic, meaning something could have been interacting with the inlet area. there were no changes in motor current leading up to these events, ruling out possible ingestion. on 10/20, an echocardiogram was performed and revealed that the pump was positioned deep and the inlet was close to the septal wall. after the pump was repositioned, suction events resolved for the remainder of the case. based on the data log and clinical details, the cause of the suction was most likely due to improper positioning. purge issues: clinical reported a spike in motor current followed by a drop in purge flows from 10-11ml/hr to 4ml/hr. tpa was started through the purge line which resolved the issue. the product was not returned. data log review confirms the spike in motor current as well as the drop in purge flows right after on 10/13. at the same time as the drop in purge flows, purge pressure also increased from ~450mmhg to 600mmhg. motor speed also deviated down at the same time motor current spiked, which is a signature of ingestion. tpa was started through the purge line, and it's seen in both clinical details and data logs that the purge issues started to resolve. the cause of the purge issues was most likely due to biomaterial ingestion, as a seen by signatures of ingestion before the purge issues occurred, and because it resolved after tpa was started through the purge line. device history review: the pump passed all post sterile inspection criteria.
2140 | 1220648-2026-00013 | 23943432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943432&pc=OZD | 10/10/2025 |  |  | 01/02/2026 |  |  | 10/10/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026699544 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Difficult to Advance | [5 empty] | Anemia | Unspecified Infection | Respiratory Failure | [10 empty] | Death | Required Intervention |  |  |  |  | PROTEK DUO. | VA ECMO. | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the patient had undergone a complex cardiothoracic surgery 10 days prior to impella 5.5 support, which included a quadruple coronary artery bypass graft procedure, aortic valve replacement, and mitral valve replacement. following surgery, the patient developed low cardiac output syndrome and was placed on veno-arterial extracorporeal membrane oxygenation (va ecmo) support. the medical team decided to bring the patient to the operating room for an impella 5.5 placement via the right axillary artery, with a protekduo cannula placed via the right internal jugular vein and remove va ecmo. during the insertion procedure for the impella 5.5, significant tortuosity was encountered, and the initial attempt to advance the device failed to cross the obstruction. the impella 5.5 was removed, a "buddy wire" was placed to facilitate access to the ascending aorta, and the device was successfully reinserted into the left ventricle (lv). the protekduo cannula was placed after the impella 5.5 was in position, and the femoral va ecmo support was subsequently removed. on day seven of support with impella 5.5 and protekduo, the patient was noted to be tachypneic, with a respiratory rate elevated to 30-40 breaths per minute. a bronchoalveolar lavage test returned positive for candida. the patient had elevated leukocytes and high lactate levels which the medical team noted as suggestive of an ongoing systemic infection and metabolic stress. the care team plan was to perform aggressive pulmonary toileting; however, due to persistent tachypnea and the risk of respiratory fatigue, the patient's clinical status was closely monitored for potential reintubation if their condition worsened. the following day the patient required reintubated and received receiving broad-spectrum antimicrobial therapy (avycaz, micafungin, and amphotericin b). the patient continued to receive dialysis and remained volume positive by approximately two liters. on day 12 of support, the patient remained intubated and was on pressure control mechanical ventilation. the patient's lactate levels had significantly improved. an antimicrobial regimen consisting of amphotericin b, avycaz, and micafungin was ongoing. the patient required a transfusion of one unit of packed red blood cells (prbcs) and one unit of platelets the previous day. continuous venovenous hemodiafiltration was in progress with a goal of achieving a net even fluid balance. a point-of-care ultrasound was performed at the bedside in response to episodes of impella suction alarms. imaging confirmed appropriate positioning of the impella device without evidence of migration or malposition. the patient continued to undergo evaluation for heart transplantation. on day 14 of support, there were reports of additional impella 5.5 suction alarms. the physician declined to make any changes to address the suction alarms. the following day, there were additional impella 5.5 suction alarms. it was noted that the patient was receiving continuous renal replacement therapy to pull fluid (crrt), and the crrt decreased slightly. the physician silenced the alarm. it was noted that the patient expired during impella support. it was noted that the patient's care was withdrawn. additional clinical details, including perioperative course, contributing factors, and cause of death, were unavailable at the time of reporting. | [5 empty] | Investigation summary: the investigation has been completed. clinical symptoms (anemia, respiratory failure): the cause of the injury was not determined due to insufficient clinical details. clinical symptoms (unspecified infection): the root cause of the injury was unable to be determined due to insufficient clinical details. difficult to deliver or advance: the patient had tortuous vessel condition and some difficulty was encountered during pump placement, but it was successfully placed on the second attempt. the cause of the difficult delivery of the pump was most likely related to patient condition, based on the given clinical details. pump suction: the pump was not returned for investigation. clinical details report multiple episodes of suction alarms, with brief events noted on 16-oct and ongoing alarms documented on 23-oct and 24-oct. suction alarms occurred while the patient was intubated, on continuous renal replacement therapy (crrt), and receiving volume management. volume was being pulled on crrt, and the care team responded by backing off ultrafiltration and silencing audio for alarms as the patient was scheduled for or. clinical datils confirmed appropriate positioning of the impella device, with no evidence of migration or malposition. data log was available for review from 10-oct to 8-nov. the data log shows several suction alarms from the start of the case on 10-oct through 24-oct, most frequently while running on higher p-levels. pump p-level was reduced to p4 after 24-oct, with resolution of suction alarms. no abnormalities were noted regarding motor current spikes, optical signal issues, or device positioning. the suction issue appeared to resolve with adjustment of volume status and supportive management. the cause of the suction issue was most likely related to patient condition, based on data log review and provided clinical details. device history review: the pump passed all post sterile inspection checks.
2141 | 1220648-2026-00014 | 23943476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943476&pc=OZD | 10/13/2025 |  |  | 01/02/2026 |  |  | 10/13/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700215 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 104 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | [12 empty] | Death | [5 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in st elevation myocardial infarction (stemi)/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was admitted on (b)(6) 2025 with a small bowel obstruction and a non-st-elevation myocardial infarction. following an uneventful small bowel resection, she suffered a stemi and required percutaneous coronary intervention to the left anterior descending coronary artery. the patient coded post-procedure in the catheterization lab, leading to the initiation of extracorporeal cardiopulmonary resuscitation. the patient was escalated to veno-arterial extracorporeal membrane oxygenation (va ecmo) with impella cp therapy via the left femoral artery on (b)(6) 2025 due to cardiogenic shock. after impella cp support was initiated, continuous suction alarms were noted from the impella cp and low flows were noted from the va ecmo. the patient was released to the intensive care unit (icu), and it was noted that va ecmo flows remained low and impella flows were low with suction alarms. the patient also experienced separate, non-impella related mediastinal bleeding requiring surgical intervention. the patient¿s neuro status was to be assessed, and the patient¿s prognosis was noted to be poor. on support day three, the patient¿s ejection fraction was noted around 20-25%. a head computed tomography scan revealed an ischemic stroke. on support day four with impella cp, the patient underwent a successful ramp and decannulation from va ecmo in the operating room, with vascular surgery repairing the right leg vascular injury caused by the va ecmo cannula. the patient was returned to the icu on impella cp support only. a subsequent neurological assessment confirmed a non-recoverable neurological exam due to an anoxic brain injury sustained during her initial cardiac arrest prior to impella support. the patient's care was transitioned to organ procurement and the patient expired after organ harvest. | [5 empty] | Although the patient's presenting condition may be more likely have impacted the event the impella pump cannot be excluded as a possible contributing factor in the patient's demise. investigation summary: the investigation has been completed. no product was returned for investigation. data logs were returned for analysis. data log analysis confirmed suction alarms throughout the case. positioning alarms were seen at the beginning of the case which included impella position unknown and impella position in aorta alarms but were resolved. a placement signal low alarm was also seen at the beginning of the case but was resolved. placement signal remained stable throughout the case with no trends or spikes. motor current showed no spikes or trends outside of p-level changes. no other abnormalities were seen. clinical details mention that the patient was found to have an ischemic stroke per a computed tomography scan during support. suction alarms were seen when the patient's heart rate elevated and on the day of implant. the patient was on extracorporeal membrane oxygenation support during impella support. pump suction: the root cause of the pump suction was not determined due to inconclusive evidence from the clinical details and data log analysis. stroke: the root cause of the stroke was unable to be determined due to insufficient clinical detail. device history review: the pump passed all post-sterile inspection checks.
2142 | 1220648-2026-00015 | 23943509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943509&pc=OZD | 10/21/2025 |  |  | 01/02/2026 |  |  | 10/21/2025 | 01/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657009 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2143 | 1220648-2026-00016 | 23943566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23943566&pc=OZD | 11/29/2025 |  |  | 01/02/2026 |  |  | 12/02/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1346265 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced an increase in purge pressure and unexpected pump stop. correct positioning of the pump was confirmed. the automated impella controller was replaced to resolved the issue and no patient harm was reported. this is one of two related reports. this report represents the automated impella controller. another report was submitted to represent the impella 5.5. refer to section h10 for the related report number. | [5 empty] | A3 and a4 are unknown. the automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2144 | 1220648-2026-00018 | 23945106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23945106&pc=OZD | 12/13/2025 |  |  | 01/02/2026 |  |  | 12/13/2025 | 01/02/2026 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026742666 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup
2145 | 1220648-2026-00019 | 23945130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23945130&pc=OZD | 12/11/2025 |  |  | 01/02/2026 |  |  | 12/11/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026787432 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. on the second day of support, the patient experienced ventricular tachycardia (vt). it was noted that the patient had a history of vt prior to impella implant. subsequent to the onset of vt, the patient was escalated to extracorporeal membrane oxygenation. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the tachycardia was unable to be determined due to insufficient clinical details.
2146 | 1220648-2026-00020 | 23945689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23945689&pc=OZD | 11/26/2025 |  |  | 01/02/2026 |  |  | 12/08/2025 | 01/02/2026 |  |  |  |  | 01/02/2026 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUM |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026655995 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A placement signal not reliable alarm occurred. at the time the alarm occurred, the 5.5 had been in use for more than 30 days. even after the alarm occurred, support continued by turning the alarm off, confirming the pump position with an ultrasound, and managing the patient's hemodynamics with a vital signs monitor.today, on (b)(6), the device was removed. note that within 30 days of the start of support for this 5.5 device, the aic in use (sn: (b)(6) had failed and was replaced with an aic (sn: (b)(6). additional information (12/10/2025) were any measures taken regarding the pump or aic (e.g., restart, replacement, etc.) when the alarm went off? no restart or replacement was performed. this pump had previously had an aic failure alarm on a different aic (reported in (b)(6) ), making it the second aic in use.also, 30 days had passed since the start of pump support and the patient was on the verge of end-of-life care, so the original plan was not to replace the pump. were there any other issues with the pump or aic operation after the alarm went off? after the alarm went off, the pump and aic operated without any problems, except for the position waveform not being displayed. the product has been discarded, but do you have any photos? no what is the patient's current condition? the reason the product itself was not replaced after 30 days was due to a care policy being adopted. the patient subsequently died and was discharged. additional information (12/16/2025) (b)(6)_fm-57609. date of death: (b)(6) final cause of death unknown: pump was not the cause. doctor's comments regarding death (if any): no comments. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2147 | 1220648-2025-49563 | 23933497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933497&pc=OZD | 08/04/2025 |  |  | 12/31/2025 |  |  | 08/22/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685699 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp and ecmo developed an oozing lump. the lump was surgically removed but the patient did not recover and expired from multi organ failure. | [5 empty] | A2 and a4 are unknown. no product was returned for investigation. the cause of the cardiac perforation was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
2148 | 1220648-2025-49564 | 23933511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933511&pc=OZD | 08/27/2025 |  |  | 12/31/2025 |  |  | 08/27/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608793 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | Yes | Device-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Description: the complainant reported a patient noted as high-risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. it was reported that after inserting the impella cp through the 14 fr 25cm peelaway introducer sheath, there was bleeding observed around the impella catheter in the 14 fr sheath. to resolve the bleeding, the repositioning sheath was used by inserting approximately 4cm into the 14 fr sheath. with the repositioning sheath in place, single access was no longer possible for the hrpci and a new access was obtained in the radial artery for the hrpci. it was noted that the patient received frequent doses of fentanyl intravenously because of pain, and that access site closure was difficult due to patient anatomy. it was reported that during closure of the impella access site the angioseal failed and a perclose device was deployed to achieve hemostasis. the patient was successfully weaned and survived the procedure. this complaint involves two impella devices: impella cp introducer 14 fr 25 cm with lot number s9158857 impella cp pump with serial number (b)(6). this mdr specifically addresses impella cp pump with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | The initial mdr was submitted timely for the introducer with mrn 1220648-2025-46855. upon review, it was determined that the pain, medication for pain, bleed, surgical intervention, additional device required, and device-device incompatibility were associated with the pump, not the introducer. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2149 | 1220648-2025-49567 | 23933645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933645&pc=OZD | 09/10/2025 |  |  | 12/31/2025 |  |  | 09/10/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699655 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 62 KG | Hispanic |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | Ischemia | Tachycardia | Respiratory Failure | Hematuria | [7 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support after decompensating during a mitral clip procedure. during the mitral clip procedure, the first clip was placed effectively, but the second clip was ineffective and led to a cord tear. the patient had already been in a tenuous hemodynamic state with an ejection fraction of 15%, and following the tear and subsequent decompensation, the patient was placed on epinephrine, levophed, and milrinone, and then the decision was made to implant the impella cp for hemodynamic support. the impella cp was inserted emergently in the setting of acute on chronic heart failure that deteriorated into cardiogenic shock. after the pump was implanted and support initiated, the patient¿s left leg was assessed for distal perfusion. it was noted there was very sluggish flow with narrow vasculature when viewed with fluoroscopy. the decision was then made to escalate the patient to an impella 5.5 and the patient was transferred to the cardiovascular intensive care unit (cvicu) while surgical consult was in progress and an operating room was prepared. during the transport to the cvicu, the patient became severely hypoxic, acidotic, tachycardia, and hypotensive due respiratory failure. the patient¿s monitor showed a heart rate of >200 bpm in the hallway during transport. of note, the impella flows maintained 3.5 l/min with a mean arterial pressure of >90 throughout the suspected ventricular tachycardia event. the patient was defibrillated three times while completing transport to the cvicu. following stabilization of the arrhythmia, it was noted that the patient was becoming hypoxic and through investigation, the physician found the endotracheal tube (ett) was no longer in the lungs. the eet was emergently replaced at the bedside and vasopressors and inotropes were utilized to stabilize the patient¿s vitals until acidosis could be corrected. a point of care ultrasound showed the impella cp was too deep, and the physician repositioned the impella cp. the patient¿s left leg was noted to be cool on arrival to the bed, and it was also noted that the patient¿s urine turned from yellow to red after arrival to the cvicu and subsequent to the hemodynamic incident. the patient was successfully escalated to an impella 5.5 for expected long duration of support, and for distal limb ischmia concerns due to the patient¿s poor vasculature. the impella 5.5 was implanted without issue, the impella cp was weaned and explanted, and the impella 5.5 support was initiated. following impella 5.5 support, the patient¿s vasopressor requirements were reduced. the following day, the patient¿s urine was noted to be yellow. on the fifth day of impella 5.5 support, the patient experienced a 50-beat run of ventricular tachycardia and it was also noted that the patient had a low hemoglobin of 6.6 g/dl and platelet count of 81,000 per microliter with no mention of cause of events. on the seventh day of support, the patient was transfused with a unit of blood, and it was noted that the patient¿s platelets had improved to 145,000 platelets per microliter., the patient was transferred to another hospital for advanced therapy evaluations. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. this complaint involves two impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). this mdr specifically addresses pump 1: impella cp, with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | A6 is unknown. investigation summary: no product was returned. hematuria, hypotension: the root cause of the injury was unable to be determined due to insufficient clinical details. tachycardia, arrhythmia, ischemia: in order to make a cause determination, additional clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of positioning issue was most likely use issue related since during transport to cardiovascular intensive care unit the endotracheal tube (ett) unknowingly became dislodged requiring emergent ett replacement as point of care ultrasound showed impella cp to be deep. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
2150 | 1220648-2025-49570 | 23933678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933678&pc=OZD | 09/11/2025 |  |  | 12/31/2025 |  |  | 09/11/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | 2000154 |  |  | ABIOMED, INC. |  |  | US |  | 2026724399 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp device experienced ventricular tachycardia. milrinone was discontinued. the patient is in stable condition. | [5 empty] | A4 is unknown. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
2151 | 1220648-2025-49571 | 23933683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933683&pc=OZD | 09/11/2025 |  |  | 12/31/2025 |  |  | 09/11/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720510 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was implanted in a shallow position. the patient was in a shallow position. | [5 empty] | A4 is unknown. no product was returned for investigation. the cause of the delivery issue was determined to be patient condition as the patient had stick which was too low for placement preventing successful delivery of impella. the device passed all post sterile inspection checks.
2152 | 1220648-2025-49573 | 23933704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933704&pc=PYX | 09/14/2025 |  |  | 12/31/2025 |  |  | 09/14/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025487237 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 135 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Multiple Organ Dysfunction Syndrome | Heart Failure/Congestive Heart Failure | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex lost pulsatility in the right and left ventricles. the patient went into renal failure and end organ failure with diffuse shock. care was withdrawn and the patient ultimately expired. | [5 empty] | Medical safety/clinical reviewed the event. the patient was admitted for presumed st elevation myocardial infarction; left/right heart catheterization demonstrated clean coronaries. the patient experienced cardiac arrest and was placed on biventricular mechanical circulatory support, with left-sided impella support implanted via the right femoral artery during cardiopulmonary resuscitation, followed by impella rp flex implantation via the right femoral vein. return of spontaneous circulation was achieved; however, there was no ventricular pulsatility. the patient was transferred to the intensive care unit on biventricular support with maximal vasopressor and inotropic therapy. the patient remained in refractory shock with anuria, no neurological response, renal failure, and progressive multisystem organ failure all known factors of patients in cardiogenic shock. after discussion with the family, further escalation of care was deferred, care was withdrawn, and the patient expired. conservatively, the death will be reported on both impella devices as the impella devices were in use at time of expiration. note: h6: device problem/component/investigation coding remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The impella passed all the post sterile inspection checks. the root cause of the cardiac failurewas unable to be determined due to insufficient clinical details. the cause of renal failure and organ failure was not determined due to insufficient clinical details.
2153 | 1220648-2025-49574 | 23933804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933804&pc=OZD | 09/18/2025 |  |  | 12/31/2025 |  |  | 09/18/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705950 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | False Alarm | Restricted Flow rate | Pumping Stopped | Malposition of Device | Device Sensing Problem |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced suction alarms because the pump was positioned shallow. volume was given and the pump was repositioned to resolve the alarms. there was also optical sensor drift and a pump stop, and the pump could not be restarted after three attempts. the pump was explanted and a new impella 5.5 was implanted to support the patient. | [5 empty] | A4 and a6 are unknown. the impella passed all post-sterile inspection checks. the cause of the positioning issue was not established as limited information regarding when and how device became mispositioned was provided. the cause of the pump suction was determined to be patient condition related as evidenced by suction resolution after volume administration. the cause of the false alarms and pump suction were as a result of the sensor wear evidenced by issue occurrence during drift down of placement signal. the cause of the optical sensor issue was determined to be due to a sensor wear as evidenced by log pattern. this occurred after 22 days of support, which is within the intended design life of 60 days per design traceability matrix 0550-9900. the cause of the pump stop was unable to be determined as no product was returned for analysis.
2154 | 1220648-2025-49565 | 23933824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933824&pc=OZD | 12/01/2025 | 02/11/2026 |  | 12/31/2025 |  |  | 12/02/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026705054 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Backflow | Failure to Pump | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that during a routine procedure, all steps were made from insertion to act measurement. the act was stated to be 250 seconds prior to insertion. after the pump was inserted and activated, two red alarms appeared on the aic screen: ¿impella stopped,¿ ¿retrograde flow,¿ and ¿motor current is high.¿ as per training guidelines, the team attempted to restart the device; however, the alarms persisted. consequently, the pump was removed and replaced with another unit. the patient was reported to have survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Corrected data: h6 updated - a141204 removed. investigation summary: no product and data logs were returned for investigation. pump did not start/retrograde pump flow: the cause of pump did not start and retrograde flow alarm issue was not determined due to product and logs being not available for investigation.
2155 | 1220648-2025-49576 | 23933829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933829&pc=OZD | 09/24/2025 | 02/17/2026 |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681694 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Restricted Flow rate | [6 empty] | Pulmonary Edema | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced suction alarms and flash pulmonary edema. in response, volume was given. later, the patient was successfully weaned from the pump and survived. | [5 empty] | Investigation summary: pulmonary edema: the cause of the injury was not determined due to insufficient clinical details. pump suction: clinical reported that the pump was unable to go above p-1 without suction alarms after placing the patient on va ecmo, and fluids were given. the product was not returned for investigation. data log review shows low ps pulsatility as well as a gradual trend down during the time of suction. after the reported suction events, ps started to increase, which was likely a result of the fluids given during suction. the cause of the suction was most likely related to the patient's condition, as suction started after the patient was put on va ecmo. additionally, ps was low and trending down at the time of suction, but resolved after fluids were given. | H6 updated. | A4 and a6 are unknown. the impella passed all post sterile inspection criteria. the cause of the pulmonary edema was not determined due to insufficient clinical details. the cause of the pump suction was most likely related to the patient's condition, as suction started after the patient was put on va ecmo. additionally, ps was low and trending down at the time of suction, but resolved after fluids were given.
2156 | 1220648-2025-49577 | 23933831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933831&pc=OZD | 09/24/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739737 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female | 54 KG |  |  | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis and shock. in response, the p level of the pump was lowered. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A5 and a6 are unknown. the impella passed all post sterile inspection checks. the pump was not returned for investigation. the cause of the hemolysis issue was not determined without returned product or sufficient clinical information including data log.
2157 | 1220648-2025-49578 | 23933837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933837&pc=OZD | 09/24/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026727149 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 82 KG | Hispanic |  | Death | Yes | No | Inaccurate Flow Rate | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced reduced purge flow and reduced purge pressure. tissue plasminogen activator was added to the purge. the patient expired while on impella support and the cause of death was not known. | [5 empty] | A6 is unknown. the impella passed all post-sterile inspection checks. no product was returned. the cause of the high purge pressure was biomaterial buildup since data log review showed a gradual increase in pp with no mc spikes. the cause of the cardiac arrest cannot be determined as insufficient clinical details were provided.
2158 | 1220648-2025-49579 | 23933843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933843&pc=OZD | 09/24/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026726044 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding. an extra suture was applied to obtain hemostasis. additionally, the patient's leg was ischemic. an additional device was added to perfuse on the opposite leg. later, the patient was successfully weaned from the device and survived. | [5 empty] | The impella passed all post sterile inspection checks. investigation summary: no product was returned for investigation. the cause of the bleeding is determined to be patient condition because of extreme calcification in vessels. the root cause of the ischemia was unable to be determined due to insufficient clinical details.
2159 | 1220648-2025-49580 | 23933844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933844&pc=OZD | 02/27/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610549 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient supported by an impella cp developed an access site bleed after the pump was explanted. two units of red blood cells were transfused and a femstop was placed to obtain hemostasis. the patient survived. | [5 empty] | The impella passed all the post sterile inspection checks. investigation summary: the cause of the bleed was not determined since product was not returned and insufficient clinical details provided.
2160 | 1220648-2025-49581 | 23933845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933845&pc=OZD | 09/24/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676982 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a paitent implanted with an impella 5.5 experienced access site bleeding. four units of red blood cells were transfused to the patient to obtain hemostasis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella passed all post sterile inspection checks. no product was returned for investigation. the root cause of the bleeding/ hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
2161 | 1220648-2025-49582 | 23933862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933862&pc=OZD | 09/24/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555054 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 92 KG |  |  | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced poor pump placement. the pump was repositioned which caused hemolysis. the pump was then explanted and replaced. | [5 empty] | A5 and a6 are unknown. the impella passed all post sterile inspection checks. product was not returned for review. the cause of the hemolysis was most likely due to pump was in improper position with the confirmation from clinical team, echo and data logs. the cause of positioning issue was use related which occurred after transfer.
2162 | 1220648-2025-49583 | 23933868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933868&pc=OZD | 09/24/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699783 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a hematoma at the access site and ischemia of the leg. the pump was explanted and replaced with an impella 5.5 pump. | [5 empty] | The impella passed all post sterile inspection checks. no product was returned for investigation. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the root cause of the ischemia was unable to be determined due to insufficient clinical details.
2163 | 1220648-2025-49569 | 23933872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933872&pc=OZD | 12/13/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/22/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026735076 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | False Alarm | Malposition of Device | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 56 year old male was supported with an impella device while in the intenisve care unit. during use a yellow placement signal not reliable alarm reportedly occured. and the aortic position waveform was no lobger displayed. the situation was monitored and the alarm reportedly resolved without intervention. no adverse patient events were reported and the aic was not replaced. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received therefore d6b was updated.
2164 | 1220648-2025-49584 | 23933878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933878&pc=OZD | 09/24/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656677 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced an access site bleed.a stitch was placed to obtain hemostasis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella passed all post sterile inspection checks. no product was returned. the cause of the bleed was most likely use issue related since adding additional stitch resolved the bleed.
2165 | 1220648-2025-49586 | 23933880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933880&pc=OZD | 09/29/2025 |  |  | 12/31/2025 |  |  | 09/29/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723751 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 116 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced worsening cardiogenic shock and arrhythmias. the patient suffered a cardiac arrest and the decision was made to escalate the patient to and impella 5.5. the patient survived. | [5 empty] | The impella passed all post sterile inspection checks. no product was returned for investigation. the root cause of the bradycardia and ventricular fibrillation was unable to be determined due to insufficient clinical details.
2166 | 1220648-2025-49585 | 23933881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933881&pc=OZD | 09/26/2025 |  |  | 12/31/2025 |  |  | 09/26/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714724 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Stroke/CVA | Fever | Hematoma | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Renal Failure | Ventricular Fibrillation | Thrombosis/Thrombus | Swelling/ Edema | Hemodynamic instability | [27 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced end-stage renal disease and was placed on dialysis. the patient had an infection and fever, and there were concerns of damage to the left ventricle with a thrombus therein. the patient then suffered a hemorrhage stroke and cardiac arrest. cardiopulmonary resuscitation was performed but ventricular tachycardia persisted. after experiencing hemolysis, a stroke, and possible disseminated intervascular coagulation care was withdrawn and the patient expired. | [5 empty] | Correction: device history review: the device passed all post sterile inspection checks. | The impella passed all post sterile inspection checks. investigation summary: data review: data logs showed pump ran without issue during support. there were positioning & suction alarms triggered during the case but resolved. there were no mc spikes, abnormal ps or purge issues observed during support. no product was returned for investigation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of hematoma was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the renal failure/fever/edema/death/hemorrhage was not determined due to insufficient clinical details. the root cause of the infection/thrombosis/stroke/cardiac arrest/ventricular fibrillation was unable to be determined due to insufficient clinical details.
2167 | 1220648-2025-49568 | 23933892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933892&pc=OZD | 12/21/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/23/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026747732 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 47 year old male patient underwent an attempted impella cp insertion. the patient reportedly had hypertrophic cardiomyopathy (hcm), and the device could not be safely or properly placed. the insertion was aborted. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction : section e1 was updated which was inadvertently left blank in the initial report. product complaint # (b)(4). investigation summary ==> the cause of the failure to advance was determined to be patient condition as the patient had hcm preventing successful delivery of impella. device history lot ==> device lot: 1980796 device history batch ==> subcomponent lot: n/a device history review ==> the pump sn (b)(6) passed all the post sterile inspection checks.
2168 | 1220648-2025-49588 | 23933893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933893&pc=OZD | 09/29/2025 |  |  | 12/31/2025 |  |  | 09/29/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700143 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 79 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp had slight ecchymosis and oozing. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella passed all post sterile inspection checks. no product was returned for investigation. the root cause of the bleeding and hematoma were unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
2169 | 1220648-2025-49589 | 23933896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23933896&pc=OZD | 09/29/2025 |  |  | 12/31/2025 |  |  | 09/29/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699631 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 56 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp was suspected of having heparin-induced thrombocytopenia and bleeding. heparin was withheld and replaced with bival. additionally, the patient coded. return of spontaneous circulation was quickly achieved. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella passed all post sterile inspection checks. no product was returned for investigation. the root cause of the cardiac arrest was unable to be determined due to insufficient clinical details. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
2170 | 1220648-2025-49591 | 23934029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934029&pc=OZD | 12/20/2025 | 12/30/2025 |  | 12/31/2025 |  |  | 12/20/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732121 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a 43 year old male was supported with an impella cp device. during use the device was reportedly noted to be kinked near he motor, and reduced flows were observed. troubleshooting was performed, and the patient was reportedly supported through the procedure. reduced flows were attributed to patient anatomy, including a short aortic arch. no additional clinical details were provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: analysis summary: the device was returned and visually inspected, cannula kink near outlet observed. data log review showed drop in flow at p9 after support initiated, followed suctions, ps spikes, mc spike at p6 observed as a result of cannula kink as reported. pd-catheter-(twist, bent, kinked): the root cause of cannula kink due to patient related due to short aortic arch. low or blocked pump flow: device kinked by motor, flows were reduced. trouble shooting but due to patient short aortic arch pump explanted. the root cause of low pump flow issue was cannula kink occurred due to patient short aortic arch. device history review: the pump passed all the post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
2171 | 1220648-2025-49590 | 23934054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934054&pc=OZD | 12/02/2025 |  |  | 12/31/2025 |  |  | 12/03/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727390 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. on support day 7, it was reported that due to frequent episodes of ventricular tachycardia (vt), the patient would undergo a vt ablation during impella cp support. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. additional clinical details were unavailable at the time of reporting. | [5 empty] | A. patient information: patient demographics are unknown. the device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. g1 added manufacturing country code as it was omitted from the initial report that was submitted.
2172 | 1220648-2025-49592 | 23934226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934226&pc=OZD | 10/16/2025 |  |  | 12/31/2025 |  |  | 12/03/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1860421 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was being implanted with an impella device for mechanical circulatory support. following a catheterization of coronary arteries and stent placement, the physician decided to implant an impella device to increase blood flow of the patient. during the setup of the automated impella controller (aic) with the impella device, the aic would not detect the impella catheter. there was no attempt to insert the impella device into the patient. the aic and the impella device were not exchanged for a new or alternate device. the physician decided to abort the impella implant as the patient¿s condition improved and was noted to be well recovered. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the patient's outcome was survived. this complaint involves two impella devices: unknown impella pump automated impella controller with serial number (b)(6) this mdr specifically addresses automated impella controller with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other device associated with this complaint. | [5 empty] | A. patient information: patient demographics are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2173 | 1220648-2025-49595 | 23934240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934240&pc=OZD | 12/18/2025 | 12/30/2025 |  | 12/31/2025 |  |  | 12/18/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694536 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 77-year-old male patient whose urine color changed from yellow to pink while in the intensive care unit. an echocardiogram confirmed the impella device was appropriately positioned. the care team elected to provide intravenous fluids. it was later reported that the impella device position was noted to be in the aorta following patient movement. the device support level was reduced, placement was confirmed by echocardiography, and the device was subsequently removed by the physician. the patient was reported to be stable with stable vital signs. no additional information was provided | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data logs confirm an impella position in aorta alarm at the end of the case, with observed flow saturation and flat motor current. device in wrong position: clinical details indicate that on 19-dec-2025, the impella position was found to be in the aorta after the patient coughed and attempted to sit up. the cause of the positioning issue was most likely related to unintentional patient movement, based on the provided clinical details. hematuria: the cause of the injury was not determined due to insufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2174 | 1220648-2025-49594 | 23934454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934454&pc=OZD | 12/18/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/18/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689041 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 89 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old female that was implanted with a impella 5.5 for support during a high-risk cardiac procedure. it was reported that the patient experienced a significant drop in hemoglobin, which was determined to be related to a gastrointestinal bleed, as evidenced by blood in the stool. all anticoagulants were discontinued, and an endoscopy was planned to be performed later that day. at the time of reporting, the patient was reported to be stable. no additional information was provided. | [5 empty] | Date of explantation, d6b, has been added. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2175 | 1220648-2025-49593 | 23934482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934482&pc=OZD | 12/18/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/18/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026743661 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 70 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Care was withdrawn and the survival outcome at explant is patient expired. | The user facility reported a 70-year-old male patient who was noted to have bleeding from the eyes, nose, and vascular access sites. the bleeding was thought to be related to disseminated intravascular coagulation (dic). the patient was on extracorporeal membrane oxygenation (ecmo), and at the time of reporting, no heparin was being administered. no additional information regarding patient outcome or further evaluation was provided. |  |  |  |  | Correction: section b5 : the patient outcome information was inadvertently left out from the initial report which had been added to this report. the cause of the injury was not determined due to insufficient clinical details. device history review: the pump sn (b)(6) passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2176 | 1220648-2025-49598 | 23934520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934520&pc=OZD | 12/18/2025 |  |  | 12/31/2025 |  |  | 12/18/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026734316 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female | 81 KG |  | White | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Tachycardia | Intraoperative Cardiac Injury | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced malpositioning of the device near the mitral valve. attempts to reposition the device were unsuccessful. the patient experienced ventricular tachycardia which resolved when repositioning of the device was abandoned. after the patient was successfully supported the pump was explanted. | [5 empty] | A5 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2177 | 1220648-2025-49599 | 23934524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934524&pc=OZD | 12/18/2025 |  |  | 12/31/2025 |  |  | 12/18/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731208 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. mechanical interaction with blood: clinical details noted patient had red tinged urine following impella placement. no additional details provided. the cause of the mechanical interaction with blood could not be determined as no product or logs were returned for evaluation. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e1 added initial reporter phone number as it was omitted from the initial report that was submitted. | A6 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2178 | 1220648-2025-49600 | 23934561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934561&pc=OZD | 12/18/2025 |  |  | 12/31/2025 |  |  | 12/18/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714020 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced bleeding in the jachson-pratt drain from the pump pocket. a pressure dressing was placed and one unit of packed red blood cells were transfused. the patient was in stable condition. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. access site adverse event: the cause of the bleed was unable to be determined as insufficient clinical details were provided. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2179 | 1220648-2025-49602 | 23934567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934567&pc=OZD | 12/18/2025 | 01/05/2026 |  | 12/31/2025 |  |  | 12/18/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026751860 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 115 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the stroke was pre-impella and was found in the emergency department with a computed tomography scan before the patient went to catheterization lab. | The user facility reported a 69-year-old male patient who was noted to have blown pupils with concern for anoxic brain injury and/or stroke. no additional clinical details, interventions, or patient outcome information were provided at the time of reporting. |  |  |  |  | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the stroke was unable to be determined due to insufficient clinical details. | B5 updated. the pump is not available for investigation. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2180 | 1220648-2025-49601 | 23934568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934568&pc=OZD | 12/02/2025 | 01/08/2026 |  | 12/31/2025 |  |  | 12/03/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026757975 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male | 109 KG |  |  | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp via the right femoral artery for mechanical circulatory support. the patient presented to the emergency room following an episode of shortness of breath and appearing diaphoretic. the patient was found to be in supraventricular tachycardia in the emergency room, and the medical team attempted to cardiovert the patient with medications. however, the patient became combative and confused. subsequently the patient was intubated, and the decision was made to place an impella cp to stabilize the patient for transfer to another hospital for further evaluation and treatment. upon arrival to the transfer hospital, the patient was found to have a mean arterial pressure of 120 mmhg and subsequently the performance level of the impella was turned down to p5 and the patient started on nipride. it was noted that the patient¿s urine output was red/pink. additionally, the patient¿s laboratory results were reported to be hemolyzed. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | Investigation summary: the pump was not returned for investigation. data log was available from 3-dec 1:14 to 3-dec 15:17. no abnormalities related to mechanical failure of the pump were noted on the available data log. mechanical interaction with blood (hemolysis): clinical details indicate that the patient developed elevated hemolysis labs and red/pink urine following impella cp c9+ implantation. by the time of explant, urine output was greater than 30 cc/hr and yellow. the cause of hemolysis was not determined without returned product and sufficient clinical information, including the full data log. clinical symptoms (hematuria): the cause of the injury was not determined due to insufficient clinical details. device history lot: device batch: 1979760. device history batch: subcomponent lot: na. device history review: the pump (b)(6) passed all post sterile inspection checks. | A5 and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2181 | 1220648-2025-49603 | 23934570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934570&pc=OZD | 12/02/2025 | 02/10/2026 |  | 12/31/2025 |  |  | 12/03/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719756 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. it was reported that the impella cp was inserted per instructions for use with no reported complications. it was reported that the pump had no flow and was subsequently removed and replaced with a new device. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. additional information regarding the replacement is unknown, and the final patient outcome is unknown. the device is reportedly available for return; however, as of the time of this report the device has not yet been returned for evaluation. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A. patient information: patient demographics are unknown. d6a and d6b are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the low pump flow has been completed. the clinical details state that the "clients complained that the pump had no flow, insertion was as per procedure. the pump was repositioned but that did not resolve the issue. there were no cannula kinks seen or related patient conditions. there was no biomaterial observed and the patients act's were between 160/180." there are no data logs returned. the returned product did not have any biomaterial noted or damages to the impeller. the was a kink in the cannula noted right above the motor housing. the pump was ran in a bucket and there was slight flow saturation as there was some blood and dried dextrose in the gap. low pump flow did not reproduce. since there are no data logs, the relation to the cannula kink and the low or blocked flow cannot be confirmed. the root cause of the low or blocked flow cannot be established.
2182 | 1220648-2025-49605 | 23934576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934576&pc=OZD | 12/18/2025 | 01/05/2026 |  | 12/31/2025 |  |  | 12/18/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026782837 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 126 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced a hematoma at the access site. manual pressure was held to resolve the issue. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. hematoma : clinical details note that hematoma began at 5.5 site of a combative patient. the hematoma was resolved after holding manual pressure for 15 minutes after the patient was restrained. pump was returned and evaluated. no abnormalities were observed. the cause of the hematoma was use related as clinical details note that the patient was combative. device history review: the pump passed all the post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2183 | 1220648-2025-49597 | 23934578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934578&pc=OZD | 12/03/2025 | 02/24/2026 |  | 12/31/2025 |  |  | 12/03/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731960 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. on support day 3, it was reported that an alarm occurred for purge pressure low. the purge cassette was replaced, but the issue persisted. on support day 8, the alarm was reported to have been resolved. it was noted that there were no issues with the pump operation, and there were no visible leaks observed. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. | [5 empty] | A2, a3, and a4 are unknown. the device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed.
2184 | 1220648-2025-49608 | 23934658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934658&pc=OZD | 12/17/2025 | 01/28/2026 |  | 12/31/2025 |  |  | 12/17/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754192 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 99 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old male patient implanted with an impella cp device. the patient was transferred from an outside facility and was noted to have a hematoma at the impella insertion site upon arrival. the doctor stated that the patient arrived from the sending facility with the hematoma already present and did not believe the hematoma had worsened since arrival. at the time of reporting, the doctor indicated he was not concerned. no additional information regarding patient outcome or further treatment was provided. | [5 empty] | Investigation summary no product returned. asae/hematoma: patient has hematoma at impella insertion site. the cause of the access site adverse event was not determined due to insufficient clinical details. device history lot device lot : 1973983 device history batch subcomponent lot : n/a device history review the pump sn (b)(6) passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2185 | 1220648-2025-49610 | 23934679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934679&pc=OZD | 11/30/2025 |  |  | 12/31/2025 |  |  | 12/05/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026778604 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 68 KG | Non Hispanic | Asian | Death | Yes | No | Air/Gas in Device | [6 empty] | Fever | Hematoma | Tachycardia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that during a cardiac surgery procedure there were air bubbles seen in aorta and lv. subsequently, it was also visualized that the air was from the impella with incorrect preparation of impella catheter. in addition, the surgeon observed that there have been cavitation on the absence of air and pressure alarms. the patient has expired. | [5 empty] | The investigation was completed and no product or data logs available. air in device: clinical details mention that there were air bubbles observed in the aorta and left ventricle during cardiac surgery procedure. after an interview with the surgeon and on-site support it was concluded that instead of air from the pump it might have been cavitation based on the absence of air and pressure alarms and comments from the surgeon. the root cause of the air being in the device was not determined due to lack of conclusive evidence from the provided clinical details, no data logs to confirm any related alarms, and unreturned product for further evaluation. hematoma: clinical details mention that heparin was put on hold as a subdural hematoma was discovered on (b)(6) 2025 which the medical team believed happened from the field when the patient collapsed. the root cause of the subdural hematoma was not determined due to lack of sufficient clinical details. fever, tachycardia: clinical details mention that the patient required medication due to being febrile and with a heart rate going into the 130s. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2186 | 1220648-2025-49613 | 23934688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934688&pc=OZD | 12/16/2025 |  |  | 12/31/2025 |  |  | 12/17/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026787348 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 61-year-old male was implanted with a impella cp for support during high-risk cardiac procedure. it was reported that the patient went for extracorporeal membrane oxygenation decannulation yesterday morning without difficulty. in the unit, postop, team weaned off epinephrine drops (epi gtt). 30 minutes later, dropped mean arterial pressure into the 40s, patient went into pulseless electrical activity arrest. pump alarming position unknown. dropped to p2, one round of cardiopulmonary resuscitation, with return of spontaneous circulation. resumed p6 and epi gtt. transthoracic echocardiogram showed pump in good position and depth. no further complications. weaned off epi and extubated this am. patient is stable this am. | [5 empty] | The investigation was completed and no device was returned. ppae( cardiac arrest): the root cause of the injury was unable to be determined due to insufficient clinical details. (hypotension): the cause of the injury was not determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2187 | 1220648-2025-49614 | 23934691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934691&pc=OZD | 12/06/2025 |  |  | 12/31/2025 |  |  | 12/06/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026760844 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male | 120 KG |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that immediately upon implantation, a spike was observed in the motor current (cardiac pulse), after which the alarm ¿placement signal not reliable¿ was triggered. the impella position was confirmed to be good, and no thrombus was noted prior to implantation. subsequently, the ps and left ventricular waveforms appeared highly atypical. a recommendation was made to exchange the pump. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product returned. data logs returned. the clinical details state that immediately upon implanting the pump there is a spike in motor current and then there were psnr alarms triggered. due to a lack of product returned and no identifiable pattern able to be established from the data logs for the ps increases, the root cause of the optical sensor issue cannot be confirmed or established. device history review: the pump sn (b)(6) passed all the post sterile inspection checks. | E1. initial reporter first and last name is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2188 | 1220648-2025-49615 | 23934700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934700&pc=OZD | 12/06/2025 |  |  | 12/31/2025 |  |  | 12/06/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026725110 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Unspecified Infection | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient had a planned left ventricular assist device (lvad) placement scheduled for monday but was deferred after new blood cultures returned positive for infection. a transesophageal echocardiogram (tee) will be performed prior to making a final decision. additionally, possible hemolysis was identified with a plasma-free hemoglobin level of 330. the heart team suspects the cause is volume-related. volume resuscitation was initiated, and the impella support level was adjusted from p-8 to p-6. additional information indicates that the patient had been on impella 5.5 support for approximately 60 days. on 12/15, a left ventricular assist device (lvad) implantation was performed, during which the impella 5.5 was explanted. the heart team observed platelet clots on the impella 5.5 device. the left ventricular assist device (lvad) procedure was completed without complications. on post-implant day 1, the patient exhibited left-sided weakness. a ct scan confirmed a stroke. the heart team suggested the event could be related to the impella 5.5 device. the patient is currently stable. | [5 empty] | B2 revised selection as it was reported incorrectly on the initial report that was submitted. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. data review: data logs showed the pump ran without issue during support. there were suction alarms triggered during the case but resolved. there were no motor current spikes, abnormal placement signal or purge issues were observed during support. the product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely patient condition related since hemolysis was resolved after giving the volume. infection: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2189 | 1220648-2025-49618 | 23934702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934702&pc=PYX | 12/07/2025 |  |  | 12/31/2025 |  |  | 12/08/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | OPTICAL AIC W/IMPELLA CONNECT, PKGD, JP |  |  | PYX |  |  | OPTICAL AIC W/IMPELLA CONNECT, PKGD, JP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1566737 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic was in use until this morning. another hospital-purchased machine was under repair, so impella connect was going to be used. when the power was turned on to replace the 5.5 aic that was being operated, a control unit abnormality alarm occurred. the hospital had purchased two machines consecutively, and now a control device malfunction alarm has appeared. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2190 | 1220648-2025-49617 | 23934705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934705&pc=OZD | 12/07/2025 |  |  | 12/31/2025 |  |  | 12/07/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026768625 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced an episode of ventricular tachycardia. an echocardiogram was performed to assess impella position, and the device was repositioned under echo guidance from 49.5 cm to 49 cm. the patient outcome is still to be determined as the patient remains on support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed. (ventricular tachycardia): the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks.
2191 | 1220648-2025-49616 | 23934706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934706&pc=OZD | 12/06/2025 |  |  | 12/31/2025 |  |  | 12/07/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026762771 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient experienced elevated purge pressure (pp) and low purge flow (pf) values of 2 and 3. the purge cassette was exchanged, and a recommendation was provided to replace the pump. the physician requested the lysis protocol, which was sent via email. additional information indicates that the case has been concluded. the physician reported no issues, and the patient is in improved condition. the outcome at the end of device support was successful weaning. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product or data logs were returned for evaluation. purge pressure high: clinical details noted that pump had high purge pressure and low purge flow. no kinks, purge cassette changed, and tpa added to purge. it was not noted if purge issue resolved before pump explant. the cause of the high purge pressure could not be determined as no product or logs were returned for evaluation and limited clinical details were provided. the pump sn (b)(6) passed all the post sterile inspection checks.
2192 | 1220648-2025-49619 | 23934711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934711&pc=OZD | 12/02/2025 |  |  | 12/31/2025 |  |  | 12/08/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026777498 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient was escalated from impella cp to impella 5.5 on december 3rd. during the impella 5.5 insertion, systolic arterial blood pressure (sabp) in the right hand decreased to approximately 30 mmhg. initially, this decrease was attributed to intraoperative manipulation; however, after returning to the icu, sabp in the left hand also decreased to approximately 30 mmhg. aortic waveforms were obtained at approximately map 80. a subsequent ct scan confirmed thrombus occlusion in three branches of the aortic arch. sheaths were inserted into both internal arteries, and perfusion was initiated via the extracorporeal membrane oxygenation (ecmo) circuit. a head ct scan revealed significant cerebral edema due to infarction. the patient subsequently expired. the physician noted that an ultrasound performed prior to impella 5.5 insertion revealed a large thrombus in the left atrial appendage. it is possible that this thrombus propagated during the impella 5.5 insertion; however, causality cannot be determined with certainty. the physician stated that the use of impella was necessary to attempt to save the patient¿s life, and there is no impact on future use of the device. additional information indicated that no issues were reported with the operational status of the device or the patient¿s condition during impella cp support. an arterial line was inserted into the right radial artery. the cerebral infarction was confirmed to be ischemic. because there was no dissociation between the arterial line inserted into the right temporal artery and the position waveform, it is unlikely that the cerebral infarction occurred during cp support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2193 | 1220648-2025-49620 | 23934714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934714&pc=OZD | 12/08/2025 | 01/22/2026 |  | 12/31/2025 |  |  | 12/08/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1379919 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the malfunction occurred at the pressure transmitter during priming. it was resolved by replacing the control device. there was no health impact. | Additional information from the complainant indicated that the patient had expired from a rupture as a complication of acute myocardial infarction that was present prior to impella treatment. the complainant stated the rupture not attributable to impella. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | B5 updated. the investigation is still ongoing.
2194 | 1220648-2025-49621 | 23934722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934722&pc=OZD | 12/08/2025 |  |  | 12/31/2025 |  |  | 12/08/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026779408 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 85 KG | Hispanic | Asian | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient's urine was still pink after reposition. a couple of hours post reposition; the urine cleared due to the reduction of afterload. the patient was reported to have survived the procedure. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the pump was not returned for investigation. data log was not provided and was not found on the remote server. mechanical interaction with blood (hemolysis): the cause of the issue was not determined without returned product investigation and sufficient clinical details, including data log review. device history review: the pump passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2195 | 1220648-2025-49622 | 23934738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934738&pc=OZD | 12/08/2025 | 01/14/2026 |  | 12/31/2025 |  |  | 12/08/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026784490 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that slight oozing was present at the insertion site, and the swan catheter was also noted to be oozing. the patient received 8,000 units of heparin and was on a cangrelor infusion. the patient survived the procedure. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the access site bleed could not be determined due to insufficient clinical details. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2196 | 1220648-2025-49625 | 23934876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934876&pc=PYX | 12/08/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1746950 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that since all available automated impella controllers (aics) were in use, this case required the use of abiomed¿s aic as a backup device. during preparation for support, an issue occurred after startup: the message ¿please connect the catheter cable to the connection cable¿ appeared, preventing further progress. the team repeatedly unplugged and reconnected the cable approximately 30 times and restarted the setup process several times, but the same message persisted. after retrying the setup using the same procedure, the device eventually proceeded normally and became operational. | [5 empty] | D6a and d6b: updated the implant and explant date. h6: updated code based on information in the complaint file. | Device status: the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. device analysis: console (ic10805) was sent back to fs for further investigation. the pump detection issue could not be reproduced despite several attempts to connect a known-good impella and power cycling the console in between attempts. root cause: the root cause of the pump detection issue could not be determined, as the issue could not be reproduced.
2197 | 1220648-2025-49612 | 23934884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934884&pc=OZD | 12/17/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/17/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634759 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 52-year-old male patient with hematuria versus amber/concentrated urine. an echocardiogram confirmed the device was in a stable position. no additional clinical details, interventions, or patient outcome information were provided at the time of reporting. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1: added product problem, this was omitted in error from the initial report. h6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary hemolysis/mechanical interaction with blood: the cause of the hemolysis was unable to be determined due to lack of product and data logs available for analysis.
2198 | 1220648-2025-49627 | 23934885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934885&pc=OZD | 12/09/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026775398 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 105 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient presented with pink-colored urine, raising concern for possible hemolysis. plasma-free hemoglobin was drawn to confirm. additional information indicated that the patient was hypovolemic; volume replacement was administered, and the hemolysis subsequently resolved. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary mechanical interaction with blood (hemolysis): the pump was not returned for investigation. data log was not provided or retrieved from the remote server. the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical details, including data log review. device history review the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2199 | 1220648-2025-49628 | 23934909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934909&pc=OZD | 12/06/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708890 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 69 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Rupture | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient had an impella cp in place. proper placement of the cross-clamp was discussed; however, the surgeon did not palpate the motor and instead clamped just proximal to the aortic cannulation site. following clamp placement, a high motor current alarm was triggered, and the impella subsequently stopped. the surgical procedure was completed without incident. the impella cp was then explanted, and closure was achieved using three perclose devices post-procedure. additional information indicated that the patient remains sedated on mechanical circulatory support at performance level 5 with flows of approximately 2 liters per minute. attempts to increase flows resulted in suction events. a swan-ganz catheter and transesophageal echocardiogram were performed earlier in the morning. the echocardiogram revealed a potential papillary muscle rupture. the patient was designated as do not attempt resuscitation (dnar), and an epinephrine infusion was initiated. | [5 empty] | H6: updated coding based on the results of the investigation. investigation summary impella returned for investigation. papillary muscle rupture: the cause of the injury was not determined due to insufficient clinical details. pump stop/low pump flow: the returned pump was inspected and dried blood was observed at the outflow cage by the impeller. the pump passed electrical testing and cannula had no kinks. the pump was soaked and run in the lab and neither low pump flow nor pump stop reproduced. no data logs were returned. the cause of the issue was not established as no defects were observed and no logs were returned for analysis device history review the pump passed all post sterile inspection checks | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2200 | 1220648-2025-49630 | 23934926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23934926&pc=PYX | 12/09/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1930376 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A nurse called to report that automated impella controller had a "controller error" alarm, and said it resolved within 45 seconds. the nurse stated that it only happened once and it did not happened again. there was no patient harm. | [5 empty] | D6a and d6b dates appeared on the initial report that was submitted and these fields should of been left blank. e1 added zip code extension as it was omitted from the initial report that was submitted. h6 revised h6 as the device was returned and were incorrectly reported on the initial report that was submitted. h11 revised additional manufacturer narrative as it was incorrectly reported on the initial report that was submitted. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2201 | 1220648-2025-49632 | 23935230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935230&pc=OZD | 12/16/2025 |  |  | 12/31/2025 |  |  | 12/16/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026769814 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was transferred from outside hospital with pulmonary embolism and ischemic stroke. the patient went to interventional radiology for clot removal that was successful. after returning patient went into ventricular fibrillation arrest which prompted impella cp insertion. later the patient was being evaluated for neuro status and was sent for computed tomography scan of head which resulted with a hemorrhagic stroke. because of this heparin was paused. the patient expired and no other information provided. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2202 | 1220648-2025-49636 | 23935240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935240&pc=OZD | 12/16/2025 |  |  | 12/31/2025 |  |  | 12/16/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732139 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after support was initiated, the patient was transferred to the surgical intensive care unit reportedly due to the patient¿s coagulopathy issues. after the transfer was completed, the patient was noted to have bleeding at the impella access site. the patient¿s hemoglobin and hematocrit dropped multiple times overnight and the patient was transfused with blood products. it was noted that the decision was made to explant the impella cp due to the bleeding around the repositioning sheath and introducer hub valve. hemostasis was noted to be achieved after the impella was removed. it was noted that the patient¿s outcome at explant was expired. additional clinical details, including perioperative course, contributing factors, and cause of death, were unavailable at the time of reporting. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. asae/ major bleed: bleeding occurred at the impella cp access site, patient has coagulopathy issue, resulting in multiple hemoglobin drops and transfusions. hemostasis was achieved after device removal. the cause of the access site adverse event was patient-related due to coagulopathy. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | A4 is unknown. the device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed.
2203 | 1220648-2025-49631 | 23935241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935241&pc=OZD | 12/09/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716155 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 68 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 79-year-old male patient supported with an impella device had a hemoccult-positive stool, indicating bleeding. the patient reportedly received a blood transfusion. a change in therapeutic response was reported. no additional clinical details were provided. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the major bleed could not be determined due to insufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2204 | 1220648-2025-49637 | 23935251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935251&pc=OZD | 12/16/2025 |  |  | 12/31/2025 |  |  | 12/16/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1079578 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that automated impella controller had a controller error. no further information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
2205 | 1220648-2025-49639 | 23935288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935288&pc=PYX | 01/24/2025 |  |  | 12/31/2025 |  |  | 12/16/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025483246 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Study | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information received that repeat venous ultrasound of upper and lower extremities (b)(6) 2025 per high risk tx team - no dvt, mild residual chronic nonocclusive bilateral ij dvt, mild residual chronic occlusive right cephalic superficial venous thrombus. | It was reported that a patient implanted with an impella rp flex experienced deep vein thrombosis and an occlusive superficial thrombus of the right cephalic vein. the patient was in stable condition. |  |  |  |  | No product was returned. the root cause of the injury was unable to be determined due to insufficient clinical details. | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2206 | 1220648-2025-49640 | 23935293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935293&pc=OZD | 12/16/2025 |  |  | 12/31/2025 |  |  | 12/16/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723044 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted in the right femoral artery with an impella cp experienced leaking below the diaphragm. the sheath was removed and pressure was held. access was successfully obtained in the left groin. the patient expired on support. | Clinical rationale: a 57-year-old male presented with acute myocardial infarction and cardiogenic shock, with a pre support clinical state consistent with scai shock stage d. during the procedure, a 14 fr 13 cm peel away sheath was inserted into the right femoral artery. upon removal of the dilator, the sheath was noted to be leaking below the diaphragm. the introducer was removed, and manual pressure was applied. vascular access was then successfully obtained via the left groin, where a 14 fr 25 cm peel away sheath was inserted, and an impella cp (s/n (b)(6)) was placed at 5:27 pm. the event involved a malfunction of the first introducer, resulting in a fluid leak and requiring device replacement. although the impella cp device was placed successfully via alternate access, the patient ultimately expired during the procedure. based on available information, the pump was not implicated in the cause of death; however, both devices will be returned for further analysis to determine whether any device related factors contributed to the clinical outcome. the impella cp will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. the patient¿s death is conservatively reported but likely unrelated to device performance but rather the patient¿s condition. |  |  |  |  | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H10 this is one of two related reports. this report represents the impella cp and another report was submitted to represent the introducer. related report was added as it was omitted from previously submitted reports. | B5 clinical rationale was added. h6 code d16 was reported incorrectly on the initial report that was submitted. the correct code was added to the investigation conclusions code. investigation summary: the investigation has been completed. no product return. analysis summary: hemodynamic instability / peri-procedural adverse event: the cause of the injury was not determined due to insufficient clinical details.
2207 | 1220648-2025-49642 | 23935328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935328&pc=OZD | 12/16/2025 | 01/05/2026 |  | 12/31/2025 |  |  | 12/16/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026780216 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Alarm System | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced a high motor current alarm during start up. the wire was removed to resolve the issue. no patient harm was reported. | [5 empty] | The investigation was completed and no product or data logs returned. temporary pump stop: clinical details state that there was a high motor current alarm during pump start. wire was not removed at the time. once wire was fully removed, the pump started and operated normally. the cause of the temporary pump stop was use-related based on clinical details noting that the wire was not removed prior to starting pump. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b updated. the impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
2208 | 1220648-2025-49645 | 23935383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935383&pc=OZD | 12/15/2025 |  |  | 12/31/2025 |  |  | 12/15/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026770795 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp via the right femoral artery for mechanical circulatory support. during the first day of support, the patient experienced bleeding at the impella access site. manual pressure was held to achieve hemostasis, but it was noted that the impella was removed later that day due to the bleeding and the patient¿s heart recovery. the device was explanted successfully after weaning. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2209 | 1220648-2025-49646 | 23935415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935415&pc=OZD | 12/15/2025 |  |  | 12/31/2025 |  |  | 12/15/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704263 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | I was called for this potential case and departed immediately but am an hour away from this remote facility. upon my arrival, impella had been placed pre pci once significant lad and lcx disease discovered. per report, impella placed without issue and both vessels wired/ballooned/stented successfully. upon my arrival to cath lab, peel-away had been removed and repositioning sheath advanced but not sutured and right heart cath being performed. ps appropriate with flows 3.5 in auto. patient slightly hypertensive with sbp 140¿s. md left to discuss with family and scrub peeled drape away to discover very large pool of blood and called md back to procedural room. large hematoma and rp bleed discovered. pressure held, per close tightened but bleed remained very large. impella weaned to p2 with no change in bp so md opted to pull impella to repair artery with covered sheath, stating he was just going to leave overnight for protection anyway and had only placed for protective pci purposes. left groin prepped and impella pulled. covered sheath successfully deployed and provided hemostasis. md felt rp bleed related to patient¿s large body habitus and pannus. the patient's outcome at the time of explant was survived. | [5 empty] | Clinical details note that peel-away was removed and repositioning sheath advanced but was not sutured and right heart cath was performed. pump was returned and evaluated. no abnormalities found. the cause of the access site adverse event was not determined due to no defect found. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2210 | 1220648-2025-49648 | 23935432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935432&pc=OZD | 12/14/2025 |  |  | 12/31/2025 |  |  | 12/15/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026769165 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp could not be implanted as the device repeatedly kinked within the aorta. an alternative pump system was used to support the patient. the patient was in critical condition. | [5 empty] | A2, a4, a5, and a6 are unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Correction : section h6: the problem code was inadvertently left out in the initial report which has been added to this report.
2211 | 1220648-2025-49649 | 23935438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935438&pc=OZD | 12/14/2025 | 01/28/2026 |  | 12/31/2025 |  |  | 12/14/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731108 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. on support day 2, it was reported that the patient became confused and pulled and removed the impella cp from their body. manual pressure was held at the impella access site to achieve hemostasis. no further information was provided. the patient's outcome at the time of explant was survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: no product returned. device in wrong position: the cause of the positioning issue was use related to patient movement as the patient became confused and pulled out his impella cp. asae/minor bleed: the cause of the access site adverse event was use related to patient movement as the issue occurred after the patient became confused and pulled out his impella cp. device history lot: device lot: 1961521. device history batch: subcomponent lot: n/a. device history review: the device (sn:(b)(6) passed all post sterile inspection checks.
2212 | 1220648-2025-49650 | 23935449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935449&pc=OZD | 12/14/2025 | 12/29/2025 |  | 12/31/2025 |  |  | 12/14/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026773979 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 93 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received that the bleeding did not occur during delivery of the device. it was after the peel away sheath was removed and peeled away. it was reported that the scrub assistant was ¿aggressive¿ when he held pressure on the groin for the physician to walk the sheath out and then peel it away. the bleeding didn¿t start until after the repositioning sheath was advanced forward. | The user facility reported that a 69-year-old male was implanted with an impella cp for support during high-risk cardiac procedure. it was reported that during removal of the 14 fr sheath from the right femoral artery, a hematoma formed above the access site shortly after a repositioning sheath was advanced. manual compression was applied. subsequent examination revealed a rapidly growing medial hematoma, prompting the operator to initiate device removal and apply a femstop for hemostasis. |  |  |  |  | The investigation of the reported failure mode has been completed. no product was received for evaluation. vascular dissection: the cause of the injury was not determined due to insufficient clinical details. access site adverse event: the cause of the hematoma was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | B5 updated with additional information received from the clinical site. correction: e4. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2213 | 1220648-2025-49651 | 23935453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935453&pc=OZD | 12/14/2025 | 02/12/2026 |  | 12/31/2025 |  |  | 12/14/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026730202 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that the hemolysis was resolved by removing extracorporeal membrane oxygenation (ecmo) and repositioning. | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. nurse reported hematuria and hemolyzed labs. impella was repositioned. the following day the pump was weaned successfully and the patient's outcome at explant is noted as survived. |  |  |  |  | B5 and h6 updated with additional information. it was reported the device was discarded therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the hemolysis and hematuria has been completed. no product or data logs returned. the cause of the hemolysis was not determined due to lack of product and data logs returned. the cause of the injury was not determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2214 | 1220648-2025-49652 | 23935454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935454&pc=OZD | 12/14/2025 |  |  | 12/31/2025 |  |  | 12/14/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026787425 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 56 YR |  | Female | 106 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support the patient had severe retroperitoneal bleed. it is unclear if caused by impella cp or femoral venoarterial extracorporeal membrane oxygenation. it was noted that the patient expired during impella cp support. no other information has been provided. | [5 empty] | B1: added serious injury h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. ppae (retroperitoneal hemorrhage): the cause of the injury was not determined due to insufficient clinical details. device history review the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2215 | 1220648-2025-49638 | 23935477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935477&pc=OZD | 12/09/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765855 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female | 86 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient coded in the intensive car unit and brought to the cath lab. physician decided to put in impella cp to support the patient. the patient expired in the cath lab, not due to impella related issues. no further information was reported. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2216 | 1220648-2025-49653 | 23935491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935491&pc=PYX | 12/09/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026651105 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was going to be implanted with an impella rp flex for mechanical circulatory support. it was reported that unable to advance the impella rp flex therefore a new device was used. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9: updated devices return information h6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary failure to advance: the cause of failure to advance cannot be determined since insufficient clinical details were provided.
2217 | 1220648-2025-49587 | 23935621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935621&pc=PYX | 09/29/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 09/29/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025600652 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp flex required escalation of pressors. the patient also experienced thrombocytopenia. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella passed all the post sterile inspection checks. the root cause of the sepsis was unable to be determined due to insufficient clinical details.
2218 | 1220648-2025-49655 | 23935626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935626&pc=OZD | 12/13/2025 | 12/29/2025 |  | 12/31/2025 |  |  | 12/14/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026752471 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Hematuria | Pericardial Effusion | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that the patient's urinary output was less than 30cc an hour and was pink red in color. the patient was later diagnosed with renal failure and required dialysis. it was further reported that the patient expired. | The user facility reported a 72 year old male was implanted with a impella cp for support during a high risk cardiac procedure. it was reported that there was minimal oozing at the access site, which managed with hemostatic dressing. hemoglobin remained stable at 14. a percardial effusion was identified on echocardiogram, and pericardiocentesis was performed. |  |  |  |  | Additional information provided and the following were updated based on the information provided by the clinical site: b2, b5, h1, h6. the exact date of death was not provided. correction: e4. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the root cause of the injuries could not determined due to insufficient clinical details.
2219 | 1220648-2025-49558 | 23935640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935640&pc=OZD | 12/04/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/04/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727420 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient died while receiving ecmella support, which involves the combined use of extracorporeal membrane oxygenation (ecmo) and the impella microaxial pump. it reported contributing factors including thrombus formation during support, intestinal obstruction, suboptimal initial vessel condition, left main (lm) stenosis, and procedural complexity. the cause is not impella related according to the physician. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the thrombosis, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2220 | 1220648-2025-49657 | 23935650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935650&pc=OZD | 12/14/2025 |  |  | 12/31/2025 |  |  | 12/14/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764102 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Arrest | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the patient has impella cp right femoral artery. the patient was having some ectopy where heart team is suspecting the impella cp may be affecting it. the ectopy occurs when the patient¿s sedation is decreased. electrophysiology consult planned for monday since cause of cardiac arrest (ventricular fibrillation arrest) is unknown. the pump was explanted the following day and the patient was noted as survived. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
2221 | 1220648-2025-49654 | 23935742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935742&pc=OZD | 12/10/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/10/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026778605 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increased Pump Speed | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a 56-year-old male patient underwent attempted impella placement via the right femoral artery. the impella pump was reportedly started prior to being positioned inside sheath and not across the aortic valve. the pump abruptly stopped with a spiked motor current while the red ez guide lumen remained within cannula. the device was set aside, and a second impella pump was opened and successfully implanted. no patient injury was reported. | [5 empty] | Cp pump, sn: (b)(6), red lumen & usb were returned. pump unintended start: clinical details noted that clinical staff accidentally started pump before pump was inserted in the sheath or into the patient. red lumen still present in pump at time of unintended start. returned product was inspected and no damage to returned pump was noted. a small kink was present on returned red lumen. pump was connected to console and was able to run without issue. data logs were reviewed and rt logs confirmed that pump was started before pump was inserted into patient. the cause of the unintended pump start was a user related issue based on clinical details provided. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | Correction: e4, h6. | The impella device was received from the customer. upon completion of the investigation, a supplemental mdr will be filed.
2222 | 1220648-2025-49660 | 23935902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935902&pc=OZD | 12/14/2025 |  |  | 12/31/2025 |  |  | 12/15/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026778559 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Cardiac Perforation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced suction alarms. the pump was penetrating the inferior wall. a patch closure was attempted but the myocardium was too fragile and the patient expired. | [5 empty] | B1: added product problem. h6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: patient device interaction problem / cardiac perforation: the cause of the injury was not determined due to insufficient clinical details. pump suction: the cause of the suction issue was not determined without returned product investigation and sufficient clinical details, including data logs. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2223 | 1220648-2025-49659 | 23935903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935903&pc=OZD | 12/14/2025 |  |  | 12/31/2025 |  |  | 12/14/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026777886 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 56 KG |  | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. upon arrival to the cath lab it was noted that the patient didn't have palpable pedal pulses. reperfusion catheter in place post pump implant. hemolysis was also noted with no additional information. it was later reported that the patient was in the operating room getting fasciotomy of left leg. the impella cp was explanted and replaced with an impella 5.5. | [5 empty] | The investigation was completed and no product was returned. mechanical interaction with blood (hemolysis): the pump was not returned for investigation. data log shows no signs of motor current spike, positioning issue, suction, or major placement signal trend during the entire case run. the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical details. clinical symptoms (ischemia): the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2224 | 1220648-2025-49661 | 23935905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935905&pc=OZD | 12/10/2025 | 02/12/2026 |  | 12/31/2025 |  |  | 12/14/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750285 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | Cardiac Perforation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient experienced cardiac tamponade requiring pericardial drainage with a subxiphoid catheter. the patient was then implanted with an impella cp. later, the patient was escalated to an impella 5.5. the patient was in stable condition. | [5 empty] | The investigation for the cardiac perforation has been completed. the pump was not returned for investigation. data log review was not required for this case run. the cause of the injury was not determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2225 | 1220648-2025-49658 | 23935907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935907&pc=OZD | 12/14/2025 | 02/10/2026 |  | 12/31/2025 |  |  | 12/14/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026774880 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria and hemolysis. the pump was confirmed to be in a good position. the patient was in stable condition. | [5 empty] | The ivnestigation for the hemolysis and hematuria has been completed. there was no product or data logs returned. without additional clinical details, the cause of hemolysis/mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the hematuria was not determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2226 | 1220648-2025-49662 | 23935909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935909&pc=OZD | 12/14/2025 |  |  | 12/31/2025 |  |  | 12/14/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026783666 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella cp and an impella rp flex for mechanical circulatory support. it was reported that the impella was alarming in ventricle and then aorta without movement. got x-ray showing impella catheter further back than that morning. unable to flow higher than p1 and getting 0.9 l/min. ic called to reposition bedside but echo windows not great and unable to flow higher than 0.9 l/min with repositioning attempts. decision to remove and leave peel away sheath in to be closed in the morning. the pump was replaced and the patient later expired. this report is for the impella cp. the death will be reported on the impella rp flex (serial number (b)(6)). | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Medical safety/clinical reviewed the event. the patient was on bipella support. the 1st impella cp had a positioning / low flow issue a couple days into support and it was replaced. bipella support continued. two days later, the patient had presumed heparin-induced thrombocytopenia (hit) attributed to the rp flex. however, before the test results could confirm, care was withdrawn and the patient then passed away. the decision to withdraw care was not attributed to the hit. in this case the patient's underlying condition contributed most to the demise outcome. note: h6 investigation type/findings/conclusion remains unchanged. related medical device reports: this is one of three devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. no physical damage was noted on the returned product. no evidence of biomaterial was found on the purge gap. the pump was tested in a bucket of water without any abnormalities. data log confirm that while running at steady p-level, pump flow dropped and became non-pulsatile, with active impella position alarms in the aorta and ventricle. low or blocked pump flow / device in wrong position: no issue was found on the returned pump. the cause of the low pump flow issue was most likely related to improper positioning, based on data log review and clinical details. however, the cause of the positioning issue was not determined without sufficient clinical details. device history review: the device passed all post sterile inspection checks.
2227 | 1220648-2025-49663 | 23935943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935943&pc=OZD | 12/14/2025 | 02/12/2026 |  | 12/31/2025 |  |  | 12/14/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711284 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male | 97 KG | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that there was "pp high" alarm. recommendations and best practices shared and local lysis administered. it was also reported the the pump was changed. patient outcome at explanted was noted as survived. | [5 empty] | The investigation for the high purge pressure has been completed. no product or data logs were returned for evaluation. the cause of the high purge pressure could not be determined as no product or logs were returned for evaluation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2228 | 1220648-2025-49665 | 23935959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935959&pc=OZD | 12/13/2025 | 02/12/2026 |  | 12/31/2025 |  |  | 12/13/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732685 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 90 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during swan securement, hematoma on left groin noted. pressure held. cp came back into the aorta. before pump alarmed, tech advanced reposheath back into the artery advancing the inlet back into the left ventricle. after 5.5 placed, cp was removed over stiff wire. 14fr cook sheath inserted with an without dilator. groin began to bleed. vascular surgery called in for repair. additional information provided that the surgery team performed a cutdown and repaired the vessel. "there is a lateral tear" was noted. there is no imaging available. | [5 empty] | The following section has been updated: d9-device return date: the investigation for the positioning issue, access site bleed, peripheral artery injury has been completed. the device was returned and visually inspected. no abnormality found. the root cause of positioning issue most likely was use related. the root cause of access site bleed/ hematoma cannot be determined due to insufficient clinical details. the cause of the injury was not determined due to insufficient clinical details. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2229 | 1220648-2025-49666 | 23935962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935962&pc=OZD | 12/13/2025 | 12/29/2025 |  | 12/31/2025 |  |  | 12/13/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026770420 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | Use of Device Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 60 a year old male patient experienced unsuccessful impella pump start while the wire had not been completely removed. a second impella 5.5 was successfully implanted. no patient injury was reported. | [5 empty] | Pump did not start / use against labeling: the pump was not returned for investigation. clinical details indicate that the wire was not completely removed before the physician asked for the pump to be started. after the wire was removed, the second and third attempts to start the pump failed. the device was unplugged from the aic and plugged back in, but another attempt was unsuccessful. data log was not provided or retrieved from the remote server. the cause of the pump not starting was not determined without returned product investigation and sufficient clinical details, including data logs. as per the ify, guidewire should be removed before pump start. the cause of the use against labeling issue was most likely use-related. the complaint pump s/n 1988143 passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2230 | 1220648-2025-49667 | 23935966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935966&pc=OZD | 12/13/2025 | 01/27/2026 |  | 12/31/2025 |  |  | 12/13/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1089162 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an automated impella controller (aic) in use for patient support with an impella pump had controller error alarms. the aic was exchanged for another aic and the alarms were resolved. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the device is reportedly available for return; however, as of the time of this report the device has not yet been returned for evaluation. additional clinical details, including the pump in use at the time of event, were unavailable at the time of reporting. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Section a: patient information: all patient demographics are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2231 | 1220648-2025-49668 | 23935967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935967&pc=OZD | 12/13/2025 | 02/03/2026 |  | 12/31/2025 |  |  | 12/13/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026774634 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 92 KG | Non Hispanic | White | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a 56-year-old male patient experienced a high purge pressure alarm while supported with an impella device. the purge flow was reportedly decreasing throughout the day. no kinks were noted, and the cassette was changed prior to contacting support. no additional clinical details were provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed. investigation summary: 5.5 pump sn615211 and a purge cassette returned. high purge pressure/purge system blocked: returned pump was visually inspected and no damage to impella cannula was observed. some small biomaterial present around impeller blade. pump was connected to console and purged in lab for approximately 5 minutes before "purge system blocked" alarms were triggered. pump was turned onto p-9 and increased purge pressure and purge alarm resolved. no data logs were returned for evaluation. clinical details noted that "high purge pressure" and "purge system blocked" alarms were triggered. no kinks noted, tpa was added to purge but did not resolve issue and pump was explanted. the cause of the high purge pressure/purge system blocked could not be determined as the purge issue did not sustain when it was reproduced in the lab, no data logs were returned for analysis, and limited clinical details were provided. device history review was completed and passed all post sterile inspection checks. | Correction to initial report: it was reported in the initial report under h6 that medication was required (f2303). however, it was inadvertently left off in event narrative in b5 the details regarding the medication administration. b5 should have stated the following: the user facility reported that a 56 year old male patient experienced a high purge pressure alarm while supported with an impella device. the purge flow was reportedly decreasing throughout the day. purge system was blocked. tissue plasminogen activator (tpa) was started with purge flows less than 1 and pressure at 1087. the patient remained on support.
2232 | 1220648-2025-49669 | 23935969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935969&pc=OZD | 12/13/2025 |  |  | 12/31/2025 |  |  | 12/13/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026782991 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The operating room team was attempting an escalation case in the cardiovascular operating room. impella 5.5 was inserted and on start up got several alarms including "impella stopped motor current high" "impella stopped controller failure" purge disc not connected" impella stopped retrograde flow" team decided to switch to another controller and pump. they proceed with the case and there was no patient harm. this report is for the pump and another report will be sent for the controller (serial number (b)(6)). | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The logs were returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
2233 | 1220648-2025-49671 | 23935976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935976&pc=OZD | 12/13/2025 | 12/26/2025 |  | 12/31/2025 |  |  | 12/13/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766720 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported on a 71-year-old female patient on an impella cp for acute myocardial infarction. it was reported that while on support, the patient developed uncontrolled access site bleeding and as a result, the patient's mean arterial pressure dropped significantly. it is unknown what type of treatment the patient was given but the issues were not resolved. it was further reported that the patient expired in the middle of the night. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary no product returned. major bleed: patient's map dropped significantly, and uncontrollable access site bleeding occurred. the cause of the access site adverse event was not determined due to insufficient clinical details. hemodynamic instability: bleeding were never resolved. the cause of the injury was not determined due to insufficient clinical details. device history review this pump passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2234 | 1220648-2025-49670 | 23935979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935979&pc=OZD | 12/13/2025 | 12/26/2025 |  | 12/31/2025 |  |  | 12/13/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713312 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 64 year old male patient underwent implantation of an impella 5.5 during a high risk cardiac procedure. within the last 24 hours, the patient reportedly developed a hematoma at the vascular access site. no additional clinical details were provided. no device malfunction was reported. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. asae/hematoma: patient developed a hematoma at access site within the last 24 hours. hemostasis achieved with manual pressure. the cause of the access site adverse event was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2235 | 1220648-2025-49672 | 23935989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935989&pc=OZD | 12/13/2025 | 12/29/2025 |  | 12/31/2025 |  |  | 12/13/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122628 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during use of the automated impella device, when turning on aic ¿ unexpected controller shut down¿. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2236 | 1220648-2025-49674 | 23935994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23935994&pc=OZD | 12/13/2025 | 12/29/2025 |  | 12/31/2025 |  |  | 12/13/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731941 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 62 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 66-year-old female patient experienced significant oozing at the access site and new onset hematuria while supported with an impella device. bleeding/ hemorrhage was reported. no additional clinical details were provided. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the hematuria was unable to be determined due to insufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2237 | 1220648-2025-49675 | 23936019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936019&pc=OZD | 12/11/2025 |  |  | 12/31/2025 |  |  | 12/12/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681982 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. on the first day of support, it was reported that the patient experienced hemolysis. the impella was repositioned and subsequently the hemolysis resolved. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | Added code a01 based on information in the complaint file. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2238 | 1220648-2025-49677 | 23936022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936022&pc=OZD | 12/12/2025 | 12/26/2025 |  | 12/31/2025 |  |  | 12/12/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026767829 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported on a 59 year old female with an impella cp for cardiomyopathy. it was reported that during support, the patient had suspected hemolysis with red urine is red. labs were hemolyzed. the patient was transferred to another facility where the pump was repositioned. urine changed to dark/amber. patient received dialysis therapy. patient remains on support. | [5 empty] | The investigation was completed and no product or data logs available. mechanical interaction with blood (hemolysis): clinical details mention that while on impella support the patient's urine output was red and their labs were reported as hemolyzed. the root cause of the hemolysis was not determined due to lack of sufficient clinical details and unreturned product and data logs for further evaluation. device history review was completed and passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2239 | 1220648-2025-49676 | 23936023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936023&pc=OZD | 12/12/2025 | 12/26/2025 |  | 12/31/2025 |  |  | 12/12/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026756979 |  | 000000000001000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 102 KG | Non Hispanic | White | Injury | Yes | Yes | Break | Material Split, Cut or Torn | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a impella 5.5 in use on a 69-year-old male patient for support during high-risk cardiac procedure. it was reported that while on the unit, the patient was turned in bed and impella in aorta alarm sounded. at the time, a large hematoma was noted at right upper chest impella insertion site. impella was turned to p2 and echo was used to attempt an advance of the impella without success. they were unable to visualize impella via tte. patient was taken to surgery for hematoma evacuation and impella repositioning. there was no patient harm reported and the patient remains on support. | [5 empty] | B1: added malfunction. h6: added code a0414. additional information: sterile sheath cover tear. ripped area cleaned with chloraprep and covered with tegraderm. | D6b explant date was added. e1 added zip code extension as it was omitted from the initial report that was submitted. e4 added selection as it was omitted from the initial report that was submitted. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2240 | 1220648-2025-49678 | 23936024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936024&pc=OZD | 12/12/2025 |  |  | 12/31/2025 |  |  | 12/12/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026730006 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Hemodynamic instability | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. on support day 5, the patient experienced episodes of ventricular tachycardia (vt). the physician attempted to reposition the device; however, the vt did not resolve. the pump was subsequently weaned and explanted, and it was noted that the device was explanted due to the complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2241 | 1220648-2025-49679 | 23936029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936029&pc=OZD | 12/12/2025 |  |  | 12/31/2025 |  |  | 12/12/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026783682 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for escalated mechanical circulatory support from an impella cp. it was reported that on day 2 of impella cp support, the patient experienced hematuria and decreased urine output. the decision was made to escalate the patient to an impella 5.5. it was reported that on the first day of impella 5.5 support, there was an alarm for placement signal not reliable. the impella 5.5 position was verified with echocardiography. it was later reported that the alarm was resolved. no further information was provided. no signs or symptoms of patient injury were reported with the impella 5.5, there was no reported harm associated with the placement signal not reliable event. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the pump was not returned for investigation. data log was not provided or retrieved from the remote server. optical signal issue: the cause of the optical signal issue was not determined due to lack of returned product and data logs for investigation. device history review: device passed all post sterile inspection checks. | The device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed.
2242 | 1220648-2025-49681 | 23936033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936033&pc=OZD | 12/11/2025 |  |  | 12/31/2025 |  |  | 12/11/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732849 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp via the right femoral artery for mechanical circulatory support during a procedure. it was reported that during the procedure, the impella pump flows were lower than expected with values of 1.4-1.7 l/min. additionally, it was noted that the impella cannula was kinked. the pump was repositioned; however, this did not resolve the issue. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the pump was weaned and explanted following the completion of the procedure. the patient's outcome at the time of explant is survived. the device is reportedly available for return; however, as of the time of this report the device has not yet been returned for evaluation. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation was completed. no product or data logs returned. low or blocked pump flow: clinical details mention that the pump flow was 1.4-1.7 throughout the case. the root cause of the low pump flow was not determined due to insufficient clinical details, and unreturned product and data logs to confirm and further investigate the issue. product damage (cannula): clinical details mention that the pump got kinked during the case. further clinical details whether the kink occurred during implant, explant, or runtime, or if the patient had any anatomical anomalies are not available. the root cause of the cannula kink was not determined due to insufficient clinical details and unreturned product for further investigation. the device history review was completed and passed all post sterile inspection checks.
2243 | 1220648-2025-49680 | 23936034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936034&pc=OZD | 12/12/2025 |  |  | 12/31/2025 |  |  | 12/12/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739121 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 69 KG | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 via the right axillary artery for mechanical circulatory support. on day 48 of support, it was reported that the anti-contamination sleeve of the pump was broken during repositioning the device. the sleeve was wrapped in tegaderm and support continued. it was noted that the pump was later explanted. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. product damage anticontamination sleeve split/torn: clinical reported sterile sleeve torn off the hub during repositioning. the cause of the anticontamination sleeve damage is not determined due to insufficient clinical details and no returned product . h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2244 | 1220648-2025-49683 | 23936040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936040&pc=OZD | 12/11/2025 | 01/08/2026 |  | 12/31/2025 |  |  | 12/11/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026775972 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 88 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 via the right axillary artery for mechanical circulatory support. on day 2 of support, the patient experienced hematoma following the impella 5.5 implant. the patient was brought to operating room, and the hematoma was evacuated, and it was reported there were no further complications. the patient was successfully weaned from the impella 5.5 during the transplant procedure and underwent a successful heart transplant. the patient survived the procedure. | [5 empty] | Investigation summary: no product returned. asae/hematoma: hematoma post op 5.5. back to or for evacuation.the cause of the access site adverse event was not determined due to insufficient clinical details. device history lot: device lot: 1994108. device history batch: subcomponent lot: n/a. device history review: device sn (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2245 | 1220648-2025-49685 | 23936042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936042&pc=OZD | 12/11/2025 |  |  | 12/31/2025 |  |  | 12/11/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758112 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 84 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. on the third day of support it was reported that the patient experienced an acute frontotemporal infarct with hemorrhage. additionally, it was noted that the patient had gastrointestinal bleeding and limb ischemia. it was noted that the patient died during impella cp support. additional clinical details, including perioperative course, contributing factors, and cause of death, were unavailable at the time of reporting. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: ppae: (ischemia & stroke): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. (major bleed): the cause of the injury was not determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2246 | 1220648-2025-49687 | 23936051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936051&pc=OZD | 12/11/2025 |  |  | 12/31/2025 |  |  | 12/11/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728852 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. during the hrpci, it was reported that the pump abruptly stopped pumping. the medical team was able to restart the pump without issue and completed the case. the pump was weaned and explanted following the procedure. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the patient's outcome at the time of explant was survived. | [5 empty] | The device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed.
2247 | 1220648-2025-49688 | 23936055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936055&pc=OZD | 12/02/2025 | 01/28/2026 |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679666 |  | 000000000001000459 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | False Alarm | Restricted Flow rate | Increase in Pressure |  |  |  |  | Great Vessel Perforation | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that an impella 5.5 was inserted for an aortic root rupture, with a purge flow (pf)of 8-10 and a purge pressure (pp) of approximately 400. on (b)(6) the aortic root ruptured again, and a reoperation was performed using surgical mode. the following day, the pp rose to 3-4 and a pp of approximately 800. by the morning of (b)(6), the pf rose to less than 1ml and the pp rose to more than 1100, triggering a purge blockage alarm. since there were no kinks in the routes and no improvement was observed, the decision was made to remove the pump. even after the alarm, the flow rate remained at 2.0l, and there were no problems with operation. a thrombus was found attached to the outlet area of the removed pump, but it was unclear whether this was from the time of graft removal. | [5 empty] | H6 corrected; problem codes a140508 and a160105 removed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2248 | 1220648-2025-49604 | 23936121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936121&pc=OZD | 12/18/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/18/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680768 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 29 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old male patient who was in septic shock and developed bacteremia. due to futility of care and lack of clinical improvement, there was a plan to withdraw care. infectious disease specialists indicated the bacteremia was not survivable. no additional information regarding patient outcome or timing of care withdrawal was provided. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. the device was not returned for evaluation. (sepsis): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. clinical review: (b)(6) 2025: impella 5.5 was implanted via right axillary/subclavian (b)(6) 2025: patient in septic shock, developed bacteremia. (b)(6) 2025: withdrew care. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2249 | 1220648-2025-49656 | 23936134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936134&pc=OZD | 12/14/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/14/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764040 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 47-year-old male was implanted with an impella cp for support during high-risk cardiac procedure. it was reported that the patient was being transferred from the cardiac cath lab table to bed. during the transfer aic alarm noted ¿in aorta¿. p- level decreased and patient placed back to cath lab table. the doctor was unable to advance impella to correct position, under fluoroscopy. the decision made to remove impella and place new one. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. device in wrong position: the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position when the patient with impella cp was being transferred from the cardiac cath lab table to bed and aic alarm was noted ¿in aorta¿. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted.
2250 | 1220648-2025-49664 | 23936137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936137&pc=OZD | 12/13/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/13/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755685 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 128 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 51-year-old male patient experienced access site bleeding while supported with an impella device. the patient reportedly received heparin and subsequently developed continued oozing despite forward tension suturing and pressure. the sheath was advanced using best practice, and additional suturing was performed. a hemostatic dressing was applied until bleeding improved as the activated clotting time decreased. no additional clinical details were provided. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. major bleed: bleeding noted at impella access site. patient received 30,000 units of heparin. the cause of the access site adverse event was user related as the activated clotting time (acts) were high and bleeding was resolved by forward tension suture and angle matching, reducing heparin. device history review: the pump passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2251 | 1220648-2025-49673 | 23936144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936144&pc=OZD | 12/13/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/13/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732580 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 78-year-old female patient supported with an impella device experienced hemolysis. the patient's hemoglobin and hematocrit reportedly decreased, and laboratory findings were consistent with hemolysis. the impella performance level was reportedly reduced to p-5, and the patient received one unit of blood. a patient device interaction problem was reported. no additional clinical details were provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2252 | 1220648-2025-49684 | 23936162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936162&pc=OZD | 12/11/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/11/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026748305 |  | 000000000001000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 79 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella 5.5 via the right axillary artery following support from an impella cp for escalation of therapy subsequent to the patient¿s lactate rising from 15 to 20 20 mmol/l despite increasing vasoactive support. the impella 5.5 was implanted, and the impella cp was explanted with no reported complications. on the 27th day of impella 5.5 support, it was reported that the placement signal values were lower than the cuff blood pressure and the medical team was concerned there was placement signal drift. there was a reported low placement signal alarm. a point of care ultrasound was ordered to review impella position. it is unknown if the ultrasound was performed, if the pump was repositioned, or if the issue was resolved. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. additional information regarding the impella cp was unavailable at the time of reporting, however, no adverse events or device malfunctions were reported. at the time of report writing, it is noted that the patient remains on impella 5.5 support. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product or data logs returned. placement signal issue: clinical reported placement signal lower than cuff blood pressure and placement signal low alarm. the cause of the issue was not established as no product or logs were returned. device history review: device passed all post sterile inspection checks. | Additional information received for d6 explant. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2253 | 1220648-2025-49689 | 23936238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936238&pc=OZD | 09/12/2025 |  |  | 12/31/2025 |  |  | 09/12/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713205 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 120 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Malposition of Device | Difficult to Advance | [5 empty] | Cardiac Arrest | Low Blood Pressure/ Hypotension | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported difficulty passing the guidewire and impella catheter across the anastomosed and across the valve when implanting an impella 5.5. after multiple attempts, the pump successfully crossed the valve. the pump faced towards the mitral valve apparatus after returning to the intensive care unit (icu). the pump was left in position by the physician¿s choice. axillary washout, ecmo decannulated, and heparin was restarted overnight. patient began showing signs of sepsis. plan to extubate and remove any lines that are not in use. became hypertensive and required levophed. went into cardiac arrest, requiring two rounds of cpr and epinephrine. ct surgery planned this morning and potentially reposition the impella as it was against the wall per echocardiogram, post arrest. intervention needed on right coronary artery but not a good candidate due to tortuosity of vessels. had suction alarms that resolved with volume. liver biopsy performed, blood was given and medications were held. the patient was ultimately explanted successfully and hemostasis was achieved. | [5 empty] | The impella passed all post sterile inspection checks. investigation summary no product was returned. the root cause of the sepsis was unable to be determined due to insufficient clinical details. the cause of the cardiac arrest and hypotension was unable to be determined due to insufficient clinical details. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy and tortuous vessels preventing successful delivery of impella. the cause of positioning issue was most likely use issue related since the pump was verified in correct position prior to leaving cath lab and was found of position after reaching icu.
2254 | 1220648-2025-49690 | 23936241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936241&pc=OZD | 09/13/2025 |  |  | 12/31/2025 |  |  | 09/13/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723866 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 62 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a hematoma so heparin was withheld. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella passed all post sterile inspection checks. no product was returned for investigation. the root cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
2255 | 1220648-2025-49691 | 23936243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936243&pc=OZD | 09/13/2025 |  |  | 12/31/2025 |  |  | 09/13/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711270 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 68 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp had concerns of heparin induced thrombocytopenia and hemolysis. after four days of support the pump was explanted. the patient survived. | [5 empty] | The impella passed all post sterile inspection checks. product was not returned for review. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the thrombocytopenia was not determined due to insufficient clinical details.
2256 | 1220648-2025-49693 | 23936248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936248&pc=OZD | 09/13/2025 |  |  | 12/31/2025 |  |  | 09/13/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615081 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced drainage and a hematoma at the insertion site. the patient went in to supraventricular tachycardia and back to spontaneous circulation without intervention. the pump was explanted and an intra-aortic balloon pump was placed. | [5 empty] | A4 and a5 are unknown. the impella passed all the post sterile inspection checks. the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details.
2257 | 1220648-2025-49694 | 23936249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936249&pc=OZD | 09/14/2025 |  |  | 12/31/2025 |  |  | 09/14/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1815943 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 102 KG | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy and cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. prior to impella insertion, balloon aortic valvuloplasty was performed. a 0.018 guidewire was inserted, and the pump was attempted to be inserted over the wire, however a second bav was required. on the second attempt, a 0.027 guidewire was placed and the impella 5.5 was successfully inserted. it was noted that the medical team was unable to rotate the pump, however the inlet was free of all mitral and wall structures. slight oozing was observed at the impella access site, and the physician noted that it was not related to the anastomosis site, but the amount of heparin given to the patient for the procedure to achieve a high activated closure time. on the 26th day of support, there were reported retrograde flow alarms when the p level was decreased to p3 and the flow was increased on the pump. three days later, retrograde flow alarms were reported to occur intermittently, and self-resolved. on the 33rd day of support, the patient¿s nurse reported that there were battery low alarms from the automated impella controller (aic). connections were all checked and verified they were all in place and appropriate. the issue was self-resolved, and the unit began recharging. the nurse was instructed to plug the aic directly into a wall outlet and to have a back up aic ready in case of failure. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complications. this complaint involves two impella devices: impella 5.5 with serial number (b)(6), automated impella controller with serial number (b)(6). this mdr specifically addresses automated impella controller with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | This mdr is being submitted under the correct device to replace mdr 1220648-2025-46934 which was previously submitted filed with the incorrect device serial number, date of event, and date of awareness due to an administrative error. the initial mdr was filed within the required reporting timeframe, therefore this replacement mdr is not considered late. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
2258 | 1220648-2025-49695 | 23936255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936255&pc=OZD | 09/14/2025 |  |  | 12/31/2025 |  |  | 09/14/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655108 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 135 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Multiple Organ Dysfunction Syndrome | Heart Failure/Congestive Heart Failure | [10 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp went into renal failure and end organ failure with diffuse shock. the pump was explanted and exchanged for an impella rp flex. | [5 empty] | The impella passed all the post sterile inspection checks. the root cause of the cardiac failure was unable to be determined due to insufficient clinical details. the cause of the renal failure and organ failure was not determined due to insufficient clinical details. | Medical safety/clinical reviewed the event. the patient was admitted for presumed st elevation myocardial infarction; left/right heart catheterization demonstrated clean coronaries. the patient experienced cardiac arrest and was placed on biventricular mechanical circulatory support, with left-sided impella support implanted via the right femoral artery during cardiopulmonary resuscitation, followed by impella rp flex implantation via the right femoral vein. return of spontaneous circulation was achieved; however, there was no ventricular pulsatility. the patient was transferred to the intensive care unit on biventricular support with maximal vasopressor and inotropic therapy. the patient remained in refractory shock with anuria, no neurological response, renal failure, and progressive multisystem organ failure all known factors of patients in cardiogenic shock. after discussion with the family, further escalation of care was deferred, care was withdrawn, and the patient expired. conservatively, the death will be reported on both impella devices as the impella devices were in use at time of expiration. date of death: the exact date of death is unknown but has been captured as the date of explant. if actual information becomes known, a supplemental report will be submitted. change in reportability: after review of the case by medical safety, it was determined the impella cp is a death type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. note: h6: device problem/component/investigation coding remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
2259 | 1220648-2025-49697 | 23936256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936256&pc=OZD | 09/19/2025 |  |  | 12/31/2025 |  |  | 09/19/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026725136 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Fluid/Blood Leak | Malposition of Device | Material Split, Cut or Torn |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had placed a 5.5 pump for a pediatric patient admitted in with a known congenital heart history. the pump had malfunctions of purge leak associated with a sleeve tear. as a result, the pump was repositioned. the patient also suffered a major bleed. there was bloody stool, hematemesis, and an aneurysm. the aneurysm was emobilized, heparin was withheld, and mass transfusion protocol was called. afterwards, the patient was febrile. the patient remains on impella support. | [5 empty] | The impella passed all post sterile inspection checks. no product was returned. the cause of the purge leak from the filter was not determined due to no product returned and insufficient clinical details. the cause of the sterile sleeve ripping was not determined due to no product returned and insufficient clinical details. the cause of the device in wrong position was not determined due to no product returned and insufficient clinical details. the cause of the bleeding was not determined due to insufficient clinical details.
2260 | 1220648-2025-49696 | 23936259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936259&pc=OZD | 09/19/2025 |  |  | 12/31/2025 |  |  | 09/19/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739021 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp migrated to the mitral valve. repositioning was not possible so the pump was explanted. no patient harm was reported. | [5 empty] | The impella passed all post sterile inspection checks. no product was returned. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the patient during the transport.
2261 | 1220648-2025-49698 | 23936264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936264&pc=OZD | 09/19/2025 |  |  | 12/31/2025 |  |  | 09/19/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673561 |  | 0048-0014 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed to support a patient with congestive heart failure. the pump was accidentally pulled out of the ventricle. the pump was exchanged for a new impella cp to continue patient support. no patient harm was reported. | [5 empty] | The impella passed all post sterile inspection checks. no product was returned. the cause of the positioning issue was determined to be an unintentional use error as the pump was accidently pulled out of position by the user after the pump was repositioned during implantation.
2262 | 1220648-2025-49699 | 23936267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936267&pc=PYX | 09/20/2025 |  |  | 12/31/2025 |  |  | 09/20/2025 | 12/31/2025 |  |  | 11/30/2025 |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024434826 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp flex experienced a fever and there were concerns of pneumonia. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella passed all post sterile inspection checks. no product was returned for investigation. the cause of the fever was not determined due to insufficient clinical details.
2263 | 1220648-2025-49700 | 23936278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936278&pc=OZD | 09/20/2025 |  |  | 12/31/2025 |  |  | 09/20/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672865 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced suction and bleeding from all sites and orifices. in response, all anticoagulation was stopped. after six days of support the patient was escalated to an impella 5.5 pump. the patient survived. | [5 empty] | A4, a5, and a6 are unknown. the impella passed all post sterile inspection criteria. the cause of the bleeding was not determined due to insufficient clinical details. the cause of the suction was most likely related to the patients condition, as placement signal was observed trending down during the suction events, and the patient was receiving volume and va ecmo support.
2264 | 1220648-2025-49701 | 23936280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936280&pc=OZD | 09/16/2025 |  |  | 12/31/2025 |  |  | 09/22/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025493943 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that implantation of an impella 5.5 was aborted due to the patient's anatomy. the pump was aborted and patient survived. | [5 empty] | A4 is unknown. the impella passed all post sterile inspection checks. no product was returned. the cause of the delivery issue was determined to be patient condition as the patient had calcification at the bifurcation of the right subclavian artery preventing successful delivery of impella.
2265 | 1220648-2025-49702 | 23936282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936282&pc=OZD | 12/18/2025 | 02/11/2026 |  | 12/31/2025 |  |  | 12/18/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026733433 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male | 84 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a hematoma at the insertion site and hematuria. the pump was repositioned and a femostop was applied. the patient is in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2266 | 1220648-2025-49703 | 23936284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936284&pc=OZD | 09/22/2025 |  |  | 12/31/2025 |  |  | 09/22/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695638 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient had presented to the hospital on (b)(6) 2025 for right and left heart catheterization. the patient was unable to lay flat for the procedure and was admitted and found to be in heart failure with atrial fibrillation with rapid ventricular response. while on the medical/surgical unit, the patient experienced an acute onset of shortness of breath and a pulse oxygenation level lower than the limit of detection. the patient was placed on a bilevel positive airway pressure machine and taken to the intensive care unit (icu). upon arrival to the icu, the patient was emergently intubated and upon induction, the patient coded. the patient was canulated for veno-arterial extracorporeal membrane oxygenation (va ecmo) and then taken to the cardiac catheterization laboratory for impella cp placement for left ventricular unloading. on the fourth day of impella cp and va ecmo support, it was noted that the patient¿s heparin was withheld because of concern for the patient¿s hematocrit and hemoglobin dropping. the patient was transfused with one unit of packed red blood cells, and there was oozing noted at the impella access site. the site was re-dressed. on the seventh day of support, the medical team decided to escalate the patient to an impella 5.5 due to the need for prolonged hemodynamic support and for weaning the patient from va ecmo. upon removal of the impella cp the access site was closed with perclose devices. it was noted after closure that the patient had bilateral hematomas at the impella access site and where the arterial ecmo cannula was removed for ecmo reconfiguration from va ecmo to veno-veno extracorporeal membrane oxygenation (vv ecmo). a vascular surgeon performed a cutdown and repair to the arteries and the patient was released to the icu for rest and recovery. on the fifth day of impella 5.5 support, it was reported that the patient spiked a fever, was tachycardiac, and had low blood pressure. no interventions were reported for the events. on the seventh day of impella 5.5 support, vv ecmo was successfully decannulated. the following day, it was noted that the patient was hemodynamically unstable, and the medical team was unable to perform a spontaneous awakening trial. the patient was administered amiodarone for arrhythmias and then cardioverted to normal sinus rhythm. additionally, it was noted that blood cultures were drawn. the patient¿s neurological status continued to be unclear and magnetic resonance imaging was recommended. the patient's code status was changed to do not resuscitate (dnr). an electroencephalogram showed generalized slowing, which was consistent with moderate diffuse encephalopathy, and it was noted that there was no evidence of devastating neurological injury. family chose to withdraw care due to poor neurological prognosis. the patient¿s care was withdrawn, and the patient expired this complaint involves two impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). this mdr specifically addresses pump 1: impella cp, with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Medical safety reviewed the event and noted the following: the patient was placed on venoarterial extracorporeal membrane oxygenation (va ecmo) for worsening heart failure and respiratory failure. an impella cp was inserted for left ventricular unloading. on day four of support, bleeding was noted at the access site, and one unit of packed red blood cells (prbc) was administered due to a drop in hemoglobin and hematocrit. on day seven, the care team escalated support to impella 5.5 with the plan to wean ecmo. impella cp was explanted, and va ecmo was converted to venovenous (vv) ecmo. post-explant, bilateral hematomas were observed at prior insertion sites. five days after impella 5.5 placement, the patient developed fever and tachycardia, consistent with infection/sepsis. native cardiac function continued to decline, and the patient¿s status was changed to do not resuscitate. care was withdrawn following electroencephalogram findings consistent with encephalitis. the hematoma at impella cp insertion site is considered a reportable event. fever and tachycardia during impella 5.5 support represent a reportable infection/sepsis event. neurologic event (encephalitis) while on impella 5.5 support is reportable. death occurred while on impella 5.5 support and is reportable, though clinical assessment suggests death was due to progressive decline rather than device malfunction. outcome: the patient expired following withdrawal of care; no evidence of device malfunction reported. investigation summary-major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5, which is associated with the demise outcome.
2267 | 1220648-2025-49704 | 23936289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936289&pc=OZD | 09/22/2025 | 02/17/2026 |  | 12/31/2025 |  |  | 09/22/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722131 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 99 KG | Non Hispanic | White | Injury | Yes | No | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella 5.5 there was a pump stop. the pump was removed, reflushed, and placed. the wire was observed to be damaged upon removal/replacement and there was bleeding that prompted red blood cells to be transfused. later, the patient was successfully weaned from the pump and survived. | The complainant reported additional information upon request: "the impella 5.5 had no issues after the initial insertion. no replacement of pump was needed. the device was disposed and no other information is available." |  |  |  |  | The impella passed all post-sterile inspection checks. no product was returned. the cause of the temporary pump stop was use related since the 0.018 guidewire was left in the pump when starting it. the cause of the difficulty delivering/advancing was patient condition related as clinical details note anatomical challenges including a sharp turn at the ascending aorta. the cause of the arrhythmia and tachycardia was unable to be determined due to insufficient clinical details. the cause of the bleeding cannot be determined as insufficient clinical details were provided.
2268 | 1220648-2025-49707 | 23936299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936299&pc=OZD | 12/05/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1451015 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the console showed a battery level below 50% after being unplugged. the patient survived the procedure, and no patient harm was reported. | [5 empty] | H3 updated due to device being investigated. h6 codes b21 and c20 were reported incorrectly on the initial report that was submitted. codes b01 and c21 should of been coded on the initial report that was submitted. type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary; the investigation has been completed. data analysis: the logs from the complaint date did not show any battery-related alarms; however, a battery level low alarm was observed on 13 november 2025, which is the case start date, and was triggered shortly after ac power disconnection. long-term power log data revealed a single, abnormal sample of battery capacity. device analysis: the console was sent back for further investigation. an unexpected battery level low alarm could not be reproduced during testing. to replicate the conditions in the clinical setting, the console was repeatedly connected to and disconnected from alternating current power along with known good pump and purge cassette running. throughout these tests, the console operated normally, and the reported battery alarm did not occur. additionally, a thorough visual inspection of the power-related circuitry was performed, and no irregularities were observed. root cause: the root cause for the console displaying a battery level below 50% after being unplugged was most likely an erroneous trigger of the invalid data sample in the smbus communication between the power battery manager and the batteries. however, the specific component responsible was not determined. product history review summary: there was one other issues which was not related to complaint discovered when reviewing the product history of the console. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2269 | 1220648-2025-49706 | 23936300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936300&pc=OZD | 01/01/2025 |  |  | 12/31/2025 |  |  | 09/22/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1855852 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | Device Difficult to Maintain | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that purge issues and a controller failure occurred on an automated impella controller. no patient use was reported with the event. | [5 empty] | There was no patient involved in the reported event. section a was left blank. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | After evaluation, the complaint does not meet the criteria for mdr reportability as defined by 21 cfr 803.50. the issue will be tracked and trended internally. a regulatory report is no longer necessary.
2270 | 1220648-2025-49710 | 23936302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936302&pc=OZD | 12/09/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026787413 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that this was the second pump used. the device was inserted without complication and was operating in the left ventricle. the physician noted that the cannula appeared kinked on itself. it was not positioned in the apex and did not appear to be caught on any structure. the physician stated that the device seemed to be in good position and that there was no resistance during insertion. the pump was subsequently removed. of note, the patient was also on veno-arterial extracorporeal membrane oxygenation (va ecmo). the patient was reported to have survived the procedure. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary device not available. pd-cannula assembly- (twist, bent, kinked): the root cause of the cannula kink cannot be determined since the complaint device not returned for analysis. device history review the pump passed all post sterile inspection checks.
2271 | 1220648-2025-49708 | 23936304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936304&pc=PYX | 09/23/2025 |  |  | 12/31/2025 |  |  | 09/23/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661972 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 127 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp flex experienced bleeding from the swan site and oozing from the pump access site. homeostasis was achieved with stitches, manual pressure, and transfusion of four units of packed red blood cells. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella passed all post sterile inspection checks. no product was returned for investigation. the cause of the bleeding is determined to be patient condition because the patient was bleeding from other location along with impella insertion site.
2272 | 1220648-2025-49712 | 23936305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936305&pc=OZD | 12/09/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026742241 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that the impella cp was plugged into the automated impella controller (aic) and priming was initiated. an alarm appeared indicating that the pump placement signal was not detected. the physician elected to proceed without waiting for an automated impella controller (aic) change-out. the patient was reported to have survived the procedure. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation was completed and the product was returned. placement signal issue: clinical details report a trigger of psnr during case start. the user opted to continue without a ps. no further details were provided. data logs were reviewed and the first rt log showed optical signal metrics aligned with an air gap (lamp maxed out, gain 2.6, snr 0, and the ps railed to 3000). the optical signal metrics and ps resolved for a matter of seconds before becoming abnormal again. imc log review showed that during initial optical sensor calibration, the calibrated g0 was set equal to the factory g0 shortly after a line indicating that the snr was below the minimum threshold calibration procedure. these two things could be related to each other and the air gap noted in the first rt log. the pump was disconnected and reconnected, and the following attempts at reading the eeprom had issues. again, this could be related to the fact that there was an air gap/poor connection of the patient cable plug to the console, though it could not be confirmed. the second rt log shows the ps generally railed to 3000. however, the raw optical signal shows a generally typical pressure trace for a short period when the values aren't railed. during this log and the remainder of the case (per lt log), snr 0, lamp 50%, gain 1.3, light 2.7, so some 5s parameters were abnormal values while others were as to be expected. this no longer aligns with the air gap characteristic. the returned pump was investigated and the optical 5s parameters were within expected range, it passed laser continuity, and when run in a bucket of water, psnr did not reproduce and optical 5s remained stable. review of imc logs confirmed pump usage data was correct and the calibrated. g0 remained equal to factory. since data logs were inconclusive and the returned product did not reproduce any of the noted issues, the cause of the placement signal issue could not be determined. device history review was completed and passed all post-sterile inspection checks.
2273 | 1220648-2025-49711 | 23936307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936307&pc=OZD | 09/24/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1786102 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a loaner automated impella controller was returned from a customer. the pump had difficulty starting up and when it did, it showed a white screen. the controller could only be switched off by emergency shutdown. there was no patient involved in this event. | [5 empty] | The reported event did not involve patient management. section a was left blank. the root cause for the white screen at boot with no bios information, observed during the functional check, was due to a defective aic lcd display.
2274 | 1220648-2025-49714 | 23936310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936310&pc=PYX | 12/09/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1484040 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a controller error notification when starting the console. the patient outcome is unknown at this time. | Additional information received that no patient was involved, the error was noticed during a training at the hospital. |  |  |  |  | E1. initial reporter first and last name is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary: data analysis: the device logs confirmed a controller error (event set #1039) alarm during startup. this alarm is triggered when the light which reaches the ccd sensor array is less than 0.75v, which indicate a malfunction of the optical lamp. the alarm was intermittently triggering. device analysis: upon device boot up, the controller error alarm appeared then cleared on its own. when the alarm was actively triggered, inspection of the optical connector fiber was performed and revealed no visible light at the ferrule face. an optical bench check was performed, and the lamp voltage was reported as 0.51v with test cavity connected. the root cause of the controller error is likely due to a defective optical lamp.
2275 | 1220648-2025-49713 | 23936311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936311&pc=OZD | 09/23/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729389 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced low suction/low flow. upon explant of the pump a large blood clot was found at the inhalation site. | [5 empty] | A4 is unknown. the impella passed all post sterile inspection checks. the root cause of the thrombosis/thrombus was unable to be determined due to insufficient clinical details. the pump was not returned for investigation. the cause of the suction issue was most likely related to biomaterial ingestion, based on the provided clinical information.
2276 | 1220648-2025-49716 | 23936318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936318&pc=OZD | 09/24/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722790 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced high purge pressure. tissue plasminogen activator was added to the purge to resolve the issue. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. high purge pressure: the cause of the high purge pressure could not be determined as no product or log were returned for evaluation. device history review: the device passed all post sterile inspection checks.
2277 | 1220648-2025-49718 | 23936324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936324&pc=PYX | 12/08/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 |  |  | 02/28/2026 |  |  |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025529169 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella device for mechanical circulatory support. it was reported that overnight the patient continued to decompensate with regard to durable lvad flows, despite the surgeon¿s efforts to increase the rp flex (rpf) pump level to improve lvad preload. hemodynamic status worsened, and the decision was made to return to the operating room and transition from rp flex to centrimag support. the surgeon commented that calculated flow readings appeared inaccurate, stating: ¿i had my hand on the pulmonary artery last night with 3 l of ¿flow¿ and no turbulence at all. hard to believe it was doing anything. after placing a central rvad at 3 l of flow, the patient immediately improved. i believe this reflects a limitation of the device.¿ the patient survived the procedure and is currently stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1: added serious injury. h1: added serious injury. h6: omit code e2403 and added code e2343. | H6: added code a0709 based on information in the complaint record. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary no product returned. no data logs available. clinical states surgeon felt no turbulence despite pump saying 3 liters of flow and once the patient got a right ventricular assist device (rvad) with 3 liters they immediately improved. surgeon states that calculated flow is fake. low/blocked pump flow: the cause of the low pump flow was not determined due to lack of product and data logs. placement signal issue: the cause of the placement signal issue was not determined due to lack of product and data logs. device history review device passed all post sterile inspection checks.
2278 | 1220648-2025-49720 | 23936334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936334&pc=OZD | 12/10/2025 | 01/05/2026 |  | 12/31/2025 |  |  | 12/10/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589654 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 lost ao signal and the p level could not be increased. no patient harm was reported. | [5 empty] | D9 updated. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2279 | 1220648-2025-49717 | 23936335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936335&pc=PYX | 09/23/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026712338 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp flex experienced alarms for placement signal not reliable, and for suction. the p level of the pump was adjusted to resolve the alarms. the patient was later successfully weaned from the pump and survived. | [5 empty] | A5 and a6 are unknown. the impella passed all post sterile inspection checks the cause of the placement signal issue was determined to be patient condition related as evidenced by issue occurrence during suction events and issue resolution following p-level reduction. the cause of the pump suction was determined to be patient condition related as evidenced by the issue occurring during high impella p-level support with concurrent ecmo support and with supporting clinical details.
2280 | 1220648-2025-49719 | 23936337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936337&pc=OZD | 09/24/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629835 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 81 KG | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 had suction remedied by fluid delivery and there were signs of infection that prompted the access site to be cut down. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella passed all post sterile inspection criteria. the cause of the access site event could not be determined due to insufficient clinical details. the cause of the suction was most likely related to the patients condition, as placement signal was initially very low and didn't start to improve until after volume was given.
2281 | 1220648-2025-49722 | 23936344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936344&pc=OZD | 12/05/2025 |  |  | 12/31/2025 |  |  | 12/11/2025 | 12/31/2025 |  |  | 12/31/2025 |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025449551 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced aortic valve insufficiency and regurgitation. the pump was explanted and the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2282 | 1220648-2025-49721 | 23936346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936346&pc=OZD | 09/23/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705502 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 became hypotensive and tachycardic. the patient was cardioverted. a washout was performed and a large number of clots were removed from the throacic cavity. bleed/thrombosis that required or washout. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella passed all post sterile inspection checks. no product was returned. the root cause of the paroxysmal atrial fibrillation, tachycardia was unable to be determined due to insufficient clinical details. the root cause of the bleed was unable to be determined due to insufficient clinical details. the root cause of the positioning issue was unable to be determined due to insufficient clinical details.
2283 | 1220648-2025-49724 | 23936347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936347&pc=OZD | 12/10/2025 |  |  | 12/31/2025 |  |  | 12/10/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026720675 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | Company representation | [8 empty] | Injury | No | Yes | Restricted Flow rate | [6 empty] | Unspecified Infection | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced decreased flow. proper pump position was confirmed with ultrasound. the pump was replaced to resolve the issue. the pump could not be retrieved due to infection. | An impella 5.5 was placed for care escalation from a prior intra-aortic balloon pump (iabp) as the iabp was not sufficient for the patient admitted in with acute coronary syndrome and underlying hepatitis c, in scai stage c shock. the age and gender as well as other underlying comorbidities were not shared and are unknown. after approximately 2 days the pump was removed and replaced due purge flow issues that could not be resolved. the purge flow was decreasing. the team did assess the purge flow by reviewing volume needs and pump positioning. a new impella 5.5 was placed and support proceeded. the purge flow and pump replacement will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however the exchange was performed with no consequence to the patient and support. |  |  |  |  | A2, a3, and a4 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Section b5 and h6 ( clinical code ) is updated with all the related information to the event which was inadvertently left out in the initial report. investigation summary: no product or logs returned. low or blocked pump flow: clinical details mention that the pump flow rate decreased suddenly three days after insertion. the pump position and volume assessment was performed using ultrasound and no issues were found. the pump was replaced. the root cause of the low pump flow was not determined due to lack of conclusive evidence from the clinical details and unreturned product and logs for further investigation. hemodynamic instability: clinical details mention that the patient was admitted to the hospital with acute coronary syndrome and was administered an iabp but was upgraded to a impella 5.5 pump due to insufficient circulatory support. the root cause of the injury was not determined due to insufficient clinical details. device history lot: device lot : 1926440. device history batch: subcomponent lot : n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
2284 | 1220648-2025-49723 | 23936348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936348&pc=OZD | 09/23/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677199 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female | 69 KG |  | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced hematuria. the pump appeared to be positioned deep. the patient survived. | [5 empty] | A5 is unknown. the impella passed all post sterile inspection checks. product was not returned for review. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
2285 | 1220648-2025-49725 | 23936359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936359&pc=OZD | 09/24/2025 |  |  | 12/31/2025 |  |  | 09/24/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740225 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Fluid/Blood Leak | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a purge leak. before the issue could be resolved the patient experienced a cardiac arrest and expired. | [5 empty] | A4 is unknown. the impella passed all post-sterile inspection checks. data logs were reviewed. since product wasn't returned for investigation, the exact source of the leak and cause could not be determined. the cause of the cardiac arrest was not determined due to insufficient clinical details.
2286 | 1220648-2025-48654 | 23936386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936386&pc=OZD | 09/29/2025 | 11/29/2025 |  | 12/31/2025 |  |  | 09/29/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026728789 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 72 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Fever | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for escalated mechanical circulatory support from an impella cp. the patient had presented to the hospital on (b)(6) 2025 for right and left heart catheterization. the patient was unable to lay flat for the procedure and was admitted and found to be in heart failure with atrial fibrillation with rapid ventricular response. while on the medical/surgical unit, the patient experienced an acute onset of shortness of breath and a pulse oxygenation level lower than the limit of detection. the patient was placed on a bilevel positive airway pressure machine and taken to the intensive care unit (icu). upon arrival to the icu, the patient was emergently intubated and upon induction, the patient coded. the patient was canulated for veno-arterial extracorporeal membrane oxygenation (va ecmo) and then taken to the cardiac catheterization laboratory for impella cp placement for left ventricular unloading. on the fourth day of impella cp and va ecmo support, it was noted that the patient¿s heparin was withheld because of concern for the patient¿s hematocrit and hemoglobin dropping. the patient was transfused with one unit of packed red blood cells, and there was oozing noted at the impella access site. the site was re-dressed. on the seventh day of support, the medical team decided to escalate the patient to an impella 5.5 due to the need for prolonged hemodynamic support and for weaning the patient from va ecmo. upon removal of the impella cp the access site was closed with perclose devices. it was noted after closure that the patient had bilateral hematomas at the impella access site and where the arterial ecmo cannula was removed for ecmo reconfiguration from va ecmo to veno-veno extracorporeal membrane oxygenation (vv ecmo). a vascular surgeon performed a cutdown and repair to the arteries and the patient was released to the icu for rest and recovery. on the fifth day of impella 5.5 support, it was reported that the patient spiked a fever, was tachycardiac, and had low blood pressure. no interventions were reported for the events. on the seventh day of impella 5.5 support, vv ecmo was successfully decannulated. the following day, it was noted that the patient was hemodynamically unstable, and the medical team was unable to perform a spontaneous awakening trial. the patient was administered amiodarone for arrhythmias and then cardioverted to normal sinus rhythm. additionally, it was noted that blood cultures were drawn. the patient¿s neurological status continued to be unclear and magnetic resonance imaging was recommended. the patient's code status was changed to do not resuscitate (dnr). an electroencephalogram showed generalized slowing, which was consistent with moderate diffuse encephalopathy, and it was noted that there was no evidence of devastating neurological injury. family chose to withdraw care due to poor neurological prognosis. the patient¿s care was withdrawn, and the patient expired. this complaint involves two impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). this mdr specifically addresses pump 2: impella 5.5, with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Medical safety reviewed the event and noted the following: the patient was placed on venoarterial extracorporeal membrane oxygenation (va ecmo) for worsening heart failure and respiratory failure. an impella cp was inserted for left ventricular unloading. on day four of support, bleeding was noted at the access site, and one unit of packed red blood cells (prbc) was administered due to a drop in hemoglobin and hematocrit. on day seven, the care team escalated support to impella 5.5 with the plan to wean ecmo. impella cp was explanted, and va ecmo was converted to venovenous (vv) ecmo. post-explant, bilateral hematomas were observed at prior insertion sites. five days after impella 5.5 placement, the patient developed fever and tachycardia, consistent with infection/sepsis. native cardiac function continued to decline, and the patient¿s status was changed to do not resuscitate. care was withdrawn following electroencephalogram findings consistent with encephalitis. the hematoma at impella cp insertion site is considered a reportable event. fever and tachycardia during impella 5.5 support represent a reportable infection/sepsis event. neurologic event (encephalitis) while on impella 5.5 support is reportable. death occurred while on impella 5.5 support and is reportable, though clinical assessment suggests death was due to progressive decline rather than device malfunction. outcome: the patient expired following withdrawal of care; no evidence of device malfunction reported. note: regarding the neurological event, the appropriate health effect clinical code for this device is unidentified at the time of this mdr writing. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding tachycardia, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. regarding the fever and hypotension, the cause of the injury was not determined due to insufficient clinical details. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp.
2287 | 1220648-2025-49727 | 23936398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936398&pc=OZD | 11/29/2025 |  |  | 12/31/2025 |  |  | 12/02/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026720674 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced an increase in purge pressure and a pump stop. the pump was replaced with another impella 5.5. the patient was in critical condition. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product or logs were returned for evaluation. purge pressure high: clinical details noted that there was an increase in purge pressure 1 day into support. sodium bicarbonate was used in the purge solution. the cause of the high purge pressure could not be determined as no product or logs were returned for analysis and limited clinical details were provided. pump stop: clinical details noted that high purge pressure followed by pump stop occurred 1 day into support. no additional details were provided. the cause of the pump stop could not be determined as no product or logs were returned for evaluation and limited clinical details were provided. device history review: the pump passed all the post sterile inspection checks.
2288 | 1220648-2025-49726 | 23936402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936402&pc=OZD | 09/30/2025 |  |  | 12/31/2025 |  |  | 09/30/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688008 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Arrhythmia | Heart Block | Cardiovascular Dysfunction/Insufficiency | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to heart transplant. on day 96 of impella 5.5 support it was reported that the patient¿s cardiac rhythm was sinus bradycardia to sinus rhythm with first degree atrioventricular block. the patient was noted as hemodynamically stable and continues to wait for a heart transplant. on day 106 of impella 5.5 support, it was reported that there was an alarm for placement signal not reliable which began two days prior. the medical team disabled the audio alarm. on day 111 of impella 5.5 support, the placement signal loss continued, and it additionally noted that there were issues with loss of impella function. the medical team monitored hemodynamic trends, laboratory results, and overall clinical status. on day 112 of impella 5.5 support, it was noted that the patient was implanted with an impella rp flex after developing right ventricular dysfunction. the patient remains on impella 5.5 and impella rp flex support as of the time of this report awaiting heart transplant. | [5 empty] | Investigation summary: optical sensor issue: clinical reported that placement signal not reliable (psnr) alarm triggered on (b)(6) 2025 without resolution for the remainder of the case. the product was not returned for investigation. the case was just over four months long, and data logs were not provided for the time which the psnr alarm reportedly triggered on (b)(6) 2025. throughout support, psnr is noisy and placement signal increases on multiple occasions, triggering impella position unknown alarms and triggering psnr once. during the psnr alarm on (b)(6) 2025, placement signal had increased to ~139mmhg before going out of range. when placement signal goes out of range, psnr and contrast drop, but gain and light were stable. motor current was also stable and pulsatile throughout the entire case. the pump run on the same console prior to this case had stable 5s parameters throughout the case. the root cause of the optical sensor issue could not be determined due to insufficient data logs during the reported event and no product returned. device history review: pump sn (b)(6) passed all post sterile inspection criteria. qn # (b)(4) was opened because pump sn (b)(6) had a white particle inside the pouch. the pump was reworked and passed all post sterile inspection criteria.
2289 | 1220648-2025-49611 | 23936410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936410&pc=OZD | 12/05/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/06/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026757994 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the impeller was located in the aorta and was pushed back into the left ventricle by 1.5 cm. following this, a ¿motor current too high¿ alarm was triggered, and the pump ceased operation. three attempts to restart the pump were made, all of which were unsuccessful. additional information indicated that the patient was stable the whole time and was not harmed by the pump stop. new impella was implanted without any problems via left axillary artery. now pump runs without any issues and act is still in range. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the cause of the pump stop was bearing wear with biomaterial potentially being a secondary cause due to the large purge gap seen on the returned pump, supported by the data logs. as this issue occurred on day 10 of support, it is beyond the design life of a cp pump. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2290 | 1220648-2025-49730 | 23936418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936418&pc=OZD | 12/10/2025 |  |  | 12/31/2025 |  |  | 12/10/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755362 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced suboptimal positioning. the pump was repositioned. hemolysis and hematuria occurred. | [5 empty] | The investigation was completed and no product returned. no data logs available. hemolysis/mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2291 | 1220648-2025-49728 | 23936424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936424&pc=OZD | 10/17/2025 |  |  | 12/31/2025 |  |  | 12/03/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026668181 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a the impella cp was not recognized by the automated impella controller. the pump was not implanted and no patient harm was reported. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product return. logs are not available. failure to detect pump: clinical details note that during pump setup, the impella could not be detected by the aic after three attempts. they did not try a new console/pump due to the patient's improving condition. the cause of the detection failure was unable to be determined since product/logs were not returned. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2292 | 1220648-2025-49731 | 23936432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936432&pc=OZD | 12/10/2025 | 02/12/2026 |  | 12/31/2025 |  |  | 12/10/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768831 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis and hematuria. the patient is in stable condition. the pump was repositioned to resolve the issue. the patient is in stable condition. | [5 empty] | The investigation for the hemolysis, hematuria, and positioning issues has been completed. the cause of hemolysis/mechanical interaction with blood was most likely due to pump was not in proper position since clinical team confirmed pump was deep & hemolysis was improved after repositioning of the pump. the cause of positioning issue was unknown. the cause of the injury was not determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2293 | 1220648-2025-49732 | 23936492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936492&pc=PYX | 12/10/2025 | 02/25/2026 |  | 12/31/2025 |  |  | 12/11/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026671817 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 64 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Hematuria | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced bleeding issues prior to implantation that continued during support, including hematuria. the patient was transfused multiple units of packed red blood cells, platelets, and fresh frozen plasma. the patient also experienced placement signal not reliable alarms. the alarm was disabled. the patient was put on continuous renal replacement therapy and burmex gtt in response to renal failure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2294 | 1220648-2025-49733 | 23936502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936502&pc=OZD | 12/10/2025 | 01/30/2026 |  | 12/31/2025 |  |  | 12/11/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026741837 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Positioning Problem | Patient Device Interaction Problem | [5 empty] | Fever | Hemolysis | Unspecified Infection | No Clinical Signs, Symptoms or Conditions | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient implanted with an impella 5.5 experienced intermittent placement signal alarms and the ao signal did not correlated with the non-invasive blood pressure readings. an echocardiogram was performed to confirm the pump was in the correct position. no patient harm was reported. | It was further reported that the patient was febrile and had blood cultures positive for gram-negative rods. additionally, it was noted that the plasma free hemoglobin was rising and the patient is intermittently alarming "placement signal low." |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b updated. the impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received: b1, b5, h1, and h6 updated based on the additional information received from the clinical site. | D9 updated. the product has been returned. the investigation is still ongoing.
2295 | 1220648-2025-49736 | 23936522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936522&pc=OZD | 12/10/2025 |  |  | 12/31/2025 |  |  | 12/10/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026757912 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male | 129 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced brown drainage at the access site. a surgical washout was performed. the patient is in stable condition. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary no product returned. asae/major bleed: serous/brown drainage was noted at the impella axillary site, with dressing saturation approximately every 24 hours. blood cultures drawn over the prior 48 hours remained negative, and the patient was planned for or axillary site washout with cardiothoracic surgery. the cause of the access site adverse event was not determined due to insufficient clinical details. device history lot device lot: 1985045 device history batch subcomponent lot: n/a device history review the pump sn (b)(6) passed all the post sterile inspection checks.
2296 | 1220648-2025-49738 | 23936534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936534&pc=OZD | 12/07/2025 |  |  | 12/31/2025 |  |  | 12/10/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026735123 |  | 000000000001000459 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support following mitral regurgitation and worsening heart failure while supported by an intra-aortic balloon pump that was placed subsequent to coronary artery bypass graft surgery. on support day 12, it was reported that there was a continuous purge fluid leak at the purge side arm of the impella 5.5. there was no fluctuation in the purge flow or purge pressure when the device was observed over 12 hours. however, purge fluid continued to leak. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the device is reportedly available for return; however, as of the time of this report the device has not yet been returned for evaluation. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This follow-up report is being submitted to correct a previously reported h6 medical device problem code. the problem code a141102 (increase in pressure) reported on the initial mdr was not applicable to the event and has been corrected to a050401 (fluid/blood leak). | H6: updated codes to reflect the results of the investigation. investigation summary: fluid leak-side arm: the cause of the sidearm leak was an insufficient radial seal of the reservoir membrane.
2297 | 1220648-2025-49740 | 23936538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936538&pc=OZD | 12/08/2025 | 01/05/2026 |  | 12/31/2025 |  |  | 12/10/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704345 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male | 81 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced an unexpected pump stop, most likely due to thrombosis. the pump was explanted and no patient harm was reported. | The complainant also reported that multiple thrombi were observed. no photos were captured. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary: the pump was not returned for investigation. data log review was not provided or retrieved from the remote server. pump stop: the cause of the pump stop issue was not determined without returned product investigation and sufficient clinical details, including data log review. thrombosis: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot : 1949026. device history batch: subcomponent lot : n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks. | B5 updated. h6 updated. the impella device has not been received from the customer; therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
2298 | 1220648-2025-49742 | 23936539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936539&pc=OZD | 12/11/2025 |  |  | 12/31/2025 |  |  | 12/11/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1131451 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) in use with an impella pump for patient support would not recognize the purge pressure disk during a cassette change and would not advance to the next step. the aic was exchanged with another unit to resolve the issue. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the patient's outcome at the time of impella explant is unknown. the device is reportedly available for return; however, as of the time of this report the device has not yet been returned for evaluation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Correction: h5 and h8 was erroneously selected on the initial manufacturer device report the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. d9 updated with device receipt date. data review: console logs were consistent with what was reported in the complaint. device analysis: the console was tested and the reported purge disc recognition issue was able to be reproduced. further inspection of the console revealed that the purge disc flag was broken (missing at blue retainer). conclusion: the root cause of the purge disc detection issue was a broken purge disc flag. device history review: the device passed all post sterile inspection checks.
2299 | 1220648-2025-49745 | 23936577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936577&pc=OZD | 12/12/2025 |  |  | 12/31/2025 |  |  | 12/13/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724581 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Vascular Dissection | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was presented emergently to cath lab due to multiple cardiac arrest. impella placed post percutaneous coronary intervention as bail out. no distal pulses to right lower extremity (rle) and leg mottled overnight. computed tomography obtained showing bleeding from fem artery. taken to operating room with vascular surgery and impella removed after wire access via side port. good hemostasis with 1 preclose perclose and 1 additional perclose. sluggish flow at common femoral artery - artery ballooned and nitroglycerin injected. pulses present but not satisfactory. fasciotomy of rle performed for concerns of compartment syndrome with improvement in rle color and pulses. 6 fr sheath left in place to support flow. medical doctor reported that the tightness from the fem stop was likely a factor in the limb ischemia along with patients large body mass index and restlessness which made it difficult to preserve angle of entry. the pump was later explanted and the patient survived. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product returned. ischemia: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. peripheral artery dissection: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of the access site adverse event was user related to patient restlessness making it difficult to preserve the angle of entry and tight femstop application. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2300 | 1220648-2025-49709 | 23936591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936591&pc=OZD | 12/08/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1291642 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic) failure alarm occurred. the patient outcome is still to be determined as the patient remains on support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2301 | 1220648-2025-49734 | 23936653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936653&pc=OZD | 12/05/2025 | 02/25/2026 |  | 12/31/2025 |  |  | 12/10/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026777402 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that the signal for position detection is not stable alarm occurred due to a temporary unstable position signal. as the unstable signal persisted, a control device abnormality alarm occurred as a secondary response. after approximately one hour, the position signal stabilized, and both alarms cleared. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2302 | 1220648-2025-49748 | 23936659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936659&pc=OZD | 12/16/2025 | 01/28/2026 |  | 12/31/2025 |  |  | 12/16/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026773647 |  | 000000000001000100 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 79 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient has been on intra-aortic balloon pump (iabp) and venoarterial extracorporeal membrane oxygenation (va ecmo) for a prolong period of time. the patient was also on continuous renal replacement therapy for renal impairment. the patient also suffered cerebrovascular accident during this hospital admission. plan of care for patient was to place impella 5.5, remove iabp, and wean va ecmo support to explant, leaving imepella 5.5 for hemodynamic support. after multiple days of heparin therapy patient was brought back for attempt at impella 5.5 placement. access was achieved through standard right axillary graft, placement of wire/catheter was difficult due to little atrioventricular (av) movement. finally access across the av was achieved. impella was loaded on biomed 0.018 and advanced into the graft and made it completely past the anastomosis, confirmed by medical doctor, but would not advance further. multiple attempts were made to advance the impella, pulling back and advancing, torquing and advancing to no advancement of the device. the medical doctor advised quite a bit of resistance and believed vessel was to small to advance. the medical doctor removed the device from the graft with intent of covering in propofol and trying again. after removing the impella there was biomaterial in the inlet and then the doctor aborted the procedure due to concern of damage to the vessel itself. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary no product was returned. thrombosis: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. failure to advance: the cause of the failure to advance was determined to be an adverse event related to the anatomy of the patient since patient had calcification and md believed vessel was too small to advance. device history lot device lot: 1991922 device history batch subcomponent lot: n/a device history review device sn: (b)(6) passed all post sterile inspection checks.
2303 | 1220648-2025-49749 | 23936664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936664&pc=PYX | 12/09/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026767500 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 68 KG |  |  | Death | No | Yes | Device Sensing Problem | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported psnr. it was also reported that the patient care was withdrawn, and the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2304 | 1220648-2025-49750 | 23936665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936665&pc=OZD | 12/09/2025 |  |  | 12/31/2025 |  |  | 12/09/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026769340 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 69 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support it was reported that the patient developed left lower extremity ischemia. unable to remove impella due to hemodynamic needs. the medical doctor did a external femoral bypass to restore flow to the leg. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2305 | 1220648-2025-49751 | 23936678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936678&pc=OZD | 12/11/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/11/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717806 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombocytopenia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male patient whose platelet count was noted to have decreased to 103. the patient was a recent cardiac surgery patient. impella waveforms and flows were noted to be adequate, with no recent suction events observed. device position had been optimized under echocardiographic guidance. no additional information was provided. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary- device in wrong position: the pump was not returned for investigation. data log was not provided or retrieved from remote server. as per the clinical details, the position was optimized under echo guidance. the cause of the positioning issue was not determined without returned product investigation and sufficient clinical details. clinical symptom (thrombocytopenia): the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot- device lot: 1951106. device history review- device sn (b)(6) passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2306 | 1220648-2025-49739 | 23936683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936683&pc=OZD | 11/02/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/10/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725941 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Damaged by Another Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an implanted impella cp was not recognized by the automated impella controller. it was suspected there was electromagnetic interference. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2307 | 1220648-2025-49735 | 23936688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936688&pc=OZD | 12/05/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/10/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1316945 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the signal for position detection is not stable alarm occurred due to a temporary instability in the position detection signal. as this condition persisted, a control device abnormality alarm occurred as a secondary safety response. after approximately one hour, the position detection signal stabilized, and as a result, both alarms cleared. | [5 empty] | Correction: d3 manufacturer contact information was erroneously entered on the initial manufacturer device report. d6a and d6b should have been left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2308 | 1220648-2025-49753 | 23936730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936730&pc=OZD | 12/10/2025 |  |  | 12/31/2025 |  |  | 12/10/2025 | 12/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758057 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. following the successful wean and explant of the device, it was reported that the perclose device failed, and a 14fr sheath was placed at the groin access site and manual pressure was held to resolve the bleeding. the patient was transferred to a bed for transport to the cardiac stepdown unit following the impella explant. subsequently, the patient became very disoriented, stopped responding to commands, and became unresponsive. the stroke team and code team were called, and the patient coded. the patient coded for 30 minutes until time of death was called, and the patient expired. the physician stated that they believed the cause of death to be a brain bleed, and they did not believe that the death was related to impella use. additional clinical details, including perioperative course and contributing factors, were unavailable at the time of reporting. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the reported event, access site bleeding, was found to be use related as bleeding occurred after failure of the perclose. | Medical safety/clinical reviewed the event. the event describes a patient undergoing high-risk percutaneous coronary intervention (pci) with successful pre-pci implantation of an impella cp device. following the procedure, perclose closure was attempted and was unsuccessful, with bleeding noted at the groin access site. the activated clotting time was reported as 277 seconds, and a 14 french sheath was placed. during transfer to a bed for transport to the step-down unit, the patient became disoriented, was not responding to commands, and subsequently became unresponsive. cardiopulmonary resuscitation was initiated, and the patient coded for approximately 30 minutes before expiring. additional information from the field clinical manager indicated that the treating team believed the cause of death was a brain bleed, and the physician did not consider the event to be impella related. the impella cp will remain a death type of reportable event. h6. investigation type/finding/conclusion remains unchanged. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2309 | 1220648-2025-49752 | 23936809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23936809&pc=OZD | 12/11/2025 | 12/31/2025 |  | 12/31/2025 |  |  | 12/11/2025 | 12/31/2025 |  |  |  |  | 12/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674621 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male patient who arrived to the intesive care unit complaining of leg pain in the limb with the impella device. a doppler ultrasound was ordered and showed no flow to the affected leg. the patient was reported to be stable and not requiring pressors. as device removal had already been planned for later that day, the decision was made to explant the impella device at that time in the cardiac catheterization lab. no additional information was provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. ppae (ischemia): in order to make a cause determination, clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
2310 | 1220648-2025-49541 | 23920962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23920962&pc=OZD | 12/14/2025 | 01/27/2026 |  | 12/30/2025 |  |  | 12/15/2025 | 12/30/2025 |  |  |  |  | 12/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754033 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 104 KG |  | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis. the p level was decreased to improve the patient's condition. | [5 empty] | Investigation summary: no product or data logs returned. hemolysis/mechanical interaction with blood: the cause of the hemolysis was not determined due to lack of product and data logs available. device history lot: device lot: 1973969. device history batch: subcomponent lot: n/a. device history review: the complaint pump s/n (b)(6) passed all post sterile inspection checks. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2311 | 1220648-2025-49542 | 23921037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23921037&pc=OZD | 12/15/2025 |  |  | 12/30/2025 |  |  | 12/15/2025 | 12/30/2025 |  |  |  |  | 12/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026775969 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 128 KG | Hispanic |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Unspecified Infection | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced elevated purge pressure and decreased purge flows. the patient was in stable condition. | Clinical review/rationale: an impella 5.5 was placed via the axillary artery graft to support the 34 year old male who had been admitted in with acute myocardial infarction and cardiogenic shock, presenting in scai stage d shock. the patient was also supported by pharmacological support and ecmo. the pump did have elevated purge pressures and decreased purge flow, which they troubleshot by adding the lytic tpa to the purge. the pump was explanted after 25 days and went home with an lvad supporting. the patient was then re-admitted to the hospital with concerns of a harm/pocket developing at the axillary site. the team sent cultures to evaluate for infection, and to date the culture results are pending. infection has not been confirmed to date. |  |  |  |  | Section d6b (explant date ) was added. investigation summary ==> no product or data logs were returned for evaluation. high purge pressure: clinical details noted that pump experienced elevated purge pressure and decreased purge flows. tpa was added to purge but it was not noted if purge issue resolved. the cause of the high purge pressure could not be determined as no product or logs were returned for evaluation and limited clinical details were provided. device history review ==> the complaint pump s/n 638234 passed all post sterile inspection checks. | A6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1 and b2 selections were added due to new information received. b5 clinical rationale was added due to new information that was received. h6 codes were updated due to new information received. h11 added investigation summary for the new codes added. infection: clinical details noted that cultures were taken due to patient having "pocket" at old impella site approximately 1 month after explant. in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. medical device site reaction: clinical details noted patient was re-admitted due to "pocket" at old impella site. the cause of the injury was not determined due to insufficient clinical details.
2312 | 1220648-2025-49545 | 23921204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23921204&pc=OZD | 11/30/2025 |  |  | 12/30/2025 |  |  | 11/30/2025 | 12/30/2025 |  |  |  |  | 12/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685384 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female | 65 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Life Threatening | [27 empty] | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. on day 4 of support, the patient experienced atrial fibrillation subsequent to the performance level being decreased. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | A5, a6 are unknown. investigation summary: ppae (arrhythmia). in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1931603. device history batch: subcomponent: n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
2313 | 1220648-2025-49546 | 23922218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23922218&pc=OZD | 11/29/2025 |  |  | 12/30/2025 |  |  | 12/01/2025 | 12/30/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026687829 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient developed abdominal bleeding and a large hematoma was found on ct scan. no surgical intervention was required, and hemoglobin levels as well as hemodynamics remained stable. the bleeding was determined to be unrelated to impella therapy. the physician reported that the patient death was not related to the device or therapy; the patient reportedly expired due to electromechanical decoupling. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. minor bleed: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2314 | 1220648-2025-49547 | 23922428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23922428&pc=OZD | 11/29/2025 |  |  | 12/30/2025 |  |  | 11/30/2025 | 12/30/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674498 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Multiple Organ Dysfunction Syndrome | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient had acute myocardial infarction/cardiogenic shock and put on impella cp support. an impella cp was inserted into the right lower limb. at that time, an antegrade sheath was used in the left lower limb to provide ischemic care. subsequently, the patient's cardiac function did not improve, and, in accordance with the family's wishes, a care policy was adopted. later the patient's right lower limb became pale, and lower limb ischemia was observed. obstruction of the antegrade sheath route was also observed. with progression of multi organ failure, the decision was made to remove impella. the doctor commented that the antegrade sheath route had become obstructed, causing lower limb ischemia. the impella was removed, and the patient's lower limbs are showing signs of improvement. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the reported injuries, ischemia and organ failure, could not be determined due to insufficient clinical details. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2315 | 1220648-2025-49549 | 23924094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23924094&pc=OZD | 09/02/2025 |  |  | 12/30/2025 |  |  | 09/02/2025 | 12/30/2025 |  |  |  |  | 12/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714432 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 61 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | Respiratory Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced right ventricle disfunction, suction alarms, and decreased flows. vasopressors and inotropes were given to support the patient, and the patient was also implanted with an impella rp flex. the p level of the pump was reduced and the patient experienced anemia, so one unit of blood was transfused. amiodarone and bivalirudin were administered. the patient experienced hemodynamic insufficiency, and could not be weaned from the pump. dobutamine was added to the treatment plan. continuous positive airway pressure was given to the patient to improve respiratory distress and poor oxygenation. norepinephrine was added to the treatment plan. the patient was weaned from the pump and is in stable condition. | [5 empty] | No product was returned for investigation. data logs were returned for analysis. the pump was run on two consoles, and the lt log for the first console was not available. rt logs for the entire case were returned. suction alarms with a drop in pump flow were confirmed at the beginning of the case around 3pm on (b)(6) 2025 based on rt log analysis. the lt log for the case run on the second console had no abnormalities. placement signal and motor current showed no abnormal trends. clinical details revealed that the patient suffered from rv disfunction shortly after the weaning of cpb and initiation of impella support, with suction alarms and low flows on the pump requiring vasopressors/inotropes to support the patient. based on these findings decision was made to proceed with rp flex insertion in the setting of bi-ventricular failure. on (b)(6) 2025, the p-level for the pump was turned down from p-7 to p-6. the patient was administered a unit of blood due to a drop in their hemoglobin. on (b)(6) 2025, it was observed that the patient was not tolerating the drop in p-level, becoming symptomatic and affecting their hemodynamics. ejection fraction was noted to be 20%, which improved to approximately 30% on (b)(6) 2025. weaning to p-5 was tried on (b)(6) 2025, but patient was unable to tolerate. cpap was started on (b)(6) 2025 to improve respiratory distress and poor oxygenation. gastrointestinal bleeding was discovered on (b)(6) 2025. pump suction: suction alarms were confirmed at the beginning of the case through data logs. clinical details mention that the patient suffered from rv disfunction, pump had suction alarms, and patient required blood products to be administered throughout the case. the root cause of the pump suction was most likely patient condition related based on the provided clinical details and data log analysis. low or blocked pump flow: low flows were confirmed when suction alarms were triggered through data logs. clinical details mention that the patient suffered from rv disfunction, pump had low flows, and patient required blood products to be administered throughout the case. the root cause of the low pump flow was most likely patient condition related based on the provided clinical details. anemia, respiratory failure, major bleed: clinical details mention that the patient was administered a unit of blood due to a drop in their hemoglobin. cpap was started on (b)(6) 2025 to improve respiratory distress and poor oxygenation. gastrointestinal bleeding was discovered on (b)(6) 2025. the root cause of the injury was not determined due to insufficient clinical details. the device passed all post-sterile inspection checks.
2316 | 1220648-2025-49550 | 23924347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23924347&pc=OZD | 10/28/2025 |  |  | 12/30/2025 |  |  | 10/28/2025 | 12/30/2025 |  |  |  |  | 12/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026735941 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed for a patient undergoing high-risk percutaneous coronary intervention. the site had a bleed that was not able to be treated by manual hold. the bleeding required intervention in the form of red blood cell infusion, pump removal and a 16fr sheath to be placed, surgeon became involved with hemostasis and patient survived. the team had to apply a stich after the 16fr placed. the patient survival outcome was noted as survived. | [5 empty] | The investigation was completed and the device was discarded. investigation summary: major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review was completed and passed all the post sterile inspection checks.
2317 | 1220648-2025-49551 | 23924898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23924898&pc=OZD | 09/15/2025 |  |  | 12/30/2025 |  |  | 09/15/2025 | 12/30/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1163521 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a positioning alarm on the supporting automated impella controller. the pump was confirmed to be in proper position so the alarm was disabled. the patient is in stable condition. | [5 empty] | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2318 | 1220648-2025-49554 | 23925327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23925327&pc=OZD | 12/07/2025 |  |  | 12/30/2025 |  |  | 12/07/2025 | 12/30/2025 |  |  |  |  | 12/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026710666 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male | 100 KG | Non Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella was repositioned overnight under echo guidance due to dark urine. urine color is now clear/yellow. additional information was received that indicated that the was patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: pump passed all post sterile inspection checks.
2319 | 1220648-2025-49478 | 23926162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23926162&pc=OZD | 11/11/2025 |  |  | 12/30/2025 |  |  | 11/11/2025 | 12/30/2025 |  |  |  |  | 12/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026770550 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced oozing at the groin. the pump was repositioned and the patient was transfused three units of packed red blood cells. the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2320 | 1220648-2025-49423 | 23926333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23926333&pc=OZD | 11/14/2025 | 12/30/2025 |  | 12/30/2025 |  |  | 11/16/2025 | 12/30/2025 |  |  |  |  | 12/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758798 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 125 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 47 yr old male was implanted with a impella cp for support during a high risk cardiac procedure. it was reported that oozing at impella site and minimal red urine was noted. pressors weaned down and pump able to wean to p2. plan to remove cp and replace with intra-aortic balloon pump. | [5 empty] | Correction: e4 and h6 medical device problem code was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. major bleed: the cause of the access site adverse event was not determined due to insufficient clinical details and no product was returned for review. device history review: the device passed all post sterile inspection checks. | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2321 | 1220648-2025-49557 | 23926475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23926475&pc=OZD | 12/08/2025 |  |  | 12/30/2025 |  |  | 12/08/2025 | 12/30/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026758621 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during a call with the md in the icu, the patient was in ct. during transport, an alarm was triggered indicating air in the purge system. after deairing, the messages ¿purge system blocked¿ and ¿purge pressure too high¿ appeared. the purge system was replaced, but this did not resolve the issue. it was recommended to monitor the motor current (mc) and replace the pump as soon as possible. the patient is now back on ecmella therapy. various options were discussed in detail, including thrombolysis, which was ruled out; the team intends to replace the pump as soon as possible. close monitoring of the mc was again recommended, and the hotline was emphasized. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety/clinical reviewed the event. a patient in cardiogenic shock secondary to an acute myocardial infarction presented and required mechanical circulatory support. the patient was classified as society for cardiovascular angiography and interventions (scai) stage e cardiogenic shock and remained scai stage e throughout support. an impella cp (pump 1) device was implanted on 11/26/2025 to provide mechanical circulatory support and was maintained until 12/8/2025, at which time it was explanted as a bridge to an impella 5.5 (pump 2). the impella 5.5 (pump 2) was implanted on 12/8/2025. later that day, while the patient was being transported to ct imaging, an alarm occurred indicating air in the purge system. after de-airing was performed, additional alarms were displayed, including ¿purge system blocked¿ and ¿purge pressure too high.¿ the purge system was replaced; however, the alarms persisted and the issue was not resolved. clinical support recommended continued monitoring of motor current and replacement of the pump as soon as possible. the patient was supported with ecmo and impella therapy during this period. on 12/9/2025 at approximately 08:02, the impella 5.5 device was explanted and replaced with an impella cp (pump 3). the impella cp was set to p-7 for left ventricular unloading, while extracorporeal life support (ecls) was maintained at approximately 3.3 l/min for systemic support, with the goal of patient stabilization and reassessment of clinical management. following implantation of the impella cp (pump 3), the device initially displayed no aortic placement signal and no left ventricular waveform. motor current remained stable and pulsatile. after the impella cp plug was fully plugged into the automated impella controller (aic), the placement signal appeared and remained stable. no sensitivity to cable movement was observed. the issue was considered corrected and resolved, and no further device performance concerns were reported for pump 3. ultimately the patient did expire on support. impella cp pump 3 is the death type of reportable event. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
2322 | 1220648-2025-49333 | 23926661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23926661&pc=OZD | 11/18/2025 |  |  | 12/30/2025 |  |  | 11/18/2025 | 12/30/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623441 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Pump | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a was implanted with an impella cp device for mechanical circulatory support. it was reported that the device was inserted into the patient's body before setup was complete. it was noted that the device would not start. the impella cp was removed and successfully replaced with a new impella cp. no patient harm was reported in relation to this event. additional information regarding the impella cp and the final patient outcome was not available in the complaint record accessible for mdr assessment at time of reporting. | [5 empty] | Additional information was received, and the h6 medical code was corrected from a141204 to a141203. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. pump did not start: clinical reported cp was inserted in body before setup process was complete. upon starting the pump it would not start. the rt logs show an alarm #13 motor current spike on pump start. the imc logs show that priming was completed. there was no placement signal reading before pump start attempt. the cause of the issue was determined to be use related as clinical details reveal pump was inserted prematurely prior to completion of setup. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2323 | 1220648-2025-49560 | 23926771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23926771&pc=OZD | 12/04/2025 |  |  | 12/30/2025 |  |  | 12/04/2025 | 12/30/2025 |  |  |  |  | 12/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED EUROPE GMBH - 3007904765 |  |  | US |  | 1379919 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during the transfer for hospital relocation, the control device ((b)(6)) from the original hospital was swapped, but the position waveform did not display on the receiving hospital¿s control device ((b)(6)). other waveforms were visible, and the pump operated without issues. the device was then swapped again to a different backup control unit at the receiving hospital, which resolved the problem, and operation continued with that unit. there were no health impacts resulting from the malfunction. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2324 | 1220648-2025-49561 | 23926870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23926870&pc=OZD | 12/07/2025 |  |  | 12/30/2025 |  |  | 12/08/2025 | 12/30/2025 |  |  |  |  | 12/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748539 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the incident was classified as high severity, confirmed as relevant with no fault present, and required additional action with ventricular extracorpeal membrane oxygenation (v-aecmo) insertion. additional information received indicated that the patient was originally supported by ecmo and intraaortic balloon pump (iabp); the ecmo was removed when the iabp was replaced with impella. the impella was subsequently removed as it was reported that the device disconnected from the aorta. after which the patient experienced cardiac arrest due to the disconnection, requiring reinsertion of ecmo. the indication for pre-op placement was for use before ventricular septal perforation (vsp) surgery. | Additional information received indicated that ventricular septal perforation (vsp) closure surgery was completed successfully, and the patient survived the procedure and was able to have the catheter removed. |  |  |  |  | Investigation summary: no product was returned. cardiac arrest: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position: the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history lot: device lot : 1977555. device history batch: subcomponent lot : n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E1. initial reporter: the reporter¿s first and last name were corrected from the initial report, which was in japanese; this report is in english.
2325 | 1220648-2025-49559 | 23927038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23927038&pc=OZD | 12/04/2025 | 12/30/2025 |  | 12/30/2025 |  |  | 12/04/2025 | 12/30/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026737677 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Low Blood Pressure/ Hypotension | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there had been a delay with impella connect, which rendered it inadequate for live support during the case. the representative was on a video call with the account to provide assistance, but impella connect did not reflect the real-time situation. the data was not current, and the team experienced impella placement signal low alarms during a hypotensive episode that were not visible on impella connect, despite the system displaying ¿live¿ and sweeping through the window. video evidence showed the patient in a crashed state with no output, while impella connect continued to display a blood pressure of 84/44, inaccurately indicating stability. additional information indicated that the patient expired. | [5 empty] | Product complaint # (b)(4). investigation summary- the cause of the optical sensor issue was determined to most likely be due to the patient's condition. the cause of the injury was not determined due to insufficient clinical details. device history lot- device lot : 1932120. device history batch- subcomponent lot : n/a. device history review- the pump sn (b)(6) passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2326 | 1220648-2025-49562 | 23930854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23930854&pc=OZD | 11/30/2025 |  |  | 12/30/2025 |  |  | 11/30/2025 | 12/30/2025 |  |  |  |  | 12/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026705593 |  | 0048-0014 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was implanted in a patient for mechanical circulatory support. while on support, the impella was in the aorta. the uci team attempted to reposition the pump. no further information was provided. there was no reported harm or injury to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2327 | 1220648-2025-49525 | 23905302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23905302&pc=OZD | 12/13/2025 | 01/05/2026 |  | 12/29/2025 |  |  | 12/13/2025 | 12/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - |  |  | US |  | 2026779271 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Pain | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 69 year old male on impella 5.5 for cardiomyopathy. it was reported that the patient was experiencing severe pain overnight at the impella axillary insertion site. the physician claimed that the sheath was compressing the brachial plexus nerve causing pain and a large hematoma wa developing at the access site. heparin was paused allow for bleeding to stop and has since been restarted. there was no repositioning or opening of surgical cutdown as patient is awaiting heart transplant. the patient remains on support. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product returned. pain: the cause of the injury was not determined due to insufficient clinical details. hematoma/major bleed: severe pain and a large hematoma developed at the impella axillary insertion site. heparin was stopped overnight to allow bleeding to stop and hemostasis was achieved. the cause of the access site adverse event was use related to anticoagulation management as the bleeding was resolved by stopping heparin. device history review: the pump passed all the post sterile inspection checks. | D6b added. the impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2328 | 1220648-2025-49526 | 23905335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23905335&pc=OZD | 12/12/2025 | 12/29/2025 |  | 12/29/2025 |  |  | 12/13/2025 | 12/29/2025 |  |  |  |  | 12/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733786 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 85 KG | Non Hispanic | Asian | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 73 year old male patient supported with an impella device experienced oozing from multiple sites with a reported decrease in hemoglobin and hematocrit. the patient received one unit of blood, and heparin was discontinued. low placement signal alarms persisted despite catheter repositioning, with imaging reportedly showing the device in appropriate position. no additional clinical details were provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1: added product problem. additional information: the bleeding was determined by the surgeon to be normal post-operative bleeding. the pump was not removed prematurely, and patient outcome was optimal.
2329 | 1220648-2025-49527 | 23905425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23905425&pc=OZD | 12/13/2025 | 12/29/2025 |  | 12/29/2025 |  |  | 12/13/2025 | 12/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026771326 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a axillary impella 5.5 attempted via right axillary graft of 10 mm vascutek gelweave. unable to advance 5.5 past sharp angle where subclavian artery turns inferiorly towards aortic arch. aborted further attempts to advance 5.5. an impella cp was opened and successfully placed via same access. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. analysis summary: difficult to advance: returned pump visually inspected. cannula kink near outlet observed. no other abnormality found. the root cause of unable to advance 5.5 pump most likely was patent condition related where subclavian artery turns inferiorly towards aortic arch and patient had an aortic valve and mitral valve replaced a week pre procedure. device history review: the complaint pump passed all post sterile inspection checks.
2330 | 1220648-2025-49528 | 23905481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23905481&pc=OZD | 12/12/2025 | 12/29/2025 |  | 12/29/2025 |  |  | 12/12/2025 | 12/29/2025 |  |  |  |  | 12/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749565 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male | 78 KG |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Decrease in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. on day 14 of impella support, it was reported that there was leaking purge solution out of the filter at the purge sidearm. patient support was not affected, and the motor current and purge flow remained within expectations. the medical team was encouraged to replace the pump due to the issues and the medical team was instructed to carefully watch the motor current, purge flow and purge pressures. subsequently, the medical team decided to wean the patient and explant the pump later the same day. the device was explanted successfully after weaning. the patient's outcome at the time of explant was survived. | [5 empty] | B3 revised date of event as it was entered incorrectly from the initial report that was submitted. | A5 is unknown. the device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed.
2331 | 1220648-2025-49529 | 23905576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23905576&pc=OZD | 12/13/2025 | 01/27/2026 |  | 12/29/2025 |  |  | 12/13/2025 | 12/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090623 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported during prep a controller was turned on for an incoming transfer. immediately received a controller error. a different controller was used and there was no patient harm. | [5 empty] | D9 device returned and date were omitted from the initial report that was submitted and were added. h6 codes b17 and d15 were reported incorrectly on the initial report that was submitted. codes were added that should of been reported on the initial report that was submitted. h11 updated additional manufacturer narrative to reflect that the device was returned. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Data review: the device logs confirmed a controller error (event set #1039) alarm during startup. this alarm is triggered when the light which reaches the ccd sensor array is less than 0.75v, which indicate a malfunction of the optical lamp. the alarm was intermittently triggering. device analysis: the reported controller error was not able to be reproduced upon console start up and testing. perform 5s parameter check, ccd distortion check, and confirmed lamp power output were all within range. conclusion: the root cause of the controller error is likely due to a defective optical bench lamp assembly. device history lot n/a device history batch subcomponent name: optical bench lamp assembly material number: 2450-0038 lot number: n/a serial number: n/a device history review q1) service history review - are there any qns associated with the complaint device¿s most recent service report? there were no nonconformances in the most recent fsr before the complaint occurrence date which were related to the failure mode. q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode? n/a q3) complaint history review - have there been any other complaints against this console? yes phr summary: there were no issues related to the complaint discovered when reviewing the product history of console ic1329. ci-29477: fm-30094 purge issue.
2332 | 1220648-2025-49530 | 23905954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23905954&pc=OZD | 12/13/2025 | 12/29/2025 |  | 12/29/2025 |  |  | 12/13/2025 | 12/29/2025 |  |  |  |  | 12/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026737753 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 95 KG |  |  | Death | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63 year old male patient on an impella cp due to cardiogenic shock. it was reported that during insertion, the red channel lumen was pulled and staff struggled to get wire through the cannula. at some point, the patient received a second cp and eventually expired. | [5 empty] | Correction: e4 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. difficult to deliver or advance: the cause of the deliver issue was determined to be use issue related since the guy pulled red cheater lumen and staff struggled to get wire through the cannula. device history review: the device passed all post sterile inspection checks. | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2333 | 1220648-2025-49531 | 23906076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23906076&pc=OZD | 12/12/2025 | 02/12/2026 |  | 12/29/2025 |  |  | 12/12/2025 | 12/29/2025 |  |  |  |  | 12/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725395 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. on the first day of support, the nurse reported that there was an issue with the pump during start up. she reported that the last start up screen never populated on the automated impella controller. the issue was resolved, but the nurse was unsure how the issue was resolved. the event logs were reviewed with the nurse, and a pump alarm ¿impella stopped. motor current high¿ alarm was found. the alarm was resolved, and the pump was successfully started. the patient was supported for a total of three days before the pump was weaned and removed after the patient¿s heart recovered. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was discarded by the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the temp pump stop has been completed. data logs were returned. the logs do not show any position related alarms. the cause of the issue was not established as limited clinical information was provided and no product was returned.
2334 | 1220648-2025-49532 | 23906343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23906343&pc=PYX | 12/12/2025 |  |  | 12/29/2025 |  |  | 12/12/2025 | 12/29/2025 |  |  | 02/28/2026 |  |  |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478743 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Difficult to Open or Close | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella rp flex via the right femoral vein for mechanical circulatory support. on support day 2, it was reported that the tuohy-borst valve would not tighten. many efforts were made to tighten the valve; however it was reported that it would only spin and would not tighten. it was noted that the impella rp was secured. on day 5 of support, the patient experienced hemolysis, and the pump was subsequently explanted. the impella rp was replaced with a protek duo. it was noted that the patient continued to experience hemolysis post impella rp explant. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. clinical review: an impella rp flex was placed via the femoral vein of the 62-year-old female post surgical patient who had been admitted in with right heart failure and in scai stage e shock. the patient has known history of prior bypass/coronary artery bypass graft (cabg) surgery for known coronary artery disease, but any other underlying comorbidities are unknown. the pump was placed, but after a day of support the team observed the tuohy borst to be malfunctioning, as it would not tighten down, however the pump was able to be secured. on day 4 of the support there was an observation made of hemolysis. the labs correlate to this as the plasma free hemoglobin was noted to be 299mg/dl and the creatinine was elevated. the team noted there was poor anticoagulation which may have contributed to the hemolysis. the team made decision to explant the rp flex and place instead a protek duo for continued support, and the hemolysis continued post rp flex explant. hemolysis is a known risk associated with impella support as described in the instructions for use and may be influenced by patient-specific factors. investigation summary the pump was not returned for investigation. usbs with data logs were returned for investigation. data log shows the pump was utilized until 16-dec-2025, 1 am. data log shows a motor current spike at around 15-dec-2025, 10 am, without impacting any other pump parameters, and later on 16-dec-2025, 00:40 am, with a noted drop in pump flow while running at p8. there were no other signs related to positioning issues or major suction during the case run. difficult to open/close catheter anchor or touhy: clinical details indicate that the tb valve was not locking down tightly and was just spinning. despite all efforts, it was unable to be tightened, although the device was secured. the cause of the difficulty in locking the tb valve was not determined without returned product investigation and sufficient clinical details. mechanical interaction with blood (hemolysis): the cause of the hemolysis issue was most likely related to biomaterial ingestion, based on data log review. device history review: the pump passed all post sterile inspection checks. | H6 was inadvertently omitted on the initial manufacturer device report number therefore it was added on this report.
2335 | 1220648-2025-49533 | 23906545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23906545&pc=OZD | 12/11/2025 | 12/29/2025 |  | 12/29/2025 |  |  | 12/11/2025 | 12/29/2025 |  |  |  |  | 12/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026621555 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 84 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Pericardial Effusion | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced low pump flows and an acute effusion. a clot was evacuated via pericardial window and the patient stabilized. the chest was opened to check for bleeding and a small puncture was found on the posterior left ventricle. a wire had perforated the ventricle in the area of the infarction. pedgetted sutures were placed to control the bleeding. the chest tube was placed and the chest was closed. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. low or blocked pump flow: the cause of the low flow issue was not determined due to no product returned and lack of clinical information. thrombosis: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. major bleeding: small puncture was found on the posterior left ventricle (lv), cannot say for sure it appeared that a wire went through the ventricle in the area of patient's infarction. pledgetted sutures were placed and the bleeding was controlled. the cause of bleeding issue was not determined due to no product returned and lack of clinical information. pericardial effusion: the cause of the injury was not determined due to insufficient clinical details. device history review the complaint pump passed all post sterile inspection checks.
2336 | 1220648-2025-49534 | 23906780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23906780&pc=OZD | 12/09/2025 |  |  | 12/29/2025 |  |  | 12/09/2025 | 12/29/2025 |  |  |  |  | 12/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026761296 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp a cannula kink. the pump was explanted and replaced with another pump. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2337 | 1220648-2025-49535 | 23907080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23907080&pc=OZD | 12/09/2025 |  |  | 12/29/2025 |  |  | 12/10/2025 | 12/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026726185 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced limb ischemia resulting in compartment syndrome. a fasciotomy was performed to resolve the ischemia. the pump was explanted and the compartment syndrome resolved. the patient is in stable condition. | [5 empty] | A4 is unknown. the impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E1 first and last name was inadvertently omitted on the initial manufacturer device report number. the investigation was completed and no product return. ppae/ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks.
2338 | 1220648-2025-49536 | 23907350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23907350&pc=OZD | 12/09/2025 | 12/29/2025 |  | 12/29/2025 |  |  | 12/09/2025 | 12/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  |  |  | 1409579 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63 year old male on impella cp for acute myocardial infarction. it was reported that 6 days after the procedure, a control unit malfunction alarm occurred. there was an automatic alarm stop after 1 minute of occurrence. there was no patient harm reported. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the return of the device. device status: the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed.
2339 | 1220648-2025-49537 | 23907947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23907947&pc=OZD | 12/08/2025 |  |  | 12/29/2025 |  |  | 12/08/2025 | 12/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1815958 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that the automated impella controller power cable separated from the plug exposing bare wire. a replacement controller was used for patient support. | [5 empty] | The console logs were not relevant and showed no errors. the console was returned to fs for further investigation. during the inspection, it was noted that all three wires were disconnected from their terminals, and the ac cord plug had detached from the cord. additionally, the outer sheath was improperly cut, in accordance with the vendor wiring instructions. also, the light clamping marks were visible on the outer sheath, which likely contributed to the wire becoming loose. the reported issue of the power cable separated from plug, bare wire showing and console inoperable was determined to be caused by defective ac power cord set. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D2a. common device name was revised as it was entered incorrectly on the initial report that was submitted. h11 added product history review as it was omitted from the previously submitted report. product history review summary: there were ten involving the same subcomponent lot as the reported issue was discovered during the review the product history of the console. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated.
2340 | 1220648-2025-49540 | 23908996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23908996&pc=OZD | 12/01/2025 |  |  | 12/29/2025 |  |  | 12/02/2025 | 12/29/2025 |  |  |  |  | 12/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766593 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male | 67 KG |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | It was reported that the "impella ¿ in aorta" alarm occurred continuously, although echocardiography confirmed the pump to be located within the ventricle. the device was operating at p2; attempts to increase the p-level immediately triggered a suction alarm and the flow rate could not be increased. motor current demonstrated mild pulsatility, and the aortic-placement signal also showed pulsatility, consistent with the patient monitor. in the parasternal long-axis echo view the pump position could be clearly assessed: the impella tip was positioned approximately 3.5 cm deep within the left ventricle and was free within the lv cavity. the pump had been in situ for 15 days. flow could not be increased and alarms persisted despite minor repositioning maneuvers. at p7 the device continued to display flows of 1.8 l/min, no lv signal could be derived, and a suction alarm occurred. purge flow and purge pressure were within normal limits. explantation of the pump was recommended because adequate pump flow could not be achieved. the patient was reported to have survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2341 | 1220648-2025-49539 | 23908999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23908999&pc=OZD | 12/04/2025 | 02/12/2026 |  | 12/29/2025 |  |  | 12/05/2025 | 12/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766621 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | Yes |  |  |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | [10 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the customer implanted the pump during a routine procedure for a cardiogenic shock patient. after activation, the performance signal (ps) disappeared. it is unclear how long the pump operated with a normal ps. the customer contacted our csc, who advised replacing the aic. they switched to an aic 3068 while keeping the same pump, but the ps issue persisted. the customer then decided to replace the pump (sn (b)(6)), and the new unit functioned without any problems. | [5 empty] | The investigation for the optical signal issue has been completed. the cause of the optical sensor issue cannot be determined since field data logs were not returned. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2342 | 1220648-2025-49397 | 23895465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895465&pc=OZD | 11/02/2025 |  |  | 12/26/2025 |  |  | 11/02/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026723067 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient presented to the hospital on (b)(6) 2025 with acute decompensated heart failure with complaints of worsening shortness of breath and had a severely reduced left ventricular ejection fraction of 25% due to ischemic cardiomyopathy, and no further revascularization options. the medical team decided to work the patient up for a left ventricular assist device (lvad) implantation or heart transplantation. the patient's condition deteriorated, requiring escalation of medical therapy to include continuous inotropic support and increased diuretic infusions. subsequent right heart catheterization (rhc) confirmed severe hemodynamic compromise, revealing elevated cardiac filling pressures and a mixed venous oxygen saturation of less than 55%. the medical team decided to place an impella 5.5 for mechanical circulatory support. the patient was brought to the operating room, was intubated, and the impella 5.5 was implanted via the right axillary artery without complications. on support day 3, the patient experienced episodes of ventricular tachycardia. it was noted that the impella flows remained stable throughout the episodes, and an echocardiogram was ordered to review impella placement. the patient remained hemodynamically stable. on support day 14, the patient experienced worsening blurry vision. an emergent computed tomography (ct) scan was performed and revealed an inclusion in the internal carotid artery. neurology was consulted to determine the etiology of the stroke with high suspicion for a cardioembolic cause. a repeat ct did not reveal a thrombus. as a result of the patient¿s complication, the impella was weaned and explanted. the patient's outcome at the time of explant is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2343 | 1220648-2025-49398 | 23895470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895470&pc=OZD | 11/03/2025 | 12/04/2025 |  | 12/26/2025 |  |  | 11/05/2025 | 12/26/2025 |  |  | 05/31/2025 |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366643 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Use of Device Problem | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Arrhythmia | Hemolysis | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 52 year female was implanted with impella cp for support during a high-risk cardiac procedure. it was reported that during support, the patient developed hemolysis with low output red concentrated urine. additionally, ectopy occured overnight, and the impella was repositioned and pulled back slightly by the doctor. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2344 | 1220648-2025-49399 | 23895473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895473&pc=OZD | 11/04/2025 |  |  | 12/26/2025 |  |  | 11/04/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026737538 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced renal failure and was placed on hemodialysis. later, the physician pulled the pump across the valve and could not recross. the pump was explanted and the patient was stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Investigation summary no product was returned. renal failure: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was being repositioned, doc pulled across valve and could not recross. device history lot: device lot: 1969394. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
2345 | 1220648-2025-49400 | 23895477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895477&pc=OZD | 11/06/2025 |  |  | 12/26/2025 |  |  | 11/06/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701729 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 71 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support lido started for arrhythmias. hemoglobin/hematocrit and platelets have gradually decreased. later it was reported new onset atrial fibrillation. additionally, heparin-induced thrombocytopenia panel sent and 1 unit overnight was given to the patient. it was also reported that the patient had an episode of orthostatic hypotension. implantable cardioverter defibrillator fired once for ventricular tachycardia. the patient remains on support. | The patient was successfully weaned and the device was explanted. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information: b5 device was explanted. d6b. explantation day, month and year.
2346 | 1220648-2025-49404 | 23895607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895607&pc=OZD | 11/06/2025 | 12/30/2025 |  | 12/26/2025 |  |  | 11/06/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026735733 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Backflow | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 80-year-old male was implanted with impella 5.5 during support for primary percutaneous coronary intervention. shortly after landing in cardiovascular intensive care unit, hematoma was noted at incision site. pressure dressing applied to site to troubleshoot. 1 unit of packed red blood cells were given. additionally, active preventing retrograde alarm and triggers discussed. will leave as is until explant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported hematoma/bleeding and retrograde pump flow has been completed. the root cause of thehematoma/bleeding and retrograde pump flow was not determined without returned product investigation and sufficient clinical details
2347 | 1220648-2025-49347 | 23895652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895652&pc=OZD | 12/04/2025 | 12/24/2025 |  | 12/26/2025 |  |  | 12/05/2025 | 12/26/2025 |  |  | 12/31/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025511993 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. left groin vascular access was attempted by the interventional cardiologists using a 6-french introducer sheath; however, the sheath could not be advanced. the cardiologist removed the sheath from the left side and achieved hemostasis using manual compression. the right groin access was used for single-access percutaneous coronary intervention (pci). at the completion of the procedure, the patient developed tachycardia and hypotension. clinical assessment identified a left-groin hematoma with active bleeding. the patient¿s hemodynamic status subsequently stabilized. a vascular surgeon was called to assist with management of the left-groin bleeding, and hemostasis was successfully achieved. the impella device was explanted once the patient was stabilized. the patient was taken to the operating room for exploration of the left groin and surgical repair. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. there are two associated devices for the reported event. the first device (impella cp) is addressed under manufacturer report number xxxxxxx-2025-xxxxx. the second device (oscor intro kit) is addressed under manufacturer report number xxxxxxx-2025-xxxxx.
2348 | 1220648-2025-49407 | 23895664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895664&pc=OZD | 05/09/2025 |  |  | 12/26/2025 |  |  | 11/07/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025593596 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Loqi database reported upon dd for the subject (b)(6). the device deficiency report in loqi states: impella cp stopped flowing. unable to reestablish flow. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. low or blocked pump flow: the data logs show pump flow was lower than expected for set p-level with pump flowing below the 2.9-3.3l/min range for p-7. there were no motor current spikes and no position related alarms. the cause of the low pump flow could not be determined as no product was returned and limited clinical information was provided. pump stop: the data logs confirm pump stop. the cause of the pump stop was determined to be biomaterial ingestion as evidenced by log pattern showing ingestion signature and saturated pump flows device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2349 | 1220648-2025-49408 | 23895710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895710&pc=OZD | 12/11/2025 | 01/02/2026 |  | 12/26/2025 |  |  | 12/11/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026738794 |  | 0550-0004 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. on day 2 of impella support, it was reported there were purge pressure high and purge system blocked alarms. the medical team added tissue plasminogen activator to the purge solution per abiomed medical office recommendation, which did not resolve the issue. subsequently, the device was removed due to the persistently high purge pressure alarms. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. patient outcome is unknown as outcome information was not available in the complaint record accessible for mdr assessment at time of reporting. the device is reportedly available for return; however, as of the time of this report the device has not yet been returned for evaluation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The pump was not returned for investigation. data log shows gradual fluctuation in purge pressure throughout the case run, with a purge pressure high and purge system block alarm noted mid-run. no motor current spike was reported throughout the case run. this resolved gradually over time. according to the clinical details, administration of tpa did not resolve the purge issue. based on the data log trend, the cause of the purge issue was most likely related to biological material buildup. | H6 has been updated and impact codes f2303 and f1903 has been added.
2350 | 1220648-2025-49409 | 23895719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895719&pc=OZD | 11/08/2025 |  |  | 12/26/2025 |  |  | 11/08/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732864 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 143 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced gastrointestinal bleeding and a hematoma at the access site. the patient is in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the hematoma at impella site after transfer to icu and was managed with manual pressure. the cause of hematoma is determined to be a use issue due to patient movement because the hematoma was noted after transfer to icu. the root cause of the gi bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
2351 | 1220648-2025-49412 | 23895738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895738&pc=OZD | 07/10/2025 |  |  | 12/26/2025 |  |  | 07/11/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691753 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing at the groin site. a suture was added to attain hemostasis. ten days after the pump was explanted the patient experienced atrial fibrillation. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: no product was returned. access site adverse event: the cause of oozing was most likely use issue related since the original suture torn and failed was corrected with another suture. paroxysmal atrial fibrillation: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history review: device passed all post sterile inspection checks.
2352 | 1220648-2025-49414 | 23895739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895739&pc=OZD | 08/12/2025 |  |  | 12/26/2025 |  |  | 08/12/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710706 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male | 117 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Fever | Unspecified Infection | Ischemia | Sepsis | Tachycardia | Cardiogenic Shock | Hematuria | Thrombosis/Thrombus | Valvular Insufficiency/ Regurgitation |  |  |  | Death | [28 empty] | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia requiring defibrillation and repositioning of the pump. the patient also exhibited vegetation, endocarditis, sepsis, and cardiogenic shock. there was concern for ischemic colitis. the patient was treated with antibiotics. clots were seen in the foley catheter. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the tachycardia, stroke, sepsis, thrombosis, infection, and ischemia was unable to be determined due to insufficient clinical details. the cause of the fever, hematuria, heart valve incompetence, and cardiogenic shock was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
2353 | 1220648-2025-49413 | 23895741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895741&pc=OZD | 07/17/2025 |  |  | 12/26/2025 |  |  | 07/17/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025536028 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 81 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient in post-cardiotomy cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support. device sensing issue was encountered, and the patient had developed hemolysis. the patient eventually escalated to an impella 5.5 and then subsequently expired. the following day after start of the impella cp device the patient was noted to have hemolysis. the patient was unable to tolerate increased chemical support. same day, low placement signal alarm triggered. troubleshooting steps were taken. performance level was dropped to p-2 and assessment of impella pump placement under echo was completed. it was noted the inlet measured 3.2 cm in free ventricular space. left ventricular cavity was rather small; however, the outlet visualized well above aortic valve. no other alarms were present after readvancing and pulling back the impella device catheter. optical sensor diastolic still read below 30; staff was encouraged to trust monitor for patient¿s blood pressures while leaving the option to disable placement monitoring. performance level was ramped back up to p-9 with appropriate flows and support continued. the following day the impella cp device was explanted and following coronary artery bypass graft surgery, mitral valve replacement (mvr), and tricuspid valve replacement (tvr), the decision was made to place an impella 5.5 directly. after 24 hours of support, the impella 5.5 generated a placement signal not reliable alarm. placement signal values and waveforms were present, and the alarm did not impact hemodynamic support. the alarm was disabled. during support, the patient developed volume overload and acute kidney injury (aki), requiring initiation of continuous renal replacement therapy (crrt). on day 9 of impella 5.5, support, the patient began to decompensate and experienced a gastrointestinal (gi) bleed. the patient was suspected to be in septic shock. after continued support, it was determined that the patient was no longer a candidate for permanent lvad therapy. the care plan shifted to continued impella 5.5 support with the goal of native heart recovery. subsequently, the impella 5.5 was repositioned, resulting in improved flows. despite ongoing support, the patient expired after 41 days of impella support. | [5 empty] | Medical safety reviewed the event and noted the following: the patient was brought to the operating room for placement of an impella 5.5 device as a bridge to coronary bypass grafting (cbg). upon evaluation, the axillary artery was determined to be too small to accommodate the impella 5.5. an impella cp was implanted without issue. post-implant, hemolysis was noted, which is considered a reportable adverse event. following cabg, mitral valve replacement (mvr), and tricuspid valve replacement (tvr), the decision was made to place an impella 5.5 directly. after 24 hours of support, the impella 5.5 generated a placement signal not reliable alarm. placement signal values and waveforms were present, and the alarm did not impact hemodynamic support. the alarm was disabled. during support, the patient developed volume overload and acute kidney injury (aki), requiring initiation of continuous renal replacement therapy (crrt). on day 9 of impella 5.5 support, the patient began to decompensate and experienced a gastrointestinal (gi) bleed. the patient was suspected to be in septic shock. after continued support, it was determined that the patient was no longer a candidate for permanent lvad therapy. the care plan shifted to continued impella 5.5 support with the goal of native heart recovery. on (b)(6) 2025, the impella 5.5 was repositioned, resulting in improved flows. despite ongoing support, the patient expired after 41 days of impella support. the impella cp hemolysis is a reportable adverse event. impella 5.5 placement signal not reliable alarm are considered device malfunction reportable events. the renal insufficiency, infection and gi bleed are reportable patient adverse events for the impella 5.5. the patient¿s death will be conservatively reported on the impella 5.5, although it is unlikely that the device contributed to the cause of death. the patient¿s deteriorating condition and non-candidacy for permanent lvad were likely contributing factors. note: the automated impella controller will remain a reportable product malfunction accordingly. device status. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella cp device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller and the impella 5.5, which is associated with the demise outcome. | Complaint coding was updated to better describe the reported event. consequently, h6 health effect: clinical coding was repopulated/updated, corrected accordingly. note: h6 health effect impact, component and investigation type/finding/conclusion coding remain unchanged as previously reported.
2354 | 1220648-2025-49416 | 23895757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895757&pc=OZD | 09/02/2025 |  |  | 12/26/2025 |  |  | 09/02/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701037 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 87 KG | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. there was a purge leak from the unit after almost 20 plus days of support, it is unknown what was dont to address the issue but the pump continued to support the patient for another 2 days before it was explanted. no patient harm was reported due to the issue. | [5 empty] | The pump was not returned for investigation. the data log was not required for the investigation. the cause of the purge leak issue was not determined, as the product was not returned and sufficient clinical information was not provided. the device passed all post sterile inspection checks.
2355 | 1220648-2025-49411 | 23895769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895769&pc=OZD | 12/12/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 12/12/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026722343 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 83 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65 year old male was implanted with a impella 5.5 for support during a high risk cardiac procedure. it was reported that there was some hemolysis noted. urine was pink tinged, and plasma free hemoglobin elevated 70. impella was noted to be shallow around 3 cm for the 5.5. it was recommended to discuss with the doctor advancing impella. patient was also started on extracorporeal membrane oxygenation (protekduo cannula) for right ventricle support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary no product returned. data logs on constellation. mechanical interaction with blood/hemolysis: clinical reported hemolysis with device shallow around 3 cm for the 5.5. the data logs do not show any motor current spikes outside p-level changes. there were a few suction alarms at case start. no information was provided on if repositioning occurred and if it resolved hemolysis. the cause of the issue was not established as no product was returned and limited clinical information was provided device in wrong position: clinical reported device shallow around 3 cm for the 5.5. no positioning alarms on the logs were noted on day of issue occurrence. no information was provided on how device became shallow. the cause of the issue was not established as no product was returned and limited clinical information was provided device history review the complaint device passed all post sterile inspection checks.
2356 | 1220648-2025-49422 | 23895832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895832&pc=OZD | 09/07/2025 |  |  | 12/26/2025 |  |  | 09/07/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716185 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 80 KG | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Malposition of Device | Difficult to Advance |  |  |  |  | Atrial Fibrillation | [12 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | It was reported that implantation of an impella 5.5 was difficult due to the patient's anatomy. the right axillary site was abandoned and the left axillary was used instead. however, the motor housing could not pass the anastomosis. implantation via the left axillary was also abandoned. the patient developed hematuria. the patient was placed on continuous renal replacement therapy. a placement signal low alarm occurred and self-resolved. the pump was positioned deep so it was repositioned. the patient experienced atrial fibrillation. amiodarone was given to resolve the fibrillation. purge pressure was unstable so the purge cassette was exchanged for a new cassette. the patient continued to have episodes of atrial fibrillation. suction events occurred so the pump was repositioned again. later, the pump was explanted. | [5 empty] | No product was returned for investigation. the cause of the difficulty in delivering the pump was patient condition related due to the noted trouble with the patients anatomy. the cause of the device in the wrong position was not determined as not enough clinical details were provided to determine how the pump became too deep. the cause of the pump suction was due to positioning issues; however, the cause of the positioning issues was not determined. the root cause of the paroxysmal atrial fibrillation was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2357 | 1220648-2025-49419 | 23895837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895837&pc=OZD | 12/12/2025 | 01/28/2026 |  | 12/26/2025 |  |  | 12/12/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554529 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Pumping Stopped | Use of Device Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a 74 year old female was implanted with a impella cp for support during a high risk cardiac procedure. as impella cross av otw the physician asked can we start, he was asked if wire had been removed to which he said yes. impella was started however distal tip of wire became wrapped around motor. and impella stopped error displayed. impella and wire removed through cp sheath and another impella device prepped and inserted w/o issue. | [5 empty] | Investigation summary: impella cp #(b)(6) and guidewire returned. no data logs were returned. pump did not start/use against labeling: as per clinical information reported pump was started before removing guidewire and the distal tip of the wire inside pump got wrapped around motor. later device was replaced after pump stop alarm. the product was returned for investigation. visual inspection of guidewire revealed deformation or kinking at distal tip of the wire. also, visual inspection of pump showed scratch or some minor dent marks at the impeller likely due to guidewire interaction when pump was accidently started with wire in place. but the pump was tested in bucket of water and ran without any issues. no other abnormalities found. as per ifu 0048-9007, the guidewire should be completely removed before pump start. the cause of the pump did not start/use against labelling issue was use related since ifu instructions were not followed during procedure. hemodynamic instability: the cause of injury is not determined due to insufficient clinical information. device history lot: device lot: 1833157. device history batch: subcomponent lot: n/a. device history review: device sn (b)(6) passed all post sterile inspection checks. | Medical safety reviewed the event. the event describes an unintentional device misuse resulting in removal and replacement of the device. the patient presented in cardiogenic shock secondary to acute myocardial infarction and underwent impella support placement. an impella cp device was selected. during device placement, a complication occurred. the physician verified that the guidewire had been removed and the impella cp device was started. subsequently, it was identified that the distal portion of the guidewire became entangled with the motor, resulting an impella stop error. the impella cp device and the guidewire were removed together through the introducer sheath. a second impella cp device was prepared and successfully implanted without complication. the replacement device provided circulatory support for approximately three days. thereafter, life-sustaining therapy was withdrawn, and the patient expired. impella cp pump 1 is a death type of reportable event. change in reportability: after review of the case by medical safety, it was determined the impella cp is a death type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. date of death: the exact date of death is unknown but has been captured as the date of explant. if actual information becomes known, a supplemental report will be submitted. note: h6: device problem/component/investigation coding remains unchanged. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2358 | 1220648-2025-49424 | 23895839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895839&pc=OZD | 09/19/2025 |  |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687534 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support the device exhibited placement signal issues, suction was above p-4 and the optical sensor was not working. additionally, an echocardiogram was performed and showed rate controlled atrial fibrillation. no further information was provided. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | The pump was not returned for investigation. according to the clinical details, pump flow could not go above p4 without triggering suction during the reported psnr event. optical sensor issue: data log shows a decreasing trend in placement signal for about 2 days. the pump flow had a completely opposite trend, followed by loss of placement signal with active psnr. flow returned to normal after the placement signal was completely lost on 20-sep, with active psnr for the remainder of the case run. all 5s optical signal parameters were within the specification range during the loss of placement signal. the cause of the optical sensor issue was determined to be sensor wear, based on data log review. pump suction: suction alarms were noted after the placement signal started to drift, and the pump was not able to achieve higher p-levels than p4 without triggering a suction alarm. suction alarm #14 characteristic curve alarms began to trigger as the aops trended down toward 0 mmhg. during this time, the motor current (mc) remained stable and within the expected range for the p-level. this occurred because the suction algorithm compares the mcmin values to the placement signal at the same time. since the placement signal was trending down due to sensor wear, without any change in the mc, the algorithm was tricked into thinking the pump was in diastolic suction. therefore, the suction alarms were determined to likely be false alarms. the cause of the suction issue was determined to be sensor wear, based on data log review. device in wrong position: the cause of the positioning issue was not determined without returned product investigation and sufficient clinical information.
2359 | 1220648-2025-49426 | 23895857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895857&pc=OZD | 09/19/2025 |  |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634406 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, bleeding was observed at the insertion site, hemostatic sutures were added, but the area continued to bleed. medication was adminstered to help control the bleeding. the customer stated that bleeding at the insertion site was controlled. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history lot device lot : 1900426 device history batch sub-component lot : n/a device history review the pump s/n: (b)(6) passed all the post sterile inspection checks.
2360 | 1220648-2025-49427 | 23895858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895858&pc=OZD | 09/11/2025 |  |  | 12/26/2025 |  |  | 09/11/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739077 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding. a forward tension suture was placed and the patient was transfused one unit of packed red blood cells. the patient is in stable condition. | [5 empty] | A4 is unknown. no product was returned for investigation. the cause of the bleeding is determined to be use issue because bleeding was resolved by one forward tension suture was placed to first eyelet and dressing angle matched. the device passed all post sterile inspection checks.
2361 | 1220648-2025-49429 | 23895861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895861&pc=OZD | 09/11/2025 |  |  | 12/26/2025 |  |  | 09/11/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712136 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant indicated that a patient diagnosed with cardiomyopathy received an impella 5.5 device for mechanical circulatory support. the patient was admitted while awaiting a heart and kidney transplant. the pump was operational when it alarmed for aortic issues due to optical signal loss. the pump was correctly positioned. additionally, the access site exhibited oozing, which necessitated the infusion of 2 units of red blood cells (rbcs). after 26 days, the patient passed away while on support. no transplant was performed before the patient's death. | [5 empty] | No product was returned. the cause of the optical issue was related to the automated impella controller. the cause of the access site bleeding was use related to failure to follow instructions for the patients act values above the recommended 180. the device passed all post sterile inspection checks.
2362 | 1220648-2025-49430 | 23895863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895863&pc=OZD | 09/19/2025 |  |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635360 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ischemia | [11 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, pedal pulses were not present in the impella leg. no further information was provided; however, it was later reported that the patient expired. | [5 empty] | Investigation summary: no product was returned for investigation. cardiac arrest/ ischemia: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1899315. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
2363 | 1220648-2025-49431 | 23895868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895868&pc=OZD | 09/11/2025 |  |  | 12/26/2025 |  |  | 09/11/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | 2000154 |  |  | ABIOMED, INC. |  |  | US |  | 2026616758 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant indicated that during the support of the impella cp, oozing at the impella site was noted. a dressing change was performed at the bedside with the registered nurse (rn), during which a small portion of the repo sheath was found to be exposed. the team was educated on the advancement of the repo sheath, and if feasible, angle matching was applied. monitoring was continued, and no hematoma was observed at this time. the plan of care (poc) was discussed at the bedside with the rn. the patient exhibited some oozing at the surgical sites, prompting a decision to postpone the restart of bivalirudin until reassessment. the cp pump was explanted, and a 5.5 pump was implanted. | [5 empty] | No product was returned for investigation. the cause of the major bleeding is determined to be use issue because the bleeding was resolved by advancing the repo sheath and matching the angle. the device passed all post sterile inspection checks.
2364 | 1220648-2025-49432 | 23895871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895871&pc=OZD | 09/19/2025 |  |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727184 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  |  | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the patient had multiple runs of ventricular tachycardia and was defibrillated 4 times. additionally, the patient experienced hemolysis, later it was noted that the patient was no longer showing signs of hemolysis. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary: data review: data logs showed pump ran without issue during support. there were suction alarms triggered in early of the case but resolved quickly. there were no mc spikes, abnormal ps or purge issues observed during support. product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. ventricular fibrillation: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot #: 1956269. device history batch: subcomponent lot#: n/a. device history review: pump sn: (b)(6) passed all post sterile inspection checks.
2365 | 1220648-2025-49433 | 23895872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895872&pc=OZD | 09/11/2025 |  |  | 12/26/2025 |  |  | 09/11/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553922 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient experienced an arterial dissection during implantation of an impella 5.5. the pump was explanted and reimplanted after the dissection was surgically repaired. the patient had ongoing suction alarms and impella in ventricle alarms. volume was given and the pump was repositioned. the patient was intubated to resolve the alarms. the patient was in critical condition. | [5 empty] | A4 is unknown. the pump was returned and no issues were found with the returned product. the cause of the vascular dissection was not determined due to insufficient clinical details. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy and resistance met axillary artery preventing successful delivery of impella. the device passed all post sterile inspection checks. | Medical safety/clinical reviewed the event. an 81-year-old patient with an acute myocardial infarction was supported with an impella cp for mechanical circulatory support. a decision was made to escalate support to an impella 5.5; however, advancement of the device was unsuccessful due to inability to pass the anastomosis, and the impella 5.5 implantation was aborted. an angiogram was initially reported as showing no dissection; however, subsequent documentation confirmed the presence of a vascular dissection, which was successfully repaired. this event is considered a serious injury attributed to the impella 5.5. the patient was returned to impella cp support. thereafter, the patient developed third-degree heart block requiring pacing. hemolysis was initially noted but later resolved, with urine reported as clear for several days. ongoing suction alarms and ¿impella in ventricle¿ alarms were reported despite volume resuscitation and repositioning attempts. echocardiography demonstrated the pump positioned across the aortic valve and free from cardiac structures. a loss-of-flow event occurred, resulting in hemodynamic decompensation and the need for intubation. following intubation, device alarms resolved and flows improved, allowing for transient stabilization. per the treating physician, the cause of the continuous suction, loss of flow, and erroneous placement alarms was unclear. later this same day, the family elected to withdraw life-sustaining care, and the patient expired. the patient¿s underlying condition and disease severity contributed most to the demise outcome. the event story involved two product complaints. impella 5.5 is a serious injury, and the impella cp is a death type of reportable event. h6: investigation type/finding/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
2366 | 1220648-2025-49434 | 23895874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895874&pc=OZD | 09/11/2025 |  |  | 12/26/2025 |  |  | 09/11/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700144 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Poor Visibility | [5 empty] | Atrial Fibrillation | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a placement signal low alarm. an ultrasound showed the pump was positioned deep. the pump was repositioned. the patient experienced atrial fibrillation and rapid ventricular response. the patient exhibited hematuria. the patient was intubated and cardioverted to normal sinus rhythm. the pump was exchanged for an impella 5.5. | [5 empty] | A5 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Complaint coding was updated to better reflect the reported event. consequently, section h6 medical device problem codes were updated/corrected and repopulated accordingly.
2367 | 1220648-2025-49435 | 23895882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895882&pc=OZD | 09/12/2025 | 11/19/2025 |  | 12/26/2025 |  |  | 09/12/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729413 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | False Alarm | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  | Hemolysis | Heart Block | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a third-degree heart block requiring pacing, and hemolysis. the pump generated suction alarms and impella in ventricle alarms. volume was given. the pump position was verified to be free from structures. the loss of flow caused the patient to decompensate and require intubation. once the patient was intubated the alarms resolved and flows improved. however, the patient expired on impella support. | [5 empty] | A4 is unknown. the cause of the hemolysis could not be determined due to insufficient clinical details and no product returned. the cause of the suction and low flow was most likely biomaterial ingestion, as the conditions didn't resolve with volume repletion or repositioning, and data logs show an mc spike and speed deviation. the cause of the positioning issues and alarms was most likely biomaterial ingestion, as the position was confirmed via echo, and data logs show an mc spike occurring just before the alarms started but an aortic pressure waveform. the pump passed all post sterile inspection criteria. | Medical safety/clinical reviewed the event. an 81-year-old patient with an acute myocardial infarction was supported with an impella cp for mechanical circulatory support. a decision was made to escalate support to an impella 5.5; however, advancement of the device was unsuccessful due to inability to pass the anastomosis, and the impella 5.5 implantation was aborted. an angiogram was initially reported as showing no dissection; however, subsequent documentation confirmed the presence of a vascular dissection, which was successfully repaired. this event is considered a serious injury attributed to the impella 5.5. the patient was returned to impella cp support. thereafter, the patient developed third-degree heart block requiring pacing. hemolysis was initially noted but later resolved, with urine reported as clear for several days. ongoing suction alarms and ¿impella in ventricle¿ alarms were reported despite volume resuscitation and repositioning attempts. echocardiography demonstrated the pump positioned across the aortic valve and free from cardiac structures. a loss-of-flow event occurred, resulting in hemodynamic decompensation and the need for intubation. following intubation, device alarms resolved and flows improved, allowing for transient stabilization. per the treating physician, the cause of the continuous suction, loss of flow, and erroneous placement alarms was unclear. later this same day, the family elected to withdraw life-sustaining care, and the patient expired. the patient¿s underlying condition and disease severity contributed most to the demise outcome. the event story involved two product complaints. impella 5.5 is a serious injury, and the impella cp is a death type of reportable event. h6: investigation type/finding/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
2368 | 1220648-2025-49436 | 23895889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895889&pc=OZD | 09/12/2025 |  |  | 12/26/2025 |  |  | 09/12/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683701 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Pump | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a pump stop. a kink in the cannula was observed near the motor housing. the pump was exchanged for a new one to continue patient support. no patient harm was reported. | [5 empty] | A4 is unknown. a review of data logs confirmed the complaint. a cannula kink was observed in the returned pump. testing of the returned pump did not reproduce the pump stop. the cause of the cannula kink was reasonably foreseeable misuse due to the kink found on the returned pump and the data logs showing the kink occurring during support. the device passed all post sterile inspection checks.
2369 | 1220648-2025-49437 | 23895905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895905&pc=OZD | 09/12/2025 |  |  | 12/26/2025 |  |  | 09/12/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026641855 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. it was reported that the patient woke up in the middle of the night with supraventricular tachycardia (svt) and moved right leg. the impella site revealed drainage and a hematoma from moving the leg aggressively. decision made to remove the pump. the patient received albumin, milrinone and epi along with. it was reported that the right ventricle began to work better; however, in the morning, there were suction alarms, so additional milrinone and epi were added. the patient was successfully weaned from the pump and the device explanted. | [5 empty] | No product was returned for investigation. the cause of hematoma was most likely use issue related since the patient moved the leg aggressively and hematoma was formed. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of positioning issue cannot be determined since insufficient clinical details were provided. the cause of pump suction was most likely patient condition related since giving additional fluid resolved the suction issue and flows were resumed. the device passed all post sterile inspection checks.
2370 | 1220648-2025-49447 | 23895997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23895997&pc=OZD | 09/19/2025 |  |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731760 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced intermittent runs of ventricular tachycardia. the device was repositioned under flouroscope, and it was reporting that the pigtail bent and readvanced after pulling back. the physician pulled back as much as possible and locked down the impella catheter. no further information was provided. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary tachycardia : in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot device lot : 1970107 device history batch sub-component lot : n/a device history review the pump s/n: (b)(6) passed all the post sterile inspection checks.
2371 | 1220648-2025-49448 | 23896000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896000&pc=OZD | 09/19/2025 |  |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724222 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the patient had several pulseless electrical activity (pea) resulting in suction alarms, the catheter was repositioned, and volume status was improved. additionally, it was noted that the patient's extremities are black, a vascular study was conducted and noted that there was no flow down either the leg or arms. the patient herniated and a brain death test was initiated; the patient was declared brain dead and died while on support. | [5 empty] | Investigation summary: no product was returned. ischemia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: no product was returned. pt had multiple peas, resulting in suction alarms. after reviewing logs, multiple suction alarms were observed along with impella in aorta alarms. the logs show suction alarms throughout the case and patient had poor prognosis and was on ecmo support as well. the cause of pump suction was most likely patient condition related since patient had poor prognosis and patient being in ecmo support. device history lot: device lot: 1974839. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
2372 | 1220648-2025-49449 | 23896009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896009&pc=OZD | 09/20/2025 |  |  | 12/26/2025 |  |  | 09/20/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026730381 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient experienced ectopy, positioning of the device was confirmed and deemed good. additionally, the patient had some thrombocytopenia and bival was withheld. also, blood remained to ooze from the patient and anticoagulation was withheld. no further information was provided. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary no product was returned for investigation. major bleed: patient remains oozy. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. arrhythmia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia: the cause of the clinical symptom was not determined due to insufficient clinical details. device history lot device lot: 1962575 device history batch subcomponent lot: n/a device history review device with sn: (b)(6) passed all post sterile inspection checks.
2373 | 1220648-2025-49428 | 23896033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896033&pc=OZD | 12/12/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 12/12/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026780982 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 28 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76 year old female was implanted with a impella cp for support during a high risk cardiac procedure. it was reported that the patient was having hemolysis due to cp position. impella was repositioned under echocardiogram. patient did require 4 units of blood to increase hgb, platelets, and help with volume status. impella sites and all access sites show no evidence of bleeding or hematoma. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product or data logs returned. hemolysis/mechanical interaction with blood: clinical details mention that the patient was having hemolysis due to cp position. the pump was repositioned under echocardiogram guidance and hemolysis resolved. the root cause of the hemolysis was use related to improper positioning as the hemolysis cleared up after repositioning the pump from the provided clinical details. device in wrong position: clinical details mention that the patient suffered from hemolysis due to the pump's position, which resolved after the pump was repositioned. the root cause of the device being in the wrong position was not determined due to insufficient clinical details and unreturned product. device history review: the complaint pump passed all post sterile inspection checks.
2374 | 1220648-2025-49425 | 23896059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896059&pc=OZD | 09/19/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669922 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during preparation for the procedure, the device exhibited a purge alarm followed by a placement signal unreliable alarm. the impella 5.5 pump was never implanted into the patient. no reported harm or injury to the patient. | [5 empty] | Mpella 5.5 sn (b)(6) + pc returned for evaluation. impella 5.5 prep while still in the tray on the or table immediate purge alarm (noted to be air in purge) followed by placement signal unreliable alarm. pump never entered patient. lt logs were the only data logs recovered but showed no recorded data. air in purge - the returned pump did not recreate air in purge alarms. the cause of the air in purge alarms was not determined due to the issue being unable to be recreated on the returned pump, no data logs with data were recovered, and insufficient clinical details. optical sensor issue - the returned pump showed no 5s abnormalities and passed the laser light continuity test. uson testing completed on returned pump had externally applied pressures of 0, 50, 100, 150, and 200 mmhg show as pump readings of 4, 55, 107, 158, 209. psnr was unable to be recreated. the differences between reading values and actual values were 4, 5, 7, 8, and 9 mmhg respectively. the cause of the optical signal issue could not be determined due to the returned pump unable to recreate psnr alarms with no optical abnormalities noted, no data logs with data were recovered, and insufficient clinical details.
2375 | 1220648-2025-49452 | 23896084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896084&pc=OZD | 09/22/2025 |  |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703644 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing high-risk percutaneous coronary intervention. the impella cp device was successfully placed. following the procedure, groin angiography demonstrated distal perfusion; however, distal pulses were neither palpable nor detectable by doppler. given minimal pulsatility and the ongoing need for hemodynamic support, the impella cp was escalated to an impella 5.5. | [5 empty] | Investigation summary: no product was returned for investigation. hematoma: heparin held overnight d/t hematoma at impella site. it was unclear whether the hemostasis was achieved as per clinical details. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. ischemia: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1951058 device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
2376 | 1220648-2025-49455 | 23896093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896093&pc=OZD | 09/22/2025 |  |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026741756 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Cardiac Arrest | Hemolysis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced hemolysis. the resolution for this issue is unknown. additionally, the patient experienced respitory arrest due to over sedation resulting in intubation. the patient was shocked twice due to ventricular tachycardia during intubation. the pump was successfully weaned and explanted. | [5 empty] | Investigation summary: data review: data logs showed pump ran without issue during support. there were suction alarms triggered in early of the case but resolved quickly. there were no mc spikes, abnormal ps or purge issues observed during support. product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. ventricular tachycardia/cardiac arrest: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot #: 1977172. device history batch: subcomponent lot#: n/a. device history review: pump sn: (b)(6) passed all post sterile inspection checks.
2377 | 1220648-2025-49451 | 23896119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896119&pc=OZD | 11/26/2025 |  |  | 12/26/2025 |  |  | 11/27/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760022 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while the patient was being moved to the operating room for ecmo removal, the impella cp was pulled into the aorta. the doctor checked and found that the catheter had been pulled all the way into the aorta, all the way to the pigtail, making reinsertion impossible. this was caused by the catheter not being properly secured because the fixation ring had not been securely fastened after position adjustment. as the ecmo was still in place, there were no problems with the patient's condition. after confirming that the patient's condition was stable, the ecmo and impella cp were removed simultaneously. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned. device in wrong position: the cause of the positioning issue was determined to be a use error as the pump was pulled out of position into aorta while the patient was being moved to the operating room and the catheter had not been properly secured. device history lot: device lot: 1988035. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6) passed all post sterile inspection checks. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2378 | 1220648-2025-49457 | 23896123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896123&pc=OZD | 09/22/2025 |  |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512574 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, a ¿placement signal not reliable¿ alarm was noted on the aic, and the medical care team was informed. no additional information is available regarding actions taken in response to the alarm. the patient remained on impella cp support via the axillary approach as a bridge to left ventricular assist device (lvad) therapy or transplant. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2379 | 1220648-2025-49459 | 23896160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896160&pc=OZD | 09/21/2025 |  |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717547 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient developed torsades de pointes an ultrasound was performed and the device position was confirmed as appropriate. the patient also went into ventricular fibrillation (vfib) and was shocked by automatic implantable cardioverter defibrillator (aicd) 4 times. it was reported that the procedure was successful and the patient was stable. | [5 empty] | Investigation summary: ventricular fibrillation/ arrhythmia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1955589. device history batch: sub-component lot: n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
2380 | 1220648-2025-49462 | 23896170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896170&pc=OZD | 09/09/2025 |  |  | 12/26/2025 |  |  | 09/09/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724564 |  | 0048-0045 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced leaking at the peel away sheath. the sheath was replaced to resolve the issue. additionally, the patient displayed hematuria. the patient was in stable condition. | [5 empty] | A4 is unknown. the pump was not returned for investigation. the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical details. the device passed all post sterile inspection checks.
2381 | 1220648-2025-49463 | 23896172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896172&pc=OZD | 09/12/2025 |  |  | 12/26/2025 |  |  | 09/12/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - |  |  | US |  | 2026727013 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the bipella patient had pulmonary hemorrhage during support, which was managed by holding off heparin. patient was successfully weaned off bipella. devices removed in cath lab. this complaint is for the impella cp. patient had impella rp (sn # 526992) which is reported under separate report. | [5 empty] | No product was returned for investigation. the root cause of the minor bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the root cause of the major bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
2382 | 1220648-2025-49466 | 23896174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896174&pc=OZD | 09/13/2025 |  |  | 12/26/2025 |  |  | 09/13/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738724 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Hispanic | White | Death | Yes | No | Device Sensing Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Hemodynamic instability | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced generalized bleeding and was transfused six units of packed red blood cells, two units of platelets, one unit of fresh frozen plasma, two units of cryoprecipitate, and factor vii. the patient's renal function also declined, and the pump generated placement signal low alarms. later, care was withdrawn and the patient expired. | [5 empty] | A4 is unknown. no product was returned for investigation. the cause of the optical signal issue was not determined due to no product returned, inconclusive data log review, and insufficient clinical details. the cause of the cardiac arrest, major bleed, renal failure, and hemodynamic instability was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | Medical safety/clinical reviewed the event. the patient is a 53-year-old female with a medical history significant for coronary artery disease status post coronary artery bypass grafting (cabg), hypertension, systemic lupus erythematosus, and kidney transplant (1998). the patient presented with chest pain and was diagnosed with a non-st elevation myocardial infarction (nstemi). cardiac catheterization demonstrated a patent left internal mammary artery (lima) to the left anterior descending artery (lad), with significant disease in the left circumflex and right coronary arteries. a subsequent abdominal ct scan identified a 5.3 cm infrarenal abdominal aortic aneurysm (aaa). cardiothoracic surgery was consulted for placement of an axillary impella 5.5 device to provide hemodynamic support during high-risk percutaneous coronary intervention (pci), with planned bilateral femoral artery cutdowns to facilitate simultaneous endovascular repair of the abdominal aortic aneurysm (evar). attempted placement of the impella 5.5 was unsuccessful due to small and tortuous vessel anatomy. the device was removed and replaced with an impella cp, which was successfully placed and initiated on auto mode with normal flows. pci was performed on the remaining conduit to the lad, and the abdominal aortic aneurysm was stented via femoral access. a sponge plug was placed in the graft, partially externalized at the axillary incision site, and secured with three-point fixation. the patient remained on impella support post-procedure. on post-implant day 1, generalized bleeding was noted to have improved; however, renal function worsened. on 9/14/2025, the patient experienced a cardiac arrest (rhythm unspecified) and received two units of packed red blood cells. the following day, the patient required dialysis and vasopressor support with norepinephrine (levophed) and epinephrine for hypotension. impella support was continued for hemodynamic recovery. despite ongoing support measures, the patient expired while on impella support. no additional information was provided. the patient had pre-existing comorbidities, including advanced coronary artery disease, prior cabg, systemic lupus erythematosus, renal dysfunction requiring dialysis, and a large abdominal aortic aneurysm requiring intervention. the initial inability to advance the impella 5.5 due to small and tortuous vessel anatomy is consistent with vascular access limitations described in the device instructions for use (ifu). the patient¿s death appears multifactorial in the setting of complex cardiovascular disease, high-risk pci, aaa repair, worsening renal failure, bleeding, and hemodynamic instability. however, the impella cp, which was in use at the time of expiration, will be reported for the death. the event involving the impella 5.5 is a malfunction type of reportable event. h6: investigation type/finding/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
2383 | 1220648-2025-49467 | 23896181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896181&pc=OZD | 09/20/2025 |  |  | 12/26/2025 |  |  | 09/20/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702592 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female | 58 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. it was reported that the impella was removed through the peel away. a clot was expelled through the hemostatic diaphragm. the physician elected not to place a wire for removal as to not push clot forward. the peel away removed in one piece and percloses deployed successful. no further information was provided. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | Product complaint # (b)(4). investigation summary: no product was returned for investigation. thrombosis: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. minor bleed: peel away remains in place, some oozing but stable overnight. the cause of the bleeding is determined to be use issue because the peel away is left in place and not removed before transfer to icu. device history lot: device lot: 1947632. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
2384 | 1220648-2025-49458 | 23896254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896254&pc=OZD | 11/28/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 11/28/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1586834 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during support with an imepella pump the automated impella controller had a failure "console error" and it stopped working. patient outcome is unknown. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2385 | 1220648-2025-49322 | 23896381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896381&pc=OZD | 09/29/2025 | 02/24/2026 |  | 12/26/2025 |  |  | 09/29/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - |  |  | US |  | 2026655189 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male | 83 KG | Non Hispanic | Asian | Injury | Yes | Yes | Difficult or Delayed Positioning | Priming Problem | [5 empty] | Air Embolism | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the device was difficult to position. it was difficult to orient the left chamber, after trimming the artificial blood vessel, torque was applied to the shaft and moved to the apical direction. additionally, the air was not properly removed from the purge line, and an ultrasound confirmed air had been trapped inside. the physician decided to continue with the surgery and there was no report of patient harm or injury. the pump was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2386 | 1220648-2025-49320 | 23896391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896391&pc=OZD | 09/29/2025 | 12/24/2025 |  | 12/26/2025 |  |  | 09/29/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731315 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. after the 5.5 was explanted, the patient showed signs of a cerebrovascular accident. the patient was transferred to a different facility for clot extraction. | [5 empty] | Investigation summary: no product was returned for investigation. stroke: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the clinical symptom was not determined due to insufficient clinical details. device history lot: device lot: 1964695. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
2387 | 1220648-2025-49319 | 23896392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896392&pc=OZD | 09/28/2025 | 12/24/2025 |  | 12/26/2025 |  |  | 09/28/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728048 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Arrest | Respiratory Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the patient endured a cardiac arrest with hypoxia and bradycardia. staff successfully resuscitated the patient and turned the cp up to p-9. the pump was eventually successfully weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary ==> cardiac arrest/ arrhythmia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. respiratory failure: the cause of the injury was not determined due to insufficient clinical details. device history lot ==> device lot : 1960369. device history batch ==> sub-component lot : n/a device history review ==> the pump s/n: (b)(6) passed all the post sterile inspection checks.
2388 | 1220648-2025-49471 | 23896457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896457&pc=OZD | 11/30/2025 |  |  | 12/26/2025 |  |  | 11/30/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605571 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. impella cp site continuously oozing and heparin stopped. oozing would not stop so decision to remove impella. patient was stable as per physician. physician states impella would have stayed in until monday but due to oozing decision is to remove impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary no product returned. major bleed: impella cp site continuously oozing. heparin stopped. oozing would not stop so decision to remove impella. the cause of the access site adverse event was not determined due to insufficient clinical details. device history review this pump passed all post sterile inspection checks.
2389 | 1220648-2025-49313 | 23896520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896520&pc=OZD | 09/24/2025 | 12/24/2025 |  | 12/26/2025 |  |  | 09/24/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722886 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 95 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing high-risk percutaneous coronary intervention. post-procedure attempts to secure the left femoral artery access site with perclose sutures were unsuccessful, as the sutures broke repeatedly. manual pressure was applied, and angiography of the left femoral artery was performed via right femoral artery access, revealing a suspected leak. vascular surgery was consulted for a surgical cutdown to close the artery due to persistent bleeding and inability to achieve hemostasis with perclose or manual pressure. the patient's condition stabilized, and the impella cp device was subsequently weaned and explanted. | [5 empty] | Investigation summary: no product was returned for investigation. mechanical interaction with tissue: there were "no" clinical details that relates to the excessive force that was used during the procedure and the cause of the issue was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. vascular dissection: after removal, md attempted to tighten down perclose sutures but they both broke. attempted to deploy multiple perclose with same result. manual pressure being held. pressure low. accessed rfa to shoot lfa angio. believes there is a leak. elected to consult vascular surgery to cutdown and close artery. the cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history lot: device lot: 1967161. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
2390 | 1220648-2025-49315 | 23896521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896521&pc=OZD | 09/24/2025 | 12/24/2025 |  | 12/26/2025 |  |  | 09/24/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661139 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 107 KG | Non Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Dyspnea | Low Blood Pressure/ Hypotension | Ischemia | Hematuria | Hemodynamic instability | [8 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, a marked change in urine color was noted following an abdominal x-ray, raising concern for hemolysis. bedside echocardiography was performed, and the impella was repositioned. following repositioning, urine color normalized and no additional signs of hemolysis were observed. despite these interventions, the patient was found to have an anoxic brain injury on computed tomography and exhibited agonal breathing. the patient¿s cardiogenic shock and overall clinical condition did not improve despite advanced circulatory and supportive care. care was ultimately withdrawn, and the patient expired. | [5 empty] | Investigation summary - data review: data logs confirmed pump was in improper position corresponded with the time hemolysis was reported per clinical description that triggered alarms impella position in ventricle. no other issues were found on the logs. product was not returned for review. mechanical interaction with blood/hematuria: the cause of mechanical interaction with blood/hematuria was most likely due to pump was in improper position since hemolysis was resolved after positioning of the pump. ischemia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension / hemodynamic instability / dyspnea: the cause of the injury was not determined due to insufficient clinical details. device history lot - device lot #: 1920161. device history batch - subcomponent lot#: n/a. device history review - pump sn: (b)(6) passed all post sterile inspection checks.
2391 | 1220648-2025-49304 | 23896534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896534&pc=OZD | 09/24/2025 | 12/24/2025 |  | 12/26/2025 |  |  | 09/24/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729385 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Embolism/Embolus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the device experienced a high motor current alarm and the pump stopped. attempts to restart the pump were unsuccessful. the impella was removed due to suspected clot, and a large calcium chunk was found lodged inside the pump. the patient survived the procedure. | [5 empty] | Investigation summary. pump stop: clinical details report that there was high motor current alarm that resulted in a pump stop. the pump was explanted due to the complication, and calcium was noted to be lodged in the pump. data logs were reviewed and the pump stop was confirmed. when the pump was first started, pump flows were variable because mc was below the transition value, so ps was calculating flows. mc was generally stable during this period. roughly 10 minutes after starting the pump, there was a spike up in mc to ~200 ma with a ms deviation down. mc remained at an elevated level, however, it still was not above the expected range for the p-level (p-4). shortly after this, mc started an accelerated trend up, leading up to a spike to 1250 ma. at this instant, the first high motor current pump stop alarm (#13) triggered. the pump was unable to be restarted, continuing to trigger alarm #13 and #115 the same console as well as a backup. product was not returned for investigation. based on the signatures noted in data logs in conjunction with the report of biomaterial (calcium) being noted in the pump upon explant, the cause of the pump stop was determined to be biomaterial ingestion. embolism: clinical details report that a chunk of calcium was noted on the pump at the time of explant. the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history lot. device lot: 1960447. device history batch. subcomponent: na. device history review. pump sn (b)(6) passed all post-sterile inspection checks.
2392 | 1220648-2025-49295 | 23896578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896578&pc=OZD | 09/23/2025 | 12/24/2025 |  | 12/26/2025 |  |  | 09/23/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712261 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | Increase in Pressure | Malposition of Device |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, a high purge pressure alarm was noted, and a small fluid leak from the purge sidearm was observed. the origin of the leak could not be determined. the local team replaced the purge cassette without improvement. the purge solution was changed sequentially from d5w to heparin; however, the high purge pressure and low purge flow persisted, as did the small leak. the cannula was retracted approximately 1 cm and continued close monitoring of the motor current. the plan was to return to a heparin-based purge solution overnight. despite these interventions, the patient¿s condition continued to deteriorate, care was withdrawn, and the patient ultimately expired. | [5 empty] | Investigation summaryno product returned. high purge pressure - less than 24 hours after implant, high purge pressure occurred. purge flows under 1 ml/hr. changed pc to no resolve. purge composition changed to no resolve. issue resolved without further troubleshooting within 24 hours. the cause of the high purge pressure was not determined due to no product returned, no data logs recovered, and insufficient clinical details. purge leak - pump - less than 24 hours after implant, small leak of fluid was noted inside the purge sidearm. it couldn't be determined as to origin of leak. purge composition changed to no resolve. the cause of the purge leak was not determined due to no product returned and insufficient clinical details. device in wrong position - less than 24 hours after implant, pump repositioned less than 1 cm back. data logs were not recovered. the cause of the device in the wrong position was not determined due to no product returned, no data logs recovered, and insufficient clinical details. device history lot: device lot: 1947655. device history batch subcomponent lot: n/a. device history review - pump s/n (b)(6) passed all post sterile inspection checks.
2393 | 1220648-2025-49472 | 23896581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896581&pc=OZD | 11/30/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 11/30/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1359283 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during assistance with impella 5.5, the automated impella controller (aic) error alarm sounded and a display problem occurred where the position waveform was not displayed, so the aic was replaced. the pump continued to operate without any problems even during the alarm and after the aic replacement, and the patient's condition was stable without any issues. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2394 | 1220648-2025-49289 | 23896590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896590&pc=OZD | 09/23/2025 | 12/24/2025 |  | 12/26/2025 |  |  | 09/23/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026730011 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during wire manipulation across the aortic valve into the left ventricle, the patient developed polymorphic ventricular tachycardia (vt), likely due to wire contact with the coronary sinus. the patient was successfully defibrillated twice. following defibrillation, the impella 5.5 was successfully positioned, ramped up, and functioned as intended. vasopressor support was subsequently weaned, and the patient was transferred to the intensive care unit for continued management. the patient later experienced another episode of vt, which required synchronized cardioversion and was successfully resolved. a small pericardial effusion was later identified on computed tomography (ct), and a pericardial window procedure was performed. the patient¿s condition stabilized, and the impella 5.5 was subsequently weaned and explanted. | [5 empty] | Investigation summary: no product was returned for investigation. pericardial effusion: the cause of the clinical symptom was not determined due to insufficient clinical details. tachycardia: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1964716. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
2395 | 1220648-2025-49279 | 23896599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896599&pc=OZD | 09/18/2025 | 12/24/2025 |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026718770 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. the patient was successfully placed on the impella cp device prior to percutaneous coronary intervention (pci). catheterization laboratory (ccl) for device removal using a manta vascular closure device. a foley catheter was in place, and the patient¿s urine appeared dark amber in color. echocardiography demonstrated that the impella catheter was positioned beneath the papillary muscle. the physician declined to reposition the device and opted for removal, with consideration of potential intra-aortic balloon pump (iabp) placement. the patient¿s hemodynamics had improved following removal, and no additional mechanical circulatory support (mcs) device was placed. the patient¿s hemodynamic status improved after device removal. no further information was provided regarding the patient¿s subsequent condition. | [5 empty] | Investigation summary: data review: data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta. there were suction alarms triggered in early of the case but resolved quickly no other issues were found on the logs. product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely due to pump was not in the proper position. the cause of positioning issue is unknown. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot #: 1955298. device history batch: subcomponent lot#: n/a. device history review: pump sn: (b)(6) passed all post sterile inspection checks.
2396 | 1220648-2025-49285 | 23896600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896600&pc=OZD | 09/22/2025 | 12/24/2025 |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656810 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Nerve Damage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. the patient was successfully escalated from an impella cp to an impella 5.5. after placement via the right axillary artery, the patient experienced increased bleeding at the graft anastomosis site and reported tingling in the right arm. bleeding was controlled with additional sutures and blood products, and hemostasis was achieved. the patient¿s condition stabilized, and the device was subsequently weaned and explanted. | [5 empty] | Investigation summary ==> no product was returned for investigation. peripheral nerve injury: the cause of the clinical symptom was not determined due to insufficient clinical details. device history lot ==> device lot: 1911452 device history batch ==> subcomponent lot: n/a device history review ==> device with sn: (b)(6) passed all post sterile inspection checks.
2397 | 1220648-2025-49274 | 23896634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896634&pc=OZD | 09/22/2025 | 12/24/2025 |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740544 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Patient became hemodynamically unstable in ccl, requiring more cpr, while actively trying to put in impella cp. d/t patient anatomy and tortuous iliac the cp was unable to advance over .018 wire. decision made to abort placement of impella cp. patient expired in ccl. | [5 empty] | Investigation summary: no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult patient anatomy and tortuous iliac preventing successful delivery of impella. device history lot: device lot: 1970191. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6) passed all post sterile inspection checks.
2398 | 1220648-2025-49269 | 23896662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896662&pc=OZD | 09/22/2025 | 12/24/2025 |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612706 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. oozing was noted around the insertion site, requiring two dressing changes. the patient was continued on broad-spectrum antibiotics, and shock laboratory parameters were trended to guide further management and assess the need for additional escalation of support. due to ongoing hemodynamic requirements, the patient was subsequently escalated from the impella cp to an impella 5.5 device. | [5 empty] | Investigation summary: no product was returned for investigation. fever: the cause of the clinical symptom was not determined due to insufficient clinical details. major bleed: oozing around the peel away sheath at the insertion site. full set of labs being sent. echo to bedside showing impella positioning itself unstable, moving with every contraction. the cause of the bleeding is determined to be use issue because of untastable and improper positioning of impella as per clinical details. device history lot: device lot: 1874053. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
2399 | 1220648-2025-49277 | 23896766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896766&pc=OZD | 09/22/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701204 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was attempted to be implanted with an impella cp for mechanical circulatory support. during the procedure, there was difficulty advancing a pigtail catheter into the left ventricle, which was eventually successful. however, the impella cp could not be advanced through the iliac artery due to challenging anatomy. multiple attempts with various wires were unsuccessful, and impella cp placement was ultimately aborted due to the patient¿s anatomical limitations. | [5 empty] | Investigation summary ==> no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy along with resistance in femoral artery preventing successful delivery of impella. device history lot ==> device lot: 1951048 device history batch ==> subcomponent lot: n/a. device history review ==> device (sn: (b)(6) passed all post sterile inspection checks.
2400 | 1220648-2025-49473 | 23896842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23896842&pc=OZD | 12/12/2025 | 02/25/2026 |  | 12/26/2025 |  |  | 12/12/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694328 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Patient Device Interaction Problem | Material Split, Cut or Torn | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported 69 year old male patient on an impella cp for acute myocardial infarction. during support, it was reported that the plasma free hemoglobin was drawn and found to be 150. the pump was found to be deep, but when repositioned with echo is now acceptable. the p level dropped from p6 to p5. it was further reported that while repositioning the impella in icu with echo, the sleeve covering the catheter ripped away. the patient survived the procedure. | Additional information was provided and the patient outcomes is unknown as patient was transferred to a different hospital. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the anticontamination sleeve product damage and the mechanical interaction with blood has been completed. the cause of the anticontamination sleeve damage was not determined due to no product returned and insufficient clinical details. the cause of the mechanical interaction with blood was not determined due to no product returned, no data logs recovered, and insufficient clinical details. | Additional information was received and b5 was updated
2401 | 1220648-2025-49438 | 23897172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897172&pc=OZD | 09/19/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714430 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Restricted Flow rate | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the device exhibited suctions issues, the patient was administered lactated ringer's and 1 unit of packed red blood cells (prbc). no further suction issues were reported. additionally, the patient experiened a gastrointestinal bleed and heparin was withheld. no further information was provided; however, it was later reported that the patient expired. | [5 empty] | Investigation summary : major bleed: heparin was held during the case due to a gi bleed and a nosebleed later on as well. the product was not returned for investigation. the cause of the bleeding could not be determined due to insufficient clinical details. stroke: in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: clinical reported that suction began to trigger multiple times when the patient was standing up. fluids were given and suction reportedly resolved. the product was not returned for investigation. data log review confirms suction events occurring on the day it was reported. at the same time suction started, an impella position unknown alarm triggered as well. the placement signal at this time loses pulsatility, and the placement signal waveforms are no longer aortic at this time. there were no motor current spikes during the case. the cause of the suction is most likely patient related, as placement signal loses pulsatility at the time of suction, and suction reportedly resolved with volume. device history lot:device lot number: 1951079. device history batch: subcomponent lot: n/a. device history review : pump sn (b)(6) passed all post sterile inspection criteria.
2402 | 1220648-2025-49439 | 23897195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897195&pc=OZD | 09/19/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026721381 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support the patient had some coagulopathy requiring multiple blood products. additionally, dialysis was started for worsening kidney and liver function. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary: no product was returned for investigation. major bleeding: pt had some coagulopathy requiring multiple blood products. the cause of the bleeding is determined to be patient condition because the patient is having coagulopathy requiring multiple blood products. acute kidney failure: the cause of the clinical symptom was not determined due to insufficient clinical details. device history lot: device lot: 1957239. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
2403 | 1220648-2025-49442 | 23897223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897223&pc=OZD | 09/19/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686766 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female |  |  |  | Death | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced hemolysis. no further information was provided; however, it was later reported that the patient expired. | [5 empty] | Investigation summary: mechanical interaction with blood (hemolysis): the product was not returned for investigation. according to the clinical details, hemolysis was reported on 19-sep based on an increase in ldh. the pump was utilized for another 3 days without any mentioned troubleshooting. during the time of reported hemolysis, data log shows a fluctuating placement signal with an opposite trend in pump flow, without any reported pump suction, positioning alarm, or motor current spike. the pump continued to be used for another 3 days, with deteriorating placement signal and pump flow trend noted at the end of the case. the cause of the hemolysis was not determined without returned product investigation and sufficient clinical details. device history lot : device batch: 1931579. device history batch : subcomponent lot: na. device history review : the pump sn (b)(6) passed all post sterile inspection checks.
2404 | 1220648-2025-49443 | 23897224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897224&pc=OZD | 09/19/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  | 10/31/2025 |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422240 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 76 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced hemolysis and the pump was repositioned. no further information was provided; however, it was later reported that care was withdrawn and the patient died. | [5 empty] | Investigation summary : data review: data logs showed pump ran without issue during support. there were no alarms triggered during the case. there were no mc spikes, abnormal ps or purge issues observed during support. mechanical interaction with blood: product was not returned for review. the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history lot : device lot #: 1760495. device history batch : subcomponent lot#: n/a. device history review : pump sn: (b)(6) passed all post sterile inspection checks.
2405 | 1220648-2025-49446 | 23897225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897225&pc=OZD | 09/19/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026748795 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced septic shock, blood bags and volume were administered. no further information was provided; however, it was later reported that care was withheld and the patient expired. | [5 empty] | Investigation summary : sepsis : in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot : 1979720. device history batch : sub-component lot : n/a. device history review : the pump s/n: passed all the post sterile inspection checks.
2406 | 1220648-2025-49461 | 23897232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897232&pc=OZD | 09/27/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729489 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Pulmonary Edema | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient underwent additional procedures including a carotid endarterectomy with hematoma evacuation. the patient experienced an episode of atrial fibrillation that was managed with an amiodarone drip without adverse blood pressure effects. the patient¿s condition stabilized, and the device was subsequently weaned and explanted. | [5 empty] | Investigation summary: hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. arrhythmia : in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. pulmonary edema : the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot : 1962539. device history batch: sub-component lot : n/a. device history review: the pump s/n:(b)(6) passed all the post sterile inspection checks.
2407 | 1220648-2025-49456 | 23897235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897235&pc=OZD | 09/22/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026720069 |  | 000000000001000459 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation | Distributor |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with postcardiotomy cardiogenic shock and low cardiac output syndrome underwent implantation of an impella 5.5 device for mechanical circulatory support. the pump was successfully positioned in the left ventricle; however, frequent pump position alarms occur once the device was activated. subsequently, alarms indicating elevated purge pressure and a possible obstruction in the purge system were triggered. suspecting a blockage, the yellow luer connector was removed to assess changes in purge pressure. although the purge pressure decreased after removal, the underlying issue persisted, and the situation was deemed unsafe. the case was therefore aborted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2408 | 1220648-2025-49453 | 23897242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897242&pc=OZD | 09/22/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750091 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with acute myocardial infarction and cardiogenic shock underwent implantation of an impella cp device for mechanical circulatory support. during preparation, the scrub technician inadvertently removed the easy guide lumen prior to backloading the impella over the abiomed 0.018" guidewire. the cardiologist elected not to proceed with backloading the impella without the easy guide lumen in place. consequently, a new impella cp device was utilized for the procedure. the replacement impella cp was successfully used to provide circulatory support. the care team planned for escalation of care, with cardiothoracic surgery (ct surgery) and the cardiologist present at the bedside. the patient remained on impella support for left ventricular (lv) unloading, with weaning of vasopressors as tolerated. the patient remained in guarded condition following the event. ultimately, the family made the decision to withdraw care later the same day. | [5 empty] | Investigation summary unable to deliver or advance: scrub tech removed easy guide lumen prior to backloading impella over abiomed 0.018 wire. the cause of the issue was determined to be use related as evidenced by clinical details citing the physician's decision not to proceed with impella insertion following the removal of easy guide lumen device history lot device lot number: 1958372. device history batch subcomponent lot: n/a. device history review the pump sn (b)(6) passed all post sterile inspection checks.
2409 | 1220648-2025-49469 | 23897246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897246&pc=OZD | 09/19/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/20/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026685977 |  | 000000000001000459 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, a leak occurred from the catheter plug. the leaking catheter plug was treated by wrapping gauze around it. there were no purge-related alarms or increased purge fluid flow rate. there was no increase in motor current consumption and the pump was operating normally. the pump was successfully weaned and explanted, no report of patient harm or injury. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2410 | 1220648-2025-49464 | 23897255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897255&pc=OZD | 09/21/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/21/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714435 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Hematoma | Pain | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient had runs of ventricular tachycardia and a hematoma in the right axillary pocket. additionally, the pump exhibited low flow. no further information was provided. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary: clinical symptoms (arrhythmia): in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms (pain): the cause of the injury was not determined due to insufficient clinical details. low or blocked pump flow: the pump was not returned for investigation. according to the clinical details, low flow was reported on 04-oct. data log shows the pump p-level was reduced from p8 to p4 on 02-oct while experiencing suction alarms. later, flow was seen to gradual decrease below 1 l/min, with some improvement noted while running on steady p4. there was no conclusive trend noted on the placement signal related to patient condition. no motor current spike or positioning issue was noted during the time of low pump flow. the cause of the low pump flow was not determined without returned product investigation and sufficient clinical details. access site adverse event (hematoma): the cause of the access site adverse event was not determined without returned product investigation and sufficient clinical details. device history lot: device batch: 1951078. device history batch: subcomponent lot: na. device history review: the pump sn (b)(6) passed all post sterile inspection checks.
2411 | 1220648-2025-49475 | 23897329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897329&pc=OZD | 11/14/2025 | 02/03/2026 |  | 12/26/2025 |  |  | 12/14/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026771475 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 88 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a 77 year old male was implanted with a impella cp for support during a high risk cardiac procedure. it was reported that there was red tinged urine due to hemolysis with last hgbf >170. patient expired and cause of death is unknown. | [5 empty] | H6: updated code a01 based on information in the complaint file. | Correction: e4 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. data log shows the pump was set to p9 throughout the case run. there were no suction alarms, positioning alarms, motor current spikes, or optical signal issues reported during the entire case run. placement signal had a fluctuating trend from the beginning, with some deteriorating trend at the end of the case run. pump flow was also seen fluctuating within specification range while running on a steady p-level. mechanical interaction with blood (hemolysis): the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical details. clinical symptom (hematuria): the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2412 | 1220648-2025-49470 | 23897333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897333&pc=OZD | 09/20/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/20/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026716146 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient had a cardiac tamponade overnight and taken to the operating room for an emergency wash out, chest tubes were exchanged, and the chest was closed. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary. no product was returned for investigation. cardiac tamponade: the cause of the clinical symptom was not determined due to insufficient clinical details. device history lot. device lot: 1954306. device history batch. subcomponent lot: n/a. device history review. device with sn:614357 passed all post sterile inspection checks.
2413 | 1220648-2025-49401 | 23897339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897339&pc=OZD | 11/06/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 11/06/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764112 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 86 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Asystole | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 83-year-old male patient was implanted with impella cp during non-st-elevation myocardial infarction suppport. it was reported that patient went asystole this morning compressions were started, and return of spontaneous circulation was obtained. patients son and care team decided to withdraw care. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2414 | 1220648-2025-49405 | 23897344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897344&pc=OZD | 11/06/2025 | 01/22/2026 |  | 12/26/2025 |  |  | 11/06/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725269 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Fever | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 66-year-old male was implanted with impella 5.5 for support for cardiac arrest. impella cp was placed. patient cardioverted with ¿wide complex tachycardia,¿ and the decision was then made to initiate transfer to facility. patient was on multiple drips on arrival, c-shock called and decision to escalate to 5.5. doctor felt pump not flowing as expected as pulmonary capillary wedge pressure is high. wants pump returned for analysis. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9 updated; the product has been returned. the investigation is still ongoing.
2415 | 1220648-2025-49403 | 23897345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897345&pc=OZD | 11/06/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 11/06/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026767598 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 57 year old male was implanted with impella cp support during acute heart failure exacerbation and ventricular tachycardia. it was reported that during support, patient had oozing from impella access site, sand bag placed to site. stick angle matched and site redressed. patient had hemolysis through entire ventricular tachycardia ablation. impella cp was explanted without any difficulty. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2416 | 1220648-2025-49332 | 23897347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897347&pc=OZD | 11/15/2025 | 02/25/2026 |  | 12/26/2025 |  |  | 11/16/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026745496 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported that a 71 yr old male was implanted with a impella 5.5 for support during high risk cardiac procedure. a hematoma was noted at incision site from cutdown to axillary artery. possible return to operating room for surgical washout. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the hematoma was not determined due to insufficient clinical details.
2417 | 1220648-2025-49338 | 23897356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897356&pc=OZD | 11/19/2025 | 01/13/2026 |  | 12/26/2025 |  |  | 11/19/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620050 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Backflow | Restricted Flow rate | Device Sensing Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support. it was noted that on day 1 of impella support the patient was also cannulated for extracorporeal membrane oxygenation (emco) support. on support day 2, the patient experienced a retroperitoneal bleed which the medical team believed was caused by the intra-aortic balloon pump puncture site. it was noted that the placement signal displayed on the automated impella controller for the pump was much lower than the patient¿s blood pressure and appeared to be causing potentially false retrograde flow alarms. the patient was taken to the operating room for bleeding in the chest cavity, and 5.5 l of blood and clot was removed from the chest. the patient was transfused with blood products as a result of blood loss. it was noted that the bleeding in the chest cavity was from prior surgery and not related to impella. on support day 3, the patient was weaned down on ecmo, and impella flows were increased. however, the patient¿s native heart function was low and was unable to compensate. the arising hypotension was treated with nitroglycerin and the patient¿s pulsatility was monitored before attempting to wean again. on support day 4, the patient was decannulated from ecmo support. additionally, the patient was transfused with a unit of platelets. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed. | H6 coding updated. the investigation is still ongoing. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical details note that the ps was much higher than bp and seemed to be causing flows displayed very low and retrograde flow alarms. it was later noted that the facility felt that impella ps issue was from the flow of the ecmo because once ecmo was removed, the ps was accurate. data log review showed a ps increase above 150 mmhg as reported on 19-nov-2025 with suction and preventing retrograde flow alarms. there was slight variation in optical parameters at this time but no indication of sensor damage. since the 5.5 was running below p-4 at this point, the pump flow is calculated with ps as a variable. around the time of the reported issue, the minimum mc values were around the expected values at the corresponding motor speed, but the mc had low pulsatility overall. there was variation in mc throughout the case as the pump attempted to maintain motor speed at the corresponding p-levels, but there were no clear motor current spikes indicating ingestion. the cause of the reported ps increase is unknown. there was intermittent low snr and increases/decreases in ps throughout the majority of the case. the correlation between motor current and snr appears inconsistent throughout the case as well as the correlation between snr and placement signal. there is no clear evidence that the console was malfunctioning. without a returned pump, the cause of the placement signal increase and low/retrograde low cannot be determined. optical sensor issue: the cause of the optical signal issue was not determined since product was not returned. low or blocked pump flow: the cause of the low pump flow was not determined since product was not returned. retrograde pump flow: the cause of the retrograde pump flow was not determined since product was not returned. major bleed: clinical details note bleeding in chest cavity from prior surgery not related to impella. the cause of the major bleed was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2418 | 1220648-2025-49480 | 23897471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897471&pc=OZD | 11/06/2025 |  |  | 12/26/2025 |  |  | 11/06/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733551 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 100 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella experienced gastrointestinal bleeding. in response, heparin was withheld. the patient survived. | [5 empty] | The cause of the major bleed was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks
2419 | 1220648-2025-49481 | 23897481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897481&pc=OZD | 10/25/2025 |  |  | 12/26/2025 |  |  | 10/25/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026743360 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was subsequently escalated to an impella 5.5 and would eventually expire. the patient presented to the emergency department approximately 6 days prior to the initial impella implant with shortness of breath, atrial flutter, and a congestive heart failure exacerbation. the patient initially improved in the cardiac intensive care unit (cicu) but decompensated over the weekend, experiencing a ventricular tachycardia (vt) episode for which his automatic implantable cardioverter-defibrillator delivered a shock. the patient required intubation and vasopressor support (phenylephrine). a decision was made to place a heart pump to stabilize the patient for transfer to an outside hospital for advanced therapy. the impella cp heart pump was successfully placed via the right femoral artery. the patient was successfully transferred to the hospital and arrived in stable condition. however, the following day on support day two, the patient¿s plasma free hemoglobin, lactate dehydrogenase, and lactate levels all increased. the impella was reduced to p-4 performance level, and vasopressors were held. a transthoracic echocardiogram (tte) showed good pump position but raised suspicion for a mobile mass on the pigtail catheter and potential pneumonia. the following day, the patient experienced multiple atrial fibrillation events requiring cardioversion and increased medication requirements. hemolysis was noted to have been resolved. the patient developed right lower lobe pneumonia on computed tomography scan. on day four of impella cp support, the patient continued to decline, and lactic acid results trended up. the clinical team determined the patient was a "bridge to nowhere" due to poor prognosis but later reversed the decision to escalate care. the medical team decided to escalate the patient to an impella 5.5. the impella 5.5 was implanted and the impella cp was explanted without complications. on day two of impella 5.5 support, the patient remained critically stable; however, the medical team was hopeful for recovery. on day six of impella 5.5 support, it was noted that the patient was making an incredible recovery with an improved ejection fraction of 30% on bedside echocardiogram, successful extubation, and intact neurological function. the following several days the patient continued to recover and the patient began physical therapy and was stable on impella p-6 and dobutamine. the team determined the patient was not a candidate for immediate advanced therapies due to extensive ischemic coronary artery disease not amenable to intervention, suggesting destination therapies might be required if improvement continued. however, the patient's condition significantly worsened after developing sepsis secondary to an ileus. the patient expired while on support. | [5 empty] | E1 added zip code extension as it was omitted from the initial report that was submitted. h6 a new health effect - clinical code was added. health effect - impact code f12 was removed. a new medical device problem code was added and code a01 was removed. | Patient-specific information for a4. weight is unknown at the time of this mdr writing. the exact event date and date of death are unknown but have been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. medical safety review. medical safety reviewed the event and noted the following: a patient with exacerbated congestive heart failure, atrial fibrillation, and ventricular tachycardia underwent successful implantation of an impella cp device for hemodynamic support. the patient was subsequently transferred to a hub hospital. approximately 48 hours post-implant, laboratory findings revealed elevated ldh, plasma-free hemoglobin, and lactate levels, raising suspicion of hemolysis. a transthoracic echocardiogram confirmed appropriate pump positioning but noted a possible mobile mass on the pigtail and potential pneumonia. in response, device support was reduced to p-4, after which hemolysis resolved. during this period, the patient experienced multiple atrial fibrillation episodes requiring pharmacologic intervention and cardioversion. on day 4, due to continued clinical decline, support was escalated to an impella 5.5 device. the patient initially demonstrated improvement; however, on day 11, the patient developed sepsis secondary to ileus, resulting in significant clinical deterioration and subsequent death. the occurrence of hemolysis during impella cp support is considered a serious injury, despite resolution following standard troubleshooting measures. the patient¿s death is conservatively attributed to the impella 5.5 device, as the patient was on support at the time of death. however, sepsis secondary to ileus is assessed as the most likely contributing factor. device involvement: impella cp associated with hemolysis (resolved). and the impella 5.5: patient was on support at time of death; no direct device malfunction identified. iin conclusion: serious injury (hemolysis) and death reported. death likely related to underlying clinical condition (sepsis from ileus) rather the impella pump. device status. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: sepsis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1976097. device history batch: subcomponent lot: n/a. device history review: the pump sn (b)(6) passed all post sterile inspection checks.
2420 | 1220648-2025-49482 | 23897482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897482&pc=OZD | 10/04/2025 | 11/18/2025 |  | 12/26/2025 |  |  | 10/04/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760663 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 89 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis and hematuria, and lab results indicative of kidney impairment. the impella was explanted and replaced with an impella 5.5. | [5 empty] | The pump was not returned for review. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the kidney impairment was not determined due to insufficient clinical details.
2421 | 1220648-2025-49476 | 23897484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897484&pc=OZD | 12/11/2025 | 02/12/2026 |  | 12/26/2025 |  |  | 12/11/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026745537 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp via the right femoral artery for mechanical circulatory support. on day 2 of impella cp support, it was reported that the patient¿s laboratory samples were hemolyzed. the performance level of the pump was turned down to attempt to alleviate the hemolysis. an echocardiogram was performed and revealed that the pump was angled towards the mitral valve. it was noted that repositioning the device did not resolve the hemolysis. subsequently, dialysis was initiated. the patient was escalated to an impella 5.5. the impella 5.5 had no reported complications or patient harm reported. the impella 5.5 was explanted successfully after weaning. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the device positioning issue, hemolysis, and organ failure has been completed. data logs were returned for review. since data logs were inconclusive, clinical details reported troubleshooting did not resolve the issue and product wasn't returned for investigation, the cause of the hemolysis could not be determined. since insufficient clinical details were provided and product wasn't returned for investigation, the cause of the device in wrong position could not be determined. the cause of the injury was not determined due to insufficient clinical details.
2422 | 1220648-2025-49483 | 23897485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897485&pc=OZD | 10/04/2025 |  |  | 12/26/2025 |  |  | 10/04/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712075 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 60 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Unspecified Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to heart transplant. on support day 13, the patient experienced a drop in hemoglobin and hematocrit and as a result, heparin was withheld. two days later, heparin was still on hold with further hemoglobin drop and the patient experienced melena. on day 17 of support, the patient received a unit of packed red blood cells from the blood loss. on support day 19, multiple bleeding spots were found on the patient¿s small intestine that were cauterized to resolve. anticoagulation therapy remained withheld and the patient¿s transplant status was changed to status 7 for inactive due to the ongoing bleeding. the following day, it was reported that the patient had melena in stool and the patient¿s blood pressure dropped and required vasopressin. the patient received two units of packed red blood cells, and the systemic heparin was withheld. the patient¿s hemoglobin was reported to be 7.7 g/dl and a hematocrit of 24.3%. the patient remained as status 7/inactive for transplant. on support day 24, the patients bleed had resolved, heparin was restarted, and the patient was relisted as status 2a. on support day 28, the patient was noted to have a urinary tract infection the patient was successfully weaned from the impella 5.5 during the transplant procedure and underwent a successful heart transplant. the patient survived the procedure. | [5 empty] | The cause of the major bleed was not determined since product was not returned and insufficient clinical details provided. the cause of the hypotension was not determined due to insufficient clinical details. the cause of the infection was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
2423 | 1220648-2025-49484 | 23897486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897486&pc=OZD | 10/02/2025 |  |  | 12/26/2025 |  |  | 10/02/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1323491 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a placement signal not reliable alarm. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2424 | 1220648-2025-49485 | 23897498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897498&pc=OZD | 10/01/2025 |  |  | 12/26/2025 |  |  | 10/01/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731032 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a ¿placement signal not reliable¿ alarm occurred during setup of an impella cp. no patient harm was reported. | [5 empty] | A4 is unknown. the root cause of the optical sensor issue was due to eeprom corruption.
2425 | 1220648-2025-49486 | 23897508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897508&pc=OZD | 10/01/2025 |  |  | 12/26/2025 |  |  | 10/01/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710457 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Pain | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced pain at the access site. pain medication was given intravenously. additionally, the anticontamination sleeve became detached from the catheter. tegaderm was used to cover the damage. | [5 empty] | The pump was not returned for investigation. data log review was not required for the investigation. the cause of the sterile sleeve damage issue was not determined without returned product investigation and sufficient clinical details. cthe cause of the access site pain was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2426 | 1220648-2025-49487 | 23897536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897536&pc=OZD | 09/30/2025 |  |  | 12/26/2025 |  |  | 09/30/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658489 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Restricted Flow rate | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | Inflammation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced suction and impella position unknown alarms. the alarms did not correlate to the non-invasive blood pressure readings. additionally, there was a kink in the purge line which resulted in alarms. the kinks were fixed and the alarms resolved. the patient also experienced inflammation requiring medication. the patient is in stable condition. | [5 empty] | No product was returned for evaluation. data logs were returned for analysis. the root cause of the positioning issue was not determined due to inconclusive evidence from the clinical details and data log analysis. the root cause of the pump suction was not determined due to inconclusive evidence from the clinical details and data log analysis. the root cause of the purge tubing kink could not be determined due to unreturned product for further evaluation. the cause of the inflammation was not determined due to insufficient clinical details. the device passed all post-sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2427 | 1220648-2025-49489 | 23897545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897545&pc=OZD | 09/26/2025 |  |  | 12/26/2025 |  |  | 09/26/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1175553 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Unintended Application Program Shut Down | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller became unresponsive to p-level changes during ecmo de-cannulation. first, the screen appeared blank and then the screen was stuck on the home screen. the controller was unresponsive to buttons or knob selections. the controller was exchanged for another controller to resolve the issue. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6a and d6b should of been left blank on the initial report that was submitted. h6 codes b17 and c20 were reported incorrectly on the initial report that was submitted as the automated impella controller was returned. a new type of investigation code and investigations findings code were added. h10 this is one of two related reports. this report represents the automated impella controller and another report was submitted to represent the impella cp. related report number was added. h11 revised additional manufacturer narrative as it was reported incorrectly on the initial report that was submitted. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | This follow up mdr is being submitted to document that the device involved in the reported event has been identified as available for return. additionally it was confirmed that replacing the automated impella controller resolve the reported issue. upon completion of the investigation, a supplemental report will be filed.
2428 | 1220648-2025-49492 | 23897562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897562&pc=OZD | 09/24/2025 |  |  | 12/26/2025 |  |  | 09/24/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - |  |  | US |  | 2026709936 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 66 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced bleeding from the left internal mammary artery. the bleed was resolved with sutures. after a post coronary artery bypass graft the patient suffered a stroke. | [5 empty] | No product was returned for investigation. the cause of the bleeding is determined to be patient condition because the bleeding is reported at non-impella site. the root cause of the stroke was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2429 | 1220648-2025-49495 | 23897566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897566&pc=OZD | 09/22/2025 |  |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026728887 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | Patient Device Interaction Problem |  |  |  |  | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. device placement was confirmed by transesophageal echocardiography. the impella 5.5 was increased to performance level p4, and cardiopulmonary bypass was successfully weaned. upon separation from bypass, the patient became hypoxic, and veno-arterial extracorporeal membrane oxygenation was initiated. later in the course, the patient experienced a low placement signal alarm with a ventricularized waveform and pink urine, suggestive of hemolysis. echocardiography revealed suboptimal device positioning, and the impella 5.5 was successfully repositioned under echocardiographic guidance. the patient¿s condition stabilized, and the device was subsequently weaned and explanted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2430 | 1220648-2025-49494 | 23897567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897567&pc=OZD | 09/22/2025 |  |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711299 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, active suction alarms were noted on the impella cp with support fluctuating between performance levels p1 and p2. the patient received three units of packed red blood cells (prbcs) and one liter of intravenous fluids (ivf). urine output was positive and yellow in color. the care team planned to initiate continuous renal replacement therapy (crrt). bleeding was observed around the repositioning sheath and cannulation sites. following plasmapheresis administration of vitamin k, the patient¿s bleeding improved. the pump was eventually weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary ==> no product was returned. thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of bleed was most likely use issue related since the act was 203 in the icu which was above the recommended range. pump suction: no product was returned. patient had suction alarms. volume resuscitation post-procedure. suction resolved. after reviewing the logs, multiple suction alarms were observed. the cause of pump suction was most likely patient condition related since suction was resolved after giving volume per clinical. device history lot ==> device lot: 1943308. device history batch ==> subcomponent lot: n/a. device history review ==> device (sn: (b)(6) passed all post sterile inspection checks.
2431 | 1220648-2025-49493 | 23897569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897569&pc=OZD | 09/22/2025 |  |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1071329 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | False Alarm | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella for mechanical circulatory support. during support, the console was experiencing an impella position unknown issue. motor currents were flat with flows dipping to 0.6l/min at p6. an echocardiogram was completed bedside and confirmed adequate positioning. the decision made to swap consoles due to low flows experienced at times. it was reported that the procedure was successful and the patient was stable. | [5 empty] | Data analysis: review of the imc log reveals 31 instances of impella position unknown (ipu) alarms, confirming the complaint of ¿console experiencing impella position unknown.¿ twenty-four of these alarms are classified as being due to low pulsatility (pulsatility.png), with the remaining 7 being attributed to ¿suction or ao-lv pressure decoupling¿ (decoupling.png). reviewing the rt logs at the time of low pulsatility alarms reveals that these alarms occurred during periods of relatively low motor current pulsatility (rtlogpulsatility.png), while the rt logs for the pressure decoupling alarms show that these occurred during periods of lower overall ventricular pressure and relatively flat differential pressure measurements (rtlogdecouplingpressure.png). reviewing all lt logs does not reveal any periods of mean pump flow falling as low as 0.6l/min at p-level 6 (ltlogplevel6all.png), but there are periods where it falls to approximately 2l/min while the mean motor current stays relatively consistent (ltlogplevel6zoom1.png and ltlogplevel6zoom2.png). reviewing the rt logs for these periods it can be seen that the flow rate does dip as low as ~0.6l/m, however, these reductions were accompanied by corresponding overall reductions in instantaneous motor current (rtlogmotorcurrent.png). an rlfs request for the logs for the second console (ic1006) used in the case was submitted, but they are unavailable for comparison. device analysis: an impella 5.5 was connected to the console and allowed to run for approximately 2 minutes at every p-level, 1 through 9. the same pump was then moved to a test console and this procedure was repeated. the rt logs for both consoles were compared side by side and the flow rates at each p-level were found to be consistent (flowcomparison.png). the analog output of the optical bench ccd array was inspected with an oscilloscope and no distortion was noted. the optical 5s parameters were tested using two different test cavities and were well within the specifications listed in pm procedure 0042-7309. the optical fiber in the blue pump receptacle was examined with a camera and was found to be extremely contaminated (receptacleprecleaning.png). it was cleaned with a one-click cleaner and re-examined and found to be much less dirty, though considerable debris/contamination remain on the cladding (receptaclepostcleaning.png). the console was opened and a general inspection for damage, loose cables/connections, etc. was performed with no issues discovered. root cause: while the complaint behavior is confirmed via log review, no problems were found with the console during investigation and it appears to have operated as designed. corrective actions: before returning the console to the customer, the following actions are to be performed: replace the front panel optical connector (0042-2750) due to the presence of contamination that persists after cleaning. perform a full functional check (0042-7316).
2432 | 1220648-2025-49497 | 23897619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897619&pc=OZD | 09/20/2025 |  |  | 12/26/2025 |  |  | 09/20/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695416 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 131 KG |  |  | Malfunction | Yes | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support the impella was pulled out of the left ventricle by the nurse accidentally. the device was replaced with another impella. it was later reported that the procedural outcome is patient expired. | [5 empty] | Investigation summary no product was returned. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the cath lab nurse accidentally. device history lot device lot: 1939789. device history batch subcomponent lot: n/a. device history review ==> device (sn: (b)(6) passed all post sterile inspection checks.
2433 | 1220648-2025-49500 | 23897665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897665&pc=OZD | 09/19/2025 |  |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701698 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Restricted Flow rate | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the device exhibited suction alarms, the patient was administered lr (lactated ringer's) and albumin. the position of the device was verified, the nurse turned the device down to p7 and this resolved the suction alarms. additionally, the optical sensor was not working, the resolution for this issue is unknown. it was also reported that the patient experienced gastrointestinal issues, the customer stated that the patient had gastrointestinal issues prior to arrival at the hospital and impella insertion. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2434 | 1220648-2025-49501 | 23897670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897670&pc=OZD | 09/18/2025 |  |  | 12/26/2025 |  |  | 09/18/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708167 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Fever | Hemolysis | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the patient experienced hemolysis during support, with reported continuous renal replacement therapy. sepsis protocol was also reported. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | A4 is unknown. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the sepsis and fever was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2435 | 1220648-2025-49504 | 23897691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897691&pc=OZD | 09/15/2025 |  |  | 12/26/2025 |  |  | 09/15/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562383 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Difficult to Open or Remove Packaging Material | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced a thrombosis causing aphasia. the clot was removed. additionally, the sterile sleeve was torn from the touhy. | [5 empty] | A4 is unknown. the returned pump confirmed the sterile sleeve was torn off the touhy. the cause of the product damage was reasonably foreseen misuse of the device due to manipulation attempt of the sleeve while the touhy was locked. the data log review showed rising purge pressure with purge flows decreasing and slight motor current increase, but no spike. the pump was returned cut but biomaterial was found in the purge gap. the cause of the purge system blocked was biomaterial build up. the device passed all post sterile inspection checks.
2436 | 1220648-2025-49488 | 23897782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897782&pc=OZD | 12/12/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 12/12/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605546 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that impella cp, being used on a 74 year old male patient for support during a high risk cardiac procedure, had an "optical sensor failure;' that occurred mid procedure. there were no aortic or left ventricular waveforms or values on aic. the troubleshooting for this event is unknown but the procedure was completed successfully. there was no patient harm reported. | [5 empty] | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the optical signal issue was not determined due to lack of product and data logs returned.
2437 | 1220648-2025-49506 | 23897918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23897918&pc=OZD | 12/12/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 12/12/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768115 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male | 65 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an 85-year-old male patient on an impella cp for support due to acute myocardial infarction. during support, the patient's hemoglobin dropped to 6.8 requiring a unit of packed red blood cells overnight and this afternoon. no other products to be given at this time per attending of record. patient is on therapeutic anticoagulation the patient's left groin, area surrounding impella insertion site, appears large/moderately taunt compared to right femoral area. vascular surgery has been following patient since insertion given implanting physician¿s concerns over vessel size and tortuosity however impella was placed in emergent circumstance. pt remains on full heparin purge, systemic heparin drip and cangrelor drip. additionally, the patient¿s urine was noted to be pink/red. pump position checked with echo and it was decided not to reposition. the pump was eventually removed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. asae/major bleed: patient was on therapeutic anticoagulation, patient has calcification and tortuosity of bilateral femoral vessels. the root cause of bleeding adverse event/major bleeding issue most likely was related to patient condition. hemolysis/mechanical interaction with blood: the cause of the hemolysis was not determined due to lack of product and data logs returned. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e1 added initial reporter phone and e4 added section as it was omitted from the initial report that was submitted.
2438 | 1220648-2025-49499 | 23898153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898153&pc=OZD | 09/19/2025 |  |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658637 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Detachment of Device or Device Component | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp developed new onset neuro changes and heparin was withheld. during explant, the physician pulled the device into the graft and during that action, the cannula of the 5.5 separated from the rest of the catheter and remained inside the graft and the axillary artery. this led to bleeding through the graft that was not controllable and thus the surgeon pushed the cannula back into the axillary artery to gain control of the bleeding through the graft. the cannula was retrieved from the axillary artery via fluoroscopy and the axillary artery was repaired with a bovine patch. a hemi-sternotomy was performed on the patient in order to get proximal control of the vessels to reduce bleeding risk. the pump was successfully weaned and explanted, and the patient survived the event. after returning from the cath lab the patient developed new neurological changes and a cerebrovascular accident. heparin was withheld. | [5 empty] | A4 is unknown. the pump and purge cassette were returned for investigation. an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2439 | 1220648-2025-49508 | 23898166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898166&pc=OZD | 12/12/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 12/12/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026737767 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 75 year old male patient on impella cp for acute myocardial infarction. it was reported that during support, the patient developed hematuria and had less than 30 cc/hr. the patient was eventually escalated to an impella 5.5 and the cp was discarded by operating room staff. the hematuria resolved. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: (hematuria): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2440 | 1220648-2025-49510 | 23898258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898258&pc=OZD | 12/11/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 12/12/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026751588 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp on a 63 year old male for support due to acute myocardial infarction. it was reported that during support, there was a placement signal issue with placement signal showing a pressure of 206/70. arterial pressure of patient was 130/70. it is unknown how the issue resolved, however, the procedure was completed successfully. there was no patient harm. | [5 empty] | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2441 | 1220648-2025-49512 | 23898316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898316&pc=OZD | 12/13/2025 |  |  | 12/26/2025 |  |  | 12/13/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The operating room team was attempting an escalation case in the cardiovascular operating room. impella 5.5 was inserted and on start up got several alarms including "impella stopped motor current high" "impella stopped controller failure" purge disc not connected" impella stopped retrograde flow" team decided to switch to another controller and pump. they proceed with the case and there was no patient harm. this report is for the controller and another report will be sent for the pump (serial number (b)(6)). | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: consoles (b)(4) and (b)(4) were not returned for investigation but will not be flagged (see below). data analysis: (b)(4) was the first console used in this case, running cp pump 624028 starting on 12/11. the clinical decision was made to switch to a 5.5 pump. ic12070 was then powered up and 5.5 pump 640949 was attempted to be run but had an immediate pump stop due to "high motor current". there were cascading alarms of "retrograde flow" and "pump speed deviation" that were symptomatic of the initial pump stop due to high motor current. the pump was attempted to be started on this console for the next 15 minutes to no avail. the cp pump had been continuously running throughout this on ic10339 but at 19:42, the cp pump was stopped and unplugged from ic10339 and the 5.5 pump was connected but the motor current high alarms persisted on this console as well. ic4039 ran a new 5.5 pump 5917792 on this patient which performed as expected. because the 5.5 pump failed on both consoles in the exact same manner, it is unlikely that either console was at fault for the pump stops and was most likely due to the pump. additionally, the complaint mentioned "purge disc not detected" alarms but there were no unexpected occurrences of these alarms that were not immediately resolved by plugging in the purge disc.
2442 | 1220648-2025-49513 | 23898429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898429&pc=OZD | 11/21/2025 | 12/19/2025 |  | 12/26/2025 |  |  | 11/20/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765779 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 113 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported bleeding from groin site. the patient partial thromboplastin time (ptt) was 150 and a redraw was sent. the patient required 1 unit of blood overnight. the patient survived the procedure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the injuries, hematuria and access site bleeding, could not be determined due to insufficient clinical details.
2443 | 1220648-2025-49511 | 23898474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898474&pc=OZD | 12/12/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 12/12/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026787530 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported on a male patient, who had an impella cp removed bedside using 14fr manta. it was reported that after removal, the femoral pulse was dopplerable, but distal pulses were not present. the patient was taken to the or for vascular exploration and an angioseal was embolized down the leg, so impella and manta ruled out. the patient survived the procedure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. post-procedural adverse event (ischemia): the returned product was investigated. dried blood was found on the sterile sleeve and suture hub. no sterile sleeve damage was seen. the catheter was damaged at the 45 cm marking, near the suture hub. no cannula kink or damage was found. the guidewire was free of kinks and no damage was seen. clinical details mention that the patient's distal pulses were not present after impella removal. after the patient was taken to the operating room for vascular exploration, it was found that the angioseal had embolized down the leg. in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: device passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2444 | 1220648-2025-49515 | 23898559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898559&pc=OZD | 11/07/2025 |  |  | 12/26/2025 |  |  | 11/07/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702651 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 113 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced malposition of the device in the mitral structure/papillary muscle. repositioning of the pump was unsuccessful. the pump was explanted and replaced with an impella cp to support the patient. this is one of two related reports. this report represents that first impella cp. another report was submitted to represent the second impella cp. refer to section h10 for the related report number. | [5 empty] | No product was returned for investigation. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2445 | 1220648-2025-49514 | 23898580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898580&pc=OZD | 12/11/2025 |  |  | 12/26/2025 |  |  | 12/11/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026780845 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 115 KG | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 via the right axillary artery for mechanical circulatory support. it was reported on support day 2, the pump had lower than expected flows for the set performance (p) level. there were no reported alarms or suction issues. the physician repositioned the device, which did not resolve the issue. an echocardiogram was performed and revealed the pump to be well positioned and measured 4.2 cm form the aortic valve. it was noted that p-6, the flows were 3.6 l/min. the decision was made to leave the pump in place. it was noted that the patient¿s blood pressure was within the physician¿s expectations and there were no impella alarms. a medical plan was made to wean the pump the following day. it was noted that the patient expired the following day while on impella support. additional clinical details, including perioperative course, contributing factors, and cause of death, were unavailable at the time of reporting. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. low or blocked pump flow: the cause of low or blocked pump flow was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2446 | 1220648-2025-49517 | 23898595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898595&pc=OZD | 09/21/2025 |  |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703560 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Inflammation | Pain | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient's right hand swelled limiting hand movement, it was reported that the right arm did start to improve. additionally, a vascular ultrasound was conducted and a small hematoma was identified, the patient was stable, and the physician decided that no intervention was required at that time. no further information is available. | [5 empty] | Investigation summary: no product was returned for investigation. hematoma: the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. inflammation/ pain: the cause of the clinical symptom was not determined due to insufficient clinical details. thrombosis: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. device history lot: device lot: 1939842. device history batch: subcomponent lot: n/a. device history review ==> device with sn: (b)(6) passed all post sterile inspection checks.
2447 | 1220648-2025-49519 | 23898606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898606&pc=OZD | 09/20/2025 |  |  | 12/26/2025 |  |  | 09/20/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666841 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Death | Yes | No | Restricted Flow rate | [6 empty] | Heart Failure/Congestive Heart Failure | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the device exhibited suction alarms and the device was pulled back 2 centimeters. it was also reported that the nurse accidently pulled out the device and it was replaced with a new pump. no further information was provided; however, it was later reported that care was withdrawn and the patient died. | [5 empty] | Investigation summary: pump suction: clinical reported the patient was in very critical condition and getting continuous suction alarms at p-8. the product was not returned for investigation. data log review confirms suction alarms triggering throughout the case. placement signal was very low throughout the case, ranging from about 40-70 mmhg and dropping to magnitudes as low as 20mmhg. clinical details stated that the patient was in very poor condition and waveforms were suggesting low volume in the lv, but no volume was given. there was also concern for bi-ventricular failure. the cause of the suction was related to patients condition, as the low placement signal indicated that there was low volume, and clinical details confirmed the patients decompensating condition. device history lot: device lot number: 1915032. device history batch: subcomponent lot: n/a. device history review: pump sn (b)(6) passed all post sterile inspection criteria. qn #(b)(4) was opened because labels were not legible or completed. the labels were reworked and all units that passed post-rework inspection were accepted.
2448 | 1220648-2025-49520 | 23898622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898622&pc=OZD | 09/19/2025 |  |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026729275 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient developed a fever. the resolution for this issue was not provided. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary no product was returned for investigation. fever: the cause of the clinical symptom was not determined due to insufficient clinical details. device history lot device lot:1960452. device history batch subcomponent lot: n/a. device history review device with sn: (b)(6) passed all post sterile inspection checks.
2449 | 1220648-2025-49522 | 23898639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898639&pc=OZD | 09/19/2025 |  |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026690486 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The patient on bipella support with impella 5.5 and rp flex. the following day, urine output was pink-red and the cause for hemolysis was being evaluated, on heparin, levophed, and vasopressin. the impella rp flex as explanted on day 4, with successful site closure with a fresh mattress suture and manual pressure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2450 | 1220648-2025-49516 | 23898720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898720&pc=OZD | 09/24/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/24/2025 | 12/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714584 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Inaccurate Flow Rate | Increase in Pressure | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | Priming Problem |  |  | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | Hematuria | Mitral Valve Insufficiency/ Regurgitation | [33 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, an air in purge system alarm and high purge pressure alarms were noted. multiple attempts to de-air the system, change purge tubing and cassette, and implement tissue plasminogen activator purge per protocol were made. despite these interventions, the purge system remained blocked, flow saturations were noted, and the pump ultimately failed. ultimately, the impella cp was explanted and replaced. it was reported that the patient developed ventricular tachycardia while on impella support. the family elected to withdraw support. the pump was pulled across the valve and care was withdrawn. the patient expired shortly after. | [5 empty] | Investigation summary ==> ischemia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria/minor bleed/mitral valve incompetence: the cause of the injury was not determined due to insufficient clinical details. access site adverse event: clinical reported some oozing present at right femoral impella site, angle matched. the cause of this issue was determined to be use related as evidenced by report of improper angle matching. the data logs confirm an air in purge alarm which triggered for over an hour during the time of patient transfer. the alarm, when triggered stops the purger as no fluid is being detected to push through to the pump. during this period, pump flows become increasingly saturated without p-level change. the logs show that the console screen displayed an unclamp line command indicating the bag line was clamped for over an hour for the duration of the air in purge alarm. the pump was de-aired, after unclamping the line, and on resolution of the air in purge alarm, the purge pressure rises over about 5 minutes to trigger high purge pressure and purge flow is restored with very low values. no motor current spike was observed outside p-level changes. high purge pressure is sustained as well as saturated flow until the pump stops with motor current high alarms. restart attempts were unsuccessful. air in purge system: the cause of the air in purge alarm was determined to be due to an use related error as logs indicated line was clamped triggering alarm. purge pressure high/purge system blocked: the cause of the high purge pressure issue was determined to be as a result of blood ingress into purge gap as evidenced by log analysis showing purge flow stopped. high or saturated pump flows: the cause of the saturated flows was determined to be as a result of blood ingress into purge gap as evidenced by log analysis showing flow rise during purge flow stop. pump stop: the cause of the pump stop was determined to be as a result of blood ingress into purge gap when there was lack of purge flow. device history lot: device lot number: 1924812. device history batch: subcomponent lot: n/a. device history review: qn 210042132 was for a containment action for winding on insertion pin. this was unrelated to the reported failure mode. | Medical safety/clinical reviewed the event. the patient presented to the emergency department on (b)(6) 2025 in an unresponsive state, requiring immediate intubation. no advance cardiopulmonary life support was required. electrocardiogram demonstrated st-elevation myocardial infarction (stemi). the patient was emergently taken to the cardiac catheterization laboratory. coronary angiography revealed calcified ostial left main (lm) disease, chronic total occlusion of the left anterior descending, and ostial right coronary artery disease. due to hemodynamic instability and critical lm disease, the physician elected to place an impella cp (pump 1) for mechanical circulatory support. after approximately three hours of intervention, one stent was placed in the left main artery. the patient remained on high-dose vasoactive medications and impella cp support at p-9. given persistent cardiogenic shock, the care team planned escalation to veno-arterial extracorporeal membrane oxygenation (va-ecmo) with impella support and transfer to a higher level of care. a prior echocardiogram performed approximately one month before the stemi showed an ejection fraction of 25%. later the same day, high purge pressure was noted and the purge system was found to be blocked. upon transfer to the receiving facility, an ¿air in purge¿ alarm was present. de-airing the purge system and exchanging the purge tubing and cassette did not resolve the alarm. tpa purge therapy was initiated per protocol, which is noted as off label use. the patient developed bilateral lower-extremity pulse deficits, and urine output became peach-pink to red in appearance. bedside echocardiography demonstrated the impella cp was positioned slightly deep at approximately 4.4 cm. the device was repositioned to a more optimal location. oozing was noted at the right femoral impella access site. the impella angle was corrected, and the device was initially weaned with plans for explant; however, echocardiography demonstrated worsening mitral regurgitation during weaning, and the decision was made to leave the pump in place despite recognized risks. the patient was extubated, vasopressors were discontinued, and afterload reduction was initiated. the access site was clean, dry, and intact, and left lower-extremity pulses were dopplerable. high purge pressure persisted, and tpa purge therapy was continued for two additional days. the patient subsequently developed frequent suction alarms with declining pressures, followed by pump stoppage. the impella cp (pump 1) was explanted due to pump stop in the setting of unresolved high purge pressure and replaced with a second impella cp device. the impella cp (pump 2) was implanted on 9/26/2025. following implantation, the patient developed ventricular tachycardia. after discussion with the care team, the family elected to withdraw care. the device was withdrawn across the aortic valve, care was withdrawn, and the patient expired shortly thereafter. impella cp pump 1 is a serious type of reportable event and impella cp pump 2 is a death. note: h6: investigation type/findings/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
2451 | 1220648-2025-49524 | 23898751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898751&pc=OZD | 11/16/2025 | 12/16/2025 |  | 12/26/2025 |  |  | 11/16/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026754168 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 61 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 72-year-old male was implanted with a impella 5.5 for support during a high-risk cardiac procedure. it was reported that cangrelor drops were started because there was an increase in bleeding from the orogastric tube.gastroenterologist performed esophagogastroduodenoscopy hemorrhagic gastritis and arteriovenous malformation treated. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2452 | 1220648-2025-49505 | 23898761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898761&pc=OZD | 09/14/2025 | 01/15/2026 |  | 12/26/2025 |  |  | 09/14/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634553 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 90 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Hematuria | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. the patient is in stable condition. | The complainant reported (upon information request) that the product is not available for return. |  |  |  |  | A6 is unknown. no product was returned for investigation. the cause of the hematuria was not determined due to insufficient clinical details. the root cause of the infection was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | B5 updated.
2453 | 1220648-2025-49287 | 23898783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898783&pc=OZD | 09/22/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/22/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722326 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 99 KG |  | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Pulmonary Embolism | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support following escalation from an impella cp and extracorporeal membrane oxygenation (ecmo) decannulation. during support, the patient remained off heparin due to an upper gastrointestinal bleed, which was subsequently resolved by the care team. the patient was later extubated, with concerns for aspiration pneumonia and leukocytosis; cultures were obtained, and antibiotic therapy was initiated. the patient was also found to have a deep vein thrombosis (dvt), with plans to initiate low-dose heparin therapy. the patient¿s condition stabilized, and the device was subsequently weaned and explanted. | [5 empty] | Investigation summary: data review: data logs showed pump ran without issue during support. there were suction alarms triggered in early of the case but resolved quickly. there were no mc spikes, abnormal ps or purge issues observed during support. product was not returned for review. pump suction: the cause of pump suction was most likely patient condition related since the suction was resolved by fluid compensation. infection/thrombosis: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pulmonary embolism/hemostasis (gi bleed): the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot #: 1958860. device history batch: subcomponent lot#: n/a. device history review: pump sn: (b)(6) passed all post sterile inspection checks.
2454 | 1220648-2025-49518 | 23898801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898801&pc=OZD | 12/11/2025 |  |  | 12/26/2025 |  |  | 12/11/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649569 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 67 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp via the right femoral artery for mechanical circulatory support. on support day one, it was reported that the sensor for the placement signal was not correlated with the aortic pressure waveform. the following morning the issue was reported to have self-resolved. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the clinical details state that the placement signal was not correlated to the a-line by 100mmhg. due to a lack of product or data logs returned, the root cause cannot be established. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2455 | 1220648-2025-49521 | 23898840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898840&pc=OZD | 09/19/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717822 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support the patient experienced hemolysis and consistent bleeding from access sites, heparin was withheld. no further information was provided. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary the pump was not returned for investigation. data log did not show any conclusive hemolysis-related abnormalities in suction, motor current spike, or major placement signal trend during the time of the reported hemolysis issue. mechanical interaction with blood (hemolysis): the cause of the hemolysis issue was not determined without returned product or sufficient clinical information. the patient was on bipella support during the reported hemolysis issue; please refer to (b)(6) for investigation information on the rp flex pump. access site adverse event (hemorrhage/bleeding): the cause of the bleeding issue was not determined without returned product and sufficient clinical information. device history lot device batch: 1952975. device history batch subcomponent lot: na. device history review the pump sn(b)(6) passed all post sterile inspection checks.
2456 | 1220648-2025-49523 | 23898847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23898847&pc=PYX | 09/19/2025 | 12/26/2025 |  | 12/26/2025 |  |  | 09/19/2025 | 12/26/2025 |  |  |  |  | 12/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2026690486 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp flex for mechanical circulatory support. during support the patient had hemolysis. no further information was provided. the pump was successfully weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary ==> mechanical interaction with blood (hemolysis): the pump was not returned for investigation. data log shows a motor current spike 20 minutes after pump start; however, pump performance was not impacted after the spike. there was no positioning alarm, suction alarm, or major placement signal trend reported during the case run. the cause of the hemolysis issue was most likely biomaterial ingestion, based on data log review. device history lot ==> device batch: 1916496. device history batch ==> subcomponent lot: na. device history review ==> the pump (b)(6) passed all post sterile inspection checks.
2457 | 1220648-2025-49378 | 23894256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894256&pc=OZD | 09/11/2025 |  |  | 12/25/2025 |  |  | 09/11/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026720880 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Restricted Flow rate | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. it was reported that the patient had hematuria. the impella was repositioned overnight. the following morning, the impella appeared to be shallow and too close to the mitral valve. the patient was successfully weaned from the pump, and the device was explanted. | [5 empty] | The device was not returned for investigation. a review of the data log confirmed the complaint. the cause of the hematuria was not determined due to insufficient clinical details. the cause of the suction was most likely patient condition, as clinical details and the trend in placement signal indicate that the patient was having volume issues. the device passed all post sterile inspection criteria.
2458 | 1220648-2025-49379 | 23894261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894261&pc=OZD | 09/11/2025 |  |  | 12/25/2025 |  |  | 09/14/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699285 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella device for mechanical circulatory support. during implant there were reduced flow alarms. the device was explanted in the procedural area. there was no reported harm or injury to the patient. | [5 empty] | A4, a5, and a6 are unknown. the cause of the optical signal issue was not determined due to no defect found on the returned pump, no data logs were recovered, and inconclusive clinical details. the device passed all post sterile inspection checks.
2459 | 1220648-2025-49380 | 23894263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894263&pc=OZD | 09/14/2025 |  |  | 12/25/2025 |  |  | 09/14/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1311804 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a controller alarm occurred on an automated impella controller. the controller was replaced with another to continue patient support. the patient expired. | [5 empty] | A3 and a4 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2460 | 1220648-2025-49381 | 23894264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894264&pc=OZD | 08/12/2025 |  |  | 12/25/2025 |  |  | 09/15/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026671831 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male | 77 KG | Non Hispanic | Asian | Injury | Yes | No | Restricted Flow rate | [6 empty] | Cardiac Arrest | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced pericardial effusion after a percutaneous coronary intervention was performed. a drain was placed, and suction alarms occurred. the correct pump position was confirmed. the patient experienced pulseless electrical activity and expired. | [5 empty] | The pump was not returned for investigation. data logs were unavailable for review. the cause of the pericardial effusion and cardiac arrest was not determined due to insufficient clinical details. the cause of the suction issue was not determined without returned product investigation and sufficient clinical details. the device passed all post sterile inspection checks.
2461 | 1220648-2025-49384 | 23894268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894268&pc=OZD | 09/18/2025 | 11/21/2025 |  | 12/25/2025 |  |  | 09/18/2025 | 12/25/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657600 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. low leakage of the purge sidearm was reported. purge system leaking was reported again, with air detection behind the purge cassette. high purge pressure has also been reported, with tissue plasminogen activator lysis as a successful remedy. no patient harm has been reported. | Clinical review/rationale: an impella 5.5 was inserted via left axillary artery in a 33-year-old female who was admitted for acute myocardial infarction with cardiogenic shock. the patient¿s comorbidities were unknown. the patient¿s clinical state prior to initiation of support was scai stage e. the patient was in the icu on support with the 5.5. on day 48 of support, it was reported there was a low leakage of the purge side arm. the pump continued to run with no changes to the values. it was also noted that there was fluid inside of the purge unit and dried glucose on the outer shell of the purge unit, both in low quantities. on day 49 of support, the amount of fluid inside the purge unit shell was unchanged, and the outside was clean. the luer connector was properly seated and a leak could not be verified. the pump was running with no alarms and all values within range. on day 71 of support, a leakage was reported and air detected behind the purge cassette, in which the cassette was replaced. on day 96 of support, purge pressure was reported high at >1000mmhg and purge fluid <1ml. the pump was running without issues or alarms. care was withdrawn, the device explanted, and the patient expired. the purge leak will be conservatively reported for hemodynamic instability due to temporary hemodynamic support interruption during the exchange; however, the exchange was performed with no consequence to the patient and support. death is a known potential adverse event of the impella 5.5 per the instructions for use. |  |  |  |  | Correction: h6 and h11 investigation results were inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. purge leak - pump: the cause of the purge leak was related to the sidearm but the location on the sidearm was unable to be determined since product was not returned for review. high purge pressure: the cause of the high purge pressure was unable to be determined since product/data were not returned for review. device history review: the device passed all post sterile inspection checks. | A4 is unknown. the device was not returned for investigation. data logs were not provided for review. the cause of the purge leak was related to the sidearm but the location on the sidearm was unable to be determined since product was not returned for review. the cause of the high purge pressure was unable to be determined since product/data were not returned for review. the device passed all post sterile inspection checks. | B5 narrative was added and malfunction was changed to death. h6 codes were updated.
2462 | 1220648-2025-49386 | 23894269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894269&pc=OZD | 09/19/2025 |  |  | 12/25/2025 |  |  | 09/19/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026679406 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with a left ventricular thrombus was implanted with an impella cp for mechanical circulatory support. a suction alarm occurred immediately after the patient coughed. flow could no longer be achieved. adjusting the pump position did not improve the situation, and it was thought that there was some kind of interference with the left ventricular thrombus, so the impella was removed. support was provided with extra corporeal membrane oxygenation (ecmo) and an intra-aortic balloon pump. the doctor considered surgical thrombus removal and impella 5.5 insertion in preparation for removal of the ecmo system. the patient survived the event. | [5 empty] | A4 is unknown. the root cause of the suction was most likely related to the patients condition as the clinical details stated that during the time of the reported event the patient was coughing, triggering the alarms. the device passed all post sterile inspection criteria. | Investigation summary: the product was received, and the investigation was completed. device history record review was completed and noted all post sterile inspection criteria passed. regarding the pump suction, clinical reported that suction triggered after the patient coughed. the patient had left ventricular thrombus prior to implant. the product was returned and there was no biomaterial observed in the pump or kinks on the cannula. data log review shows suction occurring on 9/17/2025 and on 9/19/2025 when it was reported. on 9/17, placement signal was slightly trending down leading up the suction events. the downward trend stopped soon after on 9/18/2025, and suction conditions resolved. at the time suction was reported on 9/19/2025 there were suction #73 alarms indicating the pump was in severe suction. during this time, the placement signal was spiking, which was likely due to the patient coughing, as reported in clinical details. an rt log was observed during one of the reported suction events and motor current max values dropped coinciding with pump flows and the suction alarms were accurately triggered. the root cause of the suction was most likely related to the patient's condition as the clinical details stated that during the time of the reported event the patient was coughing, triggering the alarms. correction summary: after review of the case, it was reassessed as a malfunction type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact and medical device problem codes) were repopulated for update, correcting accordingly. additionally, investigation summary was corrected to provide complete details available. consequently, h6 investigation: type, findings and conclusion codes and h6 component code were updated, repopulated accordingly based on the completed investigation.
2463 | 1220648-2025-49385 | 23894270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894270&pc=OZD | 09/19/2025 |  |  | 12/25/2025 |  |  | 09/19/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026709174 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | No | Degraded | [6 empty] | Embolism/Embolus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient showed a possible metallic foci status post impella cp removal. glasgow coma scale was 14-15, which has resolved. no patient harm has been reported. | [5 empty] | A4 and a5 are unknown. the data logs were reviewed and there were no issues seen during the case run pertaining to the motor current or pump flows. the pump was not returned. the root cause of the pump damage cannot be determined. the cause of the embolism was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2464 | 1220648-2025-49387 | 23894272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894272&pc=OZD | 09/23/2025 |  |  | 12/25/2025 |  |  | 09/24/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696131 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Female |  |  |  | Injury | Yes | No | Inaccurate Flow Rate | Device Sensing Problem | Difficult to Advance | Patient Device Interaction Problem |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that after a difficult implantation of an impella 5.5, motor current and optical signal issues occurred. additionally, the patient exhibited hemolysis. the position of the pump was confirmed to be correct. the pump was explanted. | [5 empty] | A4, a5, and a6 are unknown. the impella 5.5 pump was returned for investigation and a cannula kink was noted as well as biomaterial in the cannula around the impeller. the pump passed laser continuity testing and had expected optical 5s parameter values. the pump was connected to a console and purged overnight, where purge pressure was persistently high. the pump was able to be started, and flows were saturated at p-9 during the run. data logs were reviewed and the reports were confirmed. the cause of the saturated flow was determined to be biomaterial ingestion. the cause of the optical sensor issue was determined to be biomaterial ingestion. the cause of the hemolysis was determined to be biomaterial ingestion. the cause of the difficulty to deliver the pump could not be determined. the device passed all post-sterile inspection checks.
2465 | 1220648-2025-49388 | 23894276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894276&pc=OZD | 09/30/2025 |  |  | 12/25/2025 |  |  | 09/30/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1308458 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an automated impella controller experienced "impella stopped", "motor current high", and "contoller failure" alarms. the pump could not be restarted. the pump was explanted and no patient was in stable condition. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2466 | 1220648-2025-49389 | 23894279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894279&pc=OZD | 09/23/2025 |  |  | 12/25/2025 |  |  | 09/30/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604895 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding and a hematoma. the pump was explanted to achieve hemostasis. | [5 empty] | Patient information was not provided. section a was left blank. the cause of the injury was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks.
2467 | 1220648-2025-49390 | 23894280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894280&pc=OZD | 09/30/2025 |  |  | 12/25/2025 |  |  | 09/30/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026709093 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced an "impella in aorta" alarm. the pump could not be repositioned. the pump was explanted and the patient was in stable condition. | [5 empty] | Patient information was not provided. section a was left blank. no product was returned for investigation. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2468 | 1220648-2025-49391 | 23894281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894281&pc=OZD | 09/30/2025 |  |  | 12/25/2025 |  |  | 09/30/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690375 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced "impella stopped", "motor current high" and "contoller failure" alarms. the pump could not be restarted. no patient harm was reported. | [5 empty] | The data logs confirm impella stopped alarms and show that the first alarm trigger was an alarm #115 with a motor current spike top 30000ma. this pattern is characteristic of an electrical open failure from an open circuit. the logs show alarm #13 motor current high alarm on restart attempt. as no product was returned and limited clinical information was provided, the cause of the electrical open issue was unable to be established. the pump sn (b)(6) passed all post sterile inspection checks. | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2469 | 1220648-2025-49394 | 23894285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894285&pc=OZD | 10/02/2025 | 01/07/2026 |  | 12/25/2025 |  |  | 10/02/2025 | 12/25/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1308461 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Failure to Read Input Signal | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that an automated impella controller generated an error in the purge system. no patient harm was reported. | [5 empty] | Patient data was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The reported failure mode has been confirmed. the root cause has not been determined and no issues have been found during the investigation so far. the investigation is still ongoing.
2470 | 1220648-2025-49393 | 23894287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894287&pc=OZD | 09/27/2025 |  |  | 12/25/2025 |  |  | 10/01/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645608 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that during implantation of an impella cp the cannula kinked and a suction alarm occurred. positioning adjustments were made but the issue did not resolve. due to reduced flow, the device was explanted. no patient harm was reported. | [5 empty] | A4 is unknown. a cannula kink was observed in the returned product. the cause of the positioning issue was determined most likely to be an adverse event related to the anatomy of the patient due to heart being in a different position than usual due to anatomical reasons, and the aorta at a steep angle. the cause of pump suction/ low pump flow was most likely due to the kink in the cannula. the cause of product damage was most likely patient condition related since the patient's heart appeared to be lying flat and was in a different position than usual due to anatomical reasons, and the aorta was at a steep angle. the device passed all post sterile inspection checks.
2471 | 1220648-2025-49395 | 23894289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894289&pc=OZD | 10/05/2025 |  |  | 12/25/2025 |  |  | 10/06/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026706927 |  | 000000000001000482 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 25 YR |  | Male |  |  | White | Injury | Yes | Yes | Inaccurate Flow Rate | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Ci created for naveen nagaraj: the complainant reported that a patient with heart failure and cardiogenic shock (cgs) underwent implantation of an impella 5.5 device for mechanical circulatory support. during support, the pump was functioning without issues, but a saturated flow was observed on the automated impella controller (aic). the field team was notified for follow-up, and the physician decided to replace the pump later that afternoon. the patient also had pericarditis and endocarditis, which resulted in aortic valve perforation and associated aortic insufficiency. the patient subsequently underwent aortic valve replacement and received a new impella 5.5 device. | [5 empty] | The investigation into the reported issue has been completed. the cause of the injury was not determined due to insufficient clinical details. high or saturated pump flows: clinical reported saturated flows at 5.9l/min. the logs show a gradual motor current rise through support and a small motor current shift of about a 100ma increase on 6-oct. mean flows go from about 5.5l/min to 5.8/5.9l/min. the purge pressure and purge values also shift at that instant but trends are not sustained and soon after purge pressure values decrease and purge flows increase. the pattern suggests biomaterial interaction as evidenced by the increase in flow with an effect on motor current and purge values. placement signal values were also trending down slightly the cause of the issue was determined to be biomaterial ingestion as evidenced by log analysis the pump passed all post sterile inspection checks.
2472 | 1220648-2025-49396 | 23894291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23894291&pc=OZD | 10/08/2025 |  |  | 12/25/2025 |  |  | 10/08/2025 | 12/25/2025 |  |  |  |  | 12/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026630921 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 80 YR |  | Female | 75 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp for mechanical circulatory support (mcs) and expired with three hours on admissions to the unit. the impella cp device was placed after cardiopulmonary resuscitation was initiated, which continued through implant of the impella and post percutaneous coronary intervention (pci) of the left main, left anterior descending arteries. the patient was transferred to the unit after pci and expired within three hours. per the report, the death is unrelated to the impella device. the patient was in critical condition and impella was placed during resuscitation. | [5 empty] | Patient-specific information a5. ethnicity is unknown. additionally, the exact date of death is unknown but has been captured as the date of explant. medical safety review. medical safety review was completed and noted the following: upon prepping an impella cp device, there was noted to be a leak in the purge cassette with an air in line alarm. the cassette was replaced and set up and insertion continued without issue. impella cp was placed after cardiopulmonary resuscitation was initialed and resuscitation continued throughout insertion. post intervention, the patient went to the intensive care unit (icu) where they expired 3 hours later. due to the leak in the purge cassette, this is a device malfunction on the purge cassette. because the impella cp was in place when the patient passed away, it is being conservatively reported however is very unlikely to have contributed to the patient outcome and the death is likely the result of the patient's clinical presentation. device status. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. h6 codes were updated. e4 corrected information as it was incorrect from the initial report that was submitted. g1 corrected manufacturer contact title as it was incorrect from the initial report that was submitted.
2473 | 1220648-2025-49265 | 23887670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23887670&pc=OZD | 11/22/2025 | 12/24/2025 |  | 12/24/2025 |  |  | 11/25/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530369 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after weaning iand removing impella cp, thrombus occured and patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: ppae (thrombosis): in order to make a cause determination,the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
2474 | 1220648-2025-49267 | 23887702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23887702&pc=OZD | 09/29/2025 | 12/24/2025 |  | 12/24/2025 |  |  | 11/25/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1713329 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that once impella was placed and started ps became highly erratic with massive ps spikes. motor current appropriate, although large differential between systolic/diastolic macrophage activation syndrome. the flows were appropriate for p level once position optimized. it was discussed and encouraged for the team to use patient arterial line for hemodynamic monitoring and watch motor current and flows. the white cable was discovered that it was not fully seated into aic which once staff pushed in, placement signal unreliable alarm resolved. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2475 | 1220648-2025-49222 | 23888278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888278&pc=OZD | 12/03/2025 | 12/23/2025 |  | 12/24/2025 |  |  | 12/09/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026782190 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a thrombus in the three "branches" of the arch which was treated with perfusion, and a ct showed a cerebral infarction. the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety/clinical reviewed the event. the patient was escalated from impella cp to impella 5.5 support. at the time of impella 5.5 insertion, systolic arterial blood pressure (sabp) in the right hand decreased to 30 mmhg. this was initially attributed to intraoperative manipulation. however, upon return to the intensive care unit (icu), sabp in the left hand also decreased to approximately 30 mmhg. aortic (ao) waveforms were obtained with a mean arterial pressure (map) of approximately 80 mmhg. a subsequent computed tomography (ct) scan confirmed thrombotic occlusion in three branches of the aortic arch. sheaths were inserted into both internal arteries, and perfusion was initiated via the extracorporeal membrane oxygenation (ecmo) circuit. a head ct scan demonstrated significant cerebral edema secondary to ischemic cerebral infarction. the patient subsequently expired. physician comments indicated that ultrasound imaging performed prior to impella 5.5 insertion identified a giant thrombus in the left atrial appendage. the physician noted that the patient¿s survival would have been unlikely without impella support and indicated no anticipated impact. additional information indicates it is unclear whether a ct scan was performed at the time of escalation from impella cp to impella 5.5. an arterial line (a-line) was placed in the right radial artery. the cerebral infarction was ischemic in nature. as there was no waveform dissociation between the arterial line placed in the right temporal artery and the device position waveform, it is considered unlikely that the cerebral infarction occurred during impella cp support. based on the available information, the timing of the cerebral infarction is unknown but is considered unlikely to be associated with impella cp support per the report. the impella cp will be reported as a serious injury type of reportable event, while the impella 5.5 will be reported as a death type of reportable event. according to the physician, the patient¿s survival would have been unlikely without impella support. correction. complaint/patient coding was updated/corrected following medical safety/clinical review as follows: added embolism (e0503). therefore, section h6 (health effect ¿ clinical code, health effect ¿ impact code, and medical device problem code) has been fully repopulated and should be regarded as the final coding that accurately reflects the reported event. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The investigation of the reported failure mode has been completed. no product was received for evaluation. stroke/thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2476 | 1220648-2025-49268 | 23888294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888294&pc=OZD | 11/26/2025 | 12/23/2025 |  | 12/24/2025 |  |  | 11/26/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633441 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the procedure was aborted because of the vascular anatomy. the pump could not be advanced because the femoral artery was heavily calcified with multiple stenoses. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product was returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult vascular anatomy and severe stenosis preventing successful delivery of impella. device history review: device passed all post sterile inspection checks. | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2477 | 1220648-2025-49270 | 23888308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888308&pc=OZD | 09/22/2025 |  |  | 12/24/2025 |  |  | 09/22/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720875 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. during support, the device generated an audio alarm. the nurse disabled the audio overnight, as no further repositioning or echocardiogram (echo) evaluation was planned at that time. the pump was subsequently weaned and explanted, and the patient survived the event. | [5 empty] | No product returned. optical sensor issue : after 1 day on support, ps issue occurred. confirmed position via echo. data log review showed ps low alarms. no ps offset applied. ps monitoring disabled shortly after ps low alarms show as ceasing. there was small negative offset (-2.5 mmhg) initially applied for 9 hours and then offset of ~ -15 mmhg after 20 plus hours later. the ps low was observed only with second console after transfer. ps gradually improves throughout case. no other issue noted with the placement signal / 5s parameters besides low ps values and the ps trends up and improves by the end of the case. the ps low alarms did not resolved, but the ps monitoring was disabled. the cause of the optical signal issue was not determined with inconclusive log and clinical evidence per analysis.
2478 | 1220648-2025-49272 | 23888314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888314&pc=OZD | 09/22/2025 |  |  | 12/24/2025 |  |  | 09/22/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714078 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during placement in the cardiac catheterization laboratory, the device was improperly positioned under echocardiographic guidance. the physician attempted to reposition the impella, but the device was pulled back too far and could not be advanced across the aortic valve. the decision was made to place a new impella device. the patient was successfully weaned from impella support, and the original device was explanted. | [5 empty] | Investigation summary ==> no product was returned. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was improperly positioned via echo while in ccl. device history lot ==> device lot: 1955272 device history batch ==> subcomponent lot: . device history review ==> device (sn:(b)(6) passed all post sterile inspection checks.
2479 | 1220648-2025-49273 | 23888318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888318&pc=OZD | 09/22/2025 |  |  | 12/24/2025 |  |  | 09/22/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722079 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. on the morning of the procedure, the patient experienced runs of ventricular tachycardia. no echocardiogram had been performed at the time. the impella cp was inserted via the left femoral artery (lfa) without complication. the patient was successfully weaned from impella support, and the device was explanted without issues. | [5 empty] | Investigation summary ventricular tachycardia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot device lot : 1972075 device history batch sub-component lot : n/a device history review the pump s/n: (b)(6) passed all the post sterile inspection checks.
2480 | 1220648-2025-49275 | 23888334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888334&pc=OZD | 09/22/2025 |  |  | 12/24/2025 |  |  | 09/22/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729604 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing high-risk percutaneous coronary intervention (hrpci). vascular access was obtained, and a 13 fr peel-away sheath was inserted without complication. during support, the optical sensor malfunctioned. the physician was notified and elected to proceed with the hrpci while monitoring motor current, flows, vital signs, and fluoroscopy. the hrpci procedure was successfully completed. the pump was eventually weaned and explanted, and the patient survived the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2481 | 1220648-2025-49276 | 23888338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888338&pc=OZD | 09/22/2025 |  |  | 12/24/2025 |  |  | 09/22/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688177 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ventricular Fibrillation | [11 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing high-risk percutaneous coronary intervention. the impella cp was successfully positioned across the aortic valve, the wire was removed, and mechanical support was initiated. the patient was defibrillated once and return of spontaneous circulation was achieved. the impella cp was run at p-9 for the duration of the procedure to maintain hemodynamic stability. impella support was weaned after completion of interventions, and the device was explanted. | [5 empty] | Investigation summary: no product was returned for investigation. ventricular fibrillation/ cardiac arrest: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1934164. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
2482 | 1220648-2025-49280 | 23888354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888354&pc=OZD | 09/22/2025 |  |  | 12/24/2025 |  |  | 09/22/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501660 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 was attempted to be implanted for mechanical circulatory support. under fluoroscopic guidance, a wire exchange was performed, and the impella 5.5 catheter was advanced through the introducer but encountered resistance while passing through the anastomosis. manual manipulation allowed advancement across the aortic valve; however, transesophageal echocardiography demonstrated suboptimal positioning, with the inflow located within the mitral chordae. multiple repositioning attempts were made, but the inflow remained trapped. the procedure was aborted, and the decision was made to replace the device with a new impella 5.5. | [5 empty] | Investigation summary: no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. product damage (catheter kink): the cause of product damage was most likely use issue related since excessive force was used per clinical. device history lot: device lot: 1813984. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
2483 | 1220648-2025-49281 | 23888356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888356&pc=PYX | 09/22/2025 |  |  | 12/24/2025 |  |  | 09/22/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2026725894 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. following removal of the temporary pacing wire and placement of a swan-ganz catheter, an increase in pulmonary artery systolic pressure was observed, accompanied by a 1-liter decrease in impella flow at performance level p6, raising concern for possible thrombus ingestion. the medical team discussed potential management options and continued close monitoring. a chest x-ray confirmed stable impella rp positioning. device flow subsequently normalized later that evening, and the device appeared to function appropriately. the patient remained on systemic heparin and milrinone infusions. the pump was eventually weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary thrombosis/arrhythmia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. low or blocked pump flow: clinical reported after swan placement via femoral vein and tpm removal, there was increase in ps and drop in pump flow. the data logs confirm clinical events and show 2 motor current spikes about an hour apart with speed deviations and a shift in placement signal. the pump flows also dropped at that instant with no change in p-level. purge flows dropped after first spike and resolved after second spike. low flows were sustained for about 3 hours before issue resolved right after another motor current spike. this is likely as a result of biomaterial being displaced from the impeller. the cause of the issue was determined to be as a result of an ingestion event as evidenced by log pattern device history lot device lot number: 1945905 device history batch subcomponent lot: n/a device history review the pump sn (b)(6) passed all post sterile inspection checks.
2484 | 1220648-2025-49282 | 23888364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888364&pc=OZD | 09/22/2025 |  |  | 12/24/2025 |  |  | 09/22/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688207 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Inaccurate Flow Rate | Patient Device Interaction Problem | [5 empty] | Cardiac Perforation | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with postcardiotomy cardiogenic shock and low cardiac output syndrome underwent implantation of an impella 5.5 device for mechanical circulatory support. following implantation, the impella controller displayed multiple ¿motor current high¿ alarms, and transesophageal echocardiography demonstrated thrombus formation on the impella 5.5 inlet. shortly thereafter, the patient developed a sudden increase in chest tube output and clinical decompensation. the patient was placed on cardiopulmonary bypass, and surgical exploration revealed a small left ventricular perforation, which was surgically repaired. the reported device was explanted and replaced with a new impella 5.5 device. | [5 empty] | High or saturated pump flow: the pump was returned for investigation. data log shows the pump had a sudden spike in motor current followed by saturated pump flow during the case run. placement signal also had a fluctuating trend during this time. motor current and flow returned to normal after 30 minutes. the pump was explanted later. returned pump had biomaterial in the inlet cage, and after it was removed, the pump operated as intended. clinical details also confirmed thrombus at the inlet on tee during the period of saturated pump flow. the cause of the high or saturated pump flow was biomaterial ingestion, based on the given clinical details, data log review, and returned pump investigation. clinical symptom (thrombosis/thrombus): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. mechanical interaction with tissue (cardiac perforation): no related defect was found on the returned product. the cause of the lv perforation issue was not determined without sufficient clinical information.
2485 | 1220648-2025-49283 | 23888372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888372&pc=OZD | 09/22/2025 |  |  | 12/24/2025 |  |  | 09/22/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703253 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support during a high-risk percutaneous coronary intervention. following the procedure, the device was explanted and the site closed with a manta device. after closure, distal pulses were not initially palpable by doppler, though they had been present before explant. contralateral femoral access was obtained, and angiography revealed thrombosis of the common femoral artery. the patient underwent penumbra thrombectomy. light oozing at the access site was managed with proper dressing care. | [5 empty] | Investigation summary no product was returned for investigation. thrombosis/ ischemia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot device lot: 1951045 device history batch subcomponent lot: n/a device history review device with sn: (b)(6) passed all post sterile inspection checks.
2486 | 1220648-2025-49284 | 23888380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888380&pc=OZD | 09/22/2025 |  |  | 12/24/2025 |  |  | 09/22/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739172 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. following insertion, the 14 french peel-away sheath was removed and exchanged for a repositioning sheath. oozing and bleeding were noted at the access site, requiring manual compression for more than 15 minutes. due to continued oozing, a femostop compression device was applied at low pressure for bleeding management until anticoagulation reached therapeutic levels. the patient expired while on impella cp support. the patient¿s subsequent clinical deterioration and death were attributed to underlying cardiogenic shock and critical illness. | [5 empty] | Investigation summary minor bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history lot device lot : 1977138 device history batch sub-component lot: n/a device history review the pump s/n: (b)(6) passed all the post sterile inspection checks.
2487 | 1220648-2025-49288 | 23888390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888390&pc=OZD | 09/22/2025 | 11/20/2025 |  | 12/24/2025 |  |  | 09/22/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723020 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 88 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support due to asystole. while on support, the patient experienced an access site bleed and hematoma that was resolved by removal of the pump. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. impella cp #(b)(4) was returned for investigation. access site adverse event (hematoma/ major bleed): the patient experienced an access site bleed and hematoma that was resolved by removal of the pump. patient expired shortly thereafter. the returned pump was investigated. no sharp edges or product damage was seen. the sterile sleeve was detached from the side of the suture hub, but no damage was seen to the sleeve itself. some dried blood was seen on the suture hub and the catheter but no other abnormalities were seen. the root cause of the bleeding and hematoma were unable to be determined as insufficient clinical details were provided and inconclusive evidence was found from returned product investigation. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2488 | 1220648-2025-49291 | 23888392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888392&pc=OZD | 09/23/2025 |  |  | 12/24/2025 |  |  | 09/23/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724644 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing high-risk percutaneous coronary intervention. while using the shockwave device, the impella placement signal was temporarily lost. despite this, device flows and the patient¿s blood pressure remained stable, and impella support was maintained. the placement signal issue resolved without intervention, and support continued until the end of the case. the patient¿s condition stabilized following the procedure, and the impella cp was successfully explanted. | [5 empty] | Investigation summary ==> impella cp sn (b)(6) returned for evaluation. optical signal issue - while using shockwave in the proximal om the placement signal was lost around the 25-30th pulse. flows remained, bp stable, and support continued until completion. data log review shows ps rails to 3000 after 30 minutes of support, snr drops to zero. the returned pumped passed laser light continuity and showed an snr value of 1, indicating a broken sensor face. the cause of the optical signal issue was shockwave interaction as the clinical noted the issue occurred during shockwave use, and both the data logs and returned product show a broken sensor face. however, the distance between the shockwave and the optical sensor are not known. device history lot: device lot: 1958585. device history review: pump s/n (b)(6) passed all post sterile inspection checks.
2489 | 1220648-2025-49290 | 23888395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888395&pc=OZD | 09/23/2025 |  |  | 12/24/2025 |  |  | 09/23/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719329 |  | 0048-0014 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the local clinical team reported a high purge pressure alarm, and tissue plasminogen activator (tpa) lytic therapy was administered. the hospital site team acknowledged that the impella cp had been in use longer than the recommended duration of five days and did not attribute the event to a device malfunction. the patient¿s condition stabilized, and the impella cp device was subsequently weaned and explanted. | [5 empty] | Investigation summary ==> the product was not returned for investigation. high purge pressure: data logs were reviewed and several occurrences (>1500) of hpp alarms and purge system blocked alarms were found on 9/23 to 9/26. also, 2 instances/occurrences of hpp alarms on each days (9/17, 9/18 and 9/20) which was quick rise of 1 to 1.5 min mostly or sometimes 2 min rise as maximum. the pp resolves quickly by a drop within 1 minute or max 2 mins on few instances. these above events trend likely indicate that there was some temporary kinking in purge line causing quick hpp events. there was no mc spike observed, pump flows were within normal specification per design during hpp alarm instances from 9/23 to 9/25. gradual rise in purge pressure for 3 hrs 23 mins approx with no pump flow saturation or any mc trend, variation or spike. initial rt log showing the start of rise in purge pressure on 9/23 was missing. the next rt log shows a gradual in rise in pp and drop in purge flows with no mc, pump flows impact. the purge flows ticks take a long time ticking at 19, 27, 29 mins long approximately. the next rt log show purge pressure high alarms trigger. purge flow ticks are stalling slowly post the long purge ticking at 30, 34 mins approximately. as per clinical it was reported for high purge pressure alarms on 9/23 to 9/25. it was mentioned tpa addition resolved the purge issues. but based on logs the purge issues never resolve until 9/26 or till end of case. the cause of high purge pressure was not determined due to inconclusive evidence per data logs and also due to device not being returned for investigation. device history lot ==> device lot: 1955139 device history batch ==> subcomponent: . device history review ==> the pump s/n (b)(6) passed all post sterile inspection checks.
2490 | 1220648-2025-49292 | 23888400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888400&pc=OZD | 09/23/2025 |  |  | 12/24/2025 |  |  | 09/23/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - |  |  | US |  | 2026723852 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing high-risk percutaneous coronary intervention. during transport to the cardiac care unit, the hospital bed accidentally collided with the automated impella controller (aic), causing the white plug on the device cable to break. part of the plug remained lodged in the aic connection port. the patient was immediately returned to the cardiac catheterization laboratory, where the impella device was removed. following device removal, oozing was noted at the access site after perclose closure. manual compression was applied for approximately 20 minutes to achieve hemostasis. two perclose proglide devices had been used during the pre-close procedure. hemostasis was achieved with manual compression, and the access site remained stable. the patient¿s survived the procedure. | [5 empty] | Device received, inv in progress: the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2491 | 1220648-2025-49293 | 23888405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888405&pc=OZD | 09/23/2025 |  |  | 12/24/2025 |  |  | 09/23/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722276 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, bleeding and hematoma developed at the impella access site. the blue suture hub on the repositioning (repo) sheath appeared to have advanced fully to the wings. the site was redress and re-secure the hub with forward tension sutures, and a safeguard device was placed. when subsequent hematoma and suction alarms occurred, the patient was given fluids, the dressing was removed, and manual pressure was applied for approximately one hour. the safeguard device was replaced, and by the following morning, the access site appeared clean with no visible hematoma. the patient also experienced runs of ventricular tachycardia (vt). the patient¿s condition stabilized, and the impella cp device was subsequently weaned and explanted. | [5 empty] | Investigation summary: data review: data logs showed pump ran without issue during support. there were suction alarms triggered in early of the case but resolved quickly. there were no mc spikes, abnormal ps or purge issues observed during support. product was not returned for review. major bleed: the cause of major bleed was unable to be determined as limited clinical details were provided and no product was returned for review. pump suction: the cause of pump suction was most likely patient condition since it resolved per data log review after volume was given. device history lot, device lot #: 1956251. device history batch, subcomponent lot#: n/a. device history review, pump sn: (b)(6). passed all post sterile inspection checks.
2492 | 1220648-2025-49294 | 23888409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888409&pc=OZD | 09/23/2025 |  |  | 12/24/2025 |  |  | 09/23/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740996 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Heart Block | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. the device was successfully inserted and ramped up to provide mechanical circulatory support without any complications. the patient had an underlying complete heart block, and their heart was being fully paced with a temporary pacemaker set to 80 beats per minute. the patient¿s condition stabilized with the support of the device, which was eventually weaned off and removed successfully. | [5 empty] | Investigation summary heart block: the cause of the injury was not determined due to insufficient clinical details. device history lot device lot : 1977127. device history batch sub-component lot : n/a. device history review the pump s/n: (b)(6) passed all the post sterile inspection checks.
2493 | 1220648-2025-49296 | 23888419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888419&pc=OZD | 09/23/2025 |  |  | 12/24/2025 |  |  | 09/23/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711277 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during a subsequent coronary artery bypass procedure, the impella was positioned across the aortic valve in surgical mode. upon weaning from cardiopulmonary bypass, the impella could not be repositioned because the pigtail had wrapped around the mitral apparatus. attempts to pull it back posed a risk of mitral valve injury. the surgical team decided to arrest the heart and remove the impella manually. the device was successfully explanted, and the patient was able to come off cardiopulmonary bypass. the groin access site was closed. the patient¿s condition stabilized, and the device was subsequently weaned and explanted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2494 | 1220648-2025-47805 | 23888527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888527&pc=PYX | 09/24/2025 | 02/11/2026 |  | 12/24/2025 |  |  | 09/24/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025564025 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp and impella rp flex for bipella support. during percutaneous coronary intervention to the right coronary artery the patient coded. cardiopulmonary resuscitation was completed with return of spontaneous circulation. the circumflex artery was noted chronic total occlusion. the impella cp was placed in the right femoral artery and decision was made to place an impella rp flex via the left femoral vein. note: actual timing of impellas placement is unknown. post implant same day, the patient experienced access site bleeding and falling hemoglobin requiring the purge solution to be changed to bicarbonate. left groin hematoma was noted. massive transfusion protocol initiated, and the patient received six units of blood products. urine output was pink-red. care was withdrawn and the patient expired. | The complainant reported additional information upon request: "¿ the catheters are not available for return as the cases were done in (b)(6). ¿ the cp measuring deep at 4.2cm from (b)(6) was not repositioned." |  |  |  |  | Patient-specific information a4.weight is unknown at the time of this mdr writing. additionally, the exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. medical safety review. medical safety reviewed the event and noted the following: the patient underwent an intervention to the right coronary artery (rca) and experienced cardiac arrest during the procedure. resuscitation efforts were successful, and following the code, impella cp and impella rp flex devices were implanted for hemodynamic support. after device implantation, the patient developed significant access site bleeding and hematomas, resulting in a substantial drop in hemoglobin. this required transfusion of six units of blood products, along with plasma and platelets. additionally, urine output was noted to be pink-tinged, raising suspicion of hemolysis. despite ongoing support, care was withdrawn, and the patient expired while on device support. adverse events identified: ¿ hemolysis: suspected based on pink-tinged urine during support. ¿ major bleeding and hematoma: significant bleeding at the access site requiring multiple transfusions. ¿ patient death: occurred while on impella cp and rp flex support. reportability assessment: ¿ hemolysis is reportable for both impella cp and rp flex. ¿ major bleeding and hematoma requiring transfusion are reportable for both devices. ¿ death on support is reportable for both devices, as device contribution cannot be ruled out, although the patient¿s deteriorating condition prior to implant is considered a contributing factor. manufacturer evaluation / contributing factors: the patient¿s critical condition prior to device implantation likely contributed to the adverse outcome. however, due to the presence of suspected hemolysis and significant bleeding during device support, these events are considered reportable. device malfunction has not been confirmed at this time. device status. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella rp flex device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp. | The investigation is ongoing. | Corrected data: e0303 removed from h6. investigation summary: no product was returned. access site adverse event (bleeding, hematoma) - after rp implanted, left groin hematoma discovered with falling hgb from 12 to 7.9 to 4. uop greater than 30 cc/hr & pink-red, units rbc s. act noted as 300 several hours after insertion. the cause of the access site adverse events was use related to the patients act being 300, a failure to follow instructions failure. patient codes: hematuria: the cause of the injury was not determined due to insufficient clinical details.
2495 | 1220648-2025-49297 | 23888530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888530&pc=OZD | 11/27/2025 | 12/24/2025 |  | 12/24/2025 |  |  | 11/26/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1618084 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that the damage was discovered not during patient care but during inspection of the aic while it was in storage. aic power cord was damaged. | [5 empty] | Device status: the impella device was returned, and an evaluation/analysis is underway. upon completion of this analysis, a supplemental report will be filed. | B3 revised date as it was reported incorrectly on the initial report that was submitted. e4 selection was added as it was omitted from the initial report that was submitted. g3 date on initial report should of been 27-nov-2025. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been submitted. data analysis: the console logs were not relevant and showed no errors. device analysis: the console was returned for further investigation. during the inspection, it was found that the outer sheath of the alternating current (ac) power cord did not meet the datasheet specification. the sheath had been stripped approximately 0.44 inches more than required, which may have caused the cable to sit loose and being properly secured by the clamp inside the plug. in addition, there were no clamp marks on either the outer sheath or the internal wires. root cause: the reported issue of a damaged automated impella controller power cord was confirmed and determined to be caused by an insufficient or incorrect length of the black outer sheath (defective ac power cord). product history review summary : there were two involving the same subcomponent lot as the reported issue was discovered during the review the product history of the console. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2496 | 1220648-2025-49298 | 23888532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888532&pc=OZD | 09/24/2025 | 01/15/2026 |  | 12/24/2025 |  |  | 09/24/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676915 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp and impella rp flex for bipella support. during percutaneous coronary intervention to the right coronary artery the patient coded. cardiopulmonary resuscitation was completed with return of spontaneous circulation. the circumflex artery was noted chronic total occlusion. the impella cp was placed in the right femoral artery and decision was made to place an impella rp flex via the left femoral vein. note: actual timing of impellas placement is unknown. post implant same day, the patient experienced access site bleeding and falling hemoglobin requiring the purge solution to be changed to bicarbonate. left groin hematoma was noted. massive transfusion protocol initiated, and the patient received six units of blood products. urine output was pink-red. care was withdrawn and the patient expired. | The complainant reported additional information upon request: ¿ the catheters are not available for return as the cases were done in (b)(6). ¿ the cp measuring deep at 4.2cm from ava was not repositioned. |  |  |  |  | Patient-specific information a4.weight is unknown at the time of this mdr writing. additionally, the exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. medical safety review. medical safety reviewed the event and noted the following: the patient underwent an intervention to the right coronary artery (rca) and experienced cardiac arrest during the procedure. resuscitation efforts were successful, and following the code, impella cp and impella rp flex devices were implanted for hemodynamic support. after device implantation, the patient developed significant access site bleeding and hematomas, resulting in a substantial drop in hemoglobin. this required transfusion of six units of blood products, along with plasma and platelets. additionally, urine output was noted to be pink-tinged, raising suspicion of hemolysis. despite ongoing support, care was withdrawn, and the patient expired while on device support. adverse events identified: hemolysis: suspected based on pink-tinged urine during support. major bleeding and hematoma: significant bleeding at the access site requiring multiple transfusions. patient death: occurred while on impella cp and rp flex support. reportability assessment: hemolysis is reportable for both impella cp and rp flex. major bleeding and hematoma requiring transfusion are reportable for both devices. death on support is reportable for both devices, as device contribution cannot be ruled out, although the patient¿s deteriorating condition prior to implant is considered a contributing factor. manufacturer evaluation / contributing factors: the patient¿s critical condition prior to device implantation likely contributed to the adverse outcome. however, due to the presence of suspected hemolysis and significant bleeding during device support, these events are considered reportable. device malfunction has not been confirmed at this time. device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella rp flex. | The investigation is ongoing.
2497 | 1220648-2025-49299 | 23888533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888533&pc=OZD | 09/24/2025 |  |  | 12/24/2025 |  |  | 09/24/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725222 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male | 100 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Pulmonary Edema | Tachycardia | Hematuria | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the patient experienced multiple clinical issues, including ischemia, pulmonary edema, and tachycardia with recurrent ventricular tachycardias. laboratory values showed an elevated partial thromboplastin time, elevated lactate, and later the development of bloody urine. over time, clinical evaluation indicated no ongoing signs of hemolysis, and hemoglobin stabilized after transfusion of two units of packed red blood cells. repositioning of the impella, fluid administration, and adjustment of anticoagulation therapy were discussed. the patient was subsequently escalated from impella cp to impella 5.5 for continued hemodynamic support. ultimately, the patient was successfully weaned from impella support. | [5 empty] | Ischemia/ tachycardia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. pulmonary edema/ hematuria: the cause of the injury was not determined due to insufficient clinical details.
2498 | 1220648-2025-49302 | 23888535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888535&pc=OZD | 09/24/2025 |  |  | 12/24/2025 |  |  | 09/24/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673895 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient, who had postcardiotomy cardiogenic shock and low cardiac output syndrome. the patient was later brought to the operating room for escalation from the impella cp to an impella 5.5. during the procedure, the patient decompensated, and veno-arterial extracorporeal membrane oxygenation was initiated. due to challenging anatomy, the impella 5.5 could not be advanced into position. the physician elected to maintain arterial pressure monitoring through the impella cp re-access port. despite escalation of support and resuscitative efforts, the patient¿s condition continued to decline. the patient experienced hemodynamic collapse and cardiogenic shock that did not recover despite advanced support. care was ultimately withdrawn, and the patient expired. | [5 empty] | Investigation summary: no product was returned for investigation. hypotension: the cause of the clinical symptom was not determined due to insufficient clinical details. device history lot: device lot: 1936567 device history batch: subcomponent lot: n/a, device history review: device with sn: (b)(6) passed all post sterile inspection checks.
2499 | 1220648-2025-49300 | 23888536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888536&pc=OZD | 09/24/2025 | 02/14/2026 |  | 12/24/2025 |  |  | 09/24/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026706056 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female | 45 KG |  |  | Death | Yes | Yes | Degraded | Unintended Collision | Pumping Stopped | Device Damaged by Another Device | Material Split, Cut or Torn |  |  | Cardiac Arrest | Perforation of Vessels | Pericardial Effusion | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the indication for use was severe triple vessel disease, cardiogenic shock, and an impella cp was placed in a patient with severe triple vessel disease during a high-risk intervention. during the intervention procedure, a coronary perforation occurred, resulting in a hemodynamically significant pericardial effusion. emergency pericardial drainage was performed, stabilizing the patient. the perforation could not be treated with a covered stent due to the lesion being unrecoverable, and the patient became fully dependent on the impella pump for circulatory support. a decision was made to transfer the patient to cardiac surgery at a hub hospital. during transport to the ambulance, the hydraulic stretcher was inadvertently lowered without warning, causing the stretcher to break the impella connector cable plug, resulting in immediate device failure. the patient, who was 100% dependent on the device, unfortunately passed away. | [5 empty] | The following codes have been added: h6 clinical code: e0511, e0619. h6 medical device prob code: a0414, a0502 and removed a0405,a051205. | Patient-specific information a5. ethnicity and a6. race are unknown at the time of this mdr writing. should this information become available, a supplemental report will be submitted accordingly. medical safety review. medical safety reviewed the event and noted the following: a serious injury leading to death ¿ device-related adverse event. further noted, the pump malfunction did contribute to the cause of death, and the malfunction was caused by user. user error during patient transport led to mechanical damage of the device connector, causing loss of pump function resulting in patient death. contributing factors are patient's critical condition with dependence on hemodynamic support and lack of precautionary measures during transport. note: this mso review does not impact/change the usmdr reportability assessment. type of reportable event remains death for usmdr. device status. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2500 | 1220648-2025-49303 | 23888538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888538&pc=OZD | 09/24/2025 |  |  | 12/24/2025 |  |  | 09/24/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691909 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was attempted to be implanted to support a patient undergoing high-risk percutaneous coronary intervention. during preparation of the impella cp, the clinical team was unable to move the sterile sleeve back toward the red plug despite the tuohy borst being unlocked. there was not enough working length on the catheter, and multiple staff members attempted without success. a second impella cp was opened, prepared normally, and successfully implanted. the impella cp was later removed without complication, and the patient remained stable following the procedure. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2501 | 1220648-2025-49306 | 23888565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888565&pc=OZD | 09/24/2025 |  |  | 12/24/2025 |  |  | 09/24/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643598 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, a small right groin hematoma was noted. one unit of red blood cells was administered. the patient¿s condition stabilized, and the device was subsequently weaned and explanted. | [5 empty] | Investigation summary: no product was returned for investigation. hematoma: it was reported that right groin hematoma still present slight expansion but stable. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history lot: device lot: 1878532. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
2502 | 1220648-2025-49301 | 23888613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888613&pc=OZD | 11/26/2025 | 12/24/2025 |  | 12/24/2025 |  |  | 11/26/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655994 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 37 YR |  | Male | 188 KG |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during the atemnot to reposition the impella, it slipped out of the ventricle. the impella could not be reinserted. pump was exchanged. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product returned. device in wrong position: the cause of the positioning issue is use related as the pump slipped out of the ventricle during an attempt at repositioning. device history review: device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2503 | 1220648-2025-49308 | 23888641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888641&pc=OZD | 01/15/2025 |  |  | 12/24/2025 |  |  | 01/15/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025586494 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Female |  |  | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a kink in the cannula during repositioning of the pump. the pump was explanted. no patient harm was reported. | [5 empty] | A4 and a5 are unknown. no product was returned for investigation. the root cause for the cannula kink was most likely use issue as imaging techniques were not used while repositioning the pump. the device passed all post-sterile inspection checks.
2504 | 1220648-2025-49310 | 23888673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888673&pc=OZD | 09/05/2025 | 02/10/2026 |  | 12/24/2025 |  |  | 09/05/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1319196 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | Operating System Becomes Nonfunctional | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  | MWR-02122025-0001983438 |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that during customer training the automated impella controller screen froze. there was no patient involvement. | [5 empty] | H6: added code based on information in the complaint file. | The event did not occur with a patient. section a was left blank. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: d1, h5, and h8 was erroneously entered on the initial manufacturer device report. d4 was inadvertently omitted on the initial manufacturer device report. | The investigation for the controller/device freezes issue has been completed. the device and data logs were returned for evaluation. the logs showed that in the case start wizard at step three, the pump connection was most likely not recognized, or the white catheter plug was not completely inserted into the aic, so after 20 seconds, epm was reset automatically. after the epm reset, the pump connection remained unrecognized, confirming the reported failure mode. this could be most likely due to the crystal issue on the impellatronic (epm) board. subsequently, the soft button next was pressed, causing the console to display messages indicating ¿connect cable: detecting impella screen¿ and ¿connect cable: connect grey connectors screen.¿ the console was evaluated and during testing, the pump detection issue could not be reproduced despite several attempts to connect a known-good impella and power cycling the console in between attempts. the root cause of the aic screen freezing during customer training, where the demo case did not proceed was likely due to the rarely occurring epm crystal startup issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and has a sw mitigation in place.
2505 | 1220648-2025-49307 | 23888781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888781&pc=OZD | 10/25/2025 | 02/17/2026 |  | 12/24/2025 |  |  | 11/27/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026747662 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that just before removal, the position waveform disappeared. the motor waveform was normal, and an echocardiogram and x-ray confirmed that the position was ok. preparations were being made for removal, but there was no particular snagging or tension.nothing had been done just before the position waveform disappeared. | [5 empty] | The investigation for the optical sensor issue has been completed. based on the symptom noted in field data logs along with no defect being found on the returned pump, it was determined that the cause of the optical sensor issue was likely related to the console. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2506 | 1220648-2025-49312 | 23888907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888907&pc=OZD | 09/24/2025 |  |  | 12/24/2025 |  |  | 09/24/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596846 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 117 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during the procedure, the patient became hypotensive and required dopamine for hemodynamic support. the impella cp was inserted successfully. during subsequent imaging, the impella backed out and could not be re-crossed. the patient was taken back, where the impella cp was removed and exchanged for an intra-aortic balloon pump. the patient remained stable after the procedure. | [5 empty] | Investigation summary: no product was returned. device in wrong position (positioning issue): the cause of the positioning issue was determined most likely to be unintentional use error but unsure if the pump was pulled out of position by the user or patient during the transport to ct. device history lot: device lot: 1873885. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
2507 | 1220648-2025-49311 | 23888912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888912&pc=OZD | 11/27/2025 | 12/24/2025 |  | 12/24/2025 |  |  | 11/28/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759479 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Sepsis | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient with endocarditis, tricuspid, mitral and aortic valve have been replaced. there was bleeding post surgery, but no issues with the impella device. insertion site perfectly dry. 30 units of red blood cells were given. in addition, patient expired but it is unknown if patients' death was related to device. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. sepsis: in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed/renal failure: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2508 | 1220648-2025-49314 | 23888917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888917&pc=PYX | 09/24/2025 |  |  | 12/24/2025 |  |  | 09/24/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026671772 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 78 KG | Non Hispanic | Asian | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Pulmonary Edema | Respiratory Failure | Hematuria | Asystole | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp flex experienced hemolysis, hematuria, asystole, and pulmonary edema. the position of the pump was confirmed to be correct. the patient was treated with epinephrine, vasoactive drips, and heparin. the patient was acidotic with acute kidney injury and shock liver. the patient was cannulated with veno-arterial extracorporeal membrane oxygenation and the impella was explanted. | [5 empty] | No product was returned for investigation. the cause of the respiratory failure and pulmonary edema was not determined due to insufficient clinical details. the root cause of the asystole was unable to be determined due to insufficient clinical details. the cause of the hemolysis and hematuria was most likely patient condition-related, based on the given clinical details and data log review. the device passed all post sterile inspection checks.
2509 | 1220648-2025-49316 | 23888920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888920&pc=OZD | 09/25/2025 |  |  | 12/24/2025 |  |  | 09/25/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713911 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. the patient experienced a period of hypotension. continuous renal replacement therapy was ongoing, and vancomycin was added for infection management. during support, the impella 5.5 experienced repeated pump stop events three times and high purge pressure alarms. the patient was taken to the cardiovascular lab for pump exchange. due to extreme difficulty with the impella 5.5 placement, the decision was made to switch to the right femoral artery access and implant an impella cp. | [5 empty] | Investigation summary: impella 5.5 sn (b)(6) + purge cassette returned for evaluation. after 15 days on support, purge system blocked/high purge pressures noted. no kinks observed. pc replaced to no resolution. tpa added to purge with unknown resolve. pump stop x 3. purge pressure high + purge system blocked - data log review shows purge pressure high and purge blocked alarms occur after a 12 hour gradual rise in purge pressure with purge flows decreasing in parallel. the returned pump recreated purge pressure high alarms and had biomaterial around impeller/purge gap, blood in the purge lumen. the cause of the purge pressure high and purge system blocked issues was patient condition related to biomaterial build-up found in the purge lumen, recreating the purge issues and data logs showing biomaterial build-up from gradual pp rise. pump stop - data log review shows gradual rise in pp over 12 hours that culminate in impella stopping. pump able to resume for another 2 hours before more pump stop alarms. data logs showing biomaterial build-up from gradual pp rise. the returned pump had biomaterial wrapped around the impellor blade and had no open motor lines. also, there was blood in the purge lumen which was well past the motor housing. attempts to run the pump failed due to mc high alarms triggering. the cause of the pump stop was likely due to blood ingress, patient condition related to biomaterial build-up found wrapped around the impellor blades. patient codes - hypotension: the cause of the hypotension was not determined due to insufficient clinical details. arrythmia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the arrythmia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1951071 device history batch: subcomponent lot: n/a device history review: pump s/n (b)(6) passed all post sterile inspection checks.
2510 | 1220648-2025-49317 | 23888926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888926&pc=OZD | 09/26/2025 |  |  | 12/24/2025 |  |  | 09/26/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during clinical use of the automated impella controller, the purge disc was not detected. no additional information was available. | [5 empty] | Dhr: q1) service history review - are there any qns associated with the complaint device¿s most recent service report? no. q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode?. q3) complaint history review - have there been any other complaints against this console? yes. there were 6 other complaints made against this console: 21-05015-1: pump recognition issue. determined to be due to software. sw v8.4.1 to v8.5 should have resolved the issue. . 21-05256-1: scrolling video on impella connect? no problem found. backstrap cable replaced as a precaution. 22-11591-1: unexpected shutdown. not determined due to the inability to reproduce. 4-in-1 and compact flashcards were replaced. 23-15053-1: purge disc detection issues. broken purge disc flag. replaced to address the issue. 23-16640-1: controller failure. the purge had a damaged communication packet causing an incorrect crc resulting in a pud failure alarm. 25-42086-2: console unable to boot successfully. corrupted compact flash cards were the root cause. phr summary: there was an issue related to the complaint discovered when reviewing the product history of console (b)(6) prior complaint with the same fm/rc (23-15053-1). data review: console logs were consistent with what was reported in the complaint. purge disc not detected alarms [event set #402] were observed in the imc logs from the reported complaint date. (see (b)(6)_imclog_segment.pdf¿ in the attachments) device analysis: console was tested after arriving in field service. the reported purge disc detection issue was able to be reproduced. inspection of the console revealed that the purge disc flag was broken (missing at blue retainer). (see ¿(b)(6)_brokenpurgediscflag.png¿ in the attachments) conclusion: the root cause of the reported purge disc detection issue was the observed broken purge disc flag. repair: before sending this console back to the customer, field service was instructed the following: - replace the purge disc detect flag [3001198] and purge flag spring [3001196] - perform sections 10-12 of pm procedure (b)(6) - perform a full functional check on the console and optical system [0042-7316].