16612 | 1220648-2024-21223 | 20540483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540483&pc=OZD | 07/27/2022 | 03/28/2025 |  | 10/26/2024 |  |  | 07/27/2022 | 10/26/2024 |  |  | 12/31/2023 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023178976 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | Priming Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the device exhibited an air in line alarm. the staff de-aired the system 3 times. issues remain unresolved. pump then stopped. they attempted to restart the device which was unsuccessful. this led to the device being swapped for a new one. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the pump was returned for investigation. no logs are available for review. the purge cassette was not returned. the cause of the priming issue was not determined since product was not returned. a visual inspection of the pump revealed blood in the purge gap and blood in the motor. no blood was found in the purge lumen. the resistances between all three motor phases are within specification. the pump would not run. the cause of the pump stop was most likely use related as the purge issue likely allowed for blood ingress into the motor, stopping the pump. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This is one of two medical device reports associated with the event. refer to manufacturing report number 1220648-2024-21468 for pump 2 sn (b)(6). this report is being filed as part of a retrospective review of historical records. corrections made to initial mfr 1220648-2024-21223: b1: adverse event selected. b2: date of death added. h6: investigation findings codes 213/no device problem found and 3221 no findings available corrected to 4251 undesirable presence of endogenous materials.
16613 | 1220648-2024-21224 | 20540484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540484&pc=OZD | 01/14/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 01/14/2023 | 10/26/2024 |  |  | 08/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260863 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was needing an impella cp for mechanical circulatory support. it was noted that multiple attempts were made to place the pump, however the physician was unable to advance the device through the patient's vasculature. the implant was subsequently aborted. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of delivery issue was due patient condition. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
16614 | 1220648-2024-21225 | 20540485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540485&pc=OZD | 01/10/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 01/10/2023 | 10/26/2024 |  |  | 08/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269537 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported hematoma and subsequently delivery issue involving during impella support. during the sheath exchange following the initial delivery, a hematoma formed preventing the repositioning sheath from fully advancing into the arteriotomy. the introducer was swapped to an 18fr sheath to obtain hemostasis, and the pump was implanted without further issue. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient condition related and related to the hematoma that had formed. the cause of the access site bleeding was most likely patient condition related since oozing was also reported from the ij swan site. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
16615 | 1220648-2024-21226 | 20540486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540486&pc=OZD | 05/02/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 05/02/2023 | 10/26/2024 |  |  | 12/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321384 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was the development of a thrombus coinciding with impella support. it was noted during an echocardiogram that a thrombus was wrapped around the catheter of the impella. due to the noted thrombus, the physician was advised of the potential for a pump stop and low pulsatility. patient support continued with the implanted impella following discovery of the thrombus. care was ultimately, withdrawn and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the reported thrombus issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the patient has preexisting thrombus in the aortic root. the root cause of the thrombus issue was patient condition related. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
16616 | 1220648-2024-21227 | 20540535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540535&pc=OZD | 04/13/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 04/13/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1545509 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller failure red alarm. no patient was involved. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the console was tested but the pump stop failure was unable to be reproduced with a known-good demo pump. inspection on the impellatronic board did reveal some marks of rework fcn98 which did not affect the functionality based on in-lab testing and other 3 successful patient cases for ic8481. since these errors occurred while using the demo pump on (b)(6), it is likely that the issues stemmed from the demo pump in use. the cause of the issue was a defective demo pump. this report is being filed as part of a retrospective review of historical records.
16617 | 1220648-2024-21228 | 20540536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540536&pc=OZD | 04/05/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 04/06/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1306046 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced inability to mute suction alarms via the console's "mute" button. the console was power cycled and the issue could not be reproduced. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic5445) was sent back for further investigation. the root cause of the ¿mute¿ key unresponsive was unable to determine because the failure mode was unable to reproduce. this report is being filed as part of a retrospective review of historical records.
16618 | 1220648-2024-21229 | 20540537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540537&pc=OZD | 03/31/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 03/31/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090371 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced a battery failure red alarm that was resolved by switching to another aic. no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the failure mode was reproduced on an engineering bench. the battery failure was due to a defective battery 2. the root cause of the defective battery 2 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16619 | 1220648-2024-21230 | 20540540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540540&pc=OZD | 03/10/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 03/10/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1105128 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error alarm upon startup. it was reported the alarm was not able to be silenced or powered off, the hard power switch was used to power off the console. this aic was replaced with a backup aic and the procedure was successfully completed. there was no patient harm reported. | [5 empty] | The investigation into the controller error/loss of opm data has been completed. the impella automated controller was returned for investigation. the data analysis of the logs revealed there were no indication of the failure to confirm the complaint. the console had recently been serviced, on 02/23/2023. since the console was sent back to the hospital, it had been used for 3 patient support cases (all with cp pumps) without issue. on (b)(6) 2023, a 5.5 pump was implemented for patient support, the support was successfully without any issue, and the pump was explanted on (b)(6) 2023. through the logs, since last service to when the console was returned, there had been no issue with controller error alarm or the unable to shut down with this console ic1675. therefore, the issue could not be confirmed. the returned console was inspected and tested without major issue; glucose contamination was found near the purge door but is not related to the complaint. there was no issue found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
16620 | 1220648-2024-21231 | 20540546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540546&pc=OZD | 03/09/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 03/09/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1286797 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. it was reported that the aic did not recognize the purge cassette, reinserting the cassette did not solve the problem. the aic was replaced and there were no further issues. no report of patient harm or injury. | [5 empty] | The investigation into the purge disc/flag issues has been completed. the automated impella controller (aic) was returned for investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). the cause of the issue was broken/ missing purge flag. this report is being filed as part of a retrospective review of historical records.
16621 | 1220648-2024-21232 | 20540547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540547&pc=OZD | 03/05/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 03/05/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146946 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | Failure to Read Input Signal | [5 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. while on support, the device exhibited an air in line alarm and the purge cassette was leaking. field service went onsite to flush/ change the purge cassette and the purge disk holder found to be broken off, the purge disk was punctured. the aic pump was replaced with no further reported issue. there was no report of patient harm or injury. | [5 empty] | The investigation into the purge disc / flag issue has been completed. the automated impella controller (aic) was returned for investigation. visual inspection of the purge assembly area confirmed a ripped/ missing blue retainer. the cause of the issue was broken/ missing purge disk retainer. this report is being filed as part of a retrospective review of historical records.
16622 | 1220648-2024-21233 | 20540548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540548&pc=PYX | 07/13/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 07/13/2021 | 10/26/2024 |  |  | 03/31/2023 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2022065727 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing a coronary artery bypass was having an impella rp placed for mechanical circulatory support when the device kinked. it was reported that while advancing into the heart while being torqued, the .027 guide wire had no kinks and was straight up with no bowing into the right ventricle. reportedly as the device was being clockwise torqued through the tricuspid valve, the device kinked and was unable to be advanced. the patient was taken back to the catheterization laboratory to have a replacement impella rp inserted. there were no issues reported. the replacement impella rp was successfully placed. it was reported that the patient survived the procedure. | [5 empty] | The investigation into the product damage (kink) has been completed. the impella rp was returned for investigation. based on the clinical information provided in the complaint, the root cause of the cannula kink occurred during delivery issues is most likely patient condition. this report is being filed as part of a retrospective review of historical records.
16623 | 1220648-2024-21234 | 20540549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540549&pc=PYX | 07/25/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 07/25/2021 | 10/26/2024 |  |  | 02/28/2023 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022052390 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported the pump stopped resulting in this pump being replaced with another impella rp. reportedly, there were no issues with placement of the replacement pump and the patient remained on support. no further information was reported. | [5 empty] | The investigation into the pump stop issue has been completed. the returned data logs showed that the pump flow immediately dropped to near 0 liters per minute about three hours into support upon the pump being manually re-zeroed, which was accompanied by the occurrence of alarm 120. however, there were no "impella stopped" alarms at this time. in conclusion, it was determined that the cause of the placement signal issue was use related, specifically the offset placement signal metrics negatively affected flow calculations, thus inhibiting pump performance. the cause of the placement signal issue was use related, specifically the offset placement signal metrics negatively affected flow calculations, thus inhibiting pump performance. as noted in the impella instructions for use: impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella rp smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella rp smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16624 | 1220648-2024-21235 | 20540550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540550&pc=OZD | 03/02/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 03/02/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1204377 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the aic experienced a controller failure red alarm. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. imc logs and rtlog are consistent with the complaint. during the patient support, console alarmed for "air in purge system", purge flow ticks kept the same and purge pressure started to decline. as user acknowledged the alarm, purge wizard, de-airing, was performed and cleared the alarm. however, after the wizard, purge pressure was unable to elevate. the extremely low purge pressure eventually triggered the purge system open alarm. and as the purge was advancing to compensate the low pressure, there has no increment of purge pressure. as the result, console software algorithm regarded the situation as the malfunction of purge pressure sensor and triggered the alarm, controller failure (purge pressure sensor defective). console was tested with demo pump and purge cassette to simulate the processes had been done prior to the controller failure alarm was triggered. however, the failure mode was not reproduced. console purge pressure sensor has been tested, but the pressure reading was within the specification. the root cause of the controller failure alarm was unable to determine because the failure mode was unable to reproduce. this report is being filed as part of a retrospective review of historical records.
16625 | 1220648-2024-21236 | 20540553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540553&pc=OZD | 01/31/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 02/01/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1155603 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support, complainant reported that the aic experienced purge disc/flag issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of the issue was broken purge disk retainer. this report is being filed as part of a retrospective review of historical records.
16626 | 1220648-2024-21237 | 20540555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540555&pc=PYX | 07/25/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 07/25/2021 | 10/26/2024 |  |  | 03/31/2023 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022074843 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient had an impella rp placed for post operative mechanical circulatory support had an issue during placement. it was reported that during insertion, the catheter prolapsed and kinked while trying to advance the device through the right ventricle. this impella rp was removed and successfully replaced with another impella rp. it was reported that the patient survived the procedure. | [5 empty] | The investigation into the issue has been completed. the impella rp was not received from the customer. therefore, an evaluation of the device was not possible. per the results of the clinical review and investigation, the root cause of the kinked cannula was determined to be delivery issues and the root cause of the delivery issues was most likely a result of patient condition of diseased/small vessels. this report is being filed as part of a retrospective review of historical records.
16627 | 1220648-2024-21238 | 20540561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540561&pc=OZD | 02/23/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 02/23/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1541270 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced battery charging issues. during preventative maintenance, it was discovered that the charging current toggled between different, even negative values. it was noted that the aic was plugged into to the ac. no patient involvement. | [5 empty] | E1: the name of the initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the aic charging issue was due to a defective power supply. this report is being filed as part of a retrospective review of historical records.
16628 | 1220648-2024-21239 | 20540572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540572&pc=OZD | 02/17/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 02/17/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1380525 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's nurse reported that while preparing to implant an impella pump into a patient of unknown gender/race/ethnicity, the automated impella controller (aic) was not charged and showed no charging on the display battery icon. it was later found that the "on" button on the aic's ground had not been pressed. the patient outcome was not reported. | [5 empty] | Data analysis: the aic logs were able to confirm the complaint. both consoles were powered up after undergoing product maintenance and had an immediate "battery failure (transport connection blown/missing) - alarm issued #360" which is how the console reacts to the battery transport switch being flipped off. the failure goes away when the transport switch was corrected. root cause: the cause of the issue was the transport switch not being flipped after receipt. this report is being filed as part of a retrospective review of historical records.
16629 | 1220648-2024-21240 | 20540574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540574&pc=OZD | 02/13/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 02/13/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1135594 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the automated impella controller experienced a battery failure red alarm. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The impella device was returned and evaluated. the cause of the reported issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16630 | 1220648-2024-21241 | 20540618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540618&pc=OZD | 05/29/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 05/29/2023 | 10/26/2024 |  |  | 12/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2024323309 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced a hematoma at the impella site in the left femoral artery. manual pressure was held, and pressure dressing was applied. it was also reported that a surgical cut down was necessary for the normal removal of the impella cp. the left femoral artery was repaired and surgically closed. a wound vac was placed over the incision site. the procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported hematoma and vascular damage has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were sufficient to determine a root cause. the cause was anticoagulation management related because activated clotting time (act) was 312 and no intervention beside pressure applied at access site to solve the bleeding issue. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16631 | 1220648-2024-21242 | 20540628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540628&pc=OZD | 02/17/2023 | 10/08/2025 |  | 10/26/2024 |  |  | 02/17/2023 | 10/26/2024 |  |  | 06/30/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231681 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited suction issues indicating low flow. the resolution for this issue is unknown. the patient outcome at explant noted that the patient expired. no further information is available. omedical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | B1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21242. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21242. g1 reporting contact email revised on manufacturer device report 1220648-2024-21242 in accordance with updated procedures. h1 type of reports was erroneously selected on the initial manufacturer device report number 1220648-2024-21242. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21242 h11 investigation conclusions 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-21242.
16632 | 1220648-2024-21243 | 20540631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540631&pc=OZD | 01/28/2023 | 10/07/2025 |  | 10/26/2024 |  |  | 02/17/2023 | 10/26/2024 |  |  | 05/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219795 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient died due to respiratory arrest due to cerebral hemorrhage. it was reported that the impella cp worked as expected. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | B2 date of death was erroneously entered on the initial manufacturer device report number 1220648-2024-21243. b3 date of event was was erroneously entered on the initial manufacturer device report number 1220648-2024-21243. g1 reporting contact email revised on manufacturer device report 1220648-2024-21243 in accordance with updated procedures. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2024-21243. h6 investigation conclusions 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-21243.
16633 | 1220648-2024-21244 | 20540635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540635&pc=OZD | 02/02/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 02/02/2023 | 10/26/2024 |  |  | 10/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024286304 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a delivery issue involving a guidewire and impella cp pump. during attempted delivery of the impella pump, the guidewire was accidentally pulled from the left ventricle by the technician. as a result, the guidewire and pump were removed from the patient. however, the technician had difficulty re-wiring the pump without the easy guide lumen, and the decision was made to use another pump for insertion. there was no patient harm resulting from the delivery issue. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on the clinical data the issue occurred because the wire pulled back and because the red lumen was removed they couldn't reinsert the guide wire and the pump. root cause was use issue because of improper technique use. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16634 | 1220648-2024-21245 | 20540640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540640&pc=OZD | 05/05/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 05/08/2023 | 10/26/2024 |  |  | 11/30/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2024300586 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the original impella device was prepped and ready to implant, however it was contaminated, and a new device was prepped and inserted into the patient for support, the delay was reported to be 45 minutes. no further information is available. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported contamination issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the details provided were sufficient to determine a root cause. the root cause of the product damage was determined to be use related as the pump was contaminated during initial set up. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
16635 | 1220648-2024-21246 | 20540679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540679&pc=PYX | 08/11/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 08/11/2021 | 10/26/2024 |  |  | 03/31/2023 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022069069 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hypovolemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that the sheath kinked causing damage to the pump which was noted under fluoroscopy. this impella rp was removed and successfully was replaced with another impella rp. additional information noted the patient required one milligram of epinephrine boluses to maintain the patient's blood pressure. no further information was provided. | [5 empty] | The investigation into the issue has been completed. only the pump was returned for investigation. visual inspection of the impella revealed significant deformation of the outflow cage. no other damage or abnormalities were found. after reviewing the clinical information, it was determined that the cause of both the introducer sheath kink and the deformed outflow cage was use related. the instruction for use indicate the insertion into the left femoral vein should be avoided and the impella rp should not be used with a kinked introducer. the root cause of the event was use related. this report is being filed as part of a retrospective review of historical records.
16636 | 1220648-2024-21247 | 20540681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540681&pc=PYX | 08/13/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 08/13/2021 | 10/26/2024 |  |  | 03/31/2023 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022081693 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that the impella rp was placed at the left groin. however, a placement signal unreliable alarm generated immediately after insertion. the physician noted that a clot ingestion was possible. the placement signal stopped functioning entirely. this impella rp was removed and successfully replaced with another impella rp. there was no patient harm reported. | [5 empty] | The investigation into the placement signal issue has been completed. the impella rp was returned for investigation. the real time data log was reviewed and a placement signal not reliable alarm was confirmed on (b)(6)2021 characterized by erratic placement signal waveforms ranging from approximately -2000mmhg to 3000mmhg. additionally, the pump flow was at or near zero from approximately 13:18 onward. the returned pump was visually inspected and a deformed outflow cage as well as two catheter kinks adjacent to the motor housing were noted which likely occurred during insertion. the pump was plugged into a console and ran at p-9 with normal flow, motor current, and placement signal values, even with manipulation of the catheter at the kink location. as the placement signal and flow issues were unable to be reproduced, a specific cause for these issues could not be determined through this evaluation. however, the unreliable placement signal likely led to the reported flow issues as pump flow is calculated using the measured pressure at the differential sensor. the cause of the placement signal issue was not determined as the failure was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
16637 | 1220648-2024-21248 | 20540683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540683&pc=PYX | 08/15/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 08/15/2021 | 10/26/2024 |  |  | 01/31/2023 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022049590 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. reportedly, the pump was independently placed with phone support assistance from the local team and the clinician in route to assist. it was reported during placement of a dialysis catheter, there was an abrupt rise in the motor current from approximately 350 milliamps (ma) to 580 ma and an abrupt rise in purge pressure from 40s to 160s. also, a sudden drop in purge flow from about 30 to 18. the family declined to have the impella rp replaced. the patient was on venoarterial extracorporeal membrane oxygenation, thrombolytic therapy (systemic or in purge) and the patients activated clotting time (act) was 207 seconds at the time of the occurrence. additionally, it was noted the sheath had been indwelling since placement on (b)(6) 2021. the plan was for the patient to continue with therapy, however on (b)(6) 2021, the family made the decision to withdraw care and the patient expired. additional information from the field rep noted that the motor current, placement signal, and purge pressure all suddenly rose while the physician was placing a femoral venous dialysis catheter. the team had also been attempting left internal jugular vein (lt ij) access a few moments earlier. despite the patient's act being over 200 seconds, it was reported that there were clots present on the guidewire during the ij access attempt. the heparin induced thrombocytopenia (hit) panel was negative. the patient had been febrile over 39 degrees celsius. cultures were pending. at the time of the event, the purge solution was dextrose 5 percentin water with 25 units of heparin/milliliter (ml). upon recommendation from the field representative, the team converted it to 50 units of heparin/ml. a steady drop in motor current and purge pressure towards previous levels was then noted. it was assumed that the impella flow display was inaccurate because of the pulsatility change in the placement signal. pa pressures were the same and left-sided support was unchanged. no other troubleshooting attempts were made as the family had refused any further measures. care was withdrawn about six hours later. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the high purge pressure issue has been completed. the impella rp was not received from the customer. therefore, an evaluation of the device was not possible. the data analysis of the logs confirmed a sudden rise in motor current, placement signal, and purge pressure. the data was consistent with biomaterial being ingested into the pump. the root cause of the spike in motor current and high purge pressure was most likely ingested biomaterial resulting from the patient anatomy. as noted in the impella instructions for use: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
16638 | 1220648-2024-21249 | 20540686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540686&pc=PYX | 02/24/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 02/24/2022 | 10/26/2024 |  |  | 02/28/2022 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1469832 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. it was reported that the catheter was bending up on itself and they were not able to advance the device into proper placement due to patient anatomy. the procedure was aborted. no patient harm was reported. | [5 empty] | The investigation into the reported delivery issue has been completed. no device was received for evaluation. clinical information was sufficient in determining a root cause. the cause of the issue was a catheter kink. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
16639 | 1220648-2024-21064 | 20530153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20530153&pc=OZD | 11/22/2022 | 10/07/2025 |  | 10/25/2024 |  |  | 11/22/2022 | 10/25/2024 |  |  | 05/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223079 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Intracranial Hemorrhage | [11 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. several hours after implant the patient experienced a cerebral hemorrhage. the cerebral hemorrhage progressed to cerebral hernia, and then cardiac arrest. the patient's status was "do not resuscitate" and the patient died. no information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records. | D4 model number was erroneously entered on the initial manufacturer device report number 1220648-2024-21064. g1 reporting contact email revised on manufacturer device report 1220648-2024-21064 in accordance with updated procedures. h6 investigation conclusions code 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-21064.
16640 | 1220648-2024-21065 | 20530168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20530168&pc=OZD | 12/06/2022 | 10/07/2025 |  | 10/25/2024 |  |  | 12/06/2022 | 10/25/2024 |  |  | 07/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246476 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ventricular Fibrillation | [11 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced ventricular fibrillation, the resolution for this issue is unknown. additionally, a cerebral hemorrhage was identified and the patient died 4 days later from the cerebral hemorrhage. no additional information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | D4 model number was erroneously entered on the initial manufacturer device report number 1220648-2024-21065. g1 reporting contact email revised on manufacturer device report 1220648-2024-21065 in accordance with updated procedures.
16641 | 1220648-2024-21067 | 20530309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20530309&pc=OZD | 11/28/2022 | 10/24/2024 |  | 10/25/2024 |  |  | 11/28/2022 | 10/25/2024 |  |  | 07/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251644 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced bleeding at the access site post repositioning sheath placement. a hematoma was also noted. the patient was administered 1 unit of packed red blood cells. the physician noted that they did everything per the best practices with removal of the peel away and replacement of repositioning sheath. the physician also noted that there was not enough recoil of artery on repositioning sheath. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. | [5 empty] | B2: the date of the patient's death is unknown. the investigation for the reported access site bleeding has been completed. the product was returned. per the results of the clinical review and investigation: the root cause was determined to be complete statement with the results of the investigation. it was determined that the cause of the access site bleeding was most likely a lack of vessel recoil onto the repositioning sheath. this is attributed to the patient's condition. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16642 | 1220648-2024-21068 | 20530623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20530623&pc=OZD | 12/22/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 12/22/2022 | 10/25/2024 |  |  | 09/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023278071 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited low flow. impella started on auto mode but could never get above 2 liters of flow without suction. troubleshooting did not correlate better flow. the device was explanted and replaced with another imeplla cp. no further information is available. | [5 empty] | The investigation for the reported low flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. the cause of the low pump flow was most likely patient condition related, specifically the patient's deteriorating heart function following a perforation that occurred prior to impella implant. this report is being filed as part of a retrospective review of historical records.
16643 | 1220648-2024-21069 | 20530645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20530645&pc=OZD | 12/18/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 12/18/2022 | 10/25/2024 |  |  | 06/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237204 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | During support, it was reported that the patient did not have distal pulses. the device was going to be explanted and replaced with an impella 5.5. however, the patient died before the new device could be implanted. no further information is available. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were sufficient to determine a root cause. the cause of the limb ischemia was most likely patient condition related. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16644 | 1220648-2024-21070 | 20530660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20530660&pc=OZD | 01/13/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 01/13/2023 | 10/25/2024 |  |  | 11/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023280870 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  |  |  | Death | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device alarmed "air in the purge system". it was also reported that the pump stopped, and re-start was initially successful, however only low p-level was possible, eventually the pump stopped due to the high motor current and was unable to restart. it was also reported that the patient expired. no further information is available. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of the patient's death is unknown. the investigation for the reported purge leak and pump stop has been completed. the product was returned. per the results of the clinical review and investigation: the root cause of the cannula kink was determined to be wireless re-access which is not permitted with the use of impella cp. the pump was returned, and biomaterial was found under the impeller and in the purge gap. the pump was purged with a lab purge cassette and high purge pressure occurred. the purge cassette that was returned and was unable to be de-aired. the piston was unable to make forward strokes. the root cause of the purge cassette leak was determined to be a piston failure. the root cause of the pump stop was determined to be biomaterial ingress under the impeller. it is possible the ingress occurred during the purge issues but cannot be confirmed without the data logs. as noted in the impella ifu: impella cp with smart assist system. purge leak. section: cleaning : ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ pump stop. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
16645 | 1220648-2024-21071 | 20530680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20530680&pc=OZD | 01/14/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 01/14/2023 | 10/25/2024 |  |  | 11/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273420 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped and did not restart. procedural outcome noted the patient expired. no further information was provided. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of the patient's death is unknown. the investigation for the reported pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the pump stop and its relation to the patient outcome was not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
16646 | 1220648-2024-19777 | 20531980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20531980&pc=OZD | 02/17/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 02/17/2023 | 10/25/2024 |  |  | 12/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295684 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the purge pressure increased, and the purge flow rate decreased. reportedly, the purge cassette was replaced with no resolution. the decision was made to replace the impella. no patient harm was reported. | [5 empty] | The investigation into the high purge pressure issue has been completed. the impella pump was returned for investigation. no data logs were returned for this investigation. the pump was visually inspected, and biomaterial was found wrapped around the base of the impeller and in the purge gap. the root cause of the high purge pressure is most likely biomaterial. the confirmation of the timing of the ingestion could not be confirmed as no data logs were returned. this report is being filed as part of a retrospective review of historical records.
16647 | 1220648-2024-20583 | 20532011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532011&pc=OZD | 08/12/2024 | 05/23/2025 |  | 10/25/2024 |  |  | 08/12/2024 | 10/25/2024 |  |  | 11/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295828 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient had sepsis and gastrointestinal bleeding that led to hemorrhagic shock. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16648 | 1220648-2024-21072 | 20532189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532189&pc=OZD | 03/16/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 03/16/2023 | 10/25/2024 |  |  | 11/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024293917 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s laboratory results identified hemolysis with no pulses noted in the lower extremities, but bilateral tibials were present. a vascular study was ordered, and it was reported that the patient¿s left leg was amputated due to extracorporeal membrane oxygenation (ecmo) site and the patient received three (3) units of packed red blood cells (prbc's). the impella was in the right femoral artery and was not attributed to the amputation. the patient was successfully escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16649 | 1220648-2024-21073 | 20532202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532202&pc=OZD | 03/28/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 03/28/2023 | 10/25/2024 |  |  | 11/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024305208 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the pump abruptly stopped pumping. the impella device had been running on p-8 for 10 days at the time of the pump stop. the medical team in charge of the patient's care was aware of the possibility of potential pump stop and had been notified days prior to the pump stop to monitor the motor current. the patient had no destination therapy or advanced heart failure goals. the patient's family was not amenable to the withdrawal of the patient's care and the medical team planned on keeping the impella implanted until motor failure as there were no other options. the medical team would not offer to replace the device nor offer advanced therapies due to the patient's poor condition. the patient expired after the pump stopped and was unable to be restarted. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the long-term log shows that the pump stopped on day 10 of support. several attempts were made to re-start the pump and eventually it was restarted. the cause of the pump stop was not determined since product was not returned for investigation. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
16650 | 1220648-2024-21076 | 20532237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532237&pc=OZD | 08/25/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 09/14/2022 | 10/25/2024 |  |  | 01/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174131 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. it was noted that the patient developed a hematoma that required a blood transfusion. as a result, the emergency doctor exposed the subclavian artery and inserted the catheter using a 16f sheath via direct puncture. cardiac function was gradually recovering with the impella and extracorporeal membrane oxygenation support, but a cerebral hemorrhage occurred. additional information was provided that noted the patient expired due to multiple organ failure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16651 | 1220648-2024-21075 | 20532272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532272&pc=OZD | 01/24/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 01/24/2023 | 10/25/2024 |  |  | 07/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259439 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  |  | Death | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that the device's cannula kinked during cardio-pulmonary resuscitation (cpr). the device then exhibited low pump flow, the resolution for this issue is unknown. the device was explanted and replaced. additionally, it was reported that the patient expired. no further information was provided. | [5 empty] | The investigation for the reported cannula kink and low pump flow has been completed. the product was returned, and the cannula kink was confirmed. per the results of the clinical review and investigation: evidence of a cannula kink was found on the returned product. no other damage or abnormalities were noted. therefore, it was determined the cause of the low pump flow was a kinked cannula. this report is being filed as part of a retrospective review of historical records.
16652 | 1220648-2024-21074 | 20532278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532278&pc=OZD | 09/21/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 09/21/2022 | 10/25/2024 |  |  | 05/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218779 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support the patient had a pulmonary embolism. the pump was removed as a result. additional information was provided that noted the patient expired due to multiple organ failures. it was determined that the death was unrelated to the impella; however, it is unknown when the pulmonary embolism occurred. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
16653 | 1220648-2024-21080 | 20532449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532449&pc=OZD | 01/16/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 01/16/2023 | 10/25/2024 |  |  | 07/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254171 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, pump had low flows. under transesophageal echocardiogram (tee), it was noted that device was in proper position. the cardiologist evaluated the situation and determined the patient needed a new pump. a new pump was successfully placed with great output. however as soon as the patient got back to the unit, the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16654 | 1220648-2024-21077 | 20532452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532452&pc=OZD | 09/25/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 09/27/2022 | 10/25/2024 |  |  | 05/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223097 |  | 0048-0024-JP |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | Asian | Death | Yes | No | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with no left ventricle (lv) function and in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient's lv was "not beating" and thrombus developed which caused the pump to be removed and replaced due to flow issues. another pump was successfully placed; however, the patient expired two days later from multiple organ failure. it was reported that the thrombus may have been a contributing factor to the patient's demise. | [5 empty] | The investigation for the reported thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the thrombus was most likely patient condition related. there was left ventricular rupture prior to impella insertion, and the use of systemic heparin was stopped for bleeding control. clot observed inside the 14fr peel-away sheath that was detained. no additional treatment. no clot found on the pump body. the introducer was reported to have been flushed prior to insertion. day 1-2: heparin 10u/ml purge fluid peripheral heparin off on day 2, continuous peripheral heparin was started after half a day. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smartassist catheter insertion, section: axillary insertion of the impella cp with smartassist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
16655 | 1220648-2024-21079 | 20532515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532515&pc=OZD | 12/06/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 12/07/2022 | 10/25/2024 |  |  | 11/30/2023 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157852 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombocytopenia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after 1 day of support, the patient suffered a cerebral hemorrhage. additionally, patient experienced thrombocytopenia. no medical intervention was specified for these reported observations. when the pump was explanted, the team reported an automated impella controller (aic) red failure. the aic was removed for service as the instructions for use (ifu) was followed. post explant the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
16656 | 1220648-2024-21078 | 20532558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532558&pc=OZD | 09/25/2024 | 12/10/2024 |  | 10/25/2024 |  |  | 09/25/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025508597 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient was on impella cp support and during support, the patient had limb ischemia. the patient's impella side leg was cold and pale. no foot pulse was palpable and there was an increase in lactate. support continued. the patient was on impella support for 20.13 hours before the patient expired. the relationship between the impella and cause of death is unknown. medical safety has reviewed this complaint and determined there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21078.
16657 | 1220648-2024-21081 | 20532624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532624&pc=OZD | 09/26/2024 | 12/11/2024 |  | 10/25/2024 |  |  | 09/27/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025514290 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an unknown age and gender patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had a hemorrhagic stroke. heparin was withheld. patient expired on support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21081.
16658 | 1220648-2024-21088 | 20532719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532719&pc=OZD | 10/01/2024 | 12/06/2024 |  | 10/25/2024 |  |  | 10/01/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531362 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient was having episodes of ventricular tachycardia (vt). one episode of ventricular fibrillation (vf) during repositioning required shock and rhythm appears stable post repositioning. additionally, the patient had oral bleeding. heparin paused due to the bleed and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into vf/vt/af/at and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-21088. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16659 | 1220648-2024-21085 | 20532820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532820&pc=OZD | 09/27/2024 | 11/23/2024 |  | 10/25/2024 |  |  | 09/27/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474133 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella 5.5 support. while on support the medical doctor (md) noted a clot formation on inlet. the patient expired on support. per the md, either stents restenosis vs myocardial irritation caused the death. the md does not believe the impella was the cause of death. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-21085.
16660 | 1220648-2024-21082 | 20532821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532821&pc=OZD | 01/26/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 01/26/2023 | 10/25/2024 |  |  | 09/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023264272 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with unknow ethnicity was implanted with an impella cp for hemodynamic support. the patient experienced hemolysis the day after support began, as evident by color of urine, also noted to be related to the device. there was no suction at p-9, but hemolysis was strong; p-level lowered to p-6, but the hemolysis did not improve. the patient ultimately expired four days later. it is unknown if the death outcome is related to the impella device. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | With no product, data, or further clinical details, the cause of the hemolysis is unknown and the relationship between the device and hemolysis is unknown. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
16661 | 1220648-2024-21084 | 20532824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532824&pc=OZD | 02/09/2024 | 11/19/2024 |  | 10/25/2024 |  |  | 09/26/2024 | 10/25/2024 |  |  | 06/30/2025 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024361232 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured shock with an "unknown (undetermined)" relationship to the impella device used: patient was taken to the cath lab for urgent balloon aortic valvuloplasty (bav). patient decompensated, bav was aborted. impella cp was placed via the right groin, but patient went into ventricular fibrillation-arrest, shock x 5, down time 10-15 min, maxed dose pressor, echo bi-ventricular failure, patient transferred to cvicu post-cath lab. high dose pressors were given in cvicu, but patient was severely acidotic, and continued with high dose pressors overnight. impella cp support continued and patient was started on lasix drip for marginal urine output. patient was made dnr. remained on high dose pressors and inotropes, epi, levo, vaso are maxed. impella remains at p9, flows 4.2. the patient ultimately expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined.
16662 | 1220648-2024-21083 | 20532834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532834&pc=OZD | 09/24/2024 | 12/10/2024 |  | 10/25/2024 |  |  | 09/25/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524487 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported that during impella cp support for 74-year-old male patient, there were several suction alarms. a need for fluid was determined. fluid administration is running. partial thromboplastin time (ptt) around 30 before started with heparin. additionally, the patient became agitated and tore at the pump. pump was removed due to bad position. care was withdrawn and patient expired. medical safety review of the event noted there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the low pump flow and positioning issues has been completed. the device was not returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the positioning issue most likely was use related since the agitated patient pulled the pump out. the cause of the low pump flow issue cannot be determined since the device was not returned for analysis and data logs not available to review. corrections to the initial mfr report: d4-udi number. g1-mdr reporting contact email.
16663 | 1220648-2024-21086 | 20532838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532838&pc=OZD | 09/30/2024 | 11/26/2024 |  | 10/25/2024 |  |  | 09/30/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530643 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a bipella support ongoing for a patient admitted with multiple cardiac and systemic risk factors. the patient was being placed on impella cp while on support the patient had runs of ventricular tachycardia and required defibrillation. the family made choice to withdraw care and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-21086.
16664 | 1220648-2024-21087 | 20532844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532844&pc=OZD | 10/01/2024 | 11/24/2024 |  | 10/25/2024 |  |  | 10/01/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025539917 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 62-year-old female noted as high risk intervention was implanted with an impella cp device for mechanical circulatory support. on day 5 of the support the team observed an access site bleed that was treated by suture of the perclose and 20 minutes of pressure. the pump remains on for support currently at day 9 of support. | It was additionally reported that upon completion of support, the pump was explanted, but the patient lost distal pulses in their extremity. a thrombectomy was performed and large amounts of fibrous tissue were removed via fogarty balloon catheter. flow was restored following the intervention and the artery was repaired. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported thrombus and access site bleed has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause for the access site bleed. the root cause of access site bleeding was determined to be use issue because the bleeding was resolved by applying a suture at the site. as the necessary information was not provided, the root cause of the thrombus was not determined. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death.".
16665 | 1220648-2024-21091 | 20532849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532849&pc=OZD | 10/02/2024 | 11/20/2024 |  | 10/25/2024 |  |  | 10/02/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561577 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for support of a 66-year-old male patient admitted in an acute myocardial infarction / cardiogenic shock (ami/cgs). after a little over 1 hour of support t he pump was explanted. the pump had positioning alarms and an access site hematoma. the team was successful in replacing the pump. | [5 empty] | The investigation for the positioning issue and access site bleed has been completed. no product was returned. the root cause of the access site bleeding cannot be determined due to limited clinical information provided. there was no clinical information provided about the patients anatomy, or if the pump's position was verified with echo. due to the lack of information, the root cause of the positioning issue cannot be determined. corrections to the initial mfr report: h6 impact code 4627 changed to 4629. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16666 | 1220648-2024-21089 | 20532870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532870&pc=OZD | 10/01/2024 | 12/13/2024 |  | 10/25/2024 |  |  | 10/01/2024 | 10/25/2024 |  |  | 04/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023236577 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Cardiac Tamponade | [11 empty] | Required Intervention | Life Threatening | Disability | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Information from literature: title: a case of open-heart surgery with impella support for ventricular septal. perforation after myocardial infarction. extracorporeal circulation technology 2023; vol.50. no4,466. a male in his 70s experienced severe chest pain early in the morning and was rushed to the hospital. cag revealed a diagnosis of #6 ami and pci was performed. the patient was subsequently diagnosed with unstable hemodynamics and an echocardiogram was performed again. a swan-ganz catheter test revealed a qp/qs of 4.73 and a left-to-right shunt, leading to a diagnosis of vsp, and an impella cp placed and then on day 4, an impella 5.5 was inserted from the right subclavian artery. thereafter, qp/qs remained at 2.2-2.4. surgery was considered when the septal tissue had scarred, but hemodynamics was unstable, and surgery was performed on day 8. the surgery began of vsp closure (bovine pericardial patch) + pvi + left atrial appendage closure. the patient was able to wean from cardiopulmonary bypass without any major problems and returned to the icu. renal dysfunction was present before surgery, and the patient had no spontaneous urination, so chdf was initiated on pod 1. impella weaning was initiated on pod 5, but cardiac tamponade occurred, and hemodynamics could not be maintained, so re-thoracic hemostasis was performed on pod 7. hemodynamics stabilized thereafter, and the patient was weaned from impella 5.5 on pod 9. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ title: a case of open-heart surgery with impella support for ventricular septal perforation after myocardial infarction extracorporeal circulation technology 2023; vol.50. no4,466 | The investigation for the ventricle perforation and the renal failure has been completed. no product was returned and logs are not applicable. clinical details states that perforation causing fluid to build up around the heart. cardiac tamponade occurred after impella weaning was initiated after almost two weeks of support. rethoracic hemostasis was performed to maintain hemodynamics. no other clinical details are known about the cardiac tamponade. the cause of the cardiac tamponade was not determined since insufficient clinical details were provided and product was not returned. it was noted that the renal failure occurred sometime before the vsp closure surgery, but no other details are given. the cause of the renal failure was not determined since insufficient clinical details were provided and product/logs were not returned. corrections to the initial mfr report: g1 mdr reporting contact email.
16667 | 1220648-2024-21092 | 20532896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532896&pc=OZD | 10/04/2024 | 12/17/2024 |  | 10/25/2024 |  |  | 10/04/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533123 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp support ongoing for a patient admitted in acute myocardial infarction and cardiogenic shock. the pump support was placed via the femoral and that site had an access site bleed the day after implant. the bleed was treated by four units of blood and femstop placement, and the support continued. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report:g1 mdr reporting contact email.
16668 | 1220648-2024-21093 | 20532906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532906&pc=OZD | 10/04/2024 | 12/09/2024 |  | 10/25/2024 |  |  | 10/04/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501257 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support. the impella was placed and flows initiated on auto and turned to p8 with flows of 3.5lpm. while removing peel away sheath and placing repo sheath, suction alarm occurred. flows noted to be <1lpm. attempted to turn down p level and back up with no improvement in flows or uncoupling of waveforms. physician manipulated pump without resolution. physician opted to replace impella due to low pump flow. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the impella cp has been completed. the device was returned for evaluation. clinical details provided were sufficient to determine the root cause. the cause of the low pump flow issue was biomaterial ingestion. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21093: e4 should have been left blank. g1 reporting contact email.
16669 | 1220648-2024-21094 | 20532931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532931&pc=OZD | 10/07/2024 | 12/05/2024 |  | 10/25/2024 |  |  | 10/07/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474238 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and upon the impella explant, the pump came out smoothly, but the graft that was attached to the innominate artery was hanging off the end of the impella. the patient was bleeding and the patient's chest was opened. vascular surgery was performed. a massive transfusion was performed. the patient¿s outcome is unknown. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation into vascular damage - peripheral vessel has been completed. partial pump was returned for inspection with only the cannula. the rest of the pump was severed off - no red pump plug or patient cable was returned. there was no biomaterial in the cannula. the root cause of the vascular damage - peripheral vessel was not determined as insufficient clinical information was provided the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21094. e4 should have been left blank. g1 reporting contact email. g1 reporting contact fax number missing.
16670 | 1220648-2024-21090 | 20532956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532956&pc=OZD | 10/02/2024 | 12/13/2024 |  | 10/25/2024 |  |  | 10/02/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528819 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Female |  |  |  | Malfunction | No | Yes | Defective Device | Detachment of Device or Device Component | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant was placing a cp via the 14fr sheath for single access with 7fr pc. the 87-year-female patient had been admitted in for a high-risk coronary intervention (hrpci). the 14fr was leaking despite all best practices and so the sheath was replaced for shorter 14fr and pci sheath access made via the left femoral, and the impella remained on the right. the blood lost did not prompt any infusion back to the patient. the impella cp supported for over 2 hours and the pump was explanted. at the explant the pump fractured to 2 pieces when being pulled/removed from the sheath. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage has been completed. the impella cp and 14frx13cm 14frx25cm introducer sheaths were returned for evaluation. pump was returned with cannula detached from outflow cage. tearing and stretching was noted on both ends of separated cannula. delo bond on outflow cage remained fully intact. clinical details noted that pump got stuck in peel-away sheath when removing through peel-away sheath. it was noted that excessive force was applied when removing pump. the root cause of the product damage was most likely use related as evidence of excessive force was found on returned pump.
16671 | 1220648-2024-20215 | 20532989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532989&pc=OZD | 04/20/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 04/20/2023 | 10/25/2024 |  |  | 12/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298950 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Sense | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was blood in the foley bag. the clinician advised that the pump would likely need to be repositioned and also suggested lowering the pressure level as long as the patient could remain hemodynamically stable. additional information was not received indicating any intervention for the hemolysis. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for the reported hemolysis and placement signal issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the field long term log was reviewed and suction and positioning alarms were observed. the cause of the hemolysis was most likely improper positioning of the device. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
16672 | 1220648-2024-21095 | 20533062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20533062&pc=OZD | 09/12/2024 | 12/11/2024 |  | 10/25/2024 |  |  | 10/01/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508998 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 49-year-male patient developed an infection at their impella 5.5 access site during support. support continued with the implanted pump. additional information pertaining to potential treatment for the infection was not provided by the physician. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21095.
16673 | 1220648-2024-21096 | 20533232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20533232&pc=OZD | 09/25/2024 | 12/11/2024 |  | 10/25/2024 |  |  | 10/02/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025536165 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 60-year-old male noted as high risk intervention was implanted with an impella 5.5 device for mechanical circulatory support. during impella 5.5 support, there was leaking from near the filter of the purge side arm. there was suspected leaking from the yellow connector of the purge line, so the purge line was replaced. there was no improvement. there was no change in the purge flow or pressure, so staff decided to continue support and monitor. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the purge leak has been completed. the product was returned and evaluated. a purge cassette was connected to the pump and purge fluid started. initially, there was purge fluid sitting in retainer seal, but that was discovered to have been fluid from initial luer connection, not an active leak. the retainer was removed for closer inspection, and the pump was purged again for 10 minutes with no leaks found. log review showed no issues with purge flow or purge pressure on the day of reported leak or later. the root cause of the purge leak was most likely use-related. the instructions for use for the impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock states the following: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ corrections to the initial mfr report: h6 impact code 4629.
16674 | 1220648-2024-21097 | 20533465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20533465&pc=OZD | 09/26/2024 | 11/26/2024 |  | 10/25/2024 |  |  | 09/26/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539697 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Life Threatening | Death | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had escalated from a cp to the 5.5 pump for the support of a patient admitted in with cardiogenic shock and cardiomyopathy. while on the 5.5 for 7 days the patient developed heparin-induced thrombocytopenia. the pump remained on for 20 days and then care was withdrawn and the patient expired. the patient had developed multi-organ failure and was not a candidate for transplant. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21097. b2 death and date of death. g1 reporting contact email. h1 type of reportable event. h6 1802 added health effect - impact code.
16675 | 1220648-2024-21098 | 20533544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20533544&pc=OZD | 10/01/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 10/01/2022 | 10/25/2024 |  |  | 06/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248345 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the device exhibited low flows. the physician tried to reposition the device, the pump came across the valve, however the physician was concerned about possible clots and decided to replace the device. the original device was explanted and replaced with a new impella cp. it was noted that the patient was severely unstable prior to implanting the original pump. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported low flow and positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the low flow and positioning issue and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
16676 | 1220648-2024-21099 | 20533606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20533606&pc=OZD | 08/27/2022 | 10/07/2025 |  | 10/25/2024 |  |  | 10/12/2022 | 10/25/2024 |  |  | 04/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211695 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced a cerebral hemorrhage. it was reported that the patient experienced a cerebral hernia 11 days after the removal of the impella cp and died. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-21099. g1 reporting contact email revised on manufacturer device report 1220648-2024-21099 in accordance with updated procedures. h6 investigation conclusions code 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-21099.
16677 | 1220648-2024-21100 | 20533725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20533725&pc=OZD | 10/02/2024 | 09/22/2025 |  | 10/25/2024 |  |  | 10/02/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025517087 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a 53-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, purge flow gradually was decreasing from 9.4ml/hr to current 2ml/hr on the impella 5.5. high purge pressure alarms were occurring was well. purge cassette did not resolve the issue. tpa was initiated. it was further reported that the placement signal was not correlating to the patient¿s actual blood pressure. as the impella has been in used for over 30 days the sensor could potentially be failing. the staff will be monitoring motor current and flows. | [5 empty] | The investigation for the high purge pressure has been completed. neither data logs nor product were returned for investigation. clinical details report that tpa protocol was used to treat the high purge pressure, and within two days, the purge flows and pressures had returned to normal. that being said, since neither product nor data logs were available for investigation, the root cause of the high purge pressure could not be determined. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. neither data logs nor product were returned for investigation. the only clinical details provided suggest that the map on the console didn't match the number a blood pressure cuff was giving. the root cause of the optical signal issues could not be determined due to product and data logs not being returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. g1 updated with correct mdr reporting contact email. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-21100.
16678 | 1220648-2024-21101 | 20533773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20533773&pc=OZD | 10/21/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 10/21/2022 | 10/25/2024 |  |  | 06/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED. INC. |  |  | US |  | 2023229234 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that there was no abnormality in the impella operation, however the patient experienced severe mitral regurgitation due to papillary muscle rupture. the rupture was confirmed by echocardiography after the surgery. it was reported that the patient's circulatory status declined. the patient was managed by intubation, and a swan-ganz (s-g) catheter was also inserted and the patient was under systemic management. however, the patient died due to hypo cardiac function, including papillary muscle rupture. no further information was provided. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16679 | 1220648-2024-21104 | 20533919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20533919&pc=OZD | 12/19/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 12/19/2022 | 10/25/2024 |  |  | 09/30/2023 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022132400 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient died of cerebral hemorrhage. it was noted that the physician stated that the impella was not the direct cause of death because the patient had cardiac arrest and was in a bad condition. the physician also stated that there were no problems with the device. no further information was available. | [5 empty] | The investigation for the reported cerebral hemorrhage has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the cerebral hemorrhage and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16680 | 1220648-2024-21102 | 20533930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20533930&pc=PYX | 10/02/2024 | 11/24/2024 |  | 10/25/2024 |  |  | 10/02/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025507447 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 55-year-female patient had an impella rp placed and purge pressure low alarm occurred. purge cassette was changed, and alarm persisted. alarm went away when p-level was decreased. the surgeon is aware of the minimal flows with the impella. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation into high purge pressure has been completed. the impella device was not returned for evaluation. the data logs were reviewed and there were purge pressure low alarms seen. it it noted that the p-level was changed to try to troubleshoot these alarms but they persisted throughout the case. without the pump returned, it can not be confirmed if there are any leaks. the root cause of the purge leak - pump cannot be determined. high purge pressure is being updated to purge leak - pump as the customer explained of experiencing purge pressure low alarms.
16681 | 1220648-2024-21105 | 20533955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20533955&pc=OZD | 08/17/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 08/17/2022 | 10/25/2024 |  |  | 03/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023205761 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported upon initial activation of impella cp support, the pump had suction alarms with low flows and a dampened motor current. due to the low flows, the decision was made to remove the device and replace it with a second impella cp. upon explant, biomaterial was observed to be ingested into the pump. the patient's act was within the recommended therapeutic range. the patient was later noted to have expired on the second pump, several days later. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | No product was returned. data logs showed the "suction" alarm triggered almost immediately after starting the impella, with non-pulsatile motor current and low pump flow from the beginning. after reviewing the clinical information, it was determined that the cause of the low pump flow was ingested biomaterial. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
16682 | 1220648-2024-21106 | 20533975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20533975&pc=OZD | 11/18/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 11/18/2022 | 10/25/2024 |  |  | 05/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023218933 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient did not have an activated clotting time due to cardiogenic shock and the patient was unstable. the physician implanted the impella cp and the pump stopped. the physician decided to use an extra-corporeal membrane oxygenation device instead of the impella cp. it was reported that the patient survived explant of the device. no further information is available. | [5 empty] | The investigation for the reported pump stop has been completed. the product was returned. per the results of the clinical review and investigation: the cause of the issue was unable to be reproduced. device operated to specification. most likely an interaction with another device. evidence of scratches on the outflow cage. logs not returned. as noted in the impella ifu: impella cp with smart assist system section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16683 | 1220648-2024-21110 | 20534018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20534018&pc=OZD | 02/02/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 02/02/2023 | 10/25/2024 |  |  | 09/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272888 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported hemolysis during use of the device. attempts were made to reposition the device to no avail. the family withdrew care as the patient had an unrecoverable brain injury. the pump was not prematurely removed, nor any other intervention taken to resolve suspected hemolysis. the pump was ultimately removed after do not resuscitate and the patient expired with no allegations made toward the impella as the cause of death. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the field lt log was reviewed and pump position in the ventricle alarms and suction alarms were confirmed. ps trends were observed which correlate with the reported unrelated poor patient condition. the cause of the hemolysis was improper positioning of the device deep in the left ventricle. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
16684 | 1220648-2024-21107 | 20534091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20534091&pc=OZD | 09/26/2024 | 11/23/2024 |  | 10/25/2024 |  |  | 09/26/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025543789 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella 5.5 pump support initiated for a patient admitted for a coronary artery bypass graft and aortic valve repair and also supported by an impella rp. the patient had runs of ventricular tachycardia (vt)/ventricular fibrillation with p levels being higher. the vt prompted p level lowering and pump positioning review. the pump was appropriately placed within the left ventricle. the patient eventually expired after care was withdrawn. there is a potential relationship between the impella and cause of death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined.
16685 | 1220648-2024-21113 | 20534167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20534167&pc=OZD | 10/04/2024 | 12/05/2024 |  | 10/25/2024 |  |  | 10/04/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545050 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | Difficult to Insert | Migration or Expulsion of Device |  |  |  |  | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump delivered to support a patient admitted in with shortness of breath and atrial flutter. the cp was placed but was exhibiting signs of hemolysis on 2nd day of support. there was hematuria. the pump was repositioned and came out of the left ventricle. this prompted explant and pump replacement. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the impella cp has been completed. the device was not returned for evaluation. data logs were reviewed which showed pump ran without issue for about 24.5 hours. then pump position was in aortic area. pump then was stopped manually and disconnected from the console. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. corrections to the initial mfr report: g1 mdr reporting contact email h6 med dev prob code 1316 and 1395 changed to 1158.
16686 | 1220648-2024-21109 | 20534216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20534216&pc=OZD | 01/20/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 01/20/2023 | 10/25/2024 |  |  | 07/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258954 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. the complainant had a cp that failed delivery due to red lumen issues. the team was able to successfully place a new second pump. no harm or injury from delay in support initiation. the procedure outcome was patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. device not used because easy pass lumen removed clearly. root cause was use issue. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
16687 | 1220648-2024-21103 | 20534508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20534508&pc=OZD | 09/26/2024 | 11/24/2024 |  | 10/25/2024 |  |  | 09/26/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482263 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the peel away was exchanged for the repositioning sheath, a massive hematoma was noted. vascular surgery was called. however, the patient's wife did not want any heroic efforts taken. the wife requested the impella to be removed. the wife confirmed that they understood the patient would likely not survive without the impella. the patient was on impella support for 2.05 hours before the patient expired.there is a potential relationship between the impella and the cause of death. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
16688 | 1220648-2024-21119 | 20534681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20534681&pc=OZD | 04/19/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 04/19/2023 | 10/25/2024 |  |  | 12/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2024300704 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that there was difficulty implanting the device due to heaving kinking. however, the pump was implanted. the procedural outcome noted that the patient expired. no further information is available. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the delivery issue and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
16689 | 1220648-2024-21115 | 20534764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20534764&pc=OZD | 09/28/2024 | 11/26/2024 |  | 10/25/2024 |  |  | 09/28/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025514416 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Positioning Failure | Material Fragmentation | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and the impella cp became stuck under the papillary muscle. the impella cp was slightly pulled back and it would not budge. the impella cp was advanced and it kinked in the cannula. the doctor then pulled back to try to get the kink out and the inlet cage and pigtail popped off leaving the cannula, motor, and catheter attached. the impella cp removed with the inlet cage and pigtail left in the body. the patient was on impella support for 0.60 hours before the patient expired. there is a potential relationship between the impella and cause of death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the product damage and the positioning issues has been completed. the product was not returned. log review showed several suction alarms. the root cause of the positioning issues was most likely use-related since the impella was placed during cpr. clinical details state that the pigtail and inlet cage detached from the cannula while in the great vessel. the root cause of the cannula/pigtail detachment was not determined since product was not returned. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16690 | 1220648-2024-19934 | 20534934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20534934&pc=OZD | 06/26/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 06/26/2023 | 10/25/2024 |  |  | 01/31/2025 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312207 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a thrombus was observed on the inflow of impella. reportedly the thrombus did not cause any issues with impella flow or alarms, and the care team was aware. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the thrombus issue has been completed. per clinical details the thrombus did not cause any issues during the procedure with flows or alarms. the root cause of thrombus was most likely due to the patient¿s condition. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury. (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves as noted in the impella ifu ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ as noted in the impella ifu "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
16691 | 1220648-2024-21118 | 20534987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20534987&pc=OZD | 10/06/2024 | 12/06/2024 |  | 10/25/2024 |  |  | 10/06/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561574 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted in cardiogenic shock and with cardiomyopathy. the pump had in aorta alarms / optical signal issues and was explanted and replaced after 6 hours. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into positioning issues has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the positioning issue most likely was use related due to improper technique. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21118.
16692 | 1220648-2024-21122 | 20535032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20535032&pc=OZD | 10/06/2024 | 12/10/2024 |  | 10/25/2024 |  |  | 10/06/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493760 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed bleeding from the impella access site. fem stop was applied, and one unit of blood was administered. the patient is stable and survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16693 | 1220648-2024-21114 | 20535134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20535134&pc=OZD | 09/26/2024 | 11/24/2024 |  | 10/25/2024 |  |  | 09/26/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025539879 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp and while on support, there were ongoing suction alarms due to volume related issues. it is unknown if the suction alarms resolved before support was completed. the patient was on impella support for 64.31 hours before the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into low pump flow has been completed. the impella device was not returned for evaluation. the cause of the low pump flow issue was most likely patient condition related, based on data log review and given clinical details.
16694 | 1220648-2024-21121 | 20535168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20535168&pc=OZD | 09/25/2024 | 11/22/2024 |  | 10/25/2024 |  |  | 09/25/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522074 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and they had a right anterograde sheath placed along with left retrograde sheath for femorofemoral bypass due to loss in pulses to the right lower extremity. the right lower extremity was then dopplerable with a return in color and temperature. additionally, the patient had abdominal bleeding and massive transfusions were administered. the patient did not recover. care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the reported limb ischemia and vascular damage has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and vascular damage could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16695 | 1220648-2024-21125 | 20535226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20535226&pc=OZD | 10/06/2024 | 12/04/2024 |  | 10/25/2024 |  |  | 10/07/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1365450 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp support ongoing when there were purge alarms on the automated impella controller (aic). the aic was replaced and there was no harm or injury noted. | [5 empty] | The investigation for the console (aic) purge issue has been completed. the data logs were returned and confirmed numerous purge related alarms from the day of the occurrence including a purge system open alarm. the console was returned for evaluation. upon evaluation, the purge system open alarm was reproduced in the lab when a known good pump and purge system was run with the console. the alarm was persistent even after attempts to adjust the restrictor aperture. it was confirmed that the purge reading off of the console was significantly lower than that of a connected pressure meter at the purge pressure sensor. the purge pressure sensor was confirmed to be defective when a known good purge pressure sensor was swapped into the console and pressure meter reading matched console purge pressure readings. the root cause for the false purge pressure alarms was a defective purge pressure sensor. corrections to the initial mfr report: d9-device returned to mfr box checked. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
16696 | 1220648-2024-21117 | 20535841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20535841&pc=OZD | 10/06/2024 | 11/26/2024 |  | 10/25/2024 |  |  | 10/06/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484256 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted in acute myocardial infarction / cardiogenic shock. the pump was removed due to low pump flows after 23 minutes. the team had attempted to treat by infusion of fluids, but the pump needed explant. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. data logs were reviewed which confirmed low pump flow when pump started and remained to the rest of the case. product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely patient condition related.
16697 | 1220648-2024-21111 | 20535842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20535842&pc=OZD | 01/11/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 02/10/2023 | 10/25/2024 |  |  | 09/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278618 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient experienced distal limb ischemia. the doctor made the decision to place impella cp with small vasculature while waiting to go to the operation room for impella 5.5. an external bypass was placed to allow for distal limb perfusion. external bypass had been placed as a precaution, not as escalation. later on family decided to withdrew care and patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the limb ischemia was patient condition due to the patient's small vessels. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16698 | 1220648-2024-21120 | 20536392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20536392&pc=OZD | 09/25/2024 | 11/24/2024 |  | 10/25/2024 |  |  | 09/25/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025462087 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Ventricular Fibrillation | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support in cardiogenic shock. it was reported that while on support, hemolysis was suspected and continuous renal replacement therapy was started. the kidneys were not creating urine. therefore, four units of blood products were given to the patient. it is unknown if the impella cp was related to the hemolysis. furthermore, the patient remained on impella support for 274.75 hours before the patient had a ventricular fibrillation episode. the patient had a do not resuscitate order. therefore, the decision was made to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into hemolysis, renal failure, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. the cause of the hemolysis issue was most likely improper positioning of the device, based on given clinical details and data log review. the cause of the renal failure issue was most likely patient condition related. as the necessary information was not provided, the root cause of the vt/vf was not determined.
16699 | 1220648-2024-21108 | 20536475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20536475&pc=OZD | 12/06/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 12/06/2022 | 10/25/2024 |  |  | 08/31/2023 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022125948 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the doctor lost wire access into the left ventricle and when he tried to rewire it, he accidentally cracked the short 14 fr. peel away sheath. he was aware of what he did and held manual pressure and deployed the pre closed per close sutures. the doctor went and accessed the left side for implant with the 14 fr long sheath. no harm done to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. after reviewing the clinical information, it was determined that the cause of the introducer issue was most likely sheath kinking/damage due to improper technique. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
16700 | 1220648-2024-21128 | 20536721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20536721&pc=OZD | 06/15/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 06/15/2023 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1413761 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced purge disc/flag issues, which was resolved by changing the console. no patient harm was reported. | [5 empty] | The investigation has been completed. the console was returned for investigation. the logs were reviewed and confirmed that numerous purge pressure sensor alarms while purge cassette was being reported as inserted. visual inspection of purge assembly confirmed a broken/detached blue retainer. the cause of the issue was a defective component (broken purge disk retainer). this report is being filed as part of a retrospective review of historical records.
16701 | 1220648-2024-21127 | 20536811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20536811&pc=OZD | 06/14/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 06/14/2023 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1330939 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a failure to recognize the optical sensor on an impella cp with smartassist after being plugged in. the impella was placed in the body and was not able to display placement signal monitoring. the staff replaced the aic for a new one and issue was resolved. | [5 empty] | The investigation has been completed. the automated impella controller (aic) log and impella cp showed optical signal issue during the case start procedure which impacted calibration of the optical system. there was an unreliable placement signal alarm triggered due to an invalid placement signal and an optical sensor system error window, instructing to reconnect the impella at the pump connector. at this time, the placement signal momentarily came back before the pump was fully disconnected. despite disconnecting and reconnecting the impella pump, the unreliable placement signal could not be resolved from the customer's perspective, and the console was exchanged and there were no apparent placement signal issues. the console was tested on an engineering lab bench. the failure mode was not reproduced. upon arrival, the optical ferrule at the pump connector was inspected, and no significant contamination was found. the optical bench 5s parameter were all passing according to the functional check procedure. a known-good impella 5.5 (optical pump) was run over long-term, and no issues were observed. the optical signal issue was likely due to the pump connector not being fully connected. this report is being filed as part of a retrospective review of historical records.
16702 | 1220648-2024-21126 | 20536858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20536858&pc=OZD | 05/26/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 05/26/2023 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1438724 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced purge disc/flag issues. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console was returned. the logs showed the last case in the field before the console was sent back for investigation. the console received a purge disconnect error before shutdown was requested. the purge disc retainer was completely cracked and the purge flag was missing when the console was returned for investigation. it is likely that the failure had already started to manifest during the may 26 case wizard and then the purge retainer was completely ripped off during disconnect. once the retainer was ripped off, the purge flag could have easily fallen out. the cause of the issue was a broken purge disc retainer and missing purge flag. this report is being filed as part of a retrospective review of historical records.
16703 | 1220648-2024-21123 | 20537232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20537232&pc=OZD | 10/07/2024 | 11/21/2024 |  | 10/25/2024 |  |  | 10/07/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1739460 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the battery of an automated impella controller failed to charge despite the console being plugged in. this occurred outside of a clinical procedure. as such, there was no patient involvement. | [5 empty] | The investigation for the automated impella controller has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the reported issue was confirmed to be use related. this aic was booted while the battery transport switch was disengaged. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
16704 | 1220648-2024-21132 | 20537337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20537337&pc=OZD | 05/23/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 05/23/2023 | 10/25/2024 |  |  | 04/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228027 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported, a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the impella cp failed to start and an automated impella controller (aic) change did not resolve the issue. the clinical team mentioned that a "red line" remained in the cannula. the patient was in a heavy critical situation and the medical team decided to stop the resuscitation and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the reported pump did not start, issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that no logs were returned for investigation. the returned pump was investigated. and no red lumen was observed in the pump. no other abnormalities were observed. the pump was run in a bucket and ran within specification. the cause of the pump not starting was most likely use related and related to not removing the red lumen prior to starting the pump. this report is being filed as part of a retrospective review of historical records.
16705 | 1220648-2024-21116 | 20537877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20537877&pc=OZD | 02/08/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 02/13/2023 | 10/25/2024 |  |  | 01/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178701 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the physician reported relevant hemolysis during the first 24 hours on impella support. therapy was withdrawn in consultation with relatives. the patient expired with no allegations made toward the impella as the cause of death. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. data analysis: data log showed overall pump ran without issue during the case but pump was not in the proper position (aortic area) in the first 24 hours that triggered alarm 128 (impella position in aorta). the root cause of hemolysis was most likely due to pump was in improper position. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
16706 | 1220648-2024-21145 | 20538096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20538096&pc=OZD | 09/23/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 09/23/2022 | 10/25/2024 |  |  | 06/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238840 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that during placement of the impella the physician was unable to cross the aortic valve, this was attributed to the condition of the patient's vasculature. the device was removed and a new impella placed. additional information noted that the patient died in the catheterization laboratory. no additional information is available. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. no product was returned. the cause of the delivery issue was patient condition related and due to the tortuosity of the vessels. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
16707 | 1220648-2024-21124 | 20538208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20538208&pc=OZD | 09/29/2024 | 11/23/2024 |  | 10/25/2024 |  |  | 09/29/2024 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532993 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp because the patient was rushed to the catheterization lab for a percutaneous coronary intervention and an intra-aortic balloon pump (iabp) was placed. the iabp was causing loss of blood flow to the leg which prompted the decision to place the impella cp and remove the iabp. while on impella cp support, there was no distal pulse present in any extremities. the physician readdressed the peel away sheath the following day but the left upper extremity was noticeably cooler than the right with intermittent pulses. it was further reported that the family decided to withdrew care and the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
16708 | 1220648-2024-21112 | 20538317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20538317&pc=OZD | 01/17/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 02/09/2023 | 10/25/2024 |  |  | 10/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275240 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced right limb ischemia. it was also reported that the patient expired while on support due to fulminant cardiogenic stock after cardiopulmonary resuscitation. no allegations were made towards the impella cp as the cause of the death. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined due to insufficient clinical details. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16709 | 1220648-2024-21131 | 20538430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20538430&pc=OZD | 02/24/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 02/24/2023 | 10/25/2024 |  |  | 08/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270058 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Death | No | Yes | Deformation Due to Compressive Stress | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. the complainant reported that the catheter of their pump was found to be kinked during patient support. the pump was subsequently explanted roughly an hour after implanting due to the noted kink. an intra-aortic balloon pump (iabp) was inserted to continue with patient support. the procedural outcome noted that the patient expired three hours after admission to the unit. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. no product was returned. the cause of the delivery issue was most likely patient condition related as there was resistance crossing the aortic valve resulting in a catheter kink. this report is being filed as part of a retrospective review of historical records.
16710 | 1220648-2024-21142 | 20538749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20538749&pc=OZD | 10/11/2021 | 10/25/2024 |  | 10/25/2024 |  |  | 10/11/2021 | 10/25/2024 |  |  | 05/31/2023 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022094945 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Myocardial Infarction | Brain Injury | Heart Failure/Congestive Heart Failure | [10 empty] | Life Threatening | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient died due to myocardial infarction, heart failure, and hypoxic encephalopathy. it was reported that the impella cp was operating normally. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16711 | 1220648-2024-21136 | 20539127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539127&pc=OZD | 08/29/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 08/29/2022 | 10/25/2024 |  |  | 04/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED. INC. |  |  | US |  | 2023209990 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Unknown |  |  | Asian | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, it was noted that there was a gap in the statlock and suture pad of the repositioning sheath. bleeding was observed from the gap. the resolution for this issue is unknown. survival outcome at explant noted the patient expired, cardiac function was not recovered. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported access site bleeding and product damage has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16712 | 1220648-2024-21129 | 20539201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539201&pc=OZD | 05/07/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 05/07/2023 | 10/25/2024 |  |  | 11/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2024314625 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited suction alarms indicating low flow. the physician attempted to reposition the device several times, the pump was verified to be in the correct position, and the device was still exhibiting suction alarms. it was reported that after repositioning the heart has less arrhythmia and changes were shallower. the patient went into additionally, the patient experienced limb ischemia. no distal flow to the leg on impella side. an antegrade sheath was placed on the impella side for an external bypass. better flows to distal leg were obtained. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The investigation for the reported limb ischemia and low flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the limb ischemia and low flow were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16713 | 1220648-2024-21141 | 20539221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539221&pc=OZD | 05/29/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 05/29/2023 | 10/25/2024 |  |  | 01/31/2025 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMEPLLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2024307817 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the device exhibited suction alarms indicating low flow. the device would not go above p-1 without alarming. the device was explanted and replaced with an impella rp. the procedural outcome stated that the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The investigation for the low flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the data logs were reviewed, several suction alarms were found during the entire case at higher p-levels and the suction alarms seem to be resolving after p-level dropped to p-2. the clinical details provided were sufficient to determine a cause. the cause of the low pump flow issue was most likely patient condition related with several suction alarms observed at higher p-levels and p-level dropping to p-2 resolving the suction event per log and clinical review. this report is being filed as part of a retrospective review of historical records.
16714 | 1220648-2024-21140 | 20539288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539288&pc=OZD | 04/20/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 04/20/2023 | 10/25/2024 |  |  | 11/30/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2024300847 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 48 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced limb ischemia. no distal perfusion noted on impella cp side. the pump put in emergent setting; and no preliminary vessel evaluation was performed. the patient was reported to be very unstable, so no further intervention was performed. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the limb ischemia and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16715 | 1220648-2024-21144 | 20539297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539297&pc=OZD | 03/15/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 03/15/2022 | 10/25/2024 |  |  | 12/31/2023 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  |  |  | 2023165749 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Myocardial Contusion | Myocardial Infarction | Brain Injury | [34 empty] | A |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that the procedure was successful, and the device worked as intended. however, the patient later died due to myocardial damage and hypoxic encephalopathy due to cardiac arrest due to myocardial infarction. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | D6: device implant and explant date is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16716 | 1220648-2024-21134 | 20539433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539433&pc=OZD | 05/17/2023 | 05/30/2025 |  | 10/25/2024 |  |  | 05/17/2023 | 10/25/2024 |  |  | 01/31/2025 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2024326972 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be implanted with an impella cp device for mechanical circulatory support. it was reported that during delivery, the physician was unable to pass the impella cannula through the patient's aortic valve. the physician made several attempts; however, this was not completed due to heavy calcification. the implant was aborted. the procedural outcome noted that the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the delivery issue and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records. | H6: health effect - clinical code changed from 4445 cardiovascular dysfunction / insufficiency to 4582 no clinical signs, symptoms or conditions health effect - impact code changed from 1802 death to 2199 no apparent harm occurred in relation to the adverse event.
16717 | 1220648-2024-21139 | 20539445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539445&pc=OZD | 04/20/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 04/20/2023 | 10/25/2024 |  |  | 10/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269476 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient originally had multiple organ failure prior to impella insertion. along with the progression of multiple organ failure, renal failure also progressed. while on support, the patient and his family did not want dialysis, and the patient expired of renal failure. the impella's aic worked well until the patient expired. use of the device cannot conclusively be associated with the outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16718 | 1220648-2024-21133 | 20539488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539488&pc=OZD | 06/03/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 06/03/2023 | 10/25/2024 |  |  | 02/28/2025 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2024321024 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Injury | No | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited low flow, volume was administered to the patient. the position of the device was confirmed to be acceptable, however, to try and resolve this issue the physician decided to try and reposition the device. no improvement in flows, so the physician decided to explant the impella cp. additional information noted that the procedure was not successful, and the patient died. | [5 empty] | The investigation for the reported low flow issue has been completed. the low flow complaint was confirmed. a review of the device logs found: data showed about 30 minutes into the case, the flow decreased gradually to 1.5 l/min that triggered auto suction response and suction alarms. the low flow remained to the rest of the case. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the hemolysis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
16719 | 1220648-2024-21148 | 20539603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539603&pc=OZD | 04/14/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 04/14/2023 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1398838 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. this aic was replaced with a backup aic. there was no patient harm reported. | [5 empty] | The investigation into the issue has been completed. the console was returned for investigation. the console logs from the reported day of event were consistent with the complaint and showed that immediately after boot-up, the console presented with controller error alarm due to defective opm lamp. this was preceded by log messages showing very low value of ¿light¿ parameter (0.53 volt), indicating no light detected by the opm optoelectronics. issue was not resolved after power-cycling the console, but an attempt was made to use the pump anyway. however during the case start, the console kept showing error message ¿optical sensor system error¿. the pump was disconnected and automated impella controller power-cycled but the issue remained unresolved. the returned console was tested. initial troubleshooting indicated no light coming out of optical pump connector fiber, which was consistent with reproduced alarm. measured voltage (4.9 volt) across lamp assembly indicated that the electronics were providing sufficient power to the lamp. due to the unavailability of spare components (lamp assembly), the entire optical bench was replaced and the issue was resolved. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16720 | 1220648-2024-21130 | 20539614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539614&pc=OZD | 02/23/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 02/23/2023 | 10/25/2024 |  |  | 08/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267720 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, and post sheath removal, the patient experienced bleeding at the right femoral artery insertion site. the staff placed a femoral stop to resolve the issue. additionally, the patient experienced hemolysis, the resolution for this issue is unknown. survival outcome at explant noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were sufficient to determine a root cause. root cause is most likely the lack of vessel recoil due to the patient condition and left anterior descending (coronary artery branch) vessels treated. the investigation for access site bleeding is in process. instructions for use for the related event are as follows: hemolysis impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
16721 | 1220648-2024-21143 | 20539622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539622&pc=OZD | 03/07/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 03/07/2022 | 10/25/2024 |  |  | 11/30/2023 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  |  |  | 2022157854 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced limb ischemia, the resolution for this issue is unknown. the procedure was completed successfully with the support of the impella cp, the device was working normally. the patient was transferred to the intensive care unit and died of asystole. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16722 | 1220648-2024-21149 | 20539679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539679&pc=OZD | 05/01/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 05/11/2023 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1286944 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. this aic was replaced with a backup aic. there was no patient harm reported. | [5 empty] | The investigation into the issue has been completed. the impella automated controller was returned for investigation. the console logs from the reported day of event were consistent with the complaint and showed a controller error alarm occurring shortly after the pump was started. the real time log data indicated two instances of loss of data from opm. this was determined to have been related to a known pattern failure caused by a temporary loss of optical placement signal. the cause of the issue was not determined since the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16723 | 1220648-2024-21138 | 20539701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539701&pc=OZD | 09/21/2022 | 10/25/2024 |  | 10/25/2024 |  |  | 09/22/2022 | 10/25/2024 |  |  | 01/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176701 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the pump was not working properly and was successfully switched to another device. the survival outcome at explant noted the patient expired. use of the device cannot conclusively be associated with the outcome. no patient harm related to the device was identified. | [5 empty] | The investigation for the reported pump did not start issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the product was visually inspected, and no abnormalities were observed. cannula was visually inspected, and no kink was observed. no kink was observed using the borescope. the pump was then connected to the aic and ran at p-9. flow values were within normal limits. logs not available. clinical details showed patient was on inotropes/vasopressors 1 hour before support. the root cause of the issue was patient condition related. this report is being filed as part of a retrospective review of historical records.
16724 | 1220648-2024-21135 | 20539722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539722&pc=OZD | 06/08/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 06/08/2023 | 10/25/2024 |  |  | 08/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED. INC. |  |  | US |  | 2023257087 |  | 0048-0024-JP |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Death | No | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the device exhibited low flow. no problem was reported with the impella device; however, it was noted that the patient's organ index deteriorated. the physician decided to explant the device. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the investigation for the reported low flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the low flow and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
16725 | 1220648-2024-21146 | 20539849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539849&pc=OZD | 02/27/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 02/27/2023 | 10/25/2024 |  |  | 07/31/2024 |  | 10/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249983 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, they were unable to obtain distal pulses on both lower extremities. it appeared to be pointed to mitral valve leaflet, and an apical view was done, and the mitral valve was moving freely. no alarms on aic, good waveforms, and the decision was made to leave the impella in the current position. the patient was on multiple pressors. distal pulses were unable to be found in the coronary care unit (ccu) in both lower extremities. the patient was made a dnr the next day by the family and they withdrew care after 13.9 hours. the survival outcome at explant noted the patient expired. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The investigation for the reported limb ischemia issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the limb ischemia issue was most likely due to patient condition related with both legs being ischemic and unrelated to the impella. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16726 | 1220648-2024-21150 | 20539932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539932&pc=OZD | 05/19/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 05/19/2023 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161391 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. this aic was replaced with a backup aic and the procedure was successfully completed. there was no patient harm reported. | [5 empty] | The investigation into the issue has been completed. the console was returned for investigation. the console logs from the reported day of event are consistent with the complaint and showed controller error alarm followed by the inability to shut down the console, despite the pump being disconnected (automated impella controller software does not allow the end user to shut down if the pump is connected and, in this case, the pump was erroneously detected as connected). this issue is a known software bug (jama: gid-dft-2927090 / rasd-dft-166) that manifests in some rare scenarios when the pump is being disconnected and reconnected at certain times during case start (and often out-of-synch with the case start wizard prompts). the cause of the aic issue was a software issue. this report is being filed as part of a retrospective review of historical records.
16727 | 1220648-2024-21147 | 20539953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539953&pc=OZD | 04/03/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 04/03/2023 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1668177 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error alarm. it was reported that the white cable was secure and a backup aic was used to successfully complete the procedure. there was no patient harm reported and the patient survived the procedure. | [5 empty] | The investigation into the issue has been completed. the console was returned for investigation. the data logs aligned with the pattern failure mode, transient loss of optical data. the automated impella controller (aic) logs confirmed the reported failure mode given there was a controller error alarm triggered during the clinical procedure. he logs also reported process sampling data from optical bench: sent stop acquisition cmd ack, indicating ossiopsens receive thread gets a data sampling packet with an unexpected length. the real time logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, signal-to noise ratio (snr), contrast, lamp, gain) are lastly reported as 16384 values and then no more samples were recorded. the absence of snr samples impacted the console¿s ability to recognize the pump disconnection and required the user to perform a hard shutdown. the returned console was tested on an engineering lab bench. the failure mode was not reproduced while running an optical pump simulator. there was no issue with the console hardware. the log entry process sampling data from optical bench: sent stop acquisition cmd ack indicated the sent stop acquisition cmdflag was set when entering ossi_sampling_stopped_state which eventually led to a false communication stopped condition. the root cause of the controller error was due to an aic software issue. this report is being filed as part of a retrospective review of historical records.
16728 | 1220648-2024-21151 | 20539972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539972&pc=OZD | 05/26/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 05/26/2023 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1234572 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. this aic was replaced with a backup aic and the procedure was successfully completed. there was no patient harm reported. | [5 empty] | The investigation into the issue has been completed. the console was returned for investigation. the logs were consistent with the complaint. history logs were reviewed (since last console service), the error (loss of communication with optical bench when pump was connected) occurred and was followed by a controller error. the real time logs indicated that the loss of all ops suddenly. upon console bootup, the failure mode controller error alarm was reproduced without the pump connected. inspection of the console internal cable connection revealed a loose/misaligned ribbon cable from optical bench to 4-in-1 carrier, which led to the loss of communication condition and alarms. it was concluded from the data and device analysis, the ribbon cable was plugged in (potentially not fully since last console service), and the connection was loosened through time or as a result of console movement. the root cause of the controller error was due to a loose ribbon cable from the optical bench. this report is being filed as part of a retrospective review of historical records.
16729 | 1220648-2024-21153 | 20539982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20539982&pc=OZD | 06/01/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 06/01/2023 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1112872 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) experienced a controller error. there was no patient harm reported. | [5 empty] | The investigation into the controller error has been completed. the console was returned for investigation. the logs ended at on (b)(6) 2023 although the complaint was reported to have occurred on (b)(6) 2023. all recent available logs were viewed but none of them corresponded to the reported controller error. however, the long-term power data revealed the batteries were depleted given all battery cells we around 2.8 volt. the returned console was unable to boot up on battery power. on alternating current (ac) power, the console booted with a battery failure alarm. the batt test result indicated that both batteries were not charging despite the ac was plugged in given they were depleted and internally locked out. the investigator temporarily replaced both batteries with known good ones to ensure there was no issue with the power battery manager (pbm). the battery failure was likely due to depleted batteries. this report is being filed as part of a retrospective review of historical records.
16730 | 1220648-2024-21154 | 20540045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540045&pc=OZD | 06/04/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 06/05/2023 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1178828 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the automated impella controller (aic) had an unexpected shut down alarm. the pump was shut off and impella was stopped. the impella was switched to a new aic. there was no harm to the patient. | [5 empty] | The investigation into the power issue has been completed. the console was returned for investigation. the logs are consistent with the complaint. the console and pump were running for a few days until there was an abrupt shutdown and reboot. last time stamp on console logs were at 14:58:22, june 4th, and in no more than 90 seconds the console rebooted. later, the console issued the unexpected shutdown alarm and the pump was able to be recovered to previous running p-level. the real time logs indicated that the last sampling was 14:59:47, june 4th, which means the console was not powered down at least until this moment. rlm journal was reviewed, after synchronizing with the logs. it was found that the rlm was data streaming without any issue like restart or reboot. the console was powered on with alternating current for days. the failure mode, unexpected console reboot, was not reproduced. the console internal cable connections were inspected.no significant misalignment or cable degeneration was found, except a slight bent ribbon cable between impellatronics board and 4-in-1 (not related to the failure mode). since rlm power/communication was not interrupted, the power supplied to rlm (4-in-1 and power battery manager) was stable and excluded from the root cause of the console reboot. also, no signature of software crashed was observed from the logs. the console reboot was not a behavior to mitigate software processes abnormalities. the root cause of the unexpected controller shutdown and reboot was unable to determine because the failure mode was not reproduced. this report is being filed as part of a retrospective review of historical records.
16731 | 1220648-2024-21155 | 20540051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540051&pc=OZD | 06/07/2023 | 10/25/2024 |  | 10/25/2024 |  |  | 06/07/2023 | 10/25/2024 |  |  |  |  | 10/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1330532 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing aortic valvuloplasty was implanted with an impella device for mechanical circulatory support. prior to support, the automated impella controller (aic) experienced priming issues that were not resolved by multiple resets and replacing the purge line. the team then replaced the console to continue the procedure. there was no reported patient injury. | [5 empty] | The device was returned and evaluated, and the investigation for the reported purge disc/flag issues has been completed. the device logs from the complaint date revealed that the console issued the purge disc not detected alarm several times through a long time. re-insertion of purge cassette/purge disc was observed, yet the console still had a problem in detecting the purge disc. the console was evaluated and the issue with properly detecting the purge pressure disc was reproduced, and the purge flag was confirmed to be broken. once broken it does not react when the purge disc is attached (missing at blue disc retainer). without purge disc detected, the impella priming could not proceed. the cause of the issue was broken purge flag. this report is being filed as part of a retrospective review of historical records.
16732 | 1220648-2024-20962 | 20522126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522126&pc=OZD | 10/01/2024 | 12/12/2024 |  | 10/24/2024 |  |  | 10/01/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535206 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella cp support. it was reported that the purge filter is leaking. | [5 empty] | The investigation into purge leak ¿ pump has been completed. since data logs show the product operated within specification during the case and no leak reproduced on returned product, the root cause of the purge leak was determined to be no problem found. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-20962 | The impella device was received for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
16733 | 1220648-2024-20963 | 20522374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522374&pc=OZD | 06/05/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 10/16/2024 | 10/24/2024 |  |  | 02/29/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023203116 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there were lower than expected flows and a clot was noted. it was reported that flows were beginning to drop and suction alarms were occurring above p4. the patient had reduced cardiac function resulting in 2 minutes of cpr, bolus of epi and 1 amp of bicarb. return of spontaneous circulation was achieved. the physician did not want to replace the pump due to the patient's overall status at the time. plans were to potentially replace the pump in the early am if patient stabilized hemodynamically. despite the impella providing 2/lmin, the patient¿s labs declined overnight. the patient¿s family withdraw care and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, information provided was sufficient to determine a root cause. a clot was visualized under echo upon arrival to cvicu. due to clot formation around the cannula the impella was unable to provide maximum support. the root cause of the reported issue was biomaterial. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
16734 | 1220648-2024-20964 | 20522383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522383&pc=OZD | 10/01/2024 | 11/23/2024 |  | 10/24/2024 |  |  | 10/01/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458623 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 77-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, high purge pressure system blocked alarms occurred. purge cassette was changed but it did not resolve the issue. tpa was administered successfully. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. data logs were reviewed and found that there was a slight spike in motor current and a decrease in flows about 10 hours before the purge pressure began to rise. the purge pressure rose for about an hour before triggering high purge pressure alarms at around 1100mmhg. high purge pressure lasted for about 5 hours. there were no bag or cassette changes within a few hours of purge pressure rising. it was stated tissue plasminogen activator (tpa) was added and it resolved this issue. based on this information, the root cause of the high purge pressure is most likely due to biomaterial. added- h6 impact code 4644.
16735 | 1220648-2024-20968 | 20522447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522447&pc=OZD | 10/10/2021 | 10/24/2024 |  | 10/24/2024 |  |  | 10/10/2021 | 10/24/2024 |  |  | 10/31/2022 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021003905 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the ps alarmed and there was bleeding. the pump remained on for support despite this and the pump position was evaluated. also during support, the patient was on integerlin drip, however was bleeding all over, bilateral fem stops placed and off all heparin. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16736 | 1220648-2024-20965 | 20522462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522462&pc=OZD | 10/01/2024 | 03/06/2025 |  | 10/24/2024 |  |  | 10/01/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443371 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 34 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a 34-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during impella 5.5 support, the patient was taken back to the operating room for exploration of the pocket. additionally, the patient developed hemolysis. the device was explanted due to the hemolysis. | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured recurrent ventricular tachycardia with a "definite" relationship to the impella device used: ¿vt runs that required amiodarone. eventually started on dobutamine and vt resolved.¿. the patient recovered with no sequelae. additionally the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute left subcortical stroke with a "possible" relationship to the impella device used: ¿patient began with right arm weakness after sedation was stopped. neurology consulted. started on asa for secondary stroke prevention.¿. the patient recovered with sequelae. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | B5 revised to reflect additional information received from the clinical site. h6 revised to reflect the additional information. instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g2 report source; study was missing from manufacturer device report | The investigation for the hemolysis has been completed. data logs were reviewed finding no motor current (mc) spikes outside of p- level changes observed. no positioning issue alarms. no suction alarms. no indication observed to determine the root cause for the reported issue. the cause of the hemolysis issue cannot be determined since the complaint device not returned for analysis and insufficient clinical data provided.
16737 | 1220648-2024-20969 | 20522506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522506&pc=OZD | 10/02/2024 | 11/29/2024 |  | 10/24/2024 |  |  | 10/02/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539065 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during planned explant of an impella 5.5 device. during removal, it was documented that the implanted pump got stuck at the anastomosis and could not be explanted. due to the tension felt when pulling on the pump, the patient was taken to the operating room and their chest was opened for pump removal. the graft was ligated, and the chest was closed following a successful procedure. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into removal issues has been completed. the impella device was not returned for evaluation. the root cause of the removal issues is most likely patient condition related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20969. e4 should have been left blank. g1 reporting contact fax number missing.
16738 | 1220648-2024-20971 | 20522513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522513&pc=OZD | 10/03/2024 | 07/08/2025 |  | 10/24/2024 |  |  | 10/03/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025533300 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported the patient was on impella 5.5 support. while on support, there were decreased flows, abnormal waveform, and suspected clot per physician. the impella was explanted and the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Unwittingly, product-specific information such as lot/serial numbers, expiration/manufacturing dates, and unique device identifier number was erroneously captured on the initial medwatch report. therefore, fields in sections d4 and h4 have been updated to correctly reflect the information of complaint device and device investigated. | The investigation for the low pump flow has been completed. data logs confirmed the failure mode and clinical narrative. logs showed pump ran without issue for 190 hours, then motor current (mc) spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. corrections to the initial mfr report: h6 impact code 4629 changed to 4627. | Impella 5.5 was returned for evaluation. low pump flow: upon visual inspection, a large biomaterial was present at outflow cage. some of the biomaterial was removed and the pump was run in a bucket overnight and low pump flows reproduced. data logs were reviewed and lt logs showed pump ran without issue for 190 hours, with some suction alarms occurring. rt log of end of case showed low flows at p-8 and pump turned onto surgical mode and. no ingestion signatures observed. clinical details noted that decreased flows and abnormal waveforms were observed and pump was explanted. the root cause of the low pump flow could not be definitively determined per data log did not show definitive evidence of ingestion of biomaterial causing low pump flows. thrombus: biomaterial observed on pump explant. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
16739 | 1220648-2024-20970 | 20522516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522516&pc=OZD | 10/03/2024 | 12/04/2024 |  | 10/24/2024 |  |  | 10/03/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532360 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and there was bleeding at the access site around the repositioning sheath. vascular surgery team performed femoral cutdown and applied purse string suture. team noted calcification near impella access site. four units of red blood cells were given to the patient and two units of fresh frozen plasma were provided. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient condition related, calcification observed near impella access site as per the clinical description provided.
16740 | 1220648-2024-20973 | 20522520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522520&pc=OZD | 12/03/2021 | 10/24/2024 |  | 10/24/2024 |  |  | 11/12/2021 | 10/24/2024 |  |  | 02/28/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022043411 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the medical staff had difficult controlling site bleeding that occurred at the sheath. the patient¿s status was ¿do not attempt resuscitation¿, therefore no additional treatment was provided to manage this complication. the device was explanted, and the procedural outcome was that the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16741 | 1220648-2024-20966 | 20522625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522625&pc=OZD | 09/30/2024 | 12/03/2024 |  | 10/24/2024 |  |  | 09/30/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561578 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the patient was on impella cp support due to atrial fibrillation. while on support, purge pressure alarms were noted. purge tubing was changed, however, alarms continued. the automated impella controller (aic) was changed as well. physician made the decision to remove and replace pump. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported issue has been completed. the impella device was received from the customer and returned device was visually inspected. biomaterial was observed around the impella and purge gap. no cannula kink observed. no broken blade found. impella purged to perform the reported low purge pressure issue. leaking observed in catheter hole right before line 95. the cause of the purge leak was damaged purge tubing with catheter punctured at 95 cm marking, caused by improper technique. the cause of the optical signal issue was a damaged optical fiber line from the patient cable side inside red plug. this component failure is attributed to the manufacturing process (vendor). | Corrections to mdr mfr report # 1220648-2024-20966-01: h1-type of reportable event incorrectly stated death but should have stated malfunction.
16742 | 1220648-2024-20976 | 20522630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522630&pc=OZD | 09/25/2024 | 12/02/2024 |  | 10/24/2024 |  |  | 10/03/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528301 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella 5.5 removal, skin infection at impella tunnel site was noted. the patient is stable on wound vac and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported limb ischemia and sepsis has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and sepsis could not be determined as the device was not returned nor sufficient clinical details.
16743 | 1220648-2024-20977 | 20522637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522637&pc=OZD | 10/03/2024 | 12/04/2024 |  | 10/24/2024 |  |  | 10/03/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546058 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, there was a loss of distal pulses via impella right lower extremity (rle). the physician started an antegrade arterial line on affected impela rle. perfusion was restored and pulses were regained via doppler. external fem-fem bypass was applied the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
16744 | 1220648-2024-20974 | 20522648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522648&pc=OZD | 10/03/2024 | 11/26/2024 |  | 10/24/2024 |  |  | 10/03/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467936 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 had high purge pressure alarm. no kinks were found. the physician removed the impella successfully at bedside. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. the pump was not returned for investigation. the data logs show a 3-minute gradual rise in purge pressure with motor current spike and purge system blocked alarm while the pump was running at a steady p level. the pump flow became saturated after the purge pressure event. the pump was expanded after 1 hour of running, following a check of the purge line kick in the field. also, patient received blood products and feiba (promote clot formation and reduce blood loss) due to bleeding from surgical site. the cause of the high purge pressure issue was most likely biomaterial ingestion. the cause of the high purge pressure issue was most likely biomaterial ingestion, based on data log review.
16745 | 1220648-2024-20979 | 20522657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522657&pc=OZD | 03/17/2023 | 05/29/2025 |  | 10/24/2024 |  |  | 03/22/2023 | 10/24/2024 |  |  | 10/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269485 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient expired due to hypoxic encephalopathy. the physician noted that the impella worked fine until the end. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16746 | 1220648-2024-20967 | 20522661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522661&pc=OZD | 10/01/2024 | 12/12/2024 |  | 10/24/2024 |  |  | 10/01/2024 | 10/24/2024 |  |  | 04/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213814 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had an impella cp support ongoing in 2022 for a patient of unknown age and gender admitted in in an acute myocardial infarction / cardiogenic shock (ami/cgs). the pump was supporting for just under 3 days when it was explanted and replaced by the 5.5 pump. later in 2024 the medical team published upon the cp support and the 5.5 support. the publication noted the cp was removed due to concerns of hemolysis. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. product and data logs were not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. corrections to the initial mfr report: g1 mdr reporting contact email.
16747 | 1220648-2024-20972 | 20522674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522674&pc=OZD | 10/03/2024 | 12/04/2024 |  | 10/24/2024 |  |  | 10/03/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532414 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a patient admitted with a cardiac history. the coronary artery disease and valve disease prompted the patient to be admitted in for a high risk percutaneous coronary intervention. the team was unable to deliver the cp pump due to anatomic limitations. the team aborted the placement with a plan to return for the 5.5 pump. after the support was aborted the patient condition deteriorated and the patient expired despite all measures, cardiopulmonary resuscitation, shocks, and lucas device placement. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella unable to pass through the vasculature as per the clinical description provided.
16748 | 1220648-2024-20981 | 20522726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522726&pc=OZD | 10/01/2024 | 10/28/2025 |  | 10/24/2024 |  |  | 10/01/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531505 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Failure | Device Sensing Problem | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 66-year-old male patient was on impella cp support due to an acute myocardial infarction / cardiogenic shock (ami/cgs) and during the impella implant, the patient had ventricular fibrillation (vf). the patient was shocked and returned to baseline rhythm. additionally, there were placement signal issues with suction alarms. the pump was repositioned and the waveform was still significantly lower. audio was disabled. furthermore, the physician opted to leave the impella deep and to not reposition the pump further due to the device wanting to pop out of the left ventricle. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump and purge cassette were returned for investigation. upon investigating the pump, no abnormalities were observed; there was no biomaterial found around the impeller blades or on the sensor, no catheter or cannula kinks, or any bond failure issues. there were no broken connector pins, the optical bench parameters were within spec, and the pump passed the laser continuity test. the pump was also subject to uson pressure testing, and behaved normally, with the pump reacting to pressure changes and increasing placement signal values accordingly. however, it was observed that the placement signal reading was lower than the actual pressure inside the pressure chamber. no imc logs were returned to verify the g0 value during the case. data logs (lt and rt logs) were returned for analysis. it was seen that there were impella position in ventricle alarms around 1pm on 1 october 2024, which corresponded with abrupt rise in placement signal values. both resolved around 1:40pm. placement signal low alarms which started around 6am on 1 october 2024 got resolved around the same time as the impella position in ventricle alarm. there were suction alarms throughout the case. the 5s optical parameters did not show any abnormal trends and remained largely stable throughout the case. rt logs around the time the impella position in ventricle alarms triggered showed placement signal and lvp waveforms which followed no known pattern and were unable to be determined. clinical details revealed that the patient suffered from ventricular fibrillation once and was shocked successfully to baseline rhythm. the pump was positioned under echo post patient movement and waveforms were still significantly lower with suction alarms noted. echo on the next day of pump implant showed that the placement was deep inside the ventricle, but the decision was made to leave it at 4.1 cm inside the left ventricle owing to the device having a tendency to pop out of the left ventricle. vf/vt/af/at: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. optical signal issue: the root cause of the optical signal issue was not determined due to inconclusive evidence from the product evaluation, data logs analysis, and clinical details. positioning issues: there wasn't sufficient conclusive evidence to determine what caused the tendency of the pump to pop out of the left ventricle due to which it had to be placed deeper; whether the device was placed too deep because of the anatomy of the patient or if the pump wasn't secured properly in the left ventricle. the root cause of the positioning issue was not determined due to lack of sufficient clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
16749 | 1220648-2024-20975 | 20522728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522728&pc=OZD | 10/03/2024 | 12/06/2024 |  | 10/24/2024 |  |  | 10/03/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540218 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, there was significant bleeding at the impella access site. pressure was held and blood products were administered. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report.
16750 | 1220648-2024-20984 | 20522731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522731&pc=OZD | 03/29/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 03/29/2023 | 10/24/2024 |  |  | 11/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294912 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported after the sheath was peeled away, brisk bleeding was noted. manual pressure and femstop did not resolve the bleeding therefore systemic and purge heparin was withheld until bleeding subsided and hemostasis was achieved. the patient received one unit of blood products due to the blood loss. it was reported on (b)(6) 2023 the patient had computed tomography results that revealed a large brain hemorrhage. the patient's family decided to withdraw care and the patient was evaluated for organ donation. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16751 | 1220648-2024-20985 | 20522759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522759&pc=OZD | 02/03/2023 | 02/27/2025 |  | 10/24/2024 |  |  | 10/17/2024 | 10/24/2024 |  |  | 06/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237635 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. while on support the impella cp was in proper position and providing support; however, moved back into the aorta. the physician attempted to advance the pump back into the left ventricle but would not advance. cardiopulmonary resuscitation was initiated and the physician decided to continue pci without impella support. the physician pulled the impella and demonstrated that the cannula was kinked. the team considered placing another impella cp; however, the patient had been without a pulse or any native cardiac function for more than 1 hour and 30 min. the patient was pronounced dead. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned complaint pump visually inspected. a cannula kink was observed. the pump cannula kinked because they attempted to re-insert the pump wirelessly. the cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
16752 | 1220648-2024-20980 | 20522770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522770&pc=OZD | 09/25/2024 | 11/22/2024 |  | 10/24/2024 |  |  | 09/25/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546921 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient was on impella cp support because the patient began to rapidly decompensate becoming bradycardic and hypotensive was then taken to the central clinical laboratory where left heart catheterization revealed multi-vessel disease. the patient was intubated and temp pacer placed followed by percutaneous coronary intervention procedure of what was believed to the be culprit of the lesion in the right coronary artery. the patient continued to decompensate requiring high dose drips when the decision was made to place the impella cp. while on impella cp support the patient developed a hematoma post procedure in the unit while the patient had been moving. femstop was placed and removed, and manual pressure held. one albumin was given for hypotension; mean arterial pressure spiked into the high hundreds. the patient then became hypotensive again. the patient's hemoglobin (hgb) pre-was 13.8 just prior to removal hgb was 12.9. one unit of packed red blood cells were ordered, but not started prior to impella removal. the medical doctor discussed with the family continuation of decline in patient condition which led to the decision to pull the impella and to make the patient do not resuscitate. when care was withdrawn the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
16753 | 1220648-2024-20978 | 20522772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522772&pc=OZD | 09/25/2024 | 11/20/2024 |  | 10/24/2024 |  |  | 09/24/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532094 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Vascular Problem | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 77-year-old male patient, heparin was held due to nose and mouth bleeding. additionally, labs were hemolyzed and it was noted that the impella was caught under the papillary muscle. the physician elected to not reposition due to illness of patient. care was withdrawn and patient expired. medical safety review of the event noted there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported hemolysis has been completed. pump was not returned for investigation. data logs were returned but no motor current spikes, positioning issues or suction alarms observed. clinical information related to hemolysis is limited. the root cause of the hemolysis could not be determined without additional information.
16754 | 1220648-2024-20982 | 20522869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522869&pc=OZD | 06/08/2023 | 10/02/2025 |  | 10/24/2024 |  |  | 06/08/2023 | 10/24/2024 |  |  | 02/28/2025 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320716 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient expired due to hypoxic encephalopathy. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined b2 life threatening and required intervention was erroneously selected on the initial manufacturer device report number 1220648-2024-20982. g1 reporting contact email revised on manufacturer device report 1220648-2024-20982 in accordance with updated procedures.
16755 | 1220648-2024-20986 | 20522949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522949&pc=OZD | 04/29/2022 | 03/05/2025 |  | 10/24/2024 |  |  | 10/17/2024 | 10/24/2024 |  |  | 01/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193205 |  | 1000115 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Failure | Failure to Sense | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. while on support there was a placement signal failure that occurred immediately after dissection caused by a rotoblador procedure (not impella related). impella position was checked multiple times under direct visualization even though it was giving a constant alarm of "impella in lv." the pump was checked and confirmed in proper position, but alarms persisted. in addition, there were continuous runs of ventriculo-arterial uncoupling throughout the procedure. the pump was removed and replaced with a 2.5. additional information was provided that noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
16756 | 1220648-2024-20991 | 20522973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522973&pc=OZD | 03/25/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 03/25/2023 | 10/24/2024 |  |  | 10/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267405 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Increase in Pressure | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, device had positioning issues and high purge pressure alarms. additionally, patient had bleeding at the groin site. the groin was successfully revised and the hemorrhagic shock was under control. additional information noted that the care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported positioning issues, high purge pressure and vascular damage has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the reported issues could not be determined. this report is being filed as part of a retrospective review of historical records.
16757 | 1220648-2024-20983 | 20522977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522977&pc=OZD | 06/16/2023 | 05/23/2025 |  | 10/24/2024 |  |  | 06/16/2023 | 10/24/2024 |  |  | 08/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260680 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had signs of infection. it was indicated that it was possible that the cause of the shock was not only worsening heart failure. the patient had poor prognosis and was decided to be an end-of-life care and the patient expired from multiple organ failure. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist instructions for use for the related event are as follows: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16758 | 1220648-2024-20989 | 20522978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20522978&pc=OZD | 09/22/2021 | 10/24/2024 |  | 10/24/2024 |  |  | 09/22/2021 | 10/24/2024 |  |  | 03/31/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067366 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Pumping Stopped | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the team restarted the pump after being in surgical mode. the team attributed the stop to perhaps clamping on the motor. the pump remained on for support despite the stop/restart and not being in optimal position. they had been unable to reposition out of the papillary. care was withdrawn due to patient being poor candidate for escalation in light of overall patient condition and the patient expired minutes later. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop and positioning issues has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs were analyzed and the high motor current pump stops were confirmed. the pump was returned and was unable to start. the handle was opened and was found to be filled with blood. a hole in the catheter was found just below the motor housing. the root cause of the pump stops was determined to be an aortic cross clamp impeding the rotation of the impeller. the clamp most likely also punctured the catheter allowing blood to fill the red handle. the root cause of the positioning issues was most likely use related as imaging showed the pump in the incorrect position. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16759 | 1220648-2024-20988 | 20523038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523038&pc=OZD | 04/11/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 04/11/2022 | 10/24/2024 |  |  | 01/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183074 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, pump abruptly stopped. device was explanted. discussions were held with the patient's family to decide if support should be escalated or comfort care measures. family decided to withdraw care and patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the pump stop has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the pump stop could not be determined. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16760 | 1220648-2024-20992 | 20523119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523119&pc=OZD | 04/07/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 04/07/2023 | 10/24/2024 |  |  | 12/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024304063 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a thrombus was found in the left ventricle during impella support. it was further noted that both of the patient's feet were also ischemic due to chronic tibial disease that was found on the doppler. neither condition necessitated intervention. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. per the provided clinical information, the root cause of the limb ischemia issue was patient condition related to diseased vessels which are a result of multiple diseases patient was suffering from. the root cause of the thrombus issue was most likely patient condition related since product was not returned for investigation. as noted in the impella ifu: impella cp with smart assist system. impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
16761 | 1220648-2024-20993 | 20523140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523140&pc=OZD | 02/11/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 02/11/2023 | 10/24/2024 |  |  | 06/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270505 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 7 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. it was reported that that the patient had signs of hemolysis treated by pump repositioning. it was also reported that support was withdrawn and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was most likely patient condition related. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
16762 | 1220648-2024-20990 | 20523231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523231&pc=OZD | 05/11/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 05/11/2022 | 10/24/2024 |  |  | 12/31/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163682 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. hemolysis was identified during support. it is unknown if hemostasis was achieved. additionally, there was a lot of bleeding from the ecmo insertion site. heparin was reduced to manage this complication. weaning was successful and the pump was explanted. patient survived the explant but subsequently expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system (access site bleeding) impella cp with smart assist system (hemolysis) section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
16763 | 1220648-2024-20987 | 20523381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523381&pc=OZD | 03/06/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 03/06/2022 | 10/24/2024 |  |  | 03/31/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022063155 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the clinicians attempted to restart the pump, but their efforts were unsuccessful. subsequently, the pump was explanted and replaced with a new impella cp. patient survived the explant procedure but subsequently expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop issue has been completed. the product was returned and the pump stop was confirmed. the root cause of the pump stop was bearing wear after 9 days of use, which is beyond design life. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
16764 | 1220648-2024-20995 | 20523402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523402&pc=OZD | 03/03/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 03/03/2023 | 10/24/2024 |  |  | 11/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024293912 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was noted that there were no pulses found in both feet. physician was comfortable continuing with impella care. care was eventually withdrawn after 13.9 hours of support and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the ischemia was due to patient anatomy, most likely due to patient condition. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16765 | 1220648-2024-20996 | 20523482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523482&pc=OZD | 03/22/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 03/22/2023 | 10/24/2024 |  |  | 10/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024296916 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was noted that a younger patient was kept on impella support for 20 days. the pump was used beyond its indication and the medical team was made aware. the pump stopped on day 20. the patient suffered liver and renal failure and expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the pump stop could not be determined. this report is being filed as part of a retrospective review of historical records.
16766 | 1220648-2024-20994 | 20523495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523495&pc=PYX | 09/22/2024 | 11/22/2024 |  | 10/24/2024 |  |  | 09/22/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025545998 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella rp flex to support a patient admitted in cardiomyopathy and post left ventricular assist device placement. the impella rp flex was placed via the femoral, and two days after implant, the access site had an ooze. the team remedied the issue by placement of an extra stitch. the pump remained on for another day and was then explanted and a proteck duo was placed. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual: section: using the automated impella controller with the impella rp flex smartassist system catheter: use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be use issue because the bleeding was resolved and hemostasis was achieved by additional suture.
16767 | 1220648-2024-20998 | 20523501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523501&pc=OZD | 02/21/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 02/21/2023 | 10/24/2024 |  |  | 10/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290769 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was decided to escalate the patient to higher support using an impella 5.5. during explant, bleeding was noted and treated with one unit of blood. additionally, a thrombus was reported with no intervention specified. patient was withdrawn from care and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported thrombus and access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the access site bleeding and its relation to the impella were not determined since the product and sufficient clinical information were not provided. the femoral thrombus was determined to be most likely patient condition related. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system (access site bleeding). section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ as noted in the impella ifu: impella cp with smartassist system (thrombus). section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
16768 | 1220648-2024-20997 | 20523502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523502&pc=OZD | 03/29/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 03/29/2023 | 10/24/2024 |  |  | 11/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312785 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. pump came out of the patient's left ventricle (lv) during chest compressions and could not be re-advanced across the aortic valve (av). a new impella was used without issue. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the pump stop was use related as the pump came out of position during cpr. this report is being filed as part of a retrospective review of historical records.
16769 | 1220648-2024-21000 | 20523563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523563&pc=OZD | 10/05/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 10/05/2022 | 10/24/2024 |  |  | 06/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237221 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, access site bleeding was noted. pump was repositioned, sutures fixed, and angle dressed. soon after, patient start decompensating and expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of access site bleeding could not be determined. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16770 | 1220648-2024-21001 | 20523728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523728&pc=OZD | 11/08/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 11/09/2022 | 10/24/2024 |  |  | 08/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252049 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were suction alarms and low pump flow although volume status and positioning of the pump had been good. the patient had been in weak condition but became hemodynamically stable, therefore the physician decided to explant the impella cp device. the patient expired afterwards due to overall weak condition. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected, and no abnormalities were observed. the pump was run in a basin of water and pump flow was 3.7 l/min. the cause of the issue was most likely patient condition related. this report is being filed as part of a retrospective review of historical records.
16771 | 1220648-2024-21002 | 20523746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523746&pc=OZD | 11/17/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 11/17/2022 | 10/24/2024 |  |  | 08/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254594 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there were pump flow issues suspected to be related to ingested thrombus. the physician was advised to exchange the pump immediately and decided to leave the pump in place to transfer the patient to another hospital. the physician states that the pump was probably implanted with inadequate activated clotting time (act) due to para vasal infusion of heparin. the pump showed strange readings shortly after implant. after transfer, the patient was in bad (undefined) condition and the patient expired while on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
16772 | 1220648-2024-21009 | 20523757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523757&pc=OZD | 05/03/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 05/03/2023 | 10/24/2024 |  |  | 09/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274461 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, there was a large hematoma at the impella insertion site that could not be resolved with standard techniques. a dissection of the left main artery extending to the aorta occurred, and hemostasis was not achieved. the impella device was explanted, care was withdrawn, and the procedural outcome was that the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the reported access site bleeding and vascular damage issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was most likely use related. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records.
16773 | 1220648-2024-20999 | 20523762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523762&pc=OZD | 10/01/2024 | 12/06/2024 |  | 10/24/2024 |  |  | 10/01/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126793 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when prepping for an impella implant, the automated impella controller (aic) had an optical sensor system error in yellow alarm when the aic was plugged in. the aic was replaced and a new aic was used successfully. no patient harm has been reported. | [5 empty] | The investigation for the optical signal issue has been completed. data logs were reviewed which confirm that the optical sensor system error message was displayed by console during case start after the console failed optical bench calibration with low snr. the pump is then unplugged and re-plugged into the console eventually completing optical bench calibration before shutting down the console. the console was returned for evaluation. the front panel optical connector was inspected. there was a persistent contamination with in the connector. when a known good pump and purge line was run with console. the console operated as expected. the root cause of the optical signal issue was an air gap between the pump and console at the optical connector. the contamination was not a primary contributing factor to the failure. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
16774 | 1220648-2024-21008 | 20523829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523829&pc=OZD | 04/03/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 04/03/2023 | 10/24/2024 |  |  | 07/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251147 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during the implantation, there were positioning issues with the cp. the pump was being inserted but was caught in the mitral valve (mv) and the pump flows were not appropriate. the team made a choice to reposition the pump and this failed, and so the team explanted the impella cp and an intra-aortic balloon pump (iabp) was placed, and the patient was transferred out to another facility. the procedure outcome noted the patient expired. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the root cause of the issue was use related. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when. manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ this report is being filed as part of a retrospective review of historical records.
16775 | 1220648-2024-21005 | 20523844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523844&pc=OZD | 10/01/2024 | 11/23/2024 |  | 10/24/2024 |  |  | 10/01/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544888 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted in acute myocardial infarction / cardiogenic shock. the 2nd day of support the impella limb lost pulses. the team discussed explantation of the cp and placement of a 5.5 pump via the axillary. the cp was explanted. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: b1-should have selected adverse event.
16776 | 1220648-2024-21006 | 20523848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20523848&pc=OZD | 04/09/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 04/09/2023 | 10/24/2024 |  |  | 12/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024289490 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella purge filter broke while the patient was still in catheterization lab, shortly after implant, for a reason unknown. a leak was observed and it was reported that no retainer was used. the medical team performed a purge bypass. there was no patient harm reported. it was noted that the patient was later withdrawn from care and expired while on support. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported issue has been completed. the cause of the issue was a leak from the sidearm but the exact location of the leak was unknown since product was not returned. this report is being filed as part of a retrospective review of historical records.
16777 | 1220648-2024-21004 | 20524008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524008&pc=OZD | 03/22/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 03/28/2023 | 10/24/2024 |  |  | 11/30/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024299110 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. it was reported that the sterile sleeve came off when physician attempted to reposition impella. they covered the sterile sleeve with tegaderm. the patient was escalated to an impella 5.5. it was noted that the patient was withdrawn from care while on impella 5.5 support, and expired. medical safety review of the event noted use of the impella cp device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported product damage has been completed. the cause of the product damage, specifically damage to the sterile sleeve, was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
16778 | 1220648-2024-21003 | 20524169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524169&pc=OZD | 09/13/2022 | 05/02/2025 |  | 10/24/2024 |  |  | 09/13/2022 | 10/24/2024 |  |  | 04/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216220 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a purge leak from the cassette during impella support. the cassette was removed and replaced without further issue. there was no harm to the patient as a result of the purge leak. also, bleeding from the stomach was observed and the source of bleeding was examined, revealing gastric cancer. a gastroenterologist attempted to stop the bleeding through an endoscope, but the tumor was already at stage iv and there was nothing that could be done. heparin administration was discontinued, including the heparin in the impella purge. after informing the family of the above, they decided that no further treatment was desired and decided to proceed with dnar using only the current treatment (impella operation). it was noted that the patient later expired during support. use of the device cannot conclusively be associated with the outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. the cause of the purge leak was not determined since product was not received for evaluation. the patient death was determined to be unrelated to the impella. this report is being filed as part of a retrospective review of historical records.
16779 | 1220648-2024-21021 | 20524171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524171&pc=OZD | 09/08/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 09/12/2022 | 10/24/2024 |  |  | 03/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195386 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bleeding and a hematoma at the impella access site. as a result of the blood loss, the patient was transfused with blood products. it was noted the patient outcome was expired, and the physician stated the patient was critically ill and there was no relation of the impella with the patient's outcome. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16780 | 1220648-2024-21017 | 20524219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524219&pc=OZD | 10/02/2024 | 12/04/2024 |  | 10/24/2024 |  |  | 10/02/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025543774 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support developed hematoma at the access site. the patient was taken to the operation room for a washout of the hematoma. the patient has less swelling and pain post washout. systemic heparin held. it was further reported that the patient had ventricular tachycardia. impella was repositioned as it was close to the mitral valve. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation for the reported access site bleed and ventricular tachycardia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and ventricular tachycardia could not be determined.
16781 | 1220648-2024-21007 | 20524266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524266&pc=OZD | 02/07/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 02/07/2023 | 10/24/2024 |  |  | 09/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281416 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | Positioning Failure | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. impella cp repositioning issues due to pigtail entrapment in papillary muscle, which caused pump to flow at lower levels. pump was not replaced. patient expired three days later with several contributing factors. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The field lt was reviewed and suction with a brief pump position in the aorta alarms were observed. no pump position deep or lv waveforms were observed. the cause of the positioning issue was most likely patient condition related since patient volume and positioning in the lv anatomy was difficult. possible cannula kink since the pump pigtail could not be positioned out of the papillary. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type.
16782 | 1220648-2024-21015 | 20524294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524294&pc=OZD | 10/02/2024 | 12/06/2024 |  | 10/24/2024 |  |  | 10/02/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454530 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 and during the implant, a clot was noted in the left ventricle, but the physician placed the impella. there was suspicion that the clot was ingested by the pump due to a change in motor current and wave forms. placement monitoring was disabled. there were diminished flows and it was thought that there was a clot in the cannula. the patient remains on impella support. | [5 empty] | The investigation into low pump flow and thrombus has been completed. the failure mode (fm) thrombus removed because it is a duplicate of the low pump flow root cause. data review: data logs confirmed the failure mode & clinical narrative. logs showed after pump started & run for 4 minutes, motor current spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. d6b added. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
16783 | 1220648-2024-21013 | 20524304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524304&pc=OZD | 08/08/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 08/08/2022 | 10/24/2024 |  |  | 01/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023190900 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was going to be implanted with an impella cp for mechanical circulatory support. during delivery of the device, the doctor tried to insert the imeplla cp, however it would not pass the patient's vasculature. it was noted that the shaft of the pump catheter was kinked during delivery and could not be used. additional information noted that the patient died one day after the attempt to implant the device failed. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was most likely patient condition, as the physician noted it was a difficult insertion. as noted in the impella ifu: impella cp with smart assist system section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
16784 | 1220648-2024-21022 | 20524313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524313&pc=OZD | 10/03/2024 | 12/02/2024 |  | 10/24/2024 |  |  | 10/03/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544105 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump support placed for a patient with cardiogenic shock and cardiomyopathy. the 5.5 pump was placed as care escalation from an intra-aortic balloon pump and when transferred from community to tertiary center. the 5.5 support was successful for 12 days. the patient was noted to be therapeutic on systemic heparin and sodium bicarbonate but, with explant, an embolic cerebrovascular accident occurred. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
16785 | 1220648-2024-21012 | 20524314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524314&pc=OZD | 08/23/2024 | 11/14/2024 |  | 10/24/2024 |  |  | 08/23/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1103720 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 71-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during the initial setup, the automated impella controller (aic) failed to correctly prime and detect purge cassette. an additional cassette was used with the same result. the aic was swapped, cassette was recognized, and pump properly primed. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). patient's peri-procedural condition was critical. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the console purge issues has been completed. data logs were reviewed which partially confirmed priming issue and purge cassette detection issues. the purge kit was able to complete priming without issue but there were some imc-ahp error: 0x91 0x83 0×24 prompts observed from the reported. the console was returned for evaluation and the reported priming and purge cassette detection issues were unable to be reproduced. the aic was able to prime and detect the purge cassette as normal when it was tested in the lab. the root cause of the priming and purge cassette detection issue was likely use related. corrections to the initial mfr report: h6 impact code incorrectly coded 1802.
16786 | 1220648-2024-21019 | 20524329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524329&pc=OZD | 09/06/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 09/06/2022 | 10/24/2024 |  |  | 03/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204120 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support during an electrophysiology ablation procedure. during the support, an observation made that there was air visualized in the left ventricle and the aorta after the impella was inserted. no further information provided indicating treatment used to address the issue. additional information was provided that noted that care was withdrawn and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported embolism has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause could not be determined. as noted in the impella ifu: impella cp with smart assist system. this report is being filed as part of a retrospective review of historical records.
16787 | 1220648-2024-21010 | 20524462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524462&pc=OZD | 09/27/2021 | 10/24/2024 |  | 10/24/2024 |  |  | 09/27/2021 | 10/24/2024 |  |  | 07/31/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119443 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella was successfully placed and while preparing the patient for transport to the intensive care unit, following removal of the peel-away sheath, the flows dropped from 3.5 liters per minute (l/min) to 0.5 l/min. the decision was made to replace this pump and during removal a clot was observed in the impella cannula. this impella was removed and a new impella placed without incident. | [5 empty] | The investigation into the thrombus issue has been completed. the impella pump was returned for investigation. the data logs were returned and confirm the low flow. the product was returned, and the clot could be seen in the inflow. the pump was run in a basin of water and produced the expected amount of flow. the root cause of the low flow was determined to be ingested biomaterial. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation . section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. as noted in the impella ifu ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ as noted in the impella ifu "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
16788 | 1220648-2024-21014 | 20524484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524484&pc=OZD | 08/03/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 08/03/2022 | 10/24/2024 |  |  | 03/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211473 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited saturated flows and was replaced as a result. the patient ultimately expired, but there was no allegation that this was related to the impella. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The device was returned and evaluated, and the investigation for the reported saturated flow has been completed. visual inspection of the returned pump revealed the presence of biomaterial wrapped around the impeller and in the purge gap. no other abnormalities were found. in conclusion, it was determined that the cause of the saturated flow was ingested biomaterial. the cause of the ingested biomaterial is unable to be determined. this report is being filed as part of a retrospective review of historical records.
16789 | 1220648-2024-21011 | 20524517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524517&pc=OZD | 07/30/2022 | 09/22/2025 |  | 10/24/2024 |  |  | 08/13/2024 | 10/24/2024 |  |  | 10/31/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142122 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 13 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an unknown primary indication was implanted with an impella cp for mechanical circulatory support. during support, the patient had hemolysis from insertion to removal. the patient ultimately expired, though no allegation was made relating this to the impella. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-21011 in accordance with updated procedures. h6 health effect - impact code 1802 was erroneously entered on the initial manufacturer device report number 1220648-2024-21011.
16790 | 1220648-2024-21016 | 20524534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524534&pc=OZD | 08/09/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 08/09/2022 | 10/24/2024 |  |  | 04/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225140 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Fluid/Blood Leak | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella was found to be leaking, and the purge connection was broken. the team was planning to do a purge bypass, but the pump began to display pump saturation, so the device was explanted. additional information notes that the patient died on support. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The investigation for the reported purge leak/saturated flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the purge leak was most likely patient movement since the leak occurred during transfer and the resulting saturated flow was likely due to lack of purge. as noted in the impella ifu: purge leak - pump: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
16791 | 1220648-2024-21018 | 20524906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20524906&pc=OZD | 12/18/2021 | 10/24/2024 |  | 10/24/2024 |  |  | 08/25/2022 | 10/24/2024 |  |  | 08/31/2023 |  | 10/24/2024 |  | DSAMMARCO@ABIOMED.COM |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022117301 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp and extracorporeal membrane oxygenation for mechanical circulatory support. while on support, the patient experienced access site bleeding with hematoma. a blood transfusion was needed and it is unknown how many units of packed red blood cells the patient received. it was reported that due to the facility restrictions no further information was able to be obtained. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16792 | 1220648-2024-21033 | 20525133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20525133&pc=OZD | 10/03/2024 | 12/13/2024 |  | 10/24/2024 |  |  | 10/03/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533819 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the peel away sheath was still in place and the patient on impella cp support had no pedal pulses on the impella leg. the leg was cool and blue, thigh was still warm, but discoloration of the limb was progressing. the registered nurse wanted to wean the pump but had to talk to the medical doctor therefore treatment and outcome of the patient is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report.
16793 | 1220648-2024-21020 | 20525135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20525135&pc=OZD | 11/27/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 11/27/2022 | 10/24/2024 |  |  | 08/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265278 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that an unexpected pump stop occurred. restart was attempted but unsuccessful. it was reported to have been a flow saturation issue which had been occurring for hours prior to pump stop. also, the doctor was aware of hemolysis, no repositioning or volume given. patient was presented with ipsilateral leg ischemia. additional information was provided that noted the patient was placed on do not resuscitate status and expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. data logs showed that the purge heparin levels decreased to 0 about 24 hours prior to the pump stop. soon after, the pump flow waveform becomes saturated. the motor current and pump flow steadily increase with no p-level change (and worsening flow saturation) for 7 hours until the pump stops immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm. visual inspection of the returned pump revealed the presence of biomaterial on and around the impeller and in the purge gap. no other damage or abnormalities were found. in conclusion, it was determined that the cause of the pump stop was biomaterial buildup due to the lack of heparin in the purge solution. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
16794 | 1220648-2024-21035 | 20525205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20525205&pc=OZD | 10/04/2024 | 09/30/2025 |  | 10/24/2024 |  |  | 10/04/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510638 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Eu complainant reported the patient was on impella 5.5 support and the pump flow showed 5l, but the left ventricle was not unloading. staff suspected that the sensor was damaged. the pump was explanted the same day. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into low pump flow has been completed. no visual abnormalities were noted on returned pump. returned pump was run in the lab without issue. data logs from field showed multiple suction alarms throughout case. additionally, placement signal was noted to be noisy and trending down from start of case. the root cause of the low pump flow was most likely patient condition related. d9 date device returned to manufacturer added. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21035. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. upon visual inspection, no abnormalities of the pump were noted. optical parameters were measured in the lab, and all were within a normal range. pump was run in the lab, and no abnormalities were noted. data logs were reviewed and all optical parameters during case were found to be within range and no hard failures of the optical sensor were noted. no problem was found with the optical sensor while running in the field and with returned product. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-21035. d10 concomitant medical products and therapy dates updated.
16795 | 1220648-2024-21025 | 20525240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20525240&pc=OZD | 10/14/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 10/14/2022 | 10/24/2024 |  |  | 04/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217524 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the initial insertion attempt, the cannula kinked at the midpoint. although repositioning of the pump straightened out the kink, the component re-kinked at the same point during each subsequent attempt to place the pump. as a result, the pump was explanted and a new pump was inserted without further issue. additional information was provided that noted the patient expired. there was no reported relation of the impella with the patient's outcome. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. returned pump visually inspected. cannula kink observed. it looks like the kink happened while attempted to implant the impella. root cause of the delivery issue was improper technique of use. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
16796 | 1220648-2024-21026 | 20525252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20525252&pc=OZD | 09/26/2024 | 11/23/2024 |  | 10/24/2024 |  |  | 09/26/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505113 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support because the patient presented to the emergency department after being shocked by her implantable cardioverter defibrillator. she was found to be in atrial flutter with rapid ventricular response and started on amiodarone and heparin drips as well as diuretics. the patient continue to decompensate therefore the decision was made to place on extracorporeal membrane oxygenation and impella 5.5. while on impella 5.5 support there was increased drainage from the jackson-pratt drain. upon axillary exploration, a bleed was noted on axillary artery. the doctor was able to repair the artery which resolved the bleed. due to the complication, patient received eight packed red blood cells, five fresh frozen plasma, and six platelets. the patient was unable to maintain airway with bleeding despite multiple bronchoscopy. the patient also had other sources of bleeding. the family decided to withdraw support. extracorporeal membrane oxygenation and the impella was turned off and the patient expired immediately. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the vascular damage has been completed. pump was not returned for investigation. clinical details provided regarding the damage is very little and patient was coagulopathic and had received blood products. therefore, the root cause of the vascular damage issue was patient condition related to coagulation disorder.
16797 | 1220648-2024-21037 | 20525656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20525656&pc=OZD | 10/04/2024 | 09/24/2025 |  | 10/24/2024 |  |  | 10/04/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025504230 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Tachycardia | Embolism/Embolus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported the patient was on impella 5.5 support and while on support there were small amounts of air bubbles noted in the left ventricle. staff preemptively changed purge cassette and bag and reimaged after extubating. support continued. furthermore, there were suction events. an echocardiogram was performed and the physician reported the position was acceptable. one unit of packed red blood cells were given to the patient and the procedural outcome was the patient survived surgery. | Additionally, an impact study found that the patient had ventricular tachycardia (vt) with a possible relationship to the impella device. the vt was resolved with amiodarone drip. | The complainant also reported that there were placement signal not reliable alarms. audio was disabled and support continued. |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death.". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the embolism and low pump flow has been completed. clinical details report that small air bubbles were seen in the left ventricle. it was troubleshot by changing the purge bag and cassette. due to product not being returned and limited clinical details, the root cause of the embolism could not be determined. data logs were reviewed, and it was confirmed that pump flows were below specification or unstable throughout the majority of support. additionally, 129 suction alarms triggered throughout the case. product was not returned for investigation. based on clinical details and data log review, the root cause of the low pump flows was determined to most likely be patient condition. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information received from the clinical site. h6 revised to reflect the additional information received from the clinical site. g1 reporting contact email revised on manufacturer device report 1220648-2024-21037 in accordance with updated procedures. g2 report source; study was missing from manufacturer device report 1220648-2024-21037. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the reported placement signal not reliable (psnr) alarm was confirmed. regarding the 5s parameters, signal-to-noise ratio (snr) was below 500 for the entirety of the case, contrast was variable, and light and gain were within specification. there were many suction alarms in this case, however, the previous and next pumps used on the same console also show the snr to be consistently below specification without other contributing factors to noise. therefore, the suction is not the cause of the psnr alarms. upon review of data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5. updated to include optical signal issue description. h6. updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-21037. d.10. concomitant medical products and therapy dates updated.
16798 | 1220648-2024-21030 | 20525694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20525694&pc=OZD | 10/25/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 10/25/2022 | 10/24/2024 |  |  | 07/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255225 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the pump was measured to have a saturated flow. as a result of the noted issue, the pump was removed and replaced with a new impella cp pump. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. product was visually inspected. buildup was observed around the impeller purge gap. product was soaked in water for five minutes. biomaterial was observed around the impeller purge gap. the root cause for saturated flow is most likely biomaterial. this report is being filed as part of a retrospective review of historical records.
16799 | 1220648-2024-21024 | 20525779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20525779&pc=OZD | 10/12/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 10/12/2022 | 10/24/2024 |  |  | 04/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217240 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | Appropriate Term / Code Not Available | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implantation it was noted that when the guidewire was removed, a portion was missing from the distal end. a fluoroscopy verified that the fractured component was still lodged in the pigtail of the pump. the pump was removed with the wire intact and a new pump was implanted without further issue. patient decompensated due to left coronary artery percutaneous coronary intervention. additional information was provided that noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based off clinical details, the root cause for the guide wire fracture is most likely use related. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
16800 | 1220648-2024-21040 | 20525925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20525925&pc=OZD | 07/09/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 07/13/2022 | 10/24/2024 |  |  | 03/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209086 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was attempted to be implanted with an impella cp for mechanical circulatory support. during delivery, the 0.018 inch guidewire became kinked, preventing deployment. it was noted that the angle of insertion was 45 degrees. the guidewire was replaced, and the device was able to be deployed successfully. it was reported that the patient's cardiac function did not recover, and the patient died of multiple organ failure. the patient's physician reported that the outcome was not related to impella, nor were there any product problems. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). no patient harm related to the device was identified. | [5 empty] | The investigation for the reported delivery issue/guidewire kink has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue/kink was most likely a steep insertion angle. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16801 | 1220648-2024-21036 | 20526227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20526227&pc=OZD | 02/28/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 02/28/2022 | 10/24/2024 |  |  | 11/30/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023166779 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped, and the staff restarted the pump on p2. it was reported that the motor current and flows were very abnormal, so pump was explanted. it was noted that the patient was on support longer than the device indication. after removal, the patient was hemodynamically stable. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. | [5 empty] | B2: date of patient death is unknown the investigation for the reported pump stop has been completed. the product was returned. per the results of the clinical review and investigation: the root cause of the pump stop was determined to be bearing wear. the pump stop occurred on day 11 of support which is beyond the 8-day design life as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16802 | 1220648-2024-21023 | 20526310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20526310&pc=OZD | 09/21/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 09/21/2022 | 10/24/2024 |  |  | 07/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238707 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 MO |  | Unknown |  |  |  | Injury | Yes | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. immediately after implant, the pump stopped. restarting and changing the aic did not resolve the issue. pump was removed. the high-risk percutaneous coronary intervention was performed without impella support and afterwards, a new pump was implanted. additional information was provided that noted the patient expired. there was no reported relation of the impella with the patient's outcome. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported pump did not start issue has been completed. no logs were returned. a visual inspection of the pump revealed a fractured red lumen in the pump. the root cause of the issue was determined to be improper technique in that there was a red lumen removal issue. this report is being filed as part of a retrospective review of historical records.
16803 | 1220648-2024-21043 | 20526381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20526381&pc=OZD | 03/17/2023 | 05/23/2025 |  | 10/24/2024 |  |  | 03/17/2023 | 10/24/2024 |  |  | 12/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300830 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient did not have distal pulses in the right lower extremity as the leg was cold to the touch and slightly mottled, indicating limb ischemia. additionally, it was noted that the impella position was difficult to assess due to bleeding at the impella access site. the physician was unable to get the correct catheter measurement to place the repositioning sheath in inaccurate position. an echocardiogram (echo) was ordered, while measuring impella position, the patient coded. advanced cardiac life support was utilized however, it was reported that the patient expired soon after. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported limb ischemia and access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause for the access site bleeding issue. the root cause of the access site bleeding issue was use related to improper technique. the cause of the limb ischemia and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system. limb ischemia. impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ access site bleeding. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16804 | 1220648-2024-21027 | 20526836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20526836&pc=OZD | 10/17/2022 | 09/24/2025 |  | 10/24/2024 |  |  | 10/17/2022 | 10/24/2024 |  |  | 06/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248343 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the complainant reported that during impella device setup, the medical staff was unable to prime the purge system. because of the urgency of the case, the decision was made to open a second device to manage this complication. additional information was provided that noted the patient expired. there was no reported relation of the impella with the patient's outcome. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the investigation of priming issue has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-21027 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-21027 in accordance with updated procedures.
16805 | 1220648-2024-21041 | 20526967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20526967&pc=OZD | 06/03/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 06/03/2022 | 10/24/2024 |  |  | 01/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174073 |  | 0048-0024-JP |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced gastrointestinal bleeding, the patient received a blood transfusion, and the physician decide to explant the impella cp. it was reported that the patient died several hours after explanting the impella cp. it was reported that the patient died due to sepsis. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. no reported impella access site infection. | [5 empty] | The investigation for the reported gastrointestinal bleeding and sepsis has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the gastrointestinal bleeding and sepsis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16806 | 1220648-2024-21028 | 20527138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20527138&pc=OZD | 10/18/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 10/18/2022 | 10/24/2024 |  |  | 03/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210510 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support it was reported that there was a low flow issue. it was identified that the suboptimal flow was due to a kink in the pump. the pump was removed and replaced. additional information was provided that noted the patient expired. there was no reported relation of the impella with the patient's outcome. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the logs confirm suction alarms and a lack of motor current pulsatility. the root cause of the low flow was unable to be determined as no product was returned for analysis. this report is being filed as part of a retrospective review of historical records.
16807 | 1220648-2024-21034 | 20527208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20527208&pc=OZD | 11/10/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 11/10/2022 | 10/24/2024 |  |  | 08/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262863 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the impella was placed with the patients activated clotting time at 246. reportedly the pump was stated on auto and flows were initially 3 liters, then suction and flows went to 0.0 with motor current 0. attempts were reportedly made to change the pressure level, however, this did not resolve the flow issues. the pump was removed and replaced with an impella 5.0 for further support. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the investigation into the low pump flow has been completed. the cause of the issue was not determined since product was not returned and the clinical details provided were insufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
16808 | 1220648-2024-21031 | 20527392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20527392&pc=OZD | 10/03/2024 | 11/25/2024 |  | 10/24/2024 |  |  | 10/03/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025531894 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during support there was fluid leaking from the red plug. echocardiography showed the device malfunctioning. the impella was removed. the patient survived the event. | [5 empty] | The investigation for the product damage has been completed. data logs are not relevant to the investigation, and neither product nor photos were returned to aid the investigation. clinical details describe that fluid had gotten into the red pump plug and was leaking from there. since neither product nor photos were returned, the root cause of the product damage could not be determined. added- h6 impact code 4627. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
16809 | 1220648-2024-21038 | 20527461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20527461&pc=OZD | 04/19/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 04/19/2023 | 10/24/2024 |  |  | 12/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024310649 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced oozing noted around the impella repositioning sheath after it was swapped from introducer. the medical team attempted to angle match and secure suture hub close to the skin and they could not obtain hemostasis. after destat and other clotting agents were applied, a femstop was placed. the patient was then transferred to another hospital for additional therapy. it was noted that the following day, the patient's family decided to withdraw care from the patient, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. based on the clinical details, the root cause of the issue was lack of vessel recoil due to patient condition. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16810 | 1220648-2024-21032 | 20527500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20527500&pc=OZD | 11/08/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 11/08/2022 | 10/24/2024 |  |  | 05/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227671 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a pump stop shortly after insertion. attempts to restart the pump were unsuccessful, therefore, the device was removed and replaced. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. there are no imc logs available for the first console. the rt log from the first console (ic4750) confirms alarm 119 occurrences. the rt log from the second console (ic4749) does not show alarm 119 occurrences but shows impella stopped alarms. the pump successfully starts on the second console and is manually turned to p0 by the user which was confirmed from the imc logs. the placement signal is zero when the pump is run so it appears that the pump was run outside the body in water as a test. a visual inspection of the pump did not reveal any abnormalities. the pump was plugged into an aic and was successfully run. the 119 alarm did not reproduce. all of the motor cables were tested and found to be functioning and within specification. the pump was run in a bucket of water for 16 hours without issue. the cause of the pump not starting was not determined because it could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16811 | 1220648-2024-21052 | 20527502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20527502&pc=OZD | 07/04/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 07/04/2023 | 10/24/2024 |  |  | 03/31/2025 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343150 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | White | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the cp failed to deliver via the short 14fr sheath. the team placed a 13 cm sheath with a 25 cm sheath to insert another impella cp. the new 2nd cp was successfully placed to the patient to support till the patient expired when care was withdrawn. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that visual inspection found no issue on the pump. the introducer was not returned for review. based on clinical description, the patient had diseased vessel resulting in pump could not pass it. therefore, the root cause of delivery issue was due to patient condition. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
16812 | 1220648-2024-21050 | 20527544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20527544&pc=OZD | 07/05/2023 | 09/17/2025 |  | 10/24/2024 |  |  | 07/05/2023 | 10/24/2024 |  |  | 02/28/2025 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024321288 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while patient was on impella support had encountered a low pump flow and the patient became hypotensive with runs of ventrical fibrilations. patient was shocked 3 times this shift. impella was ruled out as cause for ectopy by echo. the patient remains on ecpella, was made a dnr. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that per the clinical information provided and pump investigation, the root cause of the low pump flow issue was biomaterial found during the motor teardown, on the magnet.
16813 | 1220648-2024-21058 | 20527724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20527724&pc=OZD | 03/16/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 03/16/2023 | 10/24/2024 |  |  | 09/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023280290 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the patient was transferred to another hospital, an impella alarm for "device defective, do not use" was received. the patient was stable and was brought to the catheterization lab to replace the impella pump. after the patient was successfully placed on the new pump, it was noted that care was withdrawn, and the procedural outcome was the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported impella defective alarm issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. data logs show that no impella defective or controller error alarms occurred throughout support. no pump stops were observed. pump performance metrics were generally within specification. therefore, was no relevant issue found during the case. this report is being filed as part of a retrospective review of historical records.
16814 | 1220648-2024-21054 | 20527862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20527862&pc=OZD | 09/05/2022 | 10/07/2025 |  | 10/24/2024 |  |  | 09/13/2022 | 10/24/2024 |  |  | 12/31/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164248 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  |  | Death | Yes | No | Restricted Flow rate | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with extracorporeal membrane oxygenation and impella cp device for mechanical circulatory support. while on support, flow decreased and could not be maintained. it was reported that the cause of low flow was most likely bleeding due to chest compression, which lead to the shortage of blood in the left ventricle and decrease in flow rate. additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-21054. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2024-21054. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-21054 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-21054 in accordance with updated procedures. h6 investigation conclusions code 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-21054.
16815 | 1220648-2024-21029 | 20527985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20527985&pc=OZD | 10/03/2024 | 12/04/2024 |  | 10/24/2024 |  |  | 10/03/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519264 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Renal Failure | Hematuria | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and the patient had red urine on support. the pump was repositioned and the patient's urine color started to clear. it was reported that there was organ damage and continuous renal replacement therapy (crrt) was started. prior to explant, the doctor stated that an mri would be done after explant of impella. upon follow up with physician, they stated that the ¿mri showed a stroke¿. treatment for the stoke is unknown however it was noted that the patient was stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the stroke, hemolysis, renal failure has been completed. data logs were reviewed which confirmed pump was in improper position corresponded with the time in clinical description hemolysis was reported that triggered alarms impella position in ventricle/aorta. no other issues were found on the logs. pump was returned for analysis. visual inspection found no issues on the pump. the root cause of hemolysis was most likely due to pump was not in the proper position as per clinical description, hemolysis was resolved after repositioning of the pump & no issue was found on the pump under visual inspection. the root cause of the renal failure and the stroke could not be determined without additional details.
16816 | 1220648-2024-21045 | 20528250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20528250&pc=OZD | 11/29/2021 | 10/24/2024 |  | 10/24/2024 |  |  | 11/29/2021 | 10/24/2024 |  |  | 07/31/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022108720 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Multiple Organ Dysfunction Syndrome | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A patient expired during impella support due to mof (multiple organ failure). the device was reportedly working as expected. there was no allegation of a relationship between the impella and the patient's death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16817 | 1220648-2024-21060 | 20528521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20528521&pc=OZD | 04/26/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 04/26/2023 | 10/24/2024 |  |  | 09/30/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152677 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implantation, the impella .018 wire was used to load and pass the cp across the aortic valve (av). once in ventricle, the impella flow started but there was no native contractility, and impella experienced ongoing suction issues throughout the course of patient support. the patient required multiple rounds of cpr, with each attempt to balloon the clotted stents, and flow was never regained. the patient eventually succumbed to their pre-existing conditions and passed away. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the reported low pump flow issue has been completed. data analysis confirmed low flow that triggered auto suction & suction alarms when pump started & remained to the rest of the case. the product was not returned for review. based on clinical description & data analysis, the root cause of low flow was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
16818 | 1220648-2024-21039 | 20528694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20528694&pc=OZD | 06/12/2022 | 05/23/2025 |  | 10/24/2024 |  |  | 06/17/2022 | 10/24/2024 |  |  | 05/31/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083758 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for cardiomyopathy. while on support, the patient had sepsis. the patient's heart recovered, but his brain was in poor condition, and 30 minutes after the impella was removed, cardiac resynchronization therapy with a defibrillator (crtd) pacing was turned off and the patient expired. the cause of death was noted to be ventricular fibrillation storm from dilated cardiomyopathy. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16819 | 1220648-2024-21057 | 20528761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20528761&pc=OZD | 12/19/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 12/19/2022 | 10/24/2024 |  |  | 08/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268444 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had chest pain and an ambulance was called to transport the patient to the hospital. upon arrival of the patient to the hospital, the patient's condition deteriorated quickly and required the patient to be intubated, an impella cp device was placed, and the patient was taken for coronary intervention procedures. it was reported while on impella cp support, that blood-tinged urine was observed after a foley catheter was inserted, although appeared to clear as the bladder emptied. on the morning of (b)(6) 2022, the patient's nurse reported that the blood samples sent for laboratory testing were hemolyzed. the physician was concerned the patient may be experiencing disseminated intravascular coagulation (dic). due to the hemolysis the patient was experiencing and the possible dic, blood products were ordered and transfused to the patient. the physician believed the impella was not involved with the dic condition. despite aggressive efforts, the patient's condition continued to decline and the patient became extremely acidotic, which caused the patient's temporary pacer to be unable to capture. the patient's pressures continued to fall, and the patient expired while on support. there were no noted impella cp device malfunctions. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | No product was returned for investigation. the long-term log shows many suction alarms at the end of the case, likely when the device was too deep in the ventricle. position in aorta alarms are then seen likely when the pump was being repositioned. the cause of the hemolysis was most likely improper positioning of the device. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
16820 | 1220648-2024-21051 | 20528934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20528934&pc=OZD | 08/02/2021 | 10/24/2024 |  | 10/24/2024 |  |  | 08/02/2021 | 10/24/2024 |  |  | 01/31/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022060056 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient died due to retroperitoneal hemorrhage. it was noted that the device worked without issue. no further information available. | [5 empty] | The investigation for the reported retroperitoneal hemorrhage has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the hemolysis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
16821 | 1220648-2024-21044 | 20528944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20528944&pc=OZD | 03/17/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 03/17/2023 | 10/24/2024 |  |  | 11/30/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024302385 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  | Black Or African American | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, effusion occurred in the unit that was unable to be controlled. it was reported that the impella cp ingested a clot after left verntricle thrombus occurred. the medical staff was unable to achieve flow utilizing the impella. the impella was restarted however, the flow remained saturated. the resolution for this issue is unknown. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The investigation for the reported saturated flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. there was no issue found related to saturated flow as per data log review. the cause of the saturation flow issue was not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
16822 | 1220648-2024-21053 | 20528954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20528954&pc=OZD | 07/26/2021 | 10/24/2024 |  | 10/24/2024 |  |  | 10/17/2024 | 10/24/2024 |  |  | 04/30/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079822 |  | 0048-00-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was bleeding from the abdominal cavity. the team attempted to control the bleeding in the patient, but the bleeding could not be stopped. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16823 | 1220648-2024-21047 | 20528976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20528976&pc=OZD | 04/26/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 04/26/2023 | 10/24/2024 |  |  | 12/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED. INC. |  |  | US |  | 2024303735 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced a cerebral hemorrhage. it was reported that the patient expired. no further information is available. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the stroke and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16824 | 1220648-2024-21046 | 20528982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20528982&pc=OZD | 04/07/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 04/07/2023 | 10/24/2024 |  |  | 01/31/2025 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2024311882 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Consumer | Health Professional |  |  |  | 63 YR |  | Female |  | Non Hispanic |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped and was able to restart but only on p1, pump stops when they go to p2. it was noted that the patient was ineligible for pump exchange due to thrombotic state. additionally, it was noted that the procedural outcome was patient expired. | [5 empty] | B2: date of patient death is unknown the investigation for the reported pump stop issue has been completed. the impella device was not received from the customer. it was noted that the pump was used beyond the impella cp design life. the clinical details provided were sufficient to determine a root cause. the cause of the pump stop issue is that the pump was used beyond impella cp design life. this report is being filed as part of a retrospective review of historical records.
16825 | 1220648-2024-21042 | 20529140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20529140&pc=OZD | 10/06/2024 | 12/06/2024 |  | 10/24/2024 |  |  | 10/06/2024 | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485258 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had the peel away sheath left in, based on the physician preference. over time as the patient swelling/volume status changed, forward tension caused the repositioning sheath to negate forward into the peel away causing oozing from the diaphragm. the repositioning sheath was pulled back slightly, the site was re-sutured and redressed. some oozing seemingly was still originating from the antegrade perfusion sheath site and hemostatic pads were applied to site. additionally, a new bleeding event occurred due to bleeding around the peel away skin entry site. there was an existing woggle suture in place that was tightened and manual pressure held for 15 min above the sheath. site ooze slowed significantly and was redressed. site appeared dry. blood products were administered. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. returned product was investigated and there were no visual abnormalities noted either on the introducer or the pump. the introducer was put through a leak test and it passed. based on clinical details provided and returned product, the root cause of the access site bleed was determined to be a use issue. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21042. d9 device returned to manufacturer: should have been marked yes. g1 reporting contact email. g1 reporting contact fax number missing.
16826 | 1220648-2024-21049 | 20529202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20529202&pc=OZD | 05/25/2023 | 10/09/2025 |  | 10/24/2024 |  |  | 05/25/2023 | 10/24/2024 |  |  | 12/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2024326308 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited high purge pressure which resulted in a pump stop. the device restarted and the patient coded. it was reported that the patient died due to organ failure. no further information was provided. | The complainant reported high purge pressure, pump stop on ic. called into unit and rn reported that they had coded the patient three times, impella noted to be running at p3. the pump stopped multiple times but was able to be restarted, patient continued to decline and expired while on support. |  |  |  |  | B2: date of patient death is unknown. the investigation for the high-pressure purge and pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the details provided were sufficient to determine a root cause. the cause of the high-pressure purge and pump was most likely biomaterial. the investigation for the cardiac arrest is in process. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the lt log shows a motor current spike. there is a brief high motor current pump stop at the time of the motor current spike. following the motor current spike and pump stop there is high purge pressure and saturated flows. there are several more high motor current pump stops but the pump is able to be restarted and the high purge pressure resolves. all of these events suggest biomaterial interaction with the pump. the cause of the pump stop was most likely biomaterial. the root cause of the high purge pressure and cardiac arrhythmia could not be determined as insufficient clinical details were provided.
16827 | 1220648-2024-21059 | 20529213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20529213&pc=OZD | 04/26/2023 | 05/22/2025 |  | 10/24/2024 |  |  | 04/26/2023 | 10/24/2024 |  |  | 10/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290754 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, limb pulses absent in right lower extremity. distal pulse not found with doppler, but popliteal pulse palpable. escalation of care to 5.5 as a bridge to decision. the impella cp was removed. later, the family decided to withdraw care of multimorbid patient, patient expired subsequently with no allegations made toward the impella cp as the cause of death. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the limb ischemia and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. instructions for use for the related event are as follows: limb ischemia impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The impella cp product was not returned for evaluation. the root cause of the limb ischemia was unable to be determined due to insufficient clinical information. this report is being filed as part of a retrospective review of historical records.
16828 | 1220648-2024-21063 | 20529442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20529442&pc=OZD | 09/07/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 09/07/2022 | 10/24/2024 |  |  | 06/30/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094862 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support the patient developed intraperitoneal bleeding which was treated by blood replacement. it was reported that the patient expired on support due to the substantial bleed. however, the cause of the intraperitoneal bleeding is unclear as there is a possibility of blood vessels being damaged by the insertion of the intra-aortic balloon pump or the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16829 | 1220648-2024-21055 | 20529519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20529519&pc=OZD | 08/18/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 08/18/2022 | 10/24/2024 |  |  | 05/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218952 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp pump stopped during weaning after the patient touched the cable with unknown force. unable to restart the pump, pump was explanted, but the patient was stable on extracorporeal membrane oxygenation (ecmo). additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | E1: the initial reporters name is unknown. the investigation for the reported issue has been completed. the pump was plugged into an aic and the alarm was reproduced. a visual inspection of the pump revealed a kink in the catheter next to the red handle. the motor phases were tested and failed specification. the cause of the pump stop was an open electrical line due to excessive force. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
16830 | 1220648-2024-21056 | 20529581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20529581&pc=OZD | 08/22/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 08/22/2022 | 10/24/2024 |  |  | 05/31/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231145 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that during advancement of the pigtail, the sheath backed out and significant bleeding occurred. the impella was loaded and placed into the ventricle. the wire was removed, and the device was initiated. suction alarms were received. it was stated that clot ingestion was the probable cause of this issue. volume was given and the position was adjusted, however, there was no change to the pump status. the decision was made to replace the impella pump with a new device. the procedural outcome was the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. data logs showed "suction" alarms triggered following a motor current spike and low pump flow upon pump activation. the motor current quickly became non-pulsatile after the alarms occurred. visual inspection of the returned pump revealed the presence of biomaterial on and around the impeller. no other abnormalities were found. therefore, it was determined that the cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
16831 | 1220648-2024-21066 | 20529647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20529647&pc=OZD | 12/07/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 12/07/2022 | 10/24/2024 |  |  | 02/29/2024 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202567 |  | 0048-0047 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Female |  | Non Hispanic | Asian | Death | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited multiple alarms, the pump stopped and there was retrograde flow when it was on auto mode. it was reported that the flow as only 0.5. unknown how this issue was resolved. additional information noted the survival outcome at explant is patient expired. no additional information is available. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown the investigation for the reported low flow has been completed. the product was returned. per the results of the clinical review and investigation: no logs were returned. a visual inspection of the pump revealed a kink in the cannula. no other abnormalities were observed. the cause of the low pump flow was most likely a kinked cannula. logs are required to confirm the cannula kink caused the low pump during the case. this report is being filed as part of a retrospective review of historical records.
16832 | 1220648-2024-21048 | 20529703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20529703&pc=OZD | 05/11/2023 | 10/24/2024 |  | 10/24/2024 |  |  | 05/11/2023 | 10/24/2024 |  |  | 01/31/2025 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2024316888 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped and was unable to be restarted. no information regarding resolution is available. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the pump stop was not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
16833 | 1220648-2024-21062 | 20529755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20529755&pc=OZD | 03/04/2022 | 10/24/2024 |  | 10/24/2024 |  |  | 03/04/2022 | 10/24/2024 |  |  | 10/31/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED. INC. |  |  | US |  | 2022147644 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during ventricular tachycardia ablation, it became ventricular fibrillation and impella and extra-corporeal membrane oxygenation (ecmo) were inserted. however, there was vascular injury that appeared when three venous sheaths for ablation were inserted, and coiling and laparotomy hemostasis were performed for intra-abdominal bleeding. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. | [5 empty] | B2: date of patient death is unknown the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16834 | 1220648-2024-21061 | 20529949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20529949&pc=OZD | 09/28/2021 | 10/24/2024 |  | 10/24/2024 |  |  | 09/28/2021 | 10/24/2024 |  |  | 02/28/2023 |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED. INC. |  |  | US |  | 2022073209 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the died from aortic dissection. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16835 | MW5161552 | 20532635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20532635&pc=OZD | 05/10/2024 | 10/21/2024 |  | 10/24/2024 |  |  |  | 10/24/2024 |  |  |  |  | 10/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP HEART PUMP |  |  | OZD | [5 empty] | ABIOMED INC. | [8 empty] | No | [6 empty] | Health Professional |  |  | Voluntary report | Not Available | Patient Family Member or Friend | [6 empty] | 81 YR |  | Female | 59 KG | Non Hispanic | White | Death | Yes | Yes | [7 empty] | Hemorrhage/Blood Loss/Bleeding | High Blood Pressure/ Hypertension | Hypoxia | Respiratory Failure | Thrombosis/Thrombus | Heart Failure/Congestive Heart Failure | Unspecified Heart Problem | Unspecified Kidney or Urinary Problem | [5 empty] | Other | Death | [27 empty] | Initial submission | [9 empty] | Impella cp heart pump failure/defective and had to be removed and replaced with balloon pump. at the beginning of may, my 81-year-old grandmother was admitted to grandview medical center for trouble breathing/out of breath walking short periods. doctors ran several tests including ecg, cardiac ct scan, heart cath, etc.. and was found to have heart failure with only 20% ejection fraction. the doctor said she would need cabg(coronary artery bypass grafting) as soon as possible. on (b)(6) 2024, a couple of days after ecg and heart cath, she was then admitted to cardiac icu where they initially attempted to insert the impella 5.5 heart pump. according to the doctor and medical records, this pump would not fit and proceeded to insert the impella cp heart pump instead. on the afternoon of (b)(6), she was taken back for surgery. she was in surgery for at least 4-5 hours. we saw them take her back to her cicu (cardiac intensive care unit) room but weren't allowed to see her. an hour after seeing her being wheeled back, the doctor came and talked to our family and said there had been complications during the surgery and she was losing a lot of blood and they needed to get the bleeding under control. the doctor also stated the impella cp pump was defective or quit working, and a balloon pump had to be installed. the doctor left and went home. nurses still could not get the bleeding under control and said she would need to be taken back to surgery. after waiting at least 2 hours for doctor and surgery team to arrive, she was wheeled back to surgery sometime after midnight. she was in surgery for a few hours to attempt to control bleeding. after surgery, doctor said bleeding was somewhat under control but had a long road ahead. the doctor stated my grandmother had lost a lot of blood, her heart was not functioning properly, and her kidneys were shutting down. she was on a ventilator and medications to help keep her heat functioning. according to medical records, she was suffering "thrombocytopenia likely related to impella", supratherapeutic inr, renal artery stenosis, uncontrolled hypertension, hyperlipidemia, acute kidney injury, acute hypoxic respiratory failure on mechanical ventilation. post op notes from the reentry procedure in the medical records state severe left ventricle disfunction, clotted blood in the anterior mediastinum, bleeding in the proximal vein graft anastomosis of the aorta, epicardial bleeding in venous areas, epicardial bleeding on the anterior surface of the heart. over the next 24 hours, her condition continued to decline and on (b)(6) 2024 she passed away. in conclusion, the impella cp heart pump was defective and had to be removed and replaced with a balloon pump during cabg. to my knowledge this has not been reported. also, i read on the fda website there was a recall of certain impella cp heart pumps, and that is what was used during my grandmother's cabg surgery. (we have copies of medical records if needed). i have medical records showing ecg, heart cath, and several other tests conducted prior to cabg surgery and insertion of impella cp heart pump.
16836 | 1220648-2024-20869 | 20513565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20513565&pc=OZD | 09/27/2024 | 11/18/2024 |  | 10/23/2024 |  |  | 09/27/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522166 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that implant of the impella cp was not successful due to kink or bend in the cannula. a new impella cp was opened and successfully inserted. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the product damage has been completed. clinical details states that during delivery of the pump, the cannula got kinked. the cannula did not kink during support. returned product was investigated and a slight kink was noted in the cannula near the outflow cage. there were no other visual abnormalities. clinical details do not provide much detail that would suggest why delivery was difficult, as the rep was not present for that part of the case. however, the second pump was able to be delivered without issue. based on returned product and clinical details, the root cause of the delivery issues was determined to most likely be a use issue. corrections to the initial mfr report: h6 impact code 4627 changed to 4627.
16837 | 1220648-2024-20871 | 20513766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20513766&pc=OZD | 09/29/2024 | 11/27/2024 |  | 10/23/2024 |  |  | 09/29/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544867 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 35-year-old female patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the patient developed a small hematoma at the access site. a pressure dressing was placed, and two units of rbcs were given. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16838 | 1220648-2024-20870 | 20513788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20513788&pc=OZD | 09/28/2024 | 12/04/2024 |  | 10/23/2024 |  |  | 09/28/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547514 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the impella cp is in the papillary muscle. unable to get a better position. care was ultimately withdrawn. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. product was not returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. corrections to the initial mfr report: g1-mfr site name and address h6 impact code 1802 was included in error.
16839 | 1220648-2024-20872 | 20514335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514335&pc=OZD | 09/23/2024 | 11/22/2024 |  | 10/23/2024 |  |  | 09/23/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025527601 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 64-year-old female patient emergently was placed on bypass due to arrest following a coronary artery bypass graft procedure, and the impella 5.5 was attempted to be implanted. due to the location of cannulas for bypass, the impella 5.5 was unable to pass via direct insertion. insertion was then attempted via the right axillary artery. impella insertion through the graft was not achievable due to the patient anatomy. a lot of force was applied when attempting to insert the pump which led to massive bleeding from anastomosis. the patient required a lot of blood products to stabilize. insertion was aborted and impella cp was placed via the femoral artery instead. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of access site bleeding and delivery issues has been completed. the impella device was not returned for evaluation. access site bleeding - the root cause of the access site bleeding was most likely due to use issue since a lot of force was applied when attempting to insert pump which led to massive bleeding from anastomosis. delivery issue - the root cause of the delivery issue was most likely patient condition since impella insertion through graft was not achievable due to patient anatomy.
16840 | 1220648-2024-20874 | 20514421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514421&pc=OZD | 09/30/2024 | 11/26/2024 |  | 10/23/2024 |  |  | 09/30/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525282 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 73-year-old male patient was on impella cp support due to st elevated myocardial infarction (stemi). while on support, initial echo for impella placement showed a small, acute ventricular septal defect (vsd) and mural thrombus. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20874.
16841 | 1220648-2024-20875 | 20514460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514460&pc=OZD | 09/30/2024 | 11/23/2024 |  | 10/23/2024 |  |  | 09/30/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522379 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 74-year-old female patient with caridomyopathy and other systemic comorbidities was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient was going to be escalated to an impella 5.5. tee revealed a large thrombus that formed around the impella cp catheter. the 5.5 implant was postponed and the patient was received an aggressive anticoagulation to dissolve the clot. the patient also received a thrombectomy to the iliac artery and 3 clots were evacuated. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-20875.
16842 | 1220648-2024-20876 | 20514462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514462&pc=OZD | 10/01/2024 | 12/04/2024 |  | 10/23/2024 |  |  | 10/01/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544898 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported that the impella 5.5 would not advance through double surgical aortic valve (av) and transcatheter aortic valve replacement (tavr) valve. decision was made to abort procedure as impella was unable to pass into left ventricle (lv) and catheter was kinked in the process. the patient¿s outcome is unknown. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the delivery issues and product damage has been completed. the impella 5.5 pump was returned and was visually inspected. catheter and cannula kinks observed. no other abnormality found. the cause of the delivery issue was patient condition related due to transcatheter aortic valve replacement (tavr) interaction and the kinks occurred during insertion process. added-d10 concomitant medical products.
16843 | 1220648-2024-20873 | 20514515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514515&pc=OZD | 09/23/2024 | 11/27/2024 |  | 10/23/2024 |  |  | 09/23/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440961 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp was bleeding continually from the impella access site. supra compresses and a pressure bandage were applied. the patient groin bleeding was controlled with anticoagulation management and sheath repositioning. the patient eventually expired due to multi-organ failure. medical safety has reviewed the event and has stated that use of the impella device cannot conclusively be associated with the outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was most likely anticoagulation management since bleeding was controlled with anticoagulation management.
16844 | 1220648-2024-20878 | 20514565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514565&pc=OZD | 09/25/2024 | 12/02/2024 |  | 10/23/2024 |  |  | 10/03/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539109 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Heart Block | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump support ongoing for a percutaneous coronary intervention and then in the intensive care unit (icu). while on support in the icu, the patient did go into heart block and a pacer was placed. the pump was explanted and upon explant, a cerebral vascular accident was noted exhibiting signs of stroke confirmed on imaging. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into stroke/cva and heart block has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
16845 | 1220648-2024-20879 | 20514575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514575&pc=OZD | 10/02/2024 | 12/11/2024 |  | 10/23/2024 |  |  | 10/02/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025537485 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and there was a leak inside the reinforced case attached to the purge side arm. there was no significant change in purge flow rate or pressure. at that stage, since there has been no significant change in the purge flow rate/pressure, it didn't seem that the pump will need to be replaced immediately, but if any changes it may be necessary to replace the pump. the staff monitored and support continues therefore explant date and patient outcome are unknown. | [5 empty] | The investigation for the purge leak has been completed. data logs were investigated, and there were no purge alarms that would suggest there was a leak. the purge flow rates were around 8 ml/hr at case start and slowly trended down toward 6 ml/hr over the course of the next 30 days. this is very much gradual and still well within the specification. product was not returned for investigation, however, photo of the purge sidearm was submitted. in that photo, there appears to be some fluid build up in the retainer around the reservoir. however, without product being returned to reproduce the leak, it cannot be confirmed that this fluid build up was due to a leak. an improper luer connection during case start could also cause this for example. due to product not being returned and no clear sign of a leak in the data logs, the root cause of the purge leak could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
16846 | 1220648-2024-20877 | 20514576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514576&pc=OZD | 10/02/2024 | 12/06/2024 |  | 10/23/2024 |  |  | 10/02/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025531908 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Difficult to Remove | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during the impella 5.5 explant there was resistance in removing the impella. the graft was cut down to remove the catheter and a hard clot on the outlet of the impella was identified. the patient is stable and no harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the thrombus issue has been completed. the product was retuned with the catheter cut. no other signs of damage were found on the product. biomaterial was seen adhered to the outside of the device. clinical details state that a clot made the removal of the pump difficult. therefore, the root cause of the removal issue was most likely patient condition related. added- h6 med dev prob code 1528.
16847 | 1220648-2024-20880 | 20514616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514616&pc=OZD | 10/02/2024 | 12/13/2024 |  | 10/23/2024 |  |  | 10/02/2024 | 10/23/2024 |  |  | 09/30/2025 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400895 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that percutaneous coronary intervention was being performed with impella cp support and there was concern of bleeding in the patient¿s groin, trying to assess it pushing with the proglides. when the patient was stable the team decided to explant impella due to continuous bleeding. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
16848 | 1220648-2024-20881 | 20514684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514684&pc=OZD | 10/02/2024 | 09/22/2025 |  | 10/23/2024 |  |  | 10/03/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546777 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was on the impella 5.5 for support. it was reported that after removing the 5.5 in the operation room the patient exhibited signs of a stroke. the stroke was confirmed on computed tomography scan and the patient was sent for a thrombectomy. the procedural outcome was the patient survived surgery. | Additional information indicated that, during patient support, the pump experienced placement signal not reliable alarm. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-20881. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. the product was not returned. the data logs were observed and this issue is related to the console not the pump. there is no problem found for the pump for the optical signal issue. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. b5 updated with additional information. g1 updated with correct mdr reporting contact email. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-20881.
16849 | 1220648-2024-20886 | 20514945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514945&pc=OZD | 09/22/2024 | 12/06/2024 |  | 10/23/2024 |  |  | 09/23/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510885 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that during impella 5.5 support the 67-year-old male patient was bleeding from chest tube. the patient received 3 units of red blood cells. it is uncertain if bleeding is somehow impella related. the patient expired while on support. medical safety review of the event noted there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: isection: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the vascular damage has been completed. no product was returned for evaluation. clinical details noted that patient was bleeding from chest tubes and experiencing mediastinum. pump was in proper position at time of bleeding. the failure mode of "vascular damage - peripheral vessel" was changed to "vascular damage - great vessel" as the pump was inserted via direct aortic access via graft. the root cause of the vascular damage - great vessel cannot be definitively determined as limited clinical details were provided.
16850 | 1220648-2024-20882 | 20514947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514947&pc=OZD | 02/06/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 02/06/2022 | 10/23/2024 |  |  | 10/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152131 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump exhibited low flows, and the patient developed a hematoma/bleeding at the access site when integrilin was introduced. a terumo radial band was placed on the radial artery in unit and the groin was redressed and resutured. the pump was replaced without further injury to the patient. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was returned. the provided data logs only contained approximately 30 minutes of data from (b)(6) 2022 while on aic ic4392. however, a transient high motor current event was observed which was followed by suction alarms, a loss of motor current pulsatility, and a reduction in pump flow. visual inspection of the returned pump revealed biomaterial around the impeller. therefore, the root cause of the low pump flow was ingested biomaterial. refer to es2020-082 for further information. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16851 | 1220648-2024-20883 | 20514985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20514985&pc=OZD | 09/20/2024 | 12/13/2024 |  | 10/23/2024 |  |  | 09/29/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025477698 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp migrated to the left ventricle (lv) while exchanging the introducer by the repositioning sheath. the impella bended in the lv and was stuck. after several maneuvers, the pump was straitened and retrieved. a new impella was successfully placed. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. the cause of the positioning issue most likely was use related due to improper technique exchanging the sheath with the repo unit. g.1 revised reporting contact email since initial manufacturer device report number 1220648-2024-20883 was submitted.
16852 | 1220648-2024-20887 | 20515074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515074&pc=OZD | 12/12/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 12/14/2021 | 10/23/2024 |  |  | 04/30/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079845 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had an intraperitoneal bleed while on support and blood transfusions were given. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the bleeding was non-obstructive intestinal ischemia and the relation to impella was determined to be unrelated. the patient died due to infection due to non-obstructive intestinal ischemia. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16853 | 1220648-2024-20889 | 20515079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515079&pc=OZD | 02/02/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 02/02/2022 | 10/23/2024 |  |  | 08/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127499 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Perforation of Vessels | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. while on support, patient experienced cardiac tamponade and hemolysis. no intervention was specified. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the field data logs were reviewed, and no abnormal alarms were captured throughout the approximately seven-hour case. the lack of provided event information regarding the reported cardiac tamponade and hemolysis indicate an unknown relation to impella. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. this report is being filed as part of a retrospective review of historical records.
16854 | 1220648-2024-20888 | 20515118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515118&pc=OZD | 09/28/2024 | 11/26/2024 |  | 10/23/2024 |  |  | 09/28/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529623 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to st elevated myocardial infarction, significant left anterior descending stenosis. while on support, bleeding was noted at impella site, arterial extracorporeal membrane oxygenation sites, and central line sites. all sites were redressed, and the patient received two unites of packed red blood cells and one unit of fresh frozen plasma. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was determined to be patient condition because of the bleeding from multiple sites. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16855 | 1220648-2024-20884 | 20515126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515126&pc=OZD | 12/14/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 12/14/2021 | 10/23/2024 |  |  | 07/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022110769 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Renal Failure | [10 empty] | Life Threatening | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, patient experienced hemolysis, limb ischemia, and renal failure. no intervention was specified. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported hemolysis, limb ischemia and renal failure has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the hemolysis was not determined. the cause of the limb ischemia was most likely the condition/size of the patient vessel. there were no reports of renal failure and the relation to impella was determined to be unknown/unrelated. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system (hemolysis) section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ as noted in the impella ifu: impella cp with smartassist system (limb ischemia and renal failure) section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
16856 | 1220648-2024-20885 | 20515167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515167&pc=OZD | 09/29/2024 | 11/27/2024 |  | 10/23/2024 |  |  | 09/29/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531601 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was placed across the valve and started without issue. however, the catheter was pulled into the aorta when the peel away sheath was removed. the impella was unable to be readvanced across the aortic valve. the decision was made to remove the impella cp. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported positioning issue has been completed. product was not returned for review. the root cause of positioning issue was use related since the pump pulled into the aorta during the peel away process.
16857 | 1220648-2024-20890 | 20515579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515579&pc=OZD | 09/30/2024 | 11/18/2024 |  | 10/23/2024 |  |  | 09/30/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548314 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and the patient was bleeding from the impella insertion site after the peel away sheath was exchanged. it was reported that there was leaking from the repositioning sheath above the blue hub. the pump was explanted. the patient survived the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported issue has been completed. the impella device was received from the customer and the returned product was visually inspected and there was a leak observed between blue butterfly and white repo hub .the root cause of the access site bleeding issue was most likely use related, repo sheath is leaking above blue hub as per the clinical description provided.
16858 | 1220648-2024-20893 | 20515661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515661&pc=OZD | 09/28/2024 | 10/23/2024 |  | 10/23/2024 |  |  | 09/28/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025525448 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella 5.5 pump for support of a 55-year-old male patient admitted in with an acute myocardial infarction / cardiogenic shock (ami/cgs). the access site had a bleed. the bleed was treated by the infusion of blood products. pump remains on. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16859 | 1220648-2024-20891 | 20515709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515709&pc=OZD | 07/01/2023 | 09/24/2025 |  | 10/23/2024 |  |  | 07/01/2023 | 10/23/2024 |  |  | 11/30/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169296 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp optical had low flows issues. there was a possible clot ingestion although the patient was anticoagulated prior to insertion. the impella was replaced successfully with a new pump. the survival outcome at explant noted the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | Us complainant reported the impella cp optical had low flows issues. placement signal low. there was a possible clot ingestion although patient was anticoagulated prior to insertion. the impella was replaced with a new pump and it was successful. the patient expired and it is uncertain if this was due to the impella. |  |  |  |  | The investigation for the reported low flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the total case runtime was 9 minutes with a reported low flow and a placement signal low alarm. data logs show the low flow from the beginning of the case. all the 5s parameters were reported to be within specification. the doctor utilized the second pump and experienced the same issue with low flow while running on p-level above p2 and placement signal low alarm with healthy 5s parameters. the cause of the low flow issue was most likely patient condition related. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The pump was not returned for investigation. the total case runtime was 9 minutes with a reported low flow and a placement signal low alarm. data logs show the low flow from the beginning of the case. all the 5s parameters were reported to be within specification. the doctor utilized the second pump and experienced the same issue with low flow while running on p-level above p2 and placement signal low alarm with healthy 5s parameters. the cause of the low flow issue was most likely patient condition related. the cause of the optical signal issue was most likely patient condition related. h6: added optical signal issue as part of a retrospective review of optical signal issues in accordance with fda recommendation.
16860 | 1220648-2024-20894 | 20515727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515727&pc=OZD | 03/24/2024 | 11/27/2024 |  | 10/23/2024 |  |  | 09/30/2024 | 10/23/2024 |  |  | 10/31/2025 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422644 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured bleeding with a probable relationship to the impella device used: patient with baseline hemoglobin/hematocrit (h/h) on (b)(6) 2024 at 22:54 of 13.0/40.0. per nurse, on (b)(6) 2024 patient had impella groin access site bleeding. impella access entry site adjusted, and resutured for hemostasis. (b)(6) 2024 h/h was 6.1/18.3. patient transfused with two units prbc. anemia was much improved by (b)(6) 2024 with h/h of 8.8/26.2. the patient recovered with sequelae. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.the root cause of access site bleeding was determined to be use issue because the bleeding was resolved by adjusting and resuturing the access site.
16861 | 1220648-2024-20892 | 20515838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515838&pc=OZD | 10/15/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 10/17/2024 | 10/23/2024 |  |  | 07/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121752 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient presented with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. after five days of support, the impella had a purge system leak. the leak was remedied by the purge bypass. it was also noted there was an ooze at the impella insertion site. manual pressure was applied to the site, the angle of the repositioning sheath was corrected and decreased heparin drip improved the ooze at the site. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported access site bleeding and purge leak has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issues could not be determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
16862 | 1220648-2024-20902 | 20515894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515894&pc=OZD | 09/30/2024 | 11/18/2024 |  | 10/23/2024 |  |  | 09/30/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563312 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and it was difficult to doppler pedis pulse and the patient's foot was cool to the touch. the doctor decided to do an exploration of the femoral artery. cutdown was performed, the impella was removed, and a thrombectomy was performed with a 4 fr fogarty with some thrombus noted. the vessel was patched with sutures and distal pulses returned. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: d4-udi number.
16863 | 1220648-2024-20897 | 20515968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515968&pc=OZD | 09/29/2024 | 12/13/2024 |  | 10/23/2024 |  |  | 09/30/2024 | 10/23/2024 |  |  | 06/30/2025 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366189 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 13 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella cp and guidewire. while preparing to insert impella, the wire was bent inside the cannula when the wire was passed through while holding the guidewire induction tube so that it would not come out. at the discretion of the physician, impella was inserted in the above state, but after placement in the left ventricle, the guidewire was stuck and could not be removed from the main impella unit. the pump and guidewire were removed once. outside the body, the guidewire was removed from the main impella unit, but there was concern that a broken part had become mixed into the cannula, so the pump was replaced. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: warnings & cautions: warnings; section: pre-support evaluation; section: direct aortic insertion; section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely use related, while preparing to insert impella, the wire was bent inside the cannula when the wire was passed through while holding the guidewire induction tube so that it would not come out, but after placement in the left ventricle, the guidewire was stuck and could not be removed from the main impella unit. this led to the removal of the guidewire and pump being replaced as per the clinical description provided. corrections to the initial mfr report: d4 udi number was missing g1 mdr reporting contact email updated.
16864 | 1220648-2024-20896 | 20515972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515972&pc=OZD | 09/29/2024 | 11/27/2024 |  | 10/23/2024 |  |  | 09/29/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533834 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 59-year-old male patient was on impella cp support due to cardiac arrest. while on support, the patient went into ventricular fibrillation (vf) arrested. cardio-pulmonary resuscitation (cpr) was required. pump placement was checked post cpr and was found to be at a left ventricle (lv) measurement of 1.5cm. pump was advanced 3cm to 96cm but remained shallow and difficult to advance into left ventricle (lv). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined.
16865 | 1220648-2024-20898 | 20515977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515977&pc=OZD | 09/30/2024 | 12/12/2024 |  | 10/23/2024 |  |  | 09/30/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510270 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support. while on support, purge-lysis ran with slow purge flow. purge pressure high / purge flow low reported. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. no product was returned. log review showed a rise in purge pressure beginning about 2 hours into support. purge flow also began just below 4 ml/hr. the increase in purge pressure was gradual from ~600 to 1100 mmhg over the course of 2-3 hours, indicating biomaterial buildup. there were no large motor current spikes, but there was a motor current downward spike that correlated with the beginning of the purge pressure rise. the high purge pressure/low purge low persisted for about 20 hours before resolving. there were no issues seen in logs with purge pressure throughout the rest of the case. purge flow remained low (below 4 ml/hr). the root cause of the high purge pressure was most likely biomaterial.
16866 | 1220648-2024-20895 | 20515999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20515999&pc=OZD | 09/29/2024 | 12/19/2024 |  | 10/23/2024 |  |  | 09/29/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513012 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was explanted due to positioning issue. after explant, the cannula was full of biomaterial. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the low pump flow, thrombus, and the positioning issues has been completed. data logs was not returned for review. pump was returned in damaged condition (catheter was cut off and missing pump plug). visual inspection found biomaterial inside cannula and pump housing. no other issues were found on the rest of the pump. the root cause of the low flow was biomaterial ingestion. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no data & images related to pump position were returned for review. corrections to the initial mfr report: b5-it should be noted that the there was initially pump flow issues in additional to the previously reported positioning issues. g1 mdr reporting contact email updated. added h6 med dev prob code 1248.
16867 | 1220648-2024-20899 | 20516129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20516129&pc=OZD | 12/03/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 10/17/2024 | 10/23/2024 |  |  | 06/30/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100658 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced hemolysis that began after the impella cp was inserted. no information was provided if any interventions were performed to treat the hemolysis. it was noted that the patient's heart function did not recover, and the patient expired. the physician determined there was no relation between impella and the patient's outcome. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
16868 | 1220648-2024-20903 | 20516143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20516143&pc=OZD | 10/14/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 11/08/2021 | 10/23/2024 |  |  | 01/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022048979 |  | 0048-0024-JP |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced bleeding with a hematoma, requiring a blood transfusion. it was also reported that the patient was progressively septic. no further information was provided as to the treatment for the sepsis. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported access site bleeding and sepsis has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16869 | 1220648-2024-20900 | 20516194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20516194&pc=OZD | 09/30/2024 | 11/18/2024 |  | 10/23/2024 |  |  | 09/30/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544106 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 64-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impella support and after the implant, flows were 2.5 l. a cannula kink was noted under fluoroscopy and the physician was unable to straighten out the kink. the doctor was happy with the subpar flows and support continued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into low pump flow has been completed. the data logs confirmed low pump flow when pump started and remained to the rest of the case. visual inspection revealed a kink on can about 15 mm from of cage and can bonding site. no other issues were found on the rest of the pump. the root cause of low flow was kinked cannula caused by user technique.
16870 | 1220648-2024-20905 | 20516665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20516665&pc=OZD | 11/24/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 10/17/2024 | 10/23/2024 |  |  | 08/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022122108 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention (pci) was implanted with ecpella (impella cp and extracorporeal membrane oxygenation) for mechanical circulatory support. after the ecpella was established, the patient's vital signs did not improve, and ventricular fibrillation occurred frequently during the pci. the impella was removed in the catheterization room, and the patient returned to the icu, where he passed away a few hours later. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
16871 | 1220648-2024-20901 | 20516699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20516699&pc=OZD | 08/09/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 10/19/2021 | 10/23/2024 |  |  | 02/28/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022073206 |  | 0048-0024-JP |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced hemolysis. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported hemolysis issues has been completed. the impella device has not been returned for evaluation. the cause of the hemolysis was not determined since product was not returned. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
16872 | 1220648-2024-20906 | 20516719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20516719&pc=OZD | 11/15/2021 | 10/09/2025 |  | 10/23/2024 |  |  | 08/17/2024 | 10/23/2024 |  |  | 08/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119831 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | Asian | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during use of the device, the patient experienced thrombocytopenia and hemolysis. no intervention or resolution was reported for these issues. the patient ultimately expired, but there was no allegation that this was related to the impella. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The complainant reported thrombocytopenia during use of the device. no intervention or premature device removal was indicated. also, cp support ongoing in japan in (b)(6) 2021. later in (b)(6) 2022 the team completed the jpvad and noted thrombocytopenia, occurred on (b)(6) 2021 however no intervention was done, and unknown if this is related to the impella at this time. patient expired |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system hemolysis: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ thrombocytopenia: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The complainant reported thrombocytopenia during use of the device. no intervention or premature device removal was indicated. also, cp support ongoing in japan in (b)(6) 2021. later in (b)(6) 2022 the team completed the jpvad and noted thrombocytopenia, occurred on (b)(6) 2021 however no intervention was done, and unknown if this is related to the impella at this time.
16873 | 1220648-2024-20907 | 20516960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20516960&pc=OZD | 11/25/2021 | 05/23/2025 |  | 10/23/2024 |  |  | 10/17/2024 | 10/23/2024 |  |  | 08/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022118814 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support hemolysis was observed and the device was explanted. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16874 | 1220648-2024-20908 | 20517001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517001&pc=OZD | 11/14/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 11/29/2021 | 10/23/2024 |  |  | 05/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022089496 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | Asian | Death | Yes | No | Defective Device | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a japanese patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had a cp support of 9 days when the patient expired due to underlying condition. the complainant noted a guidewire damage that was observed prior to the pump usage. inserted independently. 0.018 wire broke the tip or coil but was fixed by using a second one. there was no harm or injury noted as the product was replaced and not used for the deployment of the impella pump. fatal arrhythmia occurred during impella management, and ecmo was added. after that, patient died without recovering. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16875 | 1220648-2024-20909 | 20517056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517056&pc=OZD | 11/27/2021 | 05/23/2025 |  | 10/23/2024 |  |  | 10/17/2024 | 10/23/2024 |  |  | 08/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130272 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced bleeding and hematoma, requiring a blood transfusion. in addition it was reported that the patient experienced a stroke not related to the impella. the patient ultimately expired, but there was no allegation that this was related to the impella. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16876 | 1220648-2024-20904 | 20517096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517096&pc=OZD | 10/01/2024 | 12/10/2024 |  | 10/23/2024 |  |  | 10/01/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025503303 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 79-year-old male patient was on impella cp support due to occluded left main. while on support, suction alarm with 0.0 l/min flow. there were concerns for clot ingestion leading to motor failure. occasional clot noticed in tubing. | The patient also experienced hemolysis while on support. the patient family decided to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into low pump flow, thrombus, and hemolysis has been completed. the returned complaint device visually inspected. no biomaterial, no cannula kink and no broken blade observed. no other abnormality found. returned pump connected to aic to reproduce the low pump flow issue. pump run for 10 minutes at p-9. flow was within required range flow with out issue. based on the patient clinical data the low pump flow issue reproduced with the replacement pump too. the low pump flow issue most likely was patient condition related. the cause of the hemolysis issue was improper positioning related due to improper technique and impella left little shallow. as the necessary information was not provided, the root cause of the thrombus was not determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20904. b2 death and date of death. b5 additional failure mode that was inadvertently missed and patient outcome. g1 reporting contact email. h1 type of reportable event. h10 health effect - impact code 2199 was reported incorrectly. new code was added to health effect - clinical code and health effect - impact code.
16877 | 1220648-2024-20911 | 20517101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517101&pc=OZD | 03/17/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 01/05/2022 | 10/23/2024 |  |  | 10/31/2022 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021011238 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was malnourished, bedridden, and had low albumin levels even before the insertion of the impella. vascular permeability had progressed even before the impella was inserted, and even when volume was applied, it quickly escaped from the blood vessels. while on support, it was determined that it would be difficult to continue the impella in this condition, and the patient was weaned. the survival outcome at explant noted that the patient expired four (4) days after the pump was explanted due to cardiogenic shock. the sudden cardiac death and its relationship to the impella was noted as possibly related. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16878 | 1220648-2024-20910 | 20517180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517180&pc=OZD | 11/10/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 11/21/2022 | 10/23/2024 |  |  | 06/30/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023233923 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis during impella support. upon resumption of extracorporeal membrane oxygenation (ecmo) assistance, the issue was resolved and patient support continued. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported hemolysis issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
16879 | 1220648-2024-20913 | 20517525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517525&pc=OZD | 09/17/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 09/17/2022 | 10/23/2024 |  |  | 08/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022117297 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was blood leaking from the device's red plug. the physician chose to continue support as the device was operating normally. the survival outcome at explant noted that the patient expired. the relation between the impella and death could not be determined, and the cause of the death is unknown. | [5 empty] | The investigation for the reported hole in catheter issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned product was visually inspected, and blood was observed in the red handle and catheter kinks were observed at the 35 cm mark. a hole was observed in the catheter at the 33 cm mark. the cause of the issue was blood ingress due to a hole in the catheter resulting from improper use. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16880 | 1220648-2024-20915 | 20517573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517573&pc=OZD | 10/05/2024 | 12/17/2024 |  | 10/23/2024 |  |  | 10/05/2024 | 10/23/2024 |  |  |  |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540518 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient was on impella cp support and during support, there was high motor current followed by a pump stop. the pump was already being weaned when the pump stop occurred. no known patient harm. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop has been completed. product was not returned for investigation. data logs were investigated and on the 10th day of support, alarm 13 suddenly triggered. there were no motor current or purge abnormalities leading up to the complication. once the pump had stopped, alarm 115 triggered when the team was attempting to restart the pump. when alarm 115 triggered, the mc did not spike to 3000, so an electrical issue is not suspected. the data logs and fact that the pump was running for 10 days (impella cp design life is 8 days, per design trace matrix 0046-9910) suggest there is potential for bearing wear, but this cannot be confirmed without product return. since product was not returned for investigation, the root cause of the pump stop could not be determined.
16881 | 1220648-2024-20914 | 20517594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517594&pc=OZD | 02/16/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 02/16/2022 | 10/23/2024 |  |  | 11/30/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157960 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp supporting a patient along with extracorporeal membrane oxygenation (ecmo). after twelve days of support the pump stopped. it was noted that the impella cp was used beyond the indication for use. the patient later expired due to multiple organ failure. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the patient death was multiple organ failure and the relation to impella was determined to be unrelated. it was noted that the device was used beyond its indication for use. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
16882 | 1220648-2024-20916 | 20517629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517629&pc=OZD | 03/23/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 03/23/2022 | 10/23/2024 |  |  | 02/08/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067769 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient experienced hemolysis, the resolution for this device is unknown. the patient expired from an extensive cerebral infarction. it was noted that the impella cp was working as expected. no further information is available. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the investigation for the reported hemolysis has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the hemolysis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
16883 | 1220648-2024-20919 | 20517696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517696&pc=OZD | 01/31/2023 | 10/23/2024 |  | 10/23/2024 |  |  | 01/31/2023 | 10/23/2024 |  |  | 02/29/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202403 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for hemodynamic support. the patient coded prior to and during insertion of the pump. while on support, and while moving the patient, the staff accidentally pulled the pump out while the peel-away sheath was still in place. they were unable to advanced the pump further, and the decision was made to open a second pump which was successfully implanted. the patient coded again after successfully implanting the second pump but was unable to achieve return of spontaneous circulation (rosc}. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the log showed the "impella position in aorta" alarm as it became dislodged and was attempted to be advanced. the root cause of the positioning issue is most likely a use issue. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16884 | 1220648-2024-20917 | 20517832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517832&pc=OZD | 02/02/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 02/02/2022 | 10/23/2024 |  |  | 12/31/2022 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1534647 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. one day after normal explant of the impella cp, the patient experienced ventricle perforation (bleeding from an unknown location). unknown how this issue was resolved. no further information was provided, it was noted that the device operated as intended. the procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown the investigation for the ventricle perforation has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the reported cardiac tamponade was most likely patient condition and the relation to impella was determined to be unknown/unrelated. this report is being filed as part of a retrospective review of historical records.
16885 | 1220648-2024-20918 | 20517850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517850&pc=OZD | 02/27/2023 | 10/23/2024 |  | 10/23/2024 |  |  | 02/28/2023 | 10/23/2024 |  |  | 09/30/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023264158 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient expired from cardiogenic shock and cerebral hemorrhage. it was noted that the impella operated with no problem until the end. relevance of the death to the impella is unknown and there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16886 | 1220648-2024-20920 | 20517856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517856&pc=OZD | 08/29/2021 | 03/13/2025 |  | 10/23/2024 |  |  | 08/12/2024 | 10/23/2024 |  |  | 05/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022064816 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced access site bleeding and a hematoma that required a blood transfusion. the patient ultimately expired from acute coronary syndrome (acs). no further information is available. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | B2: date of patient death is unknown the impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16887 | 1220648-2024-20912 | 20517895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517895&pc=OZD | 06/24/2022 | 05/14/2025 |  | 10/23/2024 |  |  | 07/07/2022 | 10/23/2024 |  |  | 01/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174019 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Sepsis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had sepsis and a cerebrovascular accident (cva). the survival outcome at explant noted the patient expired. the physician noted that there was no relation of the death to the impella device. yet, there is not enough evidence to exclude the impella as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist instructions for use for the related event are as follows: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16888 | 1220648-2024-20922 | 20517947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20517947&pc=OZD | 12/20/2021 | 09/16/2025 |  | 10/23/2024 |  |  | 01/12/2022 | 10/23/2024 |  |  | 05/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022085953 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a hemorrhagic cerebrovascular accident (cva). it was also reported the patient experienced, bleeding including hematoma which required a blood transfusion. there was no information provided on how hemostasis was achieved. the impella was weaned and removed the same day following the cva, and the patient expired three days after explant. the outcome of the patient can not be disassociated from the use of the impella device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the site reported that the patient suffered a hemorrhagic cva on 24dec2021 after which impella cp was removed, and the patient ultimately expired. the reported cva and patient death were determined to be unrelated to impella as no imaging or evidence of the reported events was provided. additionally, no product or data logs were provided. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. corrections to the initial mfr report: g1 mdr reporting contact email.
16889 | 1220648-2024-20926 | 20518015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518015&pc=OZD | 07/04/2022 | 05/23/2025 |  | 10/23/2024 |  |  | 07/12/2022 | 10/23/2024 |  |  | 12/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023163689 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bleeding including hematoma on (b)(6) 2022. the patient required blood transfusions due to the blood loss. it was additionally reported the patient experienced sepsis on 6-july-2022. there was no additional information provided on treatments or interventions for the reported sepsis. the patient had poor recovery of cardiac function and was dependent on assisted circulation. the family made the decision to order the patient as do not resuscitate. the patient expired while on impella support. the patient experienced adverse events and given the details there is not enough information to dissociate the death outcome from the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16890 | 1220648-2024-20928 | 20518022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518022&pc=OZD | 02/11/2023 | 10/23/2024 |  | 10/23/2024 |  |  | 02/11/2023 | 10/23/2024 |  |  | 09/30/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275726 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported the impella cp pump was implanted to replace first cp used that had low pump flow. the impella inserted without issue and was ramping up when suddenly flows dropped to 0.6 l/min and were unable to increase. the medical team decided to remove the impella and replace it with an intra-aortic balloon pump. the patient continued to decompensate, blood pressure dropped to 50/30 mmhg and the patient began coding. the patient was administered epinephrine. the patient went into asystole and cardiopulmonary resuscitation was initiated. after 15 minutes of coding, the patient was pronounced as expired. | [5 empty] | The investigation for the reported delivery issue has been completed. the product was returned and a visual inspection of the pump revealed biomaterial on the impeller. the pump was run in a bucket of water and the low pump flow was reproduced. after the run in water the impeller was inspected again and biomaterial remained. no damage to the impeller was observed. the cause of the low pump flow was biomaterial. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16891 | 1220648-2024-20935 | 20518056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518056&pc=OZD | 08/08/2020 | 10/23/2024 |  | 10/23/2024 |  |  | 07/24/2024 | 10/23/2024 |  |  | 10/31/2021 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1453055 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had cerebral hemorrhage with an unknown relationship to the impella device. it was also reported that one month after explant, that the patient expired noting that it was possibly related to the impella device. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported stroke/cva issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16892 | 1220648-2024-20927 | 20518103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518103&pc=OZD | 07/25/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 07/28/2022 | 10/23/2024 |  |  | 08/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183302 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 32 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Life Threatening | Other | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a cerebral hemorrhage. there was no information provided on any treatments or interventions provided to the patient for the cerebral hemorrhage. it was reported the patient expired while on impella support. it was reported that it was unknown whether the death was related to impella, but stated the impella worked fine until the death. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16893 | 1220648-2024-20929 | 20518109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518109&pc=OZD | 02/10/2023 | 10/23/2024 |  | 10/23/2024 |  |  | 02/13/2023 | 10/23/2024 |  |  | 07/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269026 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, immediately following the decannulation of extracorporeal membrane oxygenation, the patient experienced an episode of ventricular fibrillation. direct current was administered three times, and the heartbeat resumed. after that, complete atrioventricular block occurred, and a temporary pacemaker was inserted. the impella was successfully weaned and explanted, however, the patient expired the day following explant due to multiple organ failure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16894 | 1220648-2024-20931 | 20518117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518117&pc=OZD | 02/09/2023 | 10/23/2024 |  | 10/23/2024 |  |  | 02/13/2023 | 10/23/2024 |  |  | 08/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251401 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  | Asian | Death | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient due to delirium, moved violently and and impella was pulled out of the patient's ventricle. before the position could be corrected, the patient's respiratory condition worsened and the patient required intubation. after the patient was intubated, they went into cardiac arrest. cardiopulmonary resuscitation was performed and return of spontaneous circulation was achieved. the impella was then removed and a new impella cp was inserted into the patient. the patient's condition worsened and expired the following day due to multiple organ failure. | [5 empty] | The investigation for the reported positioning issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the impella was pulled out due to the patient's delirium. the root cause of the positioning issues is due to patient movement. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ this report is being filed as part of a retrospective review of historical records.
16895 | 1220648-2024-20934 | 20518252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518252&pc=OZD | 10/13/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 10/13/2022 | 10/23/2024 |  |  | 11/30/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157939 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a sign of bleeding due to lack of volume. on the second day of support an image revealed aortic dissection. it was noted that the patient expired and was noted as circulatory failure. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported vascular damage issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the aortic dissection and its relationship to the impella were not determined since neither the product nor sufficient clinical information were provided. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16896 | 1220648-2024-20923 | 20518259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518259&pc=OZD | 01/18/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 01/18/2022 | 10/23/2024 |  |  | 07/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED. INC. |  |  | US |  | 2022114703 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed hemolysis, the physician decided to explant the device and replace it with an intra-aortic balloon pump. additional information reported that the patient expired. no further information is available. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
16897 | 1220648-2024-20924 | 20518266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518266&pc=OZD | 05/30/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 06/22/2022 | 10/23/2024 |  |  | 02/29/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023192046 |  | 0048-0024-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Cardiac Tamponade | Cardiac Perforation | [9 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with arrhythmogenic cardiomyopathy in ventricular fibrillation was implanted with an impella cp device for mechanical circulatory support. it was reported during impella cp support the patient experienced hemolysis, a serious bleed from an unknown location which required a blood transfusion, and a ventricular perforation. the patient experienced cardiac tamponade as a result of the perforation. it was unknown the exact timing or cause of the cardiac tamponade, however it was discovered during impella cp support. the patient condition was already poor, decompensated after the discovery of cardiac tamponade, and suddenly expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
16898 | 1220648-2024-20939 | 20518272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518272&pc=OZD | 10/28/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 10/28/2022 | 10/23/2024 |  |  | 02/29/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195428 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the purge side arm was broken while turning the patient. while waiting for the physician to decide on purge bypass, the pump stopped and would not restart. the device was successfully changed for a new one. further notes state that the patient coded several times, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported cassette leak and pump stop issues has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the product was visually inspected. the sidearm was observed to be damaged at the joint from filter to reservoir. buildup was observed in the outflow cage around the impeller. the product was soaked in warm water for 10 minutes. biomaterial was observed around the impeller. per clinical details, the sidearm was damaged while turning the patient. long term logs were reviewed. a sudden drop in purge pressure was observed. shortly after, there was a steady rise in motor current until eventually the pump stopped. due to the lack of purge flowing through the motor underneath the impeller, this most likely caused biomaterial to build up and caused the pump stop. the cause of the issue was biomaterial. the cause of the pump stop was equipment being placed on the sidearm causing the sidearm to break and a purge leak resulting in the biomaterial ingress pump stop. the cause of the purge leak was damage to the reservoir to filter joint. equipment was placed on top of the sidearm causing the sidearm to break and purge to leak. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
16899 | 1220648-2024-20933 | 20518294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518294&pc=OZD | 09/03/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 09/27/2022 | 10/23/2024 |  |  | 04/30/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215759 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a hemorrhagic stroke. after removal of the impella, the patient had a non-cardiac death which was noted to be possibly related to the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16900 | 1220648-2024-20938 | 20518322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518322&pc=OZD | 10/20/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 10/21/2022 | 10/23/2024 |  |  | 05/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217440 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, it was noted that the patient¿s anatomy was very difficult. tortuosity and artheriosclerosis led to a cathether kink during the impella cp insertion. a new impella was successfully inserted. once positioned, it was reported that the pump would not start, and the patient was too ill and expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump did not start issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected and biomaterial was observed on the impeller. the root cause of the issue was determined to be biomaterial. impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
16901 | 1220648-2024-20925 | 20518327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518327&pc=OZD | 06/28/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 06/28/2022 | 10/23/2024 |  |  | 02/29/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191492 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced heavy bleeding from the insertion site of the catheter inserted into the patient's neck. the patient received blood transfusions, however the blood could not be transfused fast enough and the patient expired due to multiple organ failure. it is was noted that the patient was receiving heparin for impella use at the time of the bleed, which may have contributed to the bleed and the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16902 | 1220648-2024-20941 | 20518345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518345&pc=OZD | 03/08/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 03/08/2022 | 10/23/2024 |  |  | 04/30/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022081128 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding/hematoma was noted at the access site. blood transfusion was required. impella was successfully weaned and explanted; however, cerebrovascular accident (cva)/stroke occurred 4 days post explant. patient subsequently expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular .perforation).¿
16903 | 1220648-2024-20942 | 20518348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518348&pc=OZD | 02/02/2023 | 10/23/2024 |  | 10/23/2024 |  |  | 02/02/2023 | 10/23/2024 |  |  | 08/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023271755 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient's foot on the impella side became mottled and blue during impella support. the patient was provided a set of pressors to assist with blood flow and no further intervention was performed for the condition. patient's condition subsequently deteriorated, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient's foot on the impella side became mottled and blue during impella support. the patient was provided a set of pressors to assist with blood flow and no further intervention was performed for the condition. patient's condition subsequently deteriorated, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation for the reported ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of ischemia was patient condition.
16904 | 1220648-2024-20930 | 20518353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518353&pc=OZD | 02/15/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 02/15/2022 | 10/23/2024 |  |  | 09/30/2021 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED. INC. |  |  | US |  | 1444716 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after device removal the patient experienced a cerebral hemorrhage and died. no further information is available. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16905 | 1220648-2024-20947 | 20518465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518465&pc=OZD | 08/05/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 08/05/2022 | 10/23/2024 |  |  | 03/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200393 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device along with extracorporeal membrane oxygenation for mechanical circulatory support. while on support the patient was confirmed to have a cerebral hemorrhage. heparin remained in the purge. anti-coagulation is known to be a contributing factor to cerebral hemorrhage. in addition, hemolysis was found which was resolved via placement position with an echocardiogram. additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm wheth er a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
16906 | 1220648-2024-20950 | 20518591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518591&pc=OZD | 02/08/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 02/08/2022 | 10/23/2024 |  |  | 10/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144002 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after inserting impella, echocardiogram revealed m-valve chordae tendineae rupture. regarding the m-valve chordae tendineae rupture, the vital signs did not improve even after the insertion of impella, and the physician noted that impella did not rupture from the original patient's condition, but that it was originally ruptured. the patient was not in a state of being able to undergo surgery. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16907 | 1220648-2024-20937 | 20518664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518664&pc=OZD | 06/13/2023 | 10/23/2024 |  | 10/23/2024 |  |  | 06/13/2023 | 10/23/2024 |  |  | 01/31/2025 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312731 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 92 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with severe aortic stenosis went into cardiopulmonary arrest when balloon aortic valvuloplasty was performed. cardiopulmonary resuscitation was performed and an impella cp device was inserted for mechanical circulatory support. the patient had a return of spontaneous circulation and was hemodynamically stable on impella cp support after the cardiac event. on (b)(6) 2023, computed tomography was performed for brain evaluation and hypoxic encephalopathy was noted. the causal relationship with impella is unknown, however the cerebral edema is likely secondary to hypoxic-ischemic encephalopathy which is likely secondary to the cardiac arrest which the patient's arrest occurred prior to impella insertion. no information was provided if any interventions or treatments were given to the patient for the hypoxic encephalopathy. it was reported the patient died without improvement in circulatory dynamics and the physician stated there was no causal relationship with impella. the cause of death was determined to be low cardiac function. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16908 | 1220648-2024-20940 | 20518708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518708&pc=OZD | 02/03/2023 | 10/23/2024 |  | 10/23/2024 |  |  | 02/03/2023 | 10/23/2024 |  |  | 08/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023271669 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported scant hemolysis noted so decision made to swap for impella 5.5. it was also reported that the patient's platelets dropped significantly. a heparin induced thrombocytopenia panel was sent and systemic heparin was discontinued. the patient was successfully escalated to the impella 5.5. the patient's care was later withdrawn on the 5.5 and the patient expired. medical safety review of the event noted use of the impella cp device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
16909 | 1220648-2024-20948 | 20518777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518777&pc=OZD | 08/23/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 08/23/2022 | 10/23/2024 |  |  | 09/30/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137091 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Sepsis | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device and extracorporeal membrane oxygenation for mechanical circulatory support. while on support, the patient was in poor condition with multiple comorbidities with extracorporeal membrane oxygenation support. it was reported that the patient experienced a bleed from tamponade and bleeding from the access site. this was addressed by a fixed concentration of purge fluid heparin and removal of pericardial hematoma. in addition, the patient's renal failure was addressed with continuous hemodiafiltration (chdf). the patient remained on support and the device was explanted. after the pump was removed, sepsis was noted and the patient later expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16910 | 1220648-2024-20921 | 20518782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518782&pc=OZD | 09/20/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 09/20/2021 | 10/23/2024 |  |  | 06/30/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100679 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemolysis | Sepsis | Mitral Valve Insufficiency/ Regurgitation | [10 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced aortic regurgitation, the impella was repositioned, however this did not resolve the problem. additionally, the patient developed hemolysis, adjustment of the impella pump level resolved this issue. also, the device exhibited a purge leak, and the purge cassette was successfully replaced. it was reported that two days after the device was explanted the patient developed sepsis and the medical staff administered antibacterial drugs. however, the treatment was not successful, and the patient died due to sepsis. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. purge leak : the investigation for the reported purge leak has been completed. the device was not returned; however the replaced purge cassette was returned for evaluation, and the purge leak was confirmed. per the results of the clinical review and investigation: the returned purge cassette was cut open and connected to a console. during the de-air procedure, a purge fluid leak was observed between the piston and the cylinder within the purge cassette. the root cause of the purge leak was determined to be an insufficient piston/cylinder seal. no data logs were provided. heart valve damage (mitral regurgitation), hemolysis, infection/sepsis. the impella device was not returned, and an investigation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smart assist system. hemolysis . section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction : ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ purge leak: section: cleaning : ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ mitral regurgitation: section: impella catheter in papillary muscle: ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: warnings & cautions: warnings: "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." sepsis: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16911 | 1220648-2024-20944 | 20518844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518844&pc=OZD | 06/01/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 06/03/2022 | 10/23/2024 |  |  | 03/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199437 |  | 0048-0014 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the implant, the patient suffered from an acute kidney injury and expired. there were no adverse events or pump malfunctions that occurred during the implantation of the device or during support. there is no correlation or causality noted between the outcome and impella use. | [5 empty] | The investigation for the reported renal failure and death has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16912 | 1220648-2024-20932 | 20518867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518867&pc=OZD | 01/03/2023 | 10/23/2024 |  | 10/23/2024 |  |  | 02/13/2023 | 10/23/2024 |  |  | 10/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275750 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that hemolysis was identified during the first twenty-four hours of support. there were no persistent suction alarms or positioning issues reported. it was reported the patient expired while on impella cp support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the hemolysis was not determined since insufficient clinical details were provided. relation to impella unknown. hemolysis diagnosis and troubleshooting unknown. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
16913 | 1220648-2024-20952 | 20518878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518878&pc=OZD | 07/16/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 07/16/2021 | 10/23/2024 |  |  | 03/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022066087 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a pump stop occurred while the patient was on impella cp optical support. restart attempts were unsuccessful. the pump was immediately removed. the patient was stable and needed no further mechanical circulatory support. additional information provided by the patient¿s nurse indicated that there was fluid under the purge cassette and there was a purge pressure increase prior to an increase in motor current. | [5 empty] | The investigation for the reported pump stop and purge system leak has been completed. the product was returned. the field long term and real time (rt) logs were reviewed and an impella stopped ¿ motor current high alarm was confirmed on (b)(6) 2021 followed by multiple unsuccessful restart attempts. further review of the rt logs revealed that prior to the pump stop, purge pressure had abruptly increased from approximately 500mmhg to approximately 1000mmhg at which time purge flow decreased from approximately 18ml/hr to 3ml/hr. visual inspection of the returned pump revealed some blood/biomaterial within the purge gap. the returned purge cassette was cut open and connected to a console. the purge system was then de-aired and connected to the pump with high purge pressures over 1000mmhg reproduced. examination of the purge lumen revealed blood ingress up to approximately 3cm from the motor housing which appeared to be the result of kink in this location. the root cause of the high purge pressure and subsequent pump stop was determined to be a kinked purge line. furthermore, a purge fluid leak was observed at the cartridge cap in the purge cassette. the root cause of the purge leak was determined to be a cartridge cap leak. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
16914 | 1220648-2024-20954 | 20518922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518922&pc=OZD | 08/25/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 08/26/2021 | 10/23/2024 |  |  | 04/30/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022076595 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | EXTRA-CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during wireless insertion via an 8 mm graft into the aorta, the pump could not be started. an x-ray was performed and revealed that the cannula was located completely in ventricle and kinked 180 degrees, the inflow was above outlet. after repositioning, restart was possible, but a high motor current was annotated, and the pump stopped several times while motor current increased. the device was restarted several times and verified to be in the correct position. after 12 hours, the pump stopped again and could not be restarted. the ventricle was ejecting, the aortic valve opens beside high extra-corporeal membrane oxygenation (ecmo) flows and the physician decided not to replace the device. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the device was returned, and the investigation for the reported pump stop has been completed. investigation: the device was returned for evaluation. per the results of the clinical review and investigation: confirmed to have damage to the cannula and ingested biomaterial around the impeller. the pump was not able to start until the biomaterial was removed. the cause of the pump stop was ingested biomaterial. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
16915 | 1220648-2024-20955 | 20518944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518944&pc=OZD | 05/13/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 05/13/2022 | 10/23/2024 |  |  | 02/29/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023195265 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant reported that a patient on impella cp support had a pump stop. all troubleshooting measures were taken, including aic (automated impella controller) replacement. the impella restarted and the patient continued impella support. it was further reported that the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | Product was not returned back. the clot, as visualized by the rn, most likely contributed to the pump stop. the cause of the issue was biomaterial.
16916 | 1220648-2024-20953 | 20518959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518959&pc=OZD | 08/10/2021 | 10/17/2025 |  | 10/23/2024 |  |  | 08/10/2021 | 10/23/2024 |  |  | 05/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022091818 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was placed on impella cp support for hemodynamic support. it was reported that the impella cp was inserted and running smoothly with flows of 3.9 and excellent waveforms. the pump went out of position during removal of the 14f sheath and insertion of the repo sheath. when the physician attempted repositioning, the cannula continuously kinked in one spot causing continuous suction (even in p2) and low flows. the physician decided to replace the device with another impella cp, it was reported that device was successfully implanted. additional information noted that the patient died, no further information was provided. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported production damage (cannula kink) issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the most likely root cause of the cannula kink was wireless re-access. the device is not indicated for wire re-access. this report is being filed as part of a retrospective review of historical records. | B2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2024-20953. g1 reporting contact email revised on manufacturer device report 1220648-2024-20953 in accordance with updated procedures h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2024-20953.
16917 | 1220648-2024-20957 | 20518979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20518979&pc=OZD | 03/22/2022 | 10/08/2025 |  | 10/23/2024 |  |  | 10/16/2022 | 10/23/2024 |  |  | 12/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023173120 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support pulmonary artery diastolic pressure was around 12 to 15, left ventricle diastolic pressures were very low at all times. there were also suction alarms and low purge flow. the physician suspected that the sensor was faulty and was too sensitive as positioning was good on the echocardiogram. in addition, the patient¿s right leg was mottled and there was no doppler for pedal pulses in right leg. the patient was also bleeding from various sites and was given blood products. heparin remained in the purge and systemically. the patient decompensated, return of spontaneous circulation was unable to be achieved and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | Us complainant reported that the patient (pt) was placed on impella cp for hemodynamic support. it was reported that pulmonary artery diastolic pressure was around 12 to 15, left ventricle diastolic pressures were very low at all times. there were also suction alarms and low purge flow. consistent wave form trends appearing to be diastolic function. the physician suspected that the sensor was faulty and was too sensitive as positioning was good on the echo. it was reported that the pt¿s right leg was mottled and pedal pulses in right leg unable to be dopplered. the pt was bleeding from various sites and was given blood products. heparin remained in the purge and systemically. it was reported that the patient expired. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. the returned product was visually inspected and a damaged kink protector was observed. the pump was connected to a console and no placement signal was observed and a controller failure alarm triggered. the light optics were used and damage to the optical line was observed in the red handle next to the kink protector.
16918 | 1220648-2024-20949 | 20519024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20519024&pc=OZD | 12/27/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 12/27/2022 | 10/23/2024 |  |  | 07/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252259 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was bleeding and oozing from all access sites that lasted about two days. the patient received one unit of packed red blood cells (prbcs). an additional suture was applied to the site and the oozing stopped. additional information was provided that reported the family withdrew care and the patient expired. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, there is no relation between the bleeding and the device. clinical description confirmed no further needs or concerns specific to impella. the patient had many medical conditions and process on medications. root cause most likely patient condition. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16919 | 1220648-2024-20936 | 20519034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20519034&pc=OZD | 10/17/2022 | 10/13/2025 |  | 10/23/2024 |  |  | 10/17/2022 | 10/23/2024 |  |  | 04/30/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079623 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a saturated flow issue. coinciding with the saturated flow, it was also noted that there was an intermittent 'impella in aorta' alarm. the physician repositioned under echo and the impella was pulled back 3cm, and the flows improved. also, the physician was made aware of no pulses in left foot/left leg cold. options were discussed and the physician decided they could adjust the sheath or bypass if they wanted. it was reported that the patient expired on the same day. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a saturated flow issue involving an impella device. coinciding with the saturated flow, it was also noted that there was an intermittent 'impella in aorta' alarm. physician came back to reposition under echo. impella pulled back 3cm. flows improved. it was also reported that the physician was made aware of no pulses in left foot/left leg cold. options were discussed and physician decided osu could adjust the sheath or bypass if they wanted. it was reported that the patient expired same day. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The product was returned and biomaterial was found wrapped around the impeller. the data logs show an increase in motor current and flow without a increase in p-level. there was reportedly only dextrose in the purge. the root cause of the saturated flow was determined to be biomaterial ingestion. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16920 | 1220648-2024-20958 | 20519035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20519035&pc=OZD | 03/27/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 03/26/2022 | 10/23/2024 |  |  | 08/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022133582 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Deformation Due to Compressive Stress | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient's impella leg was mottled due to limb ischemia. the family decided to withdraw care and the survival outcome at explant noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The returned product was visually inspected and a kink and fraying of the guidewire was observed. a scratch was observed on the pump outflow cage. the cause of the delivery issue was most likely a abdominal aortic aneurysm. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿. this report is being filed as part of a retrospective review of historical records.
16921 | 1220648-2024-20943 | 20519147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20519147&pc=OZD | 03/19/2023 | 02/28/2025 |  | 10/23/2024 |  |  | 03/19/2023 | 10/23/2024 |  |  | 08/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268508 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Defective Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient became agitated and moved resulting in the repositioning sheath breaking away from the suture hub. the pump was subsequently found to be in the aorta and repositioning was performed to re-advance the pump across the valve. the physician was unable to lock the impella in place due to the break, however, it was decided that patient support would continue with the current device. additionally, access site bleeding was noted. site was resutured, angle matched and tightened perclose which resolved the bleeding. survival outcome at explant noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected. repositioning sheath, suture hub and locking area inspected often and locked in different locations in the catheter no issue observed and tightened. catheter kink observed. root cause of the issue was patient movement base on the clinical data, and the catheter kink occurred while attempted to reposition due to improper technique use. this report is being filed as part of a retrospective review of historical records.
16922 | 1220648-2024-20945 | 20519152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20519152&pc=OZD | 07/22/2024 | 10/23/2024 |  | 10/23/2024 |  |  | 03/30/2022 | 10/23/2024 |  |  | 10/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022138514 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient was successfully weaned from the impella and the device explanted. after the explant, the patient was observed to have a cerebral event, a bleed in the cerebral arteries. in addition, it was noted of a bleed including hematoma, that required blood transfusions. additional information was provided that noted the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16923 | 1220648-2024-20956 | 20519156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20519156&pc=OZD | 03/10/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 03/10/2022 | 10/23/2024 |  |  | 07/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022105775 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Ischemia | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had limb ischemia. suspected cause of limb ischemia was a competitor's sheath. additionally, pump stopped during support. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system (pump stop). as noted in the impella ifu: impella cp with smartassist system. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ as noted in the impella ifu: impella cp with smartassist system (ischemia). section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
16924 | 1220648-2024-20946 | 20519169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20519169&pc=OZD | 08/03/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 08/03/2022 | 10/23/2024 |  |  | 11/30/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166521 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was suspicion of hemolysis as evident by urine color, hemolyzed labs. one unit of packed red blood cells and fluids were given to the patient. in addition, a slight adjustment of positioning was performed as intervention. unfortunately, the hemolysis remained uncontrolled. the patient was ultimately made do not resuscitate and expired due to multisystem organ failure. | [5 empty] | The investigation for the reported hemolysis has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
16925 | 1220648-2024-20959 | 20519180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20519180&pc=OZD | 12/08/2021 | 10/23/2024 |  | 10/23/2024 |  |  | 04/12/2022 | 10/23/2024 |  |  | 07/31/2023 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108707 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | [24 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced an ischemic stroke. additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported stroke/cva has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
16926 | 1220648-2024-20961 | 20519227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20519227&pc=OZD | 03/14/2023 | 03/28/2025 |  | 10/23/2024 |  |  | 10/16/2024 | 10/23/2024 |  |  | 10/31/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306112 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ischemia Stroke | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the complainant reported that the patient ¿is mottled from head to toe.¿ the medical staff was unable to obtain pulses in the upper and lower extremities. computed tomography (ct) scan revealed significant brain injury. care was withdrawn and the device explanted. additional information was provided that noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Per the clinical information provided, the root cause of the limb ischemia issue was patient condition related and not related to impella due to ischemia in both limbs. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation for stroke/cva is not determined.
16927 | 1220648-2024-20951 | 20519323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20519323&pc=OZD | 02/14/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 02/08/2022 | 10/23/2024 |  |  | 09/30/2022 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021006111 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis. no intervention was specified. the patient survived the explant procedure but subsequently expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported hemolysis has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the hemolysis could not be determined. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
16928 | 1220648-2024-20960 | 20519410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20519410&pc=OZD | 05/18/2022 | 10/23/2024 |  | 10/23/2024 |  |  | 05/20/2022 | 10/23/2024 |  |  | 02/29/2024 |  | 10/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195023 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a cp pump stop occurred after 97 hours. the patient had been transferred to the tertiary center and made care comfort only. the patient was not neuro responsive. in addition, the patient had a difficult airway and had blood pooling in posterior oropharynx. ent and surgery were consulted, and mouth was packed with moist gauze around tube. the patient also had bloody secretions/ coffee grounds from et tube and ooze from the impella site may stop heparin temporarily until bleeding under control. the survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system and impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
16929 | 1220648-2024-20746 | 20501560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20501560&pc=OZD | 07/19/2021 | 10/21/2024 |  | 10/22/2024 |  |  | 07/19/2021 | 10/22/2024 |  |  | 02/02/2023 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022054887 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced profuse oozing after the placement of the repositioning sheath from proximal end of blue suture hub and around sheath. echocardiology was performed and it was noted that the left ventricular ejection fraction was normal, but with severe mitral regurgitation. the patient was consulted for mitral valve replacement. the medical team decided to explant the impella one hour after implant and was to be replaced with an intra-aortic balloon pump. the patient went to surgery and expired. the physician's comments noted that the patient did not expire due to impella, but due to surgery complications. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the device was returned for review and no device damage was found. analysis of the video sent in shows the blue suture hub to be underneath the skin which is known to cause bleeding. the cause of the access site bleeding was improper technique of placing the blue suture up into the access site. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16930 | 1220648-2024-20748 | 20501642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20501642&pc=OZD | 07/29/2022 | 10/21/2024 |  | 10/22/2024 |  |  | 08/01/2022 | 10/22/2024 |  |  | 12/31/2022 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022039794 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Tachycardia | [11 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a percutaneous coronary intervention [pci] was performed after inserting impella cp. after the procedure, the patient experienced frequent episodes of ventricular tachycardia (vt) and therefore extracorporeal membrane oxygenation (emco) was inserted. it was later reported that the patient was taken off ecmo, but cardiopulmonary arrest (cpa) occurred and patient expired. the doctor in charge said that the patient had a cardiac rupture and that the impella was not involved, however the timing of the heart rupture seems to be after the impella insertion. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16931 | 1220648-2024-20754 | 20501649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20501649&pc=OZD | 09/13/2022 | 10/21/2024 |  | 10/22/2024 |  |  | 09/14/2022 | 10/22/2024 |  |  | 03/31/2024 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202203 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient expired due to uncontrolled gastrointestinal bleeding and multiple organ failure. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16932 | 1220648-2024-20756 | 20501806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20501806&pc=OZD | 03/14/2023 | 10/21/2024 |  | 10/22/2024 |  |  | 03/14/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1157448 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced purge disc/flag issues. the aic was unable to recognize the purge disk. no report of patient involvement. | [5 empty] | The investigation has been completed. the console (ic3079) was sent back for further investigation. the console was tested, and the reported problem was reproduced. the aic was unable to recognize the purge disk, and the purge flag was confirmed to be broken. the root cause of the issue was broken/ missing purge flag. this report is being filed as part of a retrospective review of historical records.
16933 | 1220648-2024-20745 | 20501835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20501835&pc=OZD | 06/10/2022 | 11/07/2025 |  | 10/22/2024 |  |  | 06/10/2022 | 10/22/2024 |  |  | 02/29/2024 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198982 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a loss of pulsatility following intubation and swam placement which was quickly resolved with the addition of pressors. the pump was then repositioned in the operating room on 10-jun-2022 and returned to the cardiovascular icu in stable condition. on (b)(6) 2022 it was reported that the patient was hemodynamically unstable as the patient's pressors and inotrope requirement increased. it was noted the patient was bleeding from the chest tube and as a result, the patient was massively transfused with blood products, administered methylene blue and purge heparin was reduced. the patient's renal function deteriorated and the patient was evaluated for continuous renal replacement therapy. the patient was made do not resuscitate by family and continued to deteriorate throughout the day and expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The cause of the renal failure was not determined due to insufficient clinical details. the cause of the vascular damage peripheral vessel issue was not determined due to insufficient clinical details. b1 product problem was erroneously entered on the initial manufacturer device report number 1220648-2024-20745. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-20745 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-20745 in accordance with updated procedures. h6 medical device problem code was erroneously entered on the initial manufacturer device report number 1220648-2024-20745.
16934 | 1220648-2024-20751 | 20501966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20501966&pc=OZD | 07/14/2021 | 10/31/2025 |  | 10/22/2024 |  |  | 07/14/2021 | 10/22/2024 |  |  | 03/31/2023 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022056481 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  | Black Or African American | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with a peri-valvular leak (pvl) was implanted with an impella cp device for mechanical circulatory support. it was reported the pvl was successfully contained. later in the day, it was noted that the patient was actively bleeding with a hemoglobin of 6.8 and a hematocrit of 20.1. the patient received blood transfusions and blood samples were obtained for hemolysis testing. the following day, "gross" hemolysis was noted despite good positioning. the pump stopped and was successfully replaced. the patient was noted to be stable, but in grave condition. the patient was later withdrawn from care on the replacement pump. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle: ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis. ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | E4 should have been left blank on the initial manufacturer device report number 1220648-2024-20751as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-20751 in accordance with updated procedures. the following was revised from the initial submission 1220648-2024-20751 in accordance with updated procedures. b2 death added. h1 type of report updated to death. h6 health effect: impact code 1802 added.
16935 | 1220648-2024-20757 | 20502028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502028&pc=OZD | 03/20/2023 | 10/21/2024 |  | 10/22/2024 |  |  | 03/20/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420045 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller (aic) experienced a purge disc/flag issue that was resolved by changing the console. no patient harm was reported. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. the console was tested after arriving in danvers. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the issue was broken/ missing purge flag. this report is being filed as part of a retrospective review of historical records.
16936 | 1220648-2024-20758 | 20502096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502096&pc=OZD | 03/22/2023 | 10/21/2024 |  | 10/22/2024 |  |  | 03/22/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084302 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced purge disc/flag issues, which was resolved by using a new console. there was no patient involvement. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. the blue receptacle showed a hairline fracture in the base that would have caused these errors to pop up. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16937 | 1220648-2024-20760 | 20502135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502135&pc=OZD | 03/29/2023 | 10/21/2024 |  | 10/22/2024 |  |  | 03/29/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401635 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced purge disc/flag issues. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. the root cause of the issue was a broken/missing purge flag. this report is being filed as part of a retrospective review of historical records.
16938 | 1220648-2024-20759 | 20502189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502189&pc=OZD | 03/24/2023 | 10/21/2024 |  | 10/22/2024 |  |  | 03/24/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1144656 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced purge disc/flag issues. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. the console was tested after arriving in danvers. the issue with properly detecting the purge pressure disc was reproduced. purge flag was confirmed to be broken, and purge retainer screw stand was found broken. the root cause of the issue was broken purge flag and purge retainer. this report is being filed as part of a retrospective review of historical records.
16939 | 1220648-2024-20761 | 20502197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502197&pc=OZD | 04/08/2023 | 10/21/2024 |  | 10/22/2024 |  |  | 04/08/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075487 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc/flag issue that was resolved by changing the console. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. visual inspection of the purge assembly area confirmed a ripped/missing blue retainer. the failure mode was due to broken/ missing purge retainer. this report is being filed as part of a retrospective review of historical records.
16940 | 1220648-2024-20762 | 20502220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502220&pc=OZD | 04/09/2023 | 10/21/2024 |  | 10/22/2024 |  |  | 04/09/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1168282 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc/flag issue that was resolved by changing the console. no patient harm was reported. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. visual inspection of the purge assembly area confirmed a broken blue retainer. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16941 | 1220648-2024-20744 | 20502238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502238&pc=OZD | 05/26/2023 | 10/21/2024 |  | 10/22/2024 |  |  | 05/26/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164193 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. no clinical details regarding resolution were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console was returned for investigation. the logs confirmed the reported failure. there was a battery failure alarm due to low pack voltage. the failure mode was reproduced on an engineering bench. the battery 2 pack voltage was measured to be 0 volt rather than the expected 16 volt, and the battery liquid crystal display indicator was blank (fully discharged). the batttest showed battery 2 cell 4 fully depleted. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16942 | 1220648-2024-20763 | 20502242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502242&pc=OZD | 03/09/2023 | 10/21/2024 |  | 10/22/2024 |  |  | 03/09/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1184609 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced purge disc/flag issues. there was no patient involvement, this issue occurred during a training session. when loading the purge cassette disc an error message kept showing up that there was a disc problem. if slight forward pressure was applied to the disc, the priming sequence would start however as soon as the pressure was removed it would stop priming and present an error message again. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. visual inspection of the purge assembly area confirmed a broken blue retainer, one of the screw posts was broken. the root cause of the issue was broken purge disk retainer. this report is being filed as part of a retrospective review of historical records.
16943 | 1220648-2024-20755 | 20502368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502368&pc=OZD | 01/22/2023 | 10/21/2024 |  | 10/22/2024 |  |  | 01/25/2023 | 10/22/2024 |  |  | 05/31/2024 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225194 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a cerebral hemorrhage, which caused the patient to expire. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16944 | 1220648-2024-20749 | 20502419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502419&pc=OZD | 02/13/2022 | 10/21/2024 |  | 10/22/2024 |  |  | 02/13/2022 | 10/22/2024 |  |  | 11/30/2023 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165703 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, pump stopped. the pump was removed and ultimately replaced. patient survived the explant but subsequently expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop has been completed. the product was returned, and the pump stop was confirmed. pump ran for at least 5 minutes after implant and resulted in a stoppage that could not be restarted. the cause of the pump stop was bearing wear on day 5 of support. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
16945 | 1220648-2024-20768 | 20502460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502460&pc=OZD | 02/01/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 01/31/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1232449 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. it was noted that there was no patient involvement. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. the device logs confirmed the reported failure. the failure mode was reproduced on an engineering bench. the battery 1 pack voltage was measured to be 3.5 v rather than the expected 16 v, and the battery lcd indicator was blank (fully discharged). the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16946 | 1220648-2024-20764 | 20502625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502625&pc=OZD | 03/08/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 03/08/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1418929 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller had a broken purge clip. no patient involvement was reported. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. the console was tested and the issue with properly detecting the purge pressure disc was reproduced, and purge retainer was confirmed to be broken. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16947 | 1220648-2024-20765 | 20502627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502627&pc=OZD | 03/08/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 03/08/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1207018 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a controller failure red alarm. there was no patient involvement. | [5 empty] | The investigation for the reported battery failure has been completed. the console (ic3941) was sent back for further investigation. the error was able to be reproduced in the lab repeatedly. as the console was dropped in the field, as a precautionary measure, the power battery manager will be replaced. the root cause of the aic issue was a defective power battery manager board's mismatched fw revisions. this report is being filed as part of a retrospective review of historical records.
16948 | 1220648-2024-20767 | 20502630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502630&pc=OZD | 02/03/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 02/03/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1363891 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. they noticed battery was at 50 percent and then they tried to plug in aic and the percentage would not display. they also tried to engage and disengage the battery while the alarm was active. a loaner was sent to the customer. there is no report of patient involvement. | [5 empty] | The investigation has been completed. the console (ic6395) was sent back for further investigation. aic logs confirmed the reported failure. the failure mode was reproduced on an engineering bench. the battery 1 pack voltage was measured to be 0 v rather than the expected 16 v, and the battery lcd indicator was blank (fully discharged). the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16949 | 1220648-2024-20769 | 20502634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502634&pc=OZD | 01/31/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 01/31/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1110482 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. no patient harm was reported. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) due to low pack voltage. the failure mode was reproduced on an engineering bench. the battery 1 pack voltage was measured to be 3.5 v rather than the expected 16 v, and the battery lcd indicator was blank. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16950 | 1220648-2024-20752 | 20502699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502699&pc=OZD | 11/21/2021 | 10/21/2024 |  | 10/22/2024 |  |  | 11/23/2021 | 10/22/2024 |  |  | 01/31/2022 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1473816 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device along with ecmo for mechanical circulatory support. this was the 2nd pump as the previous was explanted after reaching support of 8 days. while on support with this 2nd pump, the patient suffered from a cerebral hemorrhage and expired while on support. no product malfunction was noted. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16951 | 1220648-2024-20766 | 20502761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502761&pc=OZD | 02/06/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 02/06/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186352 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a battery failure red alarm. there was no patient involvement. | [5 empty] | The investigation for the reported battery failure has been completed. the console was sent back for further investigation. the failure mode was reproduced on an engineering bench. the battery 1 pack voltage was measured to be 3.5 v rather than the expected 16 v and battery test showed battery1 cellv1 is ~250mv less than the other cells. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16952 | 1220648-2024-20747 | 20502946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20502946&pc=OZD | 08/19/2021 | 10/21/2024 |  | 10/22/2024 |  |  | 03/23/2022 | 10/22/2024 |  |  | 02/28/2023 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022064792 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bleeding including hematoma the day after implant and 19 days post implant on (b)(6) both requiring blood transfusions of unknown quantity. additionally, the day after implant the patient experienced thrombocytopenia and it was noted the event was related to the impella cp. severity of the event was noted as serious, however there was no information provided for treatment or interventions performed for the thrombocytopenia. on (b)(6) 2021, the patient experienced a stroke. there was no information provided for treatments or interventions performed for the stroke. the outcome from these events is noted as death. the date of the patient's death is unknown. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16953 | 1220648-2024-20770 | 20504721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20504721&pc=PYX | 09/24/2024 | 11/26/2024 |  | 10/22/2024 |  |  | 09/24/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was having an attempted impella rp implant and once the pump was in the right ventricle, it would not flip towards the pulmonary atresia. the pump was removed. the patient¿s outcome is unknown. | It was reported the patient did not survive the procedure and expired on an unknown date. according to medical safety review, there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into delivery issues has been completed. the root cause of the delivery issue cannot be determined since the product was not returned and insufficient clinical details were provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20770. b2 updated to reflect adverse and death, date of death was unknown. b5 updated with patient outcome and patient outcome. g1 reporting contact email. g4 pma/510(k)number. h1 type of reportable event. h6 health effect - impact code added.
16954 | 1220648-2024-20771 | 20505055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20505055&pc=OZD | 09/25/2024 | 11/20/2024 |  | 10/22/2024 |  |  | 09/25/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1266275 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller was showing purge disk not detected alarm. no reported patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag.
16955 | 1220648-2024-20772 | 20505340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20505340&pc=OZD | 09/25/2024 | 11/26/2024 |  | 10/22/2024 |  |  | 09/25/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1119116 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during prepping the automated impella controller, it fell and the front housing needs replacement. no patient involvement has been reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into automated impella controller (aic) - damage before therapy has been completed. the console logs were analyzed with no observable anomaly. the console was visually inspected and confirmed to have cracks in the front and rear housing of the console. the root cause for aic -damage was from user interaction.
16956 | 1220648-2024-20773 | 20505480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20505480&pc=OZD | 09/24/2024 | 11/24/2024 |  | 10/22/2024 |  |  | 09/27/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534242 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 62-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the patient on impella 5.5 developed a large hematoma at the right axillary site. systemic anticoagulation was on hold due to ecmo cannula site bleeding. it was further reported that prior to cardiac transplant, 5.5 cannula was cut distal to motor per abiomed recommendation; however, surgeon used mayo scissors to cut catheter. upon removal of catheter through the axillary graft, the catheter shaft was stripped, and wire was coiled. patient is currently in icu. patient has not shown any signs of complication from this event. | [5 empty] | The investigation for the access site bleed and the removal issues has been completed. clinical details state that the customer experienced hemolysis while on impella support. it also states the customer voiced that the cannula became detached in the graft. the product was returned in two pieces; it was cut at the catheter. the red handle was not returned to confirm the serial number. there was insufficient product returned to evaluate for access site bleeding. the clinical details describe how the patient had bleeding from the ecmo site so anticoagulation was held and then the patient developed a hematoma at the impella access site. it is unclear what may have attributed to the hematoma. the root cause of the access site bleeding cannot be determined. the catheter was stripped and the cannula was detached from the catheter near the motor housing. based on the clinical details, the root cause of the product damage (detached inflow/outflow - peripheral vessels) is most likely due to use as excessive force was likely used when removing the device in the graft. data logs were reviewed finding suction and positioning alarms seen. there was no trend in the placement signal or motor current spikes present. there was no biomaterial seen or damages with the returned product impacting hemolysis. the clinical stated that the surgeon felt the device to be too close to the septal wall. repositioning was done, and the hemolysis improved. the root cause of the hemolysis issue is most likely due to improper positioning of the device. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16957 | 1220648-2024-20775 | 20505606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20505606&pc=OZD | 10/04/2024 | 12/06/2024 |  | 10/22/2024 |  |  | 10/04/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025527620 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a non-st elevated myocardial infarction patient admitted in for coronary artery bypass graft and had an impella 5.5 pump placed. the team observed when coming off the bypass and closing the chest, there was profuse bleeding. the team observed a left ventricle perforation. the cause of perforation was either the wire or pump. the patient was repaired and the patient survived. while in the intensive care unit on the pump, there were concerns of hemolysis and albumin was infused back to the patient. the patient was successfully weaned and the patient survived the explant. | [5 empty] | The investigation into ventricle perforation and hemolysis has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the pump was positioned too deep, and it could not be repositioned due to the small size of the ventricle. the root cause of the ventricle perforation was improper positioning of the device. the root cause of the hemolysis was not determined since product was not returned and insufficient clinical details were provided. g1 reporting contact email was reported incorrectly on manufacturer device report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
16958 | 1220648-2024-20774 | 20505673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20505673&pc=OZD | 09/27/2024 | 11/16/2024 |  | 10/22/2024 |  |  | 09/27/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102032 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an automated impella controller (aic) and impella 5.5 pump support ongoing for a patient admitted in with cardiogenic shock (cgs) and cardiomyopathy. the support was ongoing when the purge failure occurred. the team troubleshot by pc replacement, but this did not remedy the issue and so the aic itself was replaced. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the reported purge system failure was confirmed in the logs. the purge system failure was unable to be reproduced. inspection of the aic revealed that there were signs of ingress build up all around the purge assembly including the gaskets and the piston. the root cause of this issue was an ingress build up on the piston and around the gasket that prevented the piston from moving. g1 reporting contact fax number incorrect on manufacturer device report 1220648-2024-20774. | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
16959 | 1220648-2024-20776 | 20505735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20505735&pc=OZD | 04/27/2023 | 12/12/2024 |  | 10/22/2024 |  |  | 09/26/2024 | 10/22/2024 |  |  | 09/30/2024 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023260062 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The impact study reported that a patient on impella 5.5 support developed bleeding at the right axillary artery. the patient had hemoglobin drop from 15.1 to 7.2 and received one unit of packed red blood cells. bleeding was possibly related to the impella 5.5 study device. the patient remained stable and survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
16960 | 1220648-2024-20777 | 20505761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20505761&pc=OZD | 09/27/2024 | 11/23/2024 |  | 10/22/2024 |  |  | 09/27/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505846 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump support ongoing for a patient transferred in from the community hospital to the tertiary center. the patient had left ventricle (lv) free wall rupture post coronary artery bypass graft and planned for a re-do surgery. the 5.5 pump was placed but after 23 days there was an observation made of lv clot on the pump cannula post transient ischemic attack. the patient has survived. | [5 empty] | The investigation for the reported stroke and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of stroke and thrombus could not be determined as the device was not returned nor sufficient clinical details. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
16961 | 1220648-2024-20782 | 20505931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20505931&pc=OZD | 09/28/2024 | 11/25/2024 |  | 10/22/2024 |  |  | 09/29/2024 | 10/22/2024 |  |  | 11/30/2025 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435581 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Other | Required Intervention |  |  |  |  | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a patient admitted in with multiple cardiac comorbidities. the team placed cp-extra corporeal membrane oxygenation support for the patient despite small vessels. the team placed perfusion due to ischemia and limb pulse loss. additionally, there was an observation made during the support of a cerebrovascular accident (cva). the stroke was noted to be small and whether acute or subacute was not known. age of cva remains unknown. the patient survived the event. | [5 empty] | The investigation limb ischemia and stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 outcomes attributed to adverse event; required intervention to prevent permanent impairment/damage missing on manufacturer device report 1220648-2024-20782. h6 code 4625, 4642, and 4614 were reported incorrectly on manufacturer device report 1220648-2024-20782. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
16962 | 1220648-2024-20785 | 20505936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20505936&pc=OZD | 09/26/2024 | 12/09/2024 |  | 10/22/2024 |  |  | 09/26/2024 | 10/22/2024 |  |  | 06/30/2025 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024365582 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Abiomed received notification from the impact study team that reported a patient experienced a stroke with probable relationship to the impella 5.5 use. the event was ischemic and occurred after explant. patient info, treatment and outcome are all unknown. | [5 empty] | The investigation for the reported stroke/cva has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke/cva could not be determined. corrections to the initial manufacturer device report 1220648-2024-20785: g1-mdr reporting contact email. | The device is not expected to be returned for evaluation and review. however, the complaint investigation is not complete at this time. a supplemental and final report will be filed following the completion of the complaint investigation. we will continue to monitor for trends through the complaint system to assure patient safety.
16963 | 1220648-2024-20788 | 20505943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20505943&pc=OZD | 08/30/2024 | 10/22/2024 |  | 10/22/2024 |  |  | 09/27/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025518620 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | A notification was forwarded to the complaint team via stemi-dtu study regarding 34-year-old male patient on impella cp support that had bruising/hematoma at the access site. the clinical representative confirmed that no blood products were given, heparin was not withheld and no washout was performed. no serious intervention was performed. the patient also experienced ventricular tachycardia. the study found that there was a possible relationship between the impella, the impella procedure and the patient's ventricular tachycardia. the impella was successfully explanted after a day of support. the patient will be given a life vest upon discharge. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
16964 | 1220648-2024-20778 | 20505973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20505973&pc=OZD | 10/02/2024 | 11/27/2024 |  | 10/22/2024 |  |  | 10/02/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531361 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, suction alarms and flows dropped to 0.1 liters per minute. the physician repositioned the impella cp without resolution and stated the most likely cause was clot ingestion. an attempt was made to increase/decrease p levels without resolution of low flow state. therefore, the impella cp was explanted due to low pump flow. no harm to the patient was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. the impella was returned for evaluation and was visually inspected. biomaterial ingestion observed around impeller. slight cannula kink observed near out let. no broken blade/ no other abnormality found. the cause of the low pump flow issue was biomaterial ingestion. the cannula kink may occurred during removal.
16965 | 1220648-2024-20779 | 20505979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20505979&pc=OZD | 09/28/2024 | 11/27/2024 |  | 10/22/2024 |  |  | 09/28/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025503569 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support when the patient had a gastrointestinal bleed (gib). the clinical representative confirmed that the gib was unrelated to the impella, however the staff wanted to remove the pump early due to not being able to provide the patient proper systemic anticoagulation. one unit of packed red blood cells was provided to the patient. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿.
16966 | 1220648-2024-20780 | 20506012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506012&pc=PYX | 09/23/2024 | 11/19/2024 |  | 10/22/2024 |  |  | 09/23/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025551050 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella rp implant procedure for female patient admitted in with right heart failure and pulmonary embolism (pe), the device had been opened and primed independently. the automated impella controller (aic) did not have case start up screen and the pump failed to start (p0) at the time of delivery and possibly skipping of steps with the aic. the aic was attempted to be restarted but was unsuccessful. a new aic was used but was also unsuccessful. therefore, decision was made to use a new impella to complete the procedure. there was no harm reported to the patient. | [5 empty] | The investigation for the pump did not start issue has been completed. the pump was not returned for investigation. according to the clinical information, the doctor noted that the pump was primed prior to arrival. the purge and pump appeared to be primed with a pa pressure reading of 120; however, there was no impella startup screen. several unsuccessful attempts were made to restart the pump and switch the aic. the data log shows that the case start was initiated manually on console ic11409. the "purge cassette and pump connected" prompt was successfully executed and recorded in the imc log. the pump reached priming step 4, followed by the "purge_system_priming_complete" steps. a manual keypress was recorded, and the "case start cancel" prompt appeared. two attempts were made on this console, with the pump skipping to the "case start: exit wizard" after pump being connected to the aic. additionally, the same activity was recorded after switching to console ic3624. according to the provided imc logs, the pump was never attempted to start. the cause of the pump did not start was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
16967 | 1220648-2024-20794 | 20506108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506108&pc=OZD | 01/03/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 01/04/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1073779 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. actions taken are unknown. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. aic logs confirmed the reported failure. the failure mode was reproduced on an engineering bench. the battery 1 pack voltage was measured to be 3.5v rather than the expected 16v. the root cause of the aic issue was a single cell failure of the battery. this report is being filed as part of a retrospective review of historical records.
16968 | 1220648-2024-20783 | 20506117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506117&pc=OZD | 09/24/2024 | 11/16/2024 |  | 10/22/2024 |  |  | 09/24/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532419 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient had multiple attempts to implant the impella cp and the implant was aborted due to the pump could not cross the aortic valve. the medical doctor attributed it to tortuosity of descending aorta and inability to push impella across the valve. there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the impella was unable to cross the aortic valve due to tortuosity of the descending aorta despite multiple attempts. e4 should have been left blank on manufacturer device report 1220648-2024-20783. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
16969 | 1220648-2024-20786 | 20506130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506130&pc=OZD | 09/22/2024 | 12/12/2024 |  | 10/22/2024 |  |  | 09/26/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025517906 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 23 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intraoperative Cardiac Valve Injury | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in japan had impella cp support ongoing for a female patient admitted in cardiogenic shock (cgs). the pump was placed but the flow rate was not achieved and the chordae tendiane ruptured. the patient survived and pump was successfully explanted after two days of support. | [5 empty] | The investigation for the heart valve damage has been completed. no product or logs were returned for evaluation. clinical details noted that patient experienced a mitral valve chordae rupture while on support. it was that pump position was confirmed via echo; physician did not have difficulty inserting pump. physician noted that cardiac condition could have been a factor, however it was not certain. the root cause of the heart valve damage cannot be definitively determined as limited clinical details were provided. d4-model number updated e4 should have been left blank on manufacturer device report. g1 mdr reporting contact email updated. h5 changed to yes. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16970 | 1220648-2024-20781 | 20506152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506152&pc=OZD | 09/24/2024 | 11/16/2024 |  | 10/22/2024 |  |  | 09/24/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511310 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported an impella 5.5 was attempted to be implanted on a male patient, however upon implanting the device, there was a catheter kink noted and the implant procedure was aborted. a new impella 5.5 was used. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely use issue since the medical staff didn't attempt under fluoroscopy or any other imaging during first attempt.
16971 | 1220648-2024-20784 | 20506175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506175&pc=OZD | 09/25/2024 | 12/11/2024 |  | 10/22/2024 |  |  | 09/25/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533877 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during the impella cp placement for male patient, the introducer sheath could be placed, but the impella could not be advanced any further. the impella insertion was therefore aborted and the patient was not treated with the high risk percutaneous coronary intervention (hrpci). there was no harm to the patient and the patient is noted to be well. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature as per the clinical description provided.
16972 | 1220648-2024-20800 | 20506204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506204&pc=OZD | 09/30/2024 | 11/18/2024 |  | 10/22/2024 |  |  | 09/30/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025530540 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 access site had erythematous, macerated with some drainage. cultured was done multiple times but negative to date. iv antibiotics and cleaning site daily with betadine. the patient remains on support therefore explant date and patient outcome are unknown. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer. the necessary clinical information was not provided; therefore, the root cause of the infection/sepsis was not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & section: table 3.2 impella catheter components . ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
16973 | 1220648-2024-20789 | 20506212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506212&pc=OZD | 01/18/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 01/18/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1114582 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. there was no reported patient harm. | [5 empty] | The investigation has been completed. the console (ic1802) was sent back for further investigation. aic logs confirmed the reported failure. the root cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16974 | 1220648-2024-20787 | 20506266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506266&pc=OZD | 09/26/2024 | 11/20/2024 |  | 10/22/2024 |  |  | 09/26/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025520146 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 implant procedure for 57-year-old male patient, the pump was being prepped when it was noted that the luer lock connector near the spike became disconnected. the purge fluid flooded the automated impella controller (aic) and 'impella device defective' alarm was displayed. the pump was unable to start. the impella plug was disconnected and attempted to be dried, but the issue did not resolve. a new impella 5.5 was used. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into pump did not start has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-20787.
16975 | 1220648-2024-20799 | 20506319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506319&pc=OZD | 09/29/2024 | 11/27/2024 |  | 10/22/2024 |  |  | 09/29/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479684 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump delivered to support a patient admitted in cardiogenic shock. the patient first had venous-arterial extra corporeal membrane oxygenation placed and was cardioverted for the ventricular tachycardia. the 5.5 pump was in for two months while awaiting a transplant. the two month support was explanted due to the loss of placement signal and there was a kink observed. the 5.5 was explanted and replaced, the patient survives and still awaits the transplant. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into optical signal issue and product damage (cannula kink) has been completed. the pump was returned and evaluated. there was a significant kink found at the edge of the kink protector on the catheter side. log review showed worsening snr and contrast about ten days into the case before completely failing 60 days into the case when they lost placement signal and had placement signal not reliable alarms. the "product damage (cannula kink)" failure mode was removed since a catheter kink was reported, and the catheter kink was incorporated into the root cause of the optical signal issue. the root cause of the optical signal issue was a damaged optical fiber line caused by a catheter kink. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20799. d9 was marked no, should have been yes. e4 should have been left blank. g1 reporting contact fax number missing.
16976 | 1220648-2024-20790 | 20506321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506321&pc=OZD | 01/04/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 01/04/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146945 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. actions taken are unknown. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. aic logs confirmed the reported failure. the failure mode was reproduced on an engineering bench. the battery 2 pack voltage was measured to be 3.5 v rather than the expected 16 v, and the battery lcd indicator was blank (fully discharged). the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16977 | 1220648-2024-20796 | 20506343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506343&pc=OZD | 08/04/2024 | 11/19/2024 |  | 10/22/2024 |  |  | 09/27/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524768 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed in (b)(6) 2024 for a patient with known systemic comorbidities and in cardiogenic shock. the pump was placed via the long sheath when the anatomy was not appropriate for the shorter sheath. the support was for 15 hours when care was withdrawn and patient expired. the patient was enrolled in abiomed¿s long-term outcome and quality indicator (loqi) impella registry, and the database in 1(b)(6) 2024 noted there had been limb ischemia that was not treated due to underlying condition and plan not to escalate care and intervene. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
16978 | 1220648-2024-20791 | 20506351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506351&pc=OZD | 09/27/2024 | 11/20/2024 |  | 10/22/2024 |  |  | 09/27/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547318 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had slight oozing at the access site as the patient was moving a lot in bed. light manual pressure was held above the site. the following day, impella dressing with bloody drainage noted to gauze, angle of insertion maintained, and heparin held until partial thromboplastin time dropped. it was further reported that the patient experienced hemolysis after impella support. p-level being lowered to 6 did resolve the hemolysis. the urine cleared up in just a couple of hours. the patient survived the event. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the issue was patient movement. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20791. e4 should have been left blank. g1 reporting contact fax number missing. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16979 | 1220648-2024-20792 | 20506354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506354&pc=OZD | 09/27/2024 | 11/20/2024 |  | 10/22/2024 |  |  | 09/27/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546898 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had placed a cp for the support of a patient admitted in for a high risk percutaneous coronary intervention. the percutaneous coronary intervention was due to surgical turndown. the pump was placed but, there was an access site ooze that prompted premature explant. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16980 | 1220648-2024-20793 | 20506357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506357&pc=OZD | 09/27/2024 | 11/25/2024 |  | 10/22/2024 |  |  | 09/27/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523977 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support the patient admitted in acute myocardial infarction/cardiogenic shock. the pump support was for 14 hours due to access site bleeding. the bleed/hematoma prompted four units of blood, manual pressure, a stitch, and explant of the pump. the patient survived the event. | [5 empty] | The investigation into access site bleeding and vascular damage peripheral vessel has been completed. no product was returned for evaluation. clinical details noted that hematoma was present at impella access site and was managed with manual pressure and blood products given. patient was vomiting blood before impella placement and bleeding at access continued after patient transferred hospitals. pump was explanted and patient underwent vascular repair. the root cause of the vascular damage is most likely patient condition related. the root cause of "access site bleeding major" was removed and investigated under "vascular damage peripheral vessel". | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16981 | 1220648-2024-20795 | 20506373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506373&pc=OZD | 09/27/2024 | 12/06/2024 |  | 10/22/2024 |  |  | 09/27/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1355280 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a console seen to alarm for a system self-check error. the automated impella controller (aic) was removed from patient use and a backup console will be used. there was no patient use with the console at the time, and as such no harm or injury possible. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into automated impella controller (aic) - system self-check error has been completed. during the data analysis the console displayed the system self check failed screen during automatic reboot. this confirms the reported failure mode. this console was tested and reproduced the reported failure mode. upon initial boot up, the console rebooted automatically and displayed ¿system self check failed screen.¿ visually confirmed the pbm corrosion the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20795.
16982 | 1220648-2024-20801 | 20506435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506435&pc=OZD | 10/02/2024 | 12/04/2024 |  | 10/22/2024 |  |  | 10/02/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445892 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, an extravasation was noted from left femoral artery. the doctor opted to keep the peel away sheath in. repositioning sheath was advanced into the peel away sheath two centimeters. all sheaths were secured in place. femstop was applied over site and another angiogram was performed showing that the bleeding had resolved. additionally, a hematoma developed requiring blood transfusion. the patient survived the event. | [5 empty] | The investigation for the access site bleed and the vascular damage has been completed. no product was returned for evaluation. clinical details noted that a extravasation was noted on angiogram after insertion of the second impella cp. peel away sheath was kept in place with repositioning sheath inserted 2 cm into peel away. bleeding was managed with femstop. additionally, a hematoma was noted later in support and required pump explant and vascular repair to be performed. the root cause of the vascular damage cannot be definitively determined as no product was returned. clinical details state that the hematoma at the access site required vascular repair. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16983 | 1220648-2024-20798 | 20506463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506463&pc=OZD | 12/28/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 12/28/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1234567 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. actions taken are unknown. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic4308) was sent back for further investigation. aic logs confirmed the reported failure. the failure mode was reproduced on an engineering bench. the battery 1 pack voltage was measured to be 3.5 v rather than the expected 16 v and batt test showed battery1 cellv3 is ~375mv less than the other cells. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16984 | 1220648-2024-20797 | 20506498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506498&pc=OZD | 09/30/2024 | 11/18/2024 |  | 10/22/2024 |  |  | 09/30/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535552 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support ongoing for a patient admitted in with a st elevation myocardial infarction. it was reported that the pump malpositioning was against the mitral valve and causing minimal mitral regurgitation. the pump was removed and escalated to the 5.5. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issues most likely was use related due to improper technique.
16985 | 1220648-2024-20806 | 20506547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506547&pc=OZD | 08/01/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 08/01/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1363939 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. no patient was involved. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. when console was booted for the first time in danvers, console alarms are consistent with what was seen in the field. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16986 | 1220648-2024-20809 | 20506691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506691&pc=OZD | 07/20/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 07/20/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095656 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure after the software was upgraded that were resolved by changing the controller. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic2817) was sent back for further investigation. aic logs confirmed the reported failure. the failure mode was reproduced on an engineering bench. the battery 2 pack voltage was measured to be 3.5 v rather than the expected 16 v, and the battery lcd indicator was blank (fully discharged). battery 2 was replaced to resolve the failure alarm.the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16987 | 1220648-2024-20810 | 20506706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506706&pc=OZD | 06/16/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 06/16/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1565279 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | It was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. the battery failure alarm was due to a defective battery 1. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16988 | 1220648-2024-20807 | 20506709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506709&pc=OZD | 07/29/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 07/29/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117230 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. imc logs were reviewed. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). this report is being filed as part of a retrospective review of historical records.
16989 | 1220648-2024-20802 | 20506796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506796&pc=OZD | 09/29/2024 | 08/12/2025 |  | 10/22/2024 |  |  | 09/29/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542840 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The eu complainant had a 5.5 pump placed to support a patient admitted with known congenital disease. the pump was placed and after 32 days stopped. the pump stop occurred while the patient was walking. the team attempted to troubleshoot and restart but replaced the pump and adjusted the pharmacological support. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into pump stop has been completed. the root cause of the pump stop cannot be definitively determined as no product or logs were returned. the following have been reported incorrectly on manufacturer device report 1220648-2024-20802. | The impella device was received from the customer and an investigation regarding the returned device has been completed. the pump was returned with a cut catheter so the pump stop was unable to be reproduced. visual inspection revealed biomaterial in the outflow pre and post soaking. clinical details noted that pump had previous purge alarms two days before the pump stop. additionally, it was noted that patient was ambulating at time of the pump stop. no upward trend in motor current was noted. the cause of the pump stop was most likely biomaterial ingestion. patient ambulating at the time of the stoppage. no logs returned and pump was returned cut. thrombus failure mode was investigated also based on the cause of the pump stop most likely being biomaterial ingestion. the root cause of the thrombus was unable to be determined due to insufficient clinical details. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d.9 device returned to manufacturer date was added. a.3 added unknown selection as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-20802. h.10 instructions for use on the initial manufacturer device report number 1220648-2024-20802 contained more information then what pertained to just a pump stop. the proper pump stop instructions for use were added above.
16990 | 1220648-2024-20803 | 20506804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506804&pc=OZD | 12/25/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 12/25/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108317 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a battery failure red alarm that led to the case being aborted. | [5 empty] | The investigation has been completed. the console (ic1746) was sent back for further investigation. aic logs confirmed the reported failure. the failure mode was reproduced on an engineering bench. the battery 2 lcd indicator was blank (fully discharged). the battery failure was due to a defective battery 2. the root cause of the defective battery 2 was due to single cell failure. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16991 | 1220648-2024-20804 | 20506835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506835&pc=OZD | 09/27/2024 | 11/25/2024 |  | 10/22/2024 |  |  | 09/27/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544964 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 67-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. while on support, there was an intermittent and then continuous placement signal unreliable alarm. the representative discussed this may be an optical sensor issue. it may be a sensor or fiber so placement signals may be unreliable or possibly a kink in the catheter. the physician requested guidance for disabling the audio alarm. | [5 empty] | The investigation into placement signal issue has been completed. the root cause was changed to optical signal issue as failure is in relation to optical sensor. no product was returned for analysis. data logs investigated and no problem was found as the root cause of the optical signal issue. d6b added h6 code 765 was reported incorrectly on manufacturer device report 1220648-2024-20804 | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
16992 | 1220648-2024-20805 | 20506837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506837&pc=OZD | 07/29/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 08/01/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1615676 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic9269) was sent back for further investigation. aic logs confirmed the reported failure. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when console was booted for the first time in danvers, console alarms are consistent with what was seen in the field. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16993 | 1220648-2024-20808 | 20506857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506857&pc=OZD | 09/30/2024 | 11/20/2024 |  | 10/22/2024 |  |  | 09/30/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546737 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the pump was imaged and appeared to be under the papillary. attempts made to reposition were unsuccessful. the impella was exchanged with new cp catheter. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related.
16994 | 1220648-2024-20811 | 20506876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506876&pc=OZD | 10/02/2024 | 11/19/2024 |  | 10/22/2024 |  |  | 10/02/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515150 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp escalated to an impella 5.5 pump for a patient admitted in with cardiomyopathy and cardiogenic shock. the axillary site thrombosed and the site was evacuated and a new drain was placed. the patient additionally got fresh frozen plasma. the impella 5.5 pump remained on for support. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
16995 | 1220648-2024-20812 | 20506957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20506957&pc=OZD | 06/30/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 06/30/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm when the console was not in use. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic7560) was sent back for further investigation. aic logs confirmed the reported failure. battery switch was engaged and the console powered on without any alarms or issues. the root cause of the battery failure was due to not properly engaging the battery switch after returning from abiomed. this report is being filed as part of a retrospective review of historical records.
16996 | 1220648-2024-20814 | 20507011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507011&pc=OZD | 06/27/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 06/27/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1089158 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm, which was resolved by changing the console. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. the root cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16997 | 1220648-2024-20815 | 20507015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507015&pc=OZD | 06/20/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 06/20/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276151 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. console was powered on via battery power without issue. battery switch was disengaged and console was powered via ac without issue, triggering a battery failure as expected. while powered via ac, the battery switch was engaged and the battery failure cleared. there was no issue found with battery operation. this report is being filed as part of a retrospective review of historical records.
16998 | 1220648-2024-20820 | 20507036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507036&pc=OZD | 10/01/2024 | 11/23/2024 |  | 10/22/2024 |  |  | 10/01/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1153749 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella cp support being transferred to the secondary/tertiary medical center. when at 16,7000 feet in flight, the automated impella controller (aic) alarmed with a controller failure that was temporary. the issue resolved and when landed at the medical center, the aic was not noted to be replaced. no patient harm has been reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into automated impella controller (aic) - controller error (yellow alarm) has been completed. the data analysis showed that 5 hours into the pump running, the console alarmed controller error 1040 which resolved about six minutes later. controller error 1040 is triggered when the ambient pressure is out of range for at least two minutes, and according to the instruction for use (ifu), that range is -1000ft to 8000ft which is approximated as 550mmhg. the ambient pressure two minutes before the onset of the 1040 controller error occurred when the ambient pressure fell below 550 mmhg and resolved when the pressure came back above 550 mmhg. during the device analysis the failure mode was not reproduced as the console was not operated below 550mmhg. the root cause of the issue was the console going above 16700ft dropping ambient pressure below 550 mmhg threshold.
16999 | 1220648-2024-20816 | 20507050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507050&pc=OZD | 06/08/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 06/08/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1570900 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic8797) was sent back for further investigation. it was confirmed that persist. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. the root cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
17000 | 1220648-2024-20813 | 20507091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507091&pc=OZD | 06/27/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 06/27/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1111987 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a battery failure red alarm. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic1818) was sent back for further investigation. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. the root cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
17001 | 1220648-2024-20817 | 20507094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507094&pc=OZD | 05/24/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 05/24/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132873 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a controller error yellow alarm. the event occurred during a training setup. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic2351) was sent back for further investigation. when powering on ic2351 the controller error was reproduced. console interior was opened and checked, it was observed that the optical bench was not receiving power. the root cause of the aic issue was a defective optical bench. this report is being filed as part of a retrospective review of historical records.
17002 | 1220648-2024-20819 | 20507144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507144&pc=OZD | 09/30/2024 | 10/22/2024 |  | 10/22/2024 |  |  | 10/01/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505949 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support. approximately 7 hours post explant, the patient suffered embolic cerebral vascular accident (cva), with noted left side facial droop and left sided weakness. code stroke was called and patient was treated with tenecteplase. it was noted that the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17003 | 1220648-2024-20821 | 20507243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507243&pc=OZD | 05/13/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 05/13/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1392605 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a controller error yellow alarm, which was resolved through console replacement. there was no patient invovled. | [5 empty] | The investigation has been completed. the console (b)(6)) was sent back for further investigation. controller error and loss of optical data could not be reproduced. this is a known pattern failure, which is characterized by a loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. the failure was resolved by console power cycling. controller error and loss of optical data could not be reproduced. this report is being filed as part of a retrospective review of historical records.
17004 | 1220648-2024-20818 | 20507250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507250&pc=OZD | 09/27/2024 | 11/25/2024 |  | 10/22/2024 |  |  | 09/27/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534683 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Life Threatening | Disability | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 67-year-old male patient with cardiomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had low platelets and was heparin-induced thrombocytopenia and thrombosis positive.the staff changed from heparin to bivalrudin. support continued. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported thrombocytopenia has been completed. the device was returned for evaluation. no issues were found with the returned product. insufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia could not be determined.
17005 | 1220648-2024-20825 | 20507328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507328&pc=OZD | 03/24/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 03/24/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1360190 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a controller error yellow alarm during preventative maintenance. the sensor was replaced, and the console was run for an hour without issue. no patient was involved. | [5 empty] | The investigation has been completed. a review of the device logs confirmed the reported alarm. the root cause of the controller failure alarm was due to a defective purge pressure sensor. this report is being filed as part of a retrospective review of historical records.
17006 | 1220648-2024-20822 | 20507366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507366&pc=OZD | 09/29/2024 | 11/30/2024 |  | 10/22/2024 |  |  | 09/29/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520694 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support developed a large hematoma at the impella access site. the hematoma was present when the patient arrived at the intensive care unit from the cardiac catheterization lab. femstop was placed and heparin was held the night of insertion. the following day, the area was flat and ecchymotic. the patient survived the event. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17007 | 1220648-2024-20823 | 20507367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507367&pc=OZD | 05/12/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 05/12/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1323489 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a controller error yellow alarm, which was resolved through console replacement. | [5 empty] | The investigation has been completed. the console (ic5764) was sent back for further investigation. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. this report is being filed as part of a retrospective review of historical records.
17008 | 1220648-2024-20830 | 20507419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507419&pc=OZD | 02/15/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 02/15/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1183894 |  | 0042-0000-EU |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported that during preventative maintenance the automated impella controller experienced a controller error yellow alarm. no patient was involved. | [5 empty] | The device logs were reviewed and are consistent with the reported failure mode. a yellow controller error (pmd voltage out-of-spec) alarm was initially triggered, followed by the impella stopped alarm. although the pump stop was as expected during the short circuit preventative maintenance testing, the pump was unable to restart. in the following logs, many impellatronic communication error message appeared, as well as pmd resetting messages. eventually, the impella failure (speed deviation out-of-spec) alarm was triggered and self-resolved numerous times in the following few seconds. at the end of this power cycle, another controller error alarm appeared, indicating pbm voltages out-of-spec. the device was returned (date unknown) at abiomed in aachen germany, attempts were made to replace the impellatronic board, optical bench, and receptacle, all failed in solving this issue. the pbm board (0042-1125p) was replaced which succeeded in solving the issue. the controller error alarm was due to a defective power battery manager board (pbm) which failed to provide voltage to the impellatronic board. the exact root cause of defective pbm was unable to determine. this report is being filed as part of a retrospective review of historical records.
17009 | 1220648-2024-20831 | 20507432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507432&pc=OZD | 01/15/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 01/15/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095647 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the automated impella controller (aic) experienced a controller error yellow alarm, the resolution for this issue is unknown. no patient harm reported. | [5 empty] | The investigation has been completed. the device logs confirmed that there was a purge piston blocked event, resulting in the purge fluid not detected screen and eventually a controller error. the console was returned and the evaluation of the device found that the piston blocked event was due to build of glucose on the stepper motor gasket. when the console arrived for repair, the spindle would not move freely when the console was not powered on. this report is being filed as part of a retrospective review of historical records.
17010 | 1220648-2024-20833 | 20507434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507434&pc=OZD | 10/01/2024 | 11/25/2024 |  | 10/22/2024 |  |  | 10/01/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547602 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an impella cp implant and during the insertion, the sheath had a leak. the tip split on the sheath and a second introducer was used successfully. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into introducer damage ¿ mechanical has been completed. the root cause of the introducer damage - mechanical was not determined as no product or images were returned for analysis.
17011 | 1220648-2024-20835 | 20507440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507440&pc=OZD | 12/21/2021 | 10/22/2024 |  | 10/22/2024 |  |  | 01/05/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1422863 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the automated impella controller (aic) experienced a controller error yellow alarm, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the device logs confirmed the reported failure mode given there was a controller error alarms triggered during the clinical procedure. the console was sent back for further investigation. this known pattern failure, which is characterized by a temporary loss of optical data. triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the cause of the controller error (optical placement signal issue) was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
17012 | 1220648-2024-20827 | 20507504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507504&pc=OZD | 02/21/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 02/22/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 115417 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a controller error yellow alarm. actions taken are unknown. no patient was involved. | [5 empty] | The investigation has been completed. the device logs were reviewed and logs and are consistent with the complaint description. controller error alarm was issued upon boot-up. the cause of the aic issue was a defective kbd, the exact root cause of defective kbd was not determined. this report is being filed as part of a retrospective review of historical records.
17013 | 1220648-2024-20828 | 20507524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507524&pc=OZD | 09/29/2024 | 11/29/2024 |  | 10/22/2024 |  |  | 09/29/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1450938 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller experienced intermittent controller error alarms during patient support. the alarms self-resolved after each occurrence and there was no harm to the patient resulting from the failures. the staff was advised to swap to a back up controller to resolve the issue. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into automated impella controller (aic) - controller error (yellow alarm) has been completed. after the data log and device analysis the root cause of the controller error was determined to be a firmware issue. failures of this type will be monitored for trends.
17014 | 1220648-2024-20829 | 20507529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507529&pc=OZD | 02/16/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 02/16/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1311845 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller (aic) experienced a controller error yellow alarm, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the cause of the aic issue was a defective power battery manager board. this report is being filed as part of a retrospective review of historical records.
17015 | 1220648-2024-20826 | 20507531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507531&pc=OZD | 03/24/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 03/24/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1546473 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc/flag issue that was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. purge assembly was physically inspected, it was observed that the blue purge disc retainer had broken off. the root cause of the purge disc issue is due to a broken purge disc retainer. this report is being filed as part of a retrospective review of historical records.
17016 | 1220648-2024-20824 | 20507532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507532&pc=OZD | 05/04/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 05/04/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302148 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a controller error yellow alarm. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic5373) was sent back for further investigation. purge assembly of console was inspected. no dextrose or physical issues were found. an attempt to rotate the motor shaft by hand was unsuccessful as the shaft could not rotate freely. the motor shaft gasket was inspected, and it was observed that dextrose build up had caused the motor shaft to stall. the root cause of the inability of the motor shaft to rotate is due to dextrose build up on the motor shaft gasket. ¿ this report is being filed as part of a retrospective review of historical records.
17017 | 1220648-2024-20832 | 20507547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507547&pc=OZD | 01/03/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 01/13/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1422863 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the automated impella controller (aic) experienced a controller error yellow alarm, which was resolved by changing the console. no patient harm reported. | [5 empty] | D9: date of device return is unknown. the investigation has been completed. a review of the device logs found that there were no controller error or controller failure alarms upon boot-up. this is not consistent with the complaint. the device was returned to abiomed japan for evaluation (date of return is unknown). the root cause of the issue was not determined because issue could not be identified and/or reproduced. this report is being filed as part of a retrospective review of historical records.
17018 | 1220648-2024-20836 | 20507562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507562&pc=OZD | 09/29/2021 | 10/06/2025 |  | 10/22/2024 |  |  | 09/29/2021 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090692 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a controller error yellow alarm during maintenance. a loaner device was sent to the customer. no patient was involved. | [5 empty] | The intermittent failure exhibited in the field was most likely due to a malfunctioning pbm. although the failure was unable to be reproduced in house, the pbm was replaced out of an abundance of safety. the exact root cause of the malfunctioning pbm is undetermined. the investigation for the reported controller error has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the device logs were reviewed, and the failure could not be confirmed. the details provided indicate that the cause of system self-check failure seen in the field was likely due to a defective pbm component. however, the cause of the defective component is unknown. this report is being filed as part of a retrospective review of historical records. | The investigation for aic - controller error (yellow alarm) has been completed. data analysis: imc logs show the following, dates are likely off due to incorrect date set on the aic.05/12/11 00:02:09 sau_ser10 !! impellatronic without initialisation !! - powermod_1 communication this line entry is repeated throughout subsequent reboots. after 2 attempts at restarting the aic, the console boots into a system self check failure, as seen by the repeating entries, and remains in a failed state. this is most aachen fs received ic1341 and performed several boot cycles in an attempt to reproduce the failure .console did not reproduce the error. the pbm was inspected with no reportable findings. the intermittent failure exhibited in the field was most likely due to a malfunctioning pbm. although the failure was unable to be reproduced in house, the pbm was replaced out of an abundance of safety. the exact root cause of the malfunctioning pbm is undetermined. corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name updated.
17019 | 1220648-2024-20834 | 20507684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507684&pc=OZD | 01/13/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 01/13/2022 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 131333 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a controller error yellow alarm. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic5541) was sent back for further investigation. the console was tested on an engineering bench. it was confirmed that the console failed to boot up, which is consistent with the complaint. once console was opened, corrosion marks on power battery manager near super capacitor was observed. the console failure in boot up was due to power battery manager corrosion. this report is being filed as part of a retrospective review of historical records.
17020 | 1220648-2024-20837 | 20507722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507722&pc=OZD | 10/01/2024 | 11/27/2024 |  | 10/22/2024 |  |  | 10/01/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513204 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an attempted delivery of the impella 5.5 pump to support a patient admitted for a transcatheter aortic valve replacement. the impella 5.5 delivery failed. the team aborted the impella 5.5 delivery and instead placed an intra-aortic balloon pump. the patient¿s outcome is unknown. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature as per the clinical description provided.
17021 | 1220648-2024-20838 | 20507892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507892&pc=OZD | 09/30/2024 | 12/16/2024 |  | 10/22/2024 |  |  | 09/30/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510363 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and the patient started to bleed from their nose and the insertion site. coagulation was adjusted and the problem was resolved. the patient remained on impella support. | [5 empty] | D.6b explant date was added. the investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was determined to be anticoagulation management because the partial thromboplastin time value reported at the time of bleed is very low than the therapeutic range. bleeding was resolved and the hemostasis was achieved by adjustment of coagulation status. a.2 and a.3 revised as they were incorrect on initial manufacturer device report number 1220648-2024-20838. b.7 added as information was inadvertently omitted from initial manufacturer device report number 1220648-2024-20838. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-20838 was submitted. revised manufacturing site name and address section as several fields were reported incorrectly on the initial manufacturer device report number 1220648-2024-20838. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
17022 | 1220648-2024-20839 | 20507914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20507914&pc=OZD | 10/03/2024 | 12/06/2024 |  | 10/22/2024 |  |  | 10/03/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535361 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella cp implant and access was obtained with interventional sheath. vascular anatomy appeared would accommodate impella so proceeded to insertion. upon multiple attempts, staff were unable to advance and place impella sheath due to vascular anatomy. staff stated there was a possible issue with the sheath. the implant was aborted. the patient¿s outcome is unknown. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the delivery issue has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the sheath was unable to be inserted due to patient anatomy.
17023 | 1220648-2024-20840 | 20508297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20508297&pc=OZD | 10/03/2024 | 12/04/2024 |  | 10/22/2024 |  |  | 10/03/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547787 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support the patient experienced ventricular fibrillation, requiring several defibrillation attempts. was able to achieve return of spontaneous circulation (rosc). the patient survived the event. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
17024 | 1220648-2024-20841 | 20509135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20509135&pc=OZD | 09/27/2024 | 12/10/2024 |  | 10/22/2024 |  |  | 09/30/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529368 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient with an impella cp being inserted admitted in for a high-risk percutaneous coronary intervention. the .018 wire delivered to the left ventricle but the pump itself could not deliver due to kinking in the native tortuous anatomy. the aorta and femoral artery were known to be tortuous. a secondary impella cp was placed and supported successfully. there was no harm to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for delivery issue has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the patient has a extremely kinked vessels (aorta descendants and femoral art.), so that the implantation wasn¿t successful.
17025 | 1220648-2024-20842 | 20509465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20509465&pc=OZD | 10/04/2024 | 12/06/2024 |  | 10/22/2024 |  |  | 10/04/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534056 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, gastrointestinal (gi) bleed and bleeding at the groin were noted. femstop was applied and the patient was given one unit of red blood cells. anticoagulation was not ordered as a result of active bleed. the patient survived the event. | [5 empty] | The investigation into access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-20842. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
17026 | 1220648-2024-20845 | 20509895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20509895&pc=OZD | 09/22/2024 | 11/18/2024 |  | 10/22/2024 |  |  | 09/22/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534714 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support had a computed tomography scan and a fluid collection near the insertion sight was noted. the patient was started on zosyn and pan cultured. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported limb ischemia and sepsis has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and sepsis could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17027 | 1220648-2024-20843 | 20509967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20509967&pc=OZD | 09/22/2024 | 11/22/2024 |  | 10/22/2024 |  |  | 09/22/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530670 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed hemolysis while on impella cp support. the patients lactate dehydrogenase levels was over 2,000 and the patient had hematuria. the pump was repositioned, but the patient was ultimately made a do not resuscitate, and care was withdrawn and the patient expired prior to repeat labs and assessment to deem whether hemolysis had resolved. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the hemolysis has been completed. data log were reviewed finding many suction alarms occurred during impella support. no motor current (mc) spike observed. impella positioning alarms observed. the cause of the hemolysis most likely was improper positioning since the pump was found deep and possibly caught in a papillary muscle and the hemolysis (ldh) resolved after pulling the pump back 4cm. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
17028 | 1220648-2024-20844 | 20510036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510036&pc=OZD | 09/22/2024 | 11/22/2024 |  | 10/22/2024 |  |  | 09/22/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508856 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhagic Stroke | [11 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | Us complainant reported the patient was on impella 5.5 support. while on support, the patient developed fixed and dilated pupils. computed tomography revealed hemorrhagic stroke. the family elected to withdraw care following cerebral hemorrhagic stroke. chemical support was withdrawn. patient expired. |  |  |  |  | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. corrections to the initial mfr report: b2-selected death and date of death. b5-mso review. h1-type of reportable event changed to death. added-h6 impact code 1802. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17029 | 1220648-2024-20846 | 20510136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510136&pc=OZD | 03/17/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 03/17/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1255437 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the system self-check failed problem with the console during periodic maintenance. the console was repaired. no patient case involved. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the aic issue was contamination (corrosion). this report is being filed as part of a retrospective review of historical records.
17030 | 1220648-2024-20847 | 20510187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510187&pc=OZD | 03/17/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 03/17/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1390619 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced purge disc/flag issues. no clinical details regarding patient condition and resolution were provided. a replacement aic was shipped to the customer. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
17031 | 1220648-2024-20848 | 20510267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510267&pc=OZD | 03/17/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 03/17/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314852 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red error. no clinical details regarding patient condition and resolution were provided. a replacement aic was shipped to the customer. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to a compromised battery. this report is being filed as part of a retrospective review of historical records.
17032 | 1220648-2024-20849 | 20510293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510293&pc=OZD | 03/19/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 03/19/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1471015 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a purge issue, which was resolved by replacing the aic. no clinical details regarding patient condition were provided. a replacement aic was shipped to the customer. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the alarm of purge fluid not detected and rfid error (invisible to user) are due to defective pud. this report is being filed as part of a retrospective review of historical records.
17033 | 1220648-2024-20850 | 20510555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510555&pc=OZD | 09/22/2024 | 11/23/2024 |  | 10/22/2024 |  |  | 09/22/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529746 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support due to st elevated myocardial infarction, bradycardia and pulseless electrical activity arrest. while on support, the patient required ongoing fluid resuscitation and correction of severe lactic/metabolic acidosis. renal failure was noted and there was evidence of a new gastrointestinal bleed. care was withdrawn per family request and the patient expired. medical safety has reviewed the case and has determined that there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the gastrointestinal bleed, renal failure, and organ failure has been completed. the pump was not returned for investigation. the data log shows several suction alarms were reported during the case run, along with a deteriorating placement signal trend. no issues were noted with pump positioning or optical signal parameters. according to the clinical report, gastrointestinal bleeding was observed during the adjustment of anticoagulants during the pci procedure. signs of renal and multiple organ failure were noted during the case. arrhythmia was reported and successfully managed during case run. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the relation between gi bleed and impella was unknown. the cause of the renal failure issue was most likely related to the patient condition, as the data log justifies the deteriorating patient condition from the placement signal trend. the cause of the organ failure could not be determined without additional information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist : section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17034 | 1220648-2024-20851 | 20510629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510629&pc=OZD | 09/22/2024 | 11/20/2024 |  | 10/22/2024 |  |  | 09/22/2024 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525091 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp was inserted, the staff was unable to doppler pulses on the left foot. side port injection was performed revealing impella repositioning sheath was occlusive. external bypass was done utilizing existing right femoral artery sheath from percutaneous coronary intervention and connecting to a newly placed antegrade sheath on the left femoral artery. the patient expired after the family decided to withdraw care. medical safety has reviewed the case and has stated that there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the limb ischemia was not determined. b.2 added required intervention due to health effect - impact codes on initial manufacturer device report number 1220648-2024-20851.
17035 | 1220648-2024-20854 | 20510652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510652&pc=OZD | 06/15/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 06/15/2023 | 10/22/2024 |  |  | 02/28/2025 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024319634 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a purge system blockage and high purge pressure alarms appeared. leading up to the event, the patient also experienced oozing at their access site which only resolved after the tension sutures were readjusted. due to the patient's condition and the noted alarms, the decision was made to explant the pump. it was noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported high purge pressure and access site bleeding issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the high purge pressure issue was most likely kinked purge line based on data log review. the cause of the access site bleeding issue was most likely due to use as the sutures were fixed and tightened which resolved the issue. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17036 | 1220648-2024-20856 | 20510669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510669&pc=OZD | 10/23/2021 | 10/22/2024 |  | 10/22/2024 |  |  | 10/17/2024 | 10/22/2024 |  |  | 06/30/2023 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022105285 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/23/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with bipella, impella cp and rp, along with extracorporeal membrane oxygenationdevice for mechanical circulatory support. while on support, the impella cp was replaced due to persistent low flows. additional information was provided that noted multiple rounds of cardiopulmonary resuscitation and shocks were unsuccessful and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
17037 | 1220648-2024-20859 | 20510685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510685&pc=OZD | 08/12/2024 | 10/02/2025 |  | 10/22/2024 |  |  | 10/17/2024 | 10/22/2024 |  |  | 01/31/2025 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309579 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Brain Injury | [12 empty] | Life Threatening | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with extracorporeal membrane oxygenation-impella (ecpella) for mechanical circulatory support. from the first day of impella administration, the patient's brain prognosis was poor, and dnar was performed. no additional treatment, such as blood transfusion, was performed, and the procedure was completed. additional information was provided that noted the patient expired from hypoxic encephalopathy. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being filed as part of a retrospective review of historical records. the root cause of the stroke cannot be determined since the product was not returned and insufficient clinical details were provided. b1 required intervention was erroneously selected on the initial manufacturer device report number. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
17038 | 1220648-2024-20864 | 20510764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510764&pc=OZD | 06/22/2023 | 10/13/2025 |  | 10/22/2024 |  |  | 06/22/2023 | 10/22/2024 |  |  | 01/31/2025 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333716 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported that there was low pump flow with the impella cp. the clinical representative went on site and noted the impella flow was low and the position was unknown. patient had vt/vf for vfib arrest noted after impella was started; because positioning was lost, patient was escalated to cardiac bypass and impella was replaced after cabg was placed on cardiopulmonary bypass (cpb). impella cp restarted at p-1 and then p-2, but flows never increased past 0/0/0. impella inlet noted to be in aortic root, but would not flow. pump was turned off and the doctor explanted the device. ultimately the decision was made to insert a new cp. procedural outcome was patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
17039 | 1220648-2024-20857 | 20510767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510767&pc=OZD | 07/26/2021 | 10/22/2024 |  | 10/22/2024 |  |  | 07/26/2021 | 10/22/2024 |  |  | 02/28/2023 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022050127 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp smart assist support for hemodynamic support. while on support, the smart assist indicated there was no adequate placement signals and measurement. additional information provided that the patient expired due to maligned therapy resistant ventricular tachycardia (vt). medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that data logs were returned and many spikes in the placement signal to 3000 can be seen on the first day. no spikes were seen on the second day of support. these spikes triggered placement signal unreliable alarms. the product was returned and the placement signal functioned as expected. the placement signal spikes to 3000 when the patient cable is not fully connected to the console. the root cause of the placement signal issue was unable to be determined as the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
17040 | 1220648-2024-20855 | 20510788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510788&pc=OZD | 04/15/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 04/15/2023 | 10/22/2024 |  |  | 06/30/2024 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247903 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a large embolic cerebrovascular accident (cva) and does not have a good prognosis neuro wise. the family withdrew care due to the poor neuro prognosis. the pump was explanted and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported stroke issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the stroke issue was most likely patient condition related based on the clinical and data log review. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17041 | 1220648-2024-20860 | 20510809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510809&pc=OZD | 12/08/2021 | 10/22/2024 |  | 10/22/2024 |  |  | 12/08/2021 | 10/22/2024 |  |  | 04/30/2023 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022072707 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced low flows and suction. under imaging, the pump was observed to have a kink within the left ventricle. the physician attempted to manipulate and reposition the pump, but the kink persisted, and the device continued to produce low flows and suction alarms. the pump was explanted and replaced. procedure outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected and the pigtail kink was confirmed. no kinks in the pump cannula were confirmed. the pump was run in a basin of water and the low pump flow was unable to be reproduced. the pump operated to specification. the cause of the low pump flow was a kinked pigtail due to difficult patient anatomy. this report is being filed as part of a retrospective review of historical records.
17042 | 1220648-2024-20862 | 20510812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510812&pc=OZD | 03/22/2022 | 05/19/2025 |  | 10/22/2024 |  |  | 03/24/2022 | 10/22/2024 |  |  | 11/30/2023 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023160715 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp support in which the patient expired due to underlying cerebral hemorrhage. it was stated that the patient was being heparinized for right heart , which is when patient had stroke, and passed 4 days later. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | To make a determination as to the cause of the stroke, the following clinical information must be considered: patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation, timing of the stroke in relation to impella support (during or post-implant), detailed description of the symptoms experienced by the patient, neurological examination findings and any focal deficits observed, diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic), laboratory test results, including coagulation profiles and cardiac markers any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications, response to any previous anticoagulation or antiplatelet therapies, evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology, as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records. | The root cause of the stroke cannot be determined since the product was not returned and insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella cp ifu: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17043 | 1220648-2024-20865 | 20510815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510815&pc=OZD | 06/15/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 10/17/2024 | 10/22/2024 |  |  | 10/31/2024 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290763 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was needing an impella cp for mechanical circulatory support. during implantation it was reported that the patient¿s right iliac had tight calcific ostial lesion and the left side iliac to common femoral was <4mm. the patient¿s hemodynamics rapidly deteriorated and decision was made to attempt placement of the impella via the right femoral artery; however, the inlet was unable to advance past the lesion. an intra-aortic balloon pump was placed via the left femoral artery. the patient had deteriorated and subsequently expired following 30 plus minutes of resuscitation attempts. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was patient condition related due to patient having calcified vessels in right iliac femoral artery. this report is being filed as part of a retrospective review of historical records.
17044 | 1220648-2024-20858 | 20510836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510836&pc=OZD | 12/11/2021 | 10/22/2024 |  | 10/22/2024 |  |  | 10/17/2024 | 10/22/2024 |  |  | 05/31/2023 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080585 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there were recurrent false "impella in aorta" alarms. the alarms occurred intermittently and resolved spontaneously, and there was pulsatility in the motor current throughout the duration of support. the patient expired while on support, but no allegation was made relating the device to the patient's death. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the false impella position in aorta alarms was most likely ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
17045 | 1220648-2024-20861 | 20510841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510841&pc=OZD | 05/10/2021 | 10/22/2024 |  | 10/22/2024 |  |  | 10/16/2024 | 10/22/2024 |  |  | 01/31/2023 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022043027 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced oral and nasal bleeding. it was reported that activated clotting time control was required during the impella placement. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records.
17046 | 1220648-2024-20852 | 20510854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510854&pc=OZD | 03/20/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 03/20/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1234571 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the automated impella controller (aic) experienced a battery failure red alarm which was resolved by replacing the aic. this occurred prior to connecting to the pump. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure was due to a defective battery one. the root cause of the defective battery one was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
17047 | 1220648-2024-20853 | 20510955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510955&pc=OZD | 03/18/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 03/21/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1417541 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that that the patient was placed onto impella 5.5 support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced purge disc/flag issues that were resolved by switching to another console. the patient ultimately expired, but there was no allegation against the aic or impella related to the patient¿s death. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the cause of the issue was a broken/missing purge flag. this report is being filed as part of a retrospective review of historical records.
17048 | 1220648-2024-20863 | 20510977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20510977&pc=OZD | 03/03/2022 | 10/22/2024 |  | 10/22/2024 |  |  | 10/16/2024 | 10/22/2024 |  |  | 10/31/2023 |  | 10/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149807 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intraoperative Cardiac Valve Injury | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. during support, the sutures of the surgically-implanted aortic valve came undone. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported heart valve damage has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the heart valve damage could not be determined. however, it was noted that the impella cp cannot generate enough heat to affect sutures. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ this report is being filed as part of a retrospective review of historical records.
17049 | 1220648-2024-20866 | 20511138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20511138&pc=OZD | 05/01/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 05/01/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1432077 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a could not read disk error. the purge cassette was replaced, however the error persisted. the aic was replaced without any further reported issue. no harm or injury to patient reported. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. therefore, the root cause was due to a broken/missing purge flag. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17050 | 1220648-2024-20868 | 20511176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20511176&pc=OZD | 06/08/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 06/08/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548998 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm, which was resolved by switching to a backup console. no clinical details regarding patient involvement/condition were provided. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
17051 | 1220648-2024-20867 | 20511218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20511218&pc=OZD | 06/05/2023 | 10/22/2024 |  | 10/22/2024 |  |  | 06/05/2023 | 10/22/2024 |  |  |  |  | 10/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090369 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no patient was involved. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. there was no issue found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
17052 | 1220648-2024-20683 | 20492359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20492359&pc=OZD | 09/23/2024 | 12/12/2024 |  | 10/21/2024 |  |  | 09/26/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025470359 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and they had access site bleeding. the intervention performed for this issue was unknown. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the access site bleeding was most likely use related, the bleeding was controlled after the angle of the insertion site was adjusted as per the clinical description provided. corrections to the initial mfr report: d4 model number updated g1 mdr reporting contact email | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17053 | 1220648-2024-20684 | 20492387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20492387&pc=OZD | 09/26/2024 | 11/24/2024 |  | 10/21/2024 |  |  | 09/26/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526262 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had bloody urine noted in the foley. the impella was repositioned and the foley was replaced and hemolysis was successfully managed. it was further reported that placement signal alarms occurred with ectopy. the impella was repositioned and the ectopy since resolved with no further alarms. the patient survived the event. | [5 empty] | The investigation into vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
17054 | 1220648-2024-20685 | 20492506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20492506&pc=OZD | 09/23/2024 | 11/25/2024 |  | 10/21/2024 |  |  | 09/23/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524010 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp pump for mechanical circulatory support. while removing the peel-away sheath, the pump pulled back across the aortic valve and into the aorta. attempts to re-cross the valve were unsuccessful and the decision was ultimately made to remove the device. an intra-aortic balloon pump was subsequently placed for ongoing support. there was no patient harm. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was use related due to improper technique. added- h6 impact code 4642.
17055 | 1220648-2024-20686 | 20492810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20492810&pc=OZD | 09/26/2024 | 11/24/2024 |  | 10/21/2024 |  |  | 09/26/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024435788 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Cardiovascular Dysfunction/Insufficiency | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support starting around midnight had suction alarms occurring off and on with mean arterial pressures in the 40s per the automated impella controller (aic). no changes noted to current drips of levo 3 and dobutamine 1. continuous renal replacement therapy (crrt) was stopped earlier in the day and at this time the physician was not aware of pulmonary artery pressures. flows remained at 4.0 lpm per pump and the patient arterial line went from 40s to 10s. the patient expired. | [5 empty] | The investigation for the low pump flow has been completed. no product was returned. log review showed suction alarms throughout the case, but a larger quantity on the last day of support as reported. placement signal began to trend down in last 8 hours of case. there were no motor current spikes suggesting biomaterial ingestion. there were positioning alarms throughout the case, but no clinical details about positioning. pump flow was about 4.0 l/min at p-7, sometimes dropping lower, during the last 8 hours. the root cause of the low pump flow was most likely patient condition related. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17056 | 1220648-2024-20687 | 20493013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493013&pc=OZD | 09/25/2024 | 11/22/2024 |  | 10/21/2024 |  |  | 09/25/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539603 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support due to uncontrolled diabetes mellitus, multi vessel coronary artery disease, and acute hypoxia failure. while on support, the patient went into ventricular fibrillation. the patient was shocked back into normal sinus rhythm and regained full consciousness. the following day, the patient had four episodes of ventricular tachycardia requiring shock. additionally, the impella stopped. attempts were made to restart the pump and were unsuccessful. the pump was explanted. the patient survived the event. | [5 empty] | The investigation of vf/vt/af/at and pump stop has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the vt/vf was not determined. the cause of the pump stop issue was most likely open electrical line, based on data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17057 | 1220648-2024-20690 | 20493078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493078&pc=OZD | 09/18/2024 | 12/09/2024 |  | 10/21/2024 |  |  | 09/24/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528370 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was admitted in acute myocardial infarction/cardiogenic shock and was placed on impella cp support. the pump stopped after nine days. the patient was fortunate to be able to be weaned. the pump stop did not cause any harm or injury and the pump was not replaced. the patient survived the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into pump stop has been completed. the product was returned and evaluated. a large purge gap indicating bearing wear was found. no other abnormalities were found. log review showed a motor current spike before several "impella stopped - motor current high" alarms. the pump was able to be restarted, but then the same alarms occurred with a pump stop and it could not be restarted after multiple attempts. the initial pump stop occurred almost exactly 8 days into support. the final pump stop was on the 9th day of support. the root cause of the pump stop was bearing wear after almost 8 days of use, which is right at the 8-day design life. the following have been reported incorrectly or missing from manufacturer device report.
17058 | 1220648-2024-20689 | 20493086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493086&pc=OZD | 09/22/2024 | 11/21/2024 |  | 10/21/2024 |  |  | 09/24/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1359283 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that the automated impella controller (aic) displayed a red alarm, purge line click-on not detected. after comparing an unused aic with the aic that had caused the alarm, it was determined that the cause was a broken nail ¿ damaged assembly. the aic was replaced. the patient remained on impella support. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
17059 | 1220648-2024-20688 | 20493099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493099&pc=OZD | 09/26/2024 | 01/28/2025 |  | 10/21/2024 |  |  | 09/26/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095115 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient on impella 5.5 support starting around midnight had suction alarms occurring off and on with mean arterial pressures in the 40s per the automated impella controller (aic). there was brief period of controller failure. the patient expired. the patient¿s death is captured under the pump (5.5) complaint. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the automated impella controller failure - red alarm has been completed. there was no evidence of a controller failure from the logs but a brief period of transient loss of optical data which is not related to the suction alarms. the root cause for the transient loss of optical data was an optical bench fw communication protocol anomaly, resulting in the misalignment of data communication packets. | The automated impella controller (aic) was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. d.9 date device returned to manufacturerwas aadded. h.3 updated due to device being returned. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-20688 was submitted in accordance with updated procedures. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-20688 was submitted in accordance with updated procedures.
17060 | 1220648-2024-20691 | 20493262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493262&pc=OZD | 08/06/2024 | 11/25/2024 |  | 10/21/2024 |  |  | 09/23/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Other | [5 empty] | 72 YR |  | Male | 38 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Myocardial Infarction | Pleural Effusion | Tachycardia | Cardiogenic Shock | Tissue Breakdown | Unspecified Kidney or Urinary Problem | [7 empty] | Life Threatening | Required Intervention | Disability | [22 empty] | Invalid/Incomplete | Other |  | Yes | Initial submission | Followup | [8 empty] | The family member of a cardiac patient in california reported a user medwatch (ref# mw5159126). the daughter reported upon "potential medical negligence" and multiple adverse events that occurred to the patient supported by a 5.5 pump for 3 weeks of his hospitalization. there was reported ami, pleural effusion, tachycardia, cgs, kidney disease, possible systemic complications. no direct allegation against the impella causing the harms is made, but rather the medwatch speaks to the use of the 5.5 despite the recall is noted. | [5 empty] | There is no complaint console or serial number attached to this complaint. due to lack of case details, no data logs, and no product returned the root cause cannot be determined. the following have been reported incorrectly or missing from manufacturer device report . e4 should have been marked yes h7 medwatch number missing medwatch report attachment missing. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17061 | 1220648-2024-20694 | 20493279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493279&pc=OZD | 09/25/2024 | 11/24/2024 |  | 10/21/2024 |  |  | 09/25/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495517 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRA-CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp placed to support a patient admitted in post coronary artery bypass graft the week prior. the patient had the cp and extra-corporeal membrane oxygenation to support the patient's hemodynamics. the access site of the cp had an ooze that was treated by pressure and purse string sutures. the ooze resolved with those measures and gauze/dressing. the cp supported for 16 days until the pump repeatedly stopped. the pump was explanted. the patient survived the event. | [5 empty] | The investigation into pump stop and access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the pump stop issue was most likely biomaterial. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20694: e4 should have been left blank. g1 reporting contact fax number missing. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17062 | 1220648-2024-20693 | 20493333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493333&pc=OZD | 09/24/2024 | 11/22/2024 |  | 10/21/2024 |  |  | 09/24/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1271667 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that when the impella catheter was plugged in, an error message of optical sensor not detected occurred. staff unplugged and re-plugged the catheter in a few times before switching out consoles (automated impella controller). the same error code came up when it was plugged in to the same console. a new console worked successfully with no error message. no patient harm has been reported. | [5 empty] | The investigation for the optical issue has been completed. console logs from the reported day of event are consistent with complaint and show three attempts to use pump disconnecting and reconnecting, following case start wizard prompts - each time resulting in psnr condition, and eventually an alarm ¿placement signal not reliable¿ (#1036, optical signal invalid). the backup console sn had not been identified, therefore absence of psnr alarms on backup console could not be verified. rtlog data shows that when pump was connected snr was very low, and placement signal values could not be determined, resulting in psnr condition. console was returned for evaluation and contamination was found on the optical pump connector fiber ferrule. the cause of the aic issue was a defective optical bench. added-h6 impact code 4629. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
17063 | 1220648-2024-20692 | 20493413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493413&pc=OZD | 09/24/2024 | 11/24/2024 |  | 10/21/2024 |  |  | 09/24/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539578 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 pump for mechanical circulatory support. upon initiation of flows, a placement signal unreliable alarm appeared. the pump was subsequently replaced to resolve the noted failure. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into optical signal issue has been completed. no visual abnormalities were noted on returned pump. pump passed laser continuity testing, no biomaterial was present on pump. returned pump optical parameters were within a normal range. pump was run in a bucket of water for approximately 20 minutes and no psnr alarm was reproduced. no data logs were returned for evaluation. clinical details noted a psnr alarm was noted in the field after pump was initiated. the root cause of the optical signal issues cannot be determined as limited clinical details were provided and no data logs were returned. e4 should have been left blank on manufacturer device report 1220648-2024-20692.
17064 | 1220648-2024-20696 | 20493416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493416&pc=OZD | 09/25/2024 | 11/20/2024 |  | 10/21/2024 |  |  | 09/25/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | GM |  | 2025524408 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and they had oozing from the access site. angle matching was performed, dressing was changed and two units of red blood cells were provided to the patient. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was determined to be use issue because the bleeding was resolved and hemostasis was achieved by angle matching and dressing change. e4 should have been left blank on manufacturer device report: 1220648-2024-20696. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17065 | 1220648-2024-20695 | 20493421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493421&pc=OZD | 09/25/2024 | 11/20/2024 |  | 10/21/2024 |  |  | 09/25/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544991 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump being placed for delivery for the patient admitted in acute myocardial infarction / cardiogenic shock. the team was unable to deliver the 5.5 and so the team instead placed a cp. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass 5.5 beyond the subclavian into innominate. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17066 | 1220648-2024-20697 | 20493460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493460&pc=OZD | 09/26/2024 | 12/12/2024 |  | 10/21/2024 |  |  | 09/27/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525146 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support for 16 days before there was a pump stop. no motor current issues were reported prior to the pump stop. the pump was explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported pump stop has been completed. the device was not returned for evaluation. data logs were reviewed which revealed a motor current spike before the pump stopped and the "impella stopped - motor current high" alarms. the logs also showed there was no purge trend before the pump stop. therefore, the root cause of the pump stop could not be determined.
17067 | 1220648-2024-20698 | 20493564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493564&pc=OZD | 09/30/2024 | 12/20/2024 |  | 10/21/2024 |  |  | 09/30/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205056 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following initiation of impella 5.5 support, the staff was unable to toggle the support level on the automated impella controller (aic) above p2. the aic was swapped to a backup console to resolve the noted issue. there was no patient harm. | [5 empty] | The investigation for the kbd/rotary knob issues has been completed. data logs were reviewed which show that after pump start at p1, user couldn¿t choose any p-level above p2. this confirms the reported failure mode. the console was returned for evaluation. functional evaluation was able to reproduce the reported failure mode, the push button allows to toggle between only two selections, p0 and p1. upon unplugging the pump while running at p1 and re-plugging back, the console resumed running the pump at p1, and in this case, the push button allows to toggle only between p1 and p2. after replacing the original push button with a known good, the reported failure mode still exists. replacing the original impellatronic with known good resolved the reported failure mode. upon mounting the original impellatronic board on the known good test fixture, reproduce the reported failure mode. the root cause for unselectable menu options was a defective impellatronic board. added h6 impact code 4629 corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
17068 | 1220648-2024-20701 | 20493640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493640&pc=OZD | 09/24/2024 | 12/10/2024 |  | 10/21/2024 |  |  | 09/24/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025517285 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient admitted in acute myocardial infarction/cardiogenic shock on impella cp support. on the day of implant, the team observed a purge break and leak. the team troubleshot the pump by placing a tie at the red plug and using bone wax. the pump remained on for support despite the malfunction. the patient surived the reported issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the purge leak has been completed. no product was returned for evaluation. clinical details noted that red plug came apart where the red plug meets the clear purge sidearm when inserting the pump. as the clinical details noted that the pump came apart at the red plug which likely led to the purge leak. the root cause of the product damage (bond failure) cannot be definitively determined as limited clinical details were provided and no product was returned. corrections to the initial mfr report: d4 udi number g1-mdr reporting contact email.
17069 | 1220648-2024-20700 | 20493789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493789&pc=OZD | 09/26/2024 | 12/10/2024 |  | 10/21/2024 |  |  | 09/26/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495723 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp and extracorporeal membrane oxygenation (ecmo) support ongoing for a patient admitted in with acute myocardial infarction. the ecmo and impella cp access sites were bleeding and treated by red blood cell infusion. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely patient condition since bleeding was occurring from multiple sites. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17070 | 1220648-2024-20704 | 20493804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493804&pc=OZD | 09/26/2024 | 11/20/2024 |  | 10/21/2024 |  |  | 09/26/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547356 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | RANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the pump was implanted the flows were 3 liters per minute (lpm) in auto. staff had difficulty engaging left main through transcatheter aortic valve replacement (tvar) frame and with impella in place. during the process of adjusting various guides, notable decrease in flows to just about 1 lpm. physician felt the impella was creating more difficulty with engaging on and decided to explant at that time. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into low pump flow has been completed. the pump was returned for investigation. data logs show several motor spikes followed by low pump flow at the start of the case run. the impeller blade was found broken on the returned pump, and low pump flow was observed during the test. as per the clinical. patient had transcatheter aortic valve replacement (tavr) valve. pump flow was reported low from the start of the case and it drop further during adjustment. the cause of the low pump flow issue was fractured impeller blades. the fracture was characteristic of an interaction with another device most likely tavr but that was unable to be definitively determined since fluoroscopic imaging and/or additional clinical details were not provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20704.
17071 | 1220648-2024-20702 | 20493814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493814&pc=OZD | 09/24/2024 | 11/08/2024 |  | 10/21/2024 |  |  | 09/24/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1154073 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had a controller error (yellow alarm). there was a banner stating to switch to a back up controller. the aic was replaced. this issue was discovered during preparation. no harm has been reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the console (aic) controller alarm has been completed. data logs were returned finding that during the bootup, the log recorded the communication error ¿imc-kbd error: 0xa4 0x01 0x00. after bootup, the console immediately displayed ¿controller error (loss of comm with keypad or between the kbd and i2c for kbd hw revision 1) - alarm¿, event #24. this confirms the reported failure mode. the console was returned for evaluation. upon power cycling the console 200 times, it was unable to reproduce the reported failure mode - controller error. the root cause for the intermittent keyboard communication issue could not be determined due to the inability to reproduce. on further investigating for data streaming intermittently turning on and off and by going through the rlm logs it was evident that the root cause of this the data streaming was intermittently turning on and off was a corrupted microsd card in the impella connect module/rlm.
17072 | 1220648-2024-20705 | 20493888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493888&pc=OZD | 09/26/2024 | 12/12/2024 |  | 10/21/2024 |  |  | 09/26/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025496834 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant had a impella cp placed to support an acute myocardial infarction and cardiogenic shock patient. the pump was placed and due to groin site bleeding, noted to be "heavily" the team placed a 16fr medikit sheath to achieve hemostasis and a new impella cp was delivered. the procedural outcome was the patient survived surgery. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: b2-should have selected death and added date of death. b5-mso review. d4 model number updated. g1 mdr reporting contact email. h1-changed to death. h6 impact code 1802 added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17073 | 1220648-2024-20706 | 20493904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493904&pc=OZD | 09/26/2024 | 12/12/2024 |  | 10/21/2024 |  |  | 09/26/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025538946 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Eu complainant reported that two impella cp pumps were explanted around the same time ((b)(6)) and one of them was explanted due to patient bleeding. the clinical rep does not know which pump was explanted due to the bleeding issue. the issue will be documented against sn: (b)(6) however, there is a chance that the issue occurred with sn: (b)(6). the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17074 | 1220648-2024-20709 | 20493905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493905&pc=OZD | 09/27/2024 | 11/25/2024 |  | 10/21/2024 |  |  | 09/27/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544879 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to support a patient admitted in with known cardiomyopathy and heart failure. the patient suffered a bleed and was treated by the use of reversing protamine. the pump remains on for support therefore explant date and patient outcome are unknown. | [5 empty] | The investigation into access site bleeding has been completed. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis e4 should have been left blank on manufacturer device report 1220648-2024-20709. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
17075 | 1220648-2024-20707 | 20493908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493908&pc=OZD | 09/12/2024 | 11/23/2024 |  | 10/21/2024 |  |  | 09/26/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025517052 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Renal Failure | Sepsis | Tachycardia | Thrombocytopenia | [8 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 support in (b)(6) 2024 for a patient admitted in acute myocardial infarction / cardiogenic shock and the support duration was just over 8 days when the patient expired. there was no allegation made at the time. later in (b)(6) 2024 abiomed¿s long-term outcome and quality indicator (loqi) impella registry database reported upon four adverse events with relationship to the use of the impella. thrombocytopenia was possible relationship to the impella, sepsis also was possible, worsening of hepatic failure also was possible, and ventricular tachycardia also was possible. | [5 empty] | The investigation thrombocytopenia, infection/sepsis, renal failure, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20707. b3 changed to reflect the earliest event date. e4 should have been left blank. g1 reporting contact fax number missing. g2 report source; study missing. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17076 | 1220648-2024-20712 | 20493922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493922&pc=OZD | 10/03/2024 | 12/04/2024 |  | 10/21/2024 |  |  | 10/03/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547409 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed for a high risk percutaneous coronary intervention patient with multiple cardiac comorbidities. the patient had atrial fibrillation, mitral regurgitation, hypertension, hyperlipidemia, peripheral vascular disease, chronic obstructive pulmonary disease, and planned percutaneous coronary intervention. the pump was placed and perclose planned for hemostasis. after the explant and perclose adjusted/closed the patient condition deteriorated. the team believed a rp bleed to have occurred and they placed a femostop, a balloon tamponade, and covered stenting. multiple units of blood were also administered. the patient survived the event. | [5 empty] | The investigation for the retroperitoneal bleed has been completed. clinical details states that due to the retroperitoneal bleed, a 7mm covered stent was used to stop the bleeding. the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17077 | 1220648-2024-20699 | 20493940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493940&pc=OZD | 09/26/2024 | 11/19/2024 |  | 10/21/2024 |  |  | 09/26/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495399 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and the patient had hemolysis. the p level was decreased however, labs were still showing hemolysis. the impella was showing deep and the staff did not have plans to reposition. the p level was then increased and the patients urine was turning pink. the impella was explanted shortly after the events. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the hemolysis and positioning issues has been completed. data logs showed pump ran without issue during support. there were multiple suction alarms triggered near the end of the case when pump was run on p-1. no positioning issue alarm was triggered on the logs. product was not returned for review. per clinical description, echo confirmed pump was deep in ventricle but no reposition was done. the root cause of hemolysis was most likely due to pump was not in the proper position. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
17078 | 1220648-2024-20713 | 20493941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493941&pc=OZD | 10/03/2024 | 11/27/2024 |  | 10/21/2024 |  |  | 10/03/2024 | 10/21/2024 |  |  | 03/31/2025 |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024327479 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient with impella cp had a hematoma at the impella access site. the impella was tightened with perclose, manual pressure, and femstop was applied. the patient survived the event. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely use issue since hemostasis was achieved by tightening the perclose. the following have been reported incorrectly or missing from manufacturer device report: 1220648-2024-20713. e4: should have been left blank. g1: reporting contact fax number missing. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17079 | 1220648-2024-20703 | 20493986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20493986&pc=OZD | 09/26/2024 | 11/23/2024 |  | 10/21/2024 |  |  | 09/26/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546619 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had oozing at the access site and the labs were hemolyzing. a femstop was placed to assist with the bleeding and no known blood products were given. the impella position was confirmed to be in the mid ventricular position. it is unknown if the hemolysis resolved. the patient was transferred for more mechanical circulatory support as the impella cp was not enough. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation low pump flow and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20703.
17080 | 1220648-2024-20708 | 20494061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20494061&pc=OZD | 10/05/2024 | 11/21/2024 |  | 10/21/2024 |  |  | 10/05/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466805 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support. while on support, there were multiple suction alarms. the impella kinked. the impella was explanted due to product malfunction. the patient survived the event. | [5 empty] | The investigation for the impella cp has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of low flow was most likely due to a kinked cannula. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
17081 | 1220648-2024-20710 | 20494064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20494064&pc=OZD | 09/28/2024 | 11/27/2024 |  | 10/21/2024 |  |  | 09/28/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025491161 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella cp could not be inserted due to vascular anatomy. there was no reported patient harm. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20710 e4 should have been left blank h6 code 4627 was reported incorrectly. new code was added to health effect - impact code. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
17082 | 1220648-2024-20711 | 20494073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20494073&pc=PYX | 09/28/2024 | 11/19/2024 |  | 10/21/2024 |  |  | 09/28/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025477201 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support and there was slight bleeding at the access site. an additional suture was applied and the bleeding resolved. furthermore, biomaterial passed through the device causing a motor current spike and elevated pulmonary artery pressures. heparin was put in the purge. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation thrombus, access site bleeding, and saturated flow has been completed. the impella device was not returned for evaluation. due to lacking clinical detail and product not being returned, the root cause of the access site bleed could not be determined. product was not returned for investigation. based on clinical details and data logs, the root cause of the low pump flows was determined to be biomaterial. since the thrombus reported by the voice of the customer was determined to be the root cause of the low pump flows, the thrombus failure mode is captured under this. e4 should have been left blank on manufacturer device report: 1220648-2024-20711. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
17083 | 1220648-2024-20714 | 20494312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20494312&pc=PYX | 09/24/2024 | 11/22/2024 |  | 10/21/2024 |  |  | 09/24/2024 | 10/21/2024 |  |  | 12/31/2024 |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024296442 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic |  | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella cp support ongoing for a patient admitted in with cardiac history and suffering from acute myocardial infarction/cardiogenic shock. the team questioned the patient to be suffering from a pulmonary embolism as the coronary arteries were clean, and so the team made the decision to place an impella rp pump. the impella rp failed to deliver due to patient anatomy and the patient expired. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the implantation was aborted due to patient anatomy. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella rp smart assist system with automated impella controller: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp smartassist system with automated impella controller: section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
17084 | 1220648-2024-20716 | 20494470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20494470&pc=OZD | 09/27/2024 | 01/06/2025 |  | 10/21/2024 |  |  | 09/30/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1410591 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that automated impella controller (aic) was making an ¿odd grinding noise¿. no alarms or issues otherwise noted. the aic was exchanged. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) damage during therapy has been completed. data logs from the complaint date don't reveal any major purge related alarms on the date of occurrence. review of rt logs show no observable anomaly in the purge related anomaly. the console was returned for evaluation. the purge assembly unit and internal circuitry was visually inspected for anomalies without success. a known good pump and purge line was run on the console and the no unusual sound was heard. the console could not be confirmed for any failure as reported. the console logs showed no irregularities with the purge system on the date of occurrence up until the console was shutdown. there was no problem found with the console.
17085 | 1220648-2024-20715 | 20494523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20494523&pc=OZD | 09/29/2024 | 11/17/2024 |  | 10/21/2024 |  |  | 09/30/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1727322 |  | 0042-0040-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a automated impella controller (aic) support for a patient admitted in acute myocardial infarction and cardiogenic shock. the team noted a yellow aic alarm that prompted the team to replace the aic. the patient remains on support therefore the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) controller failure yellow alarm has been completed. data logs were reviewed which show two commands "process's sampling data from optical bench: sent stop acquisition cmd ack" then "clear all optical data shared memory: 0x4000". this results in"handle optical error conditions withv alid sample: calculator placement signal 130" which is indicative of the pattern failure mode loss of optical data due to a software defect. the console was tested on an engineering lab bench. the failure mode was not reproduced while running an optical pump simulator. there was no issue with the console hardware. the root cause of the controller error was due to an aic software issue. corrections to the initial mfr report: g1-mfr site name and address.
17086 | 1220648-2024-20717 | 20494591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20494591&pc=OZD | 09/24/2024 | 11/20/2024 |  | 10/21/2024 |  |  | 09/24/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525251 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had distal extremity cold and no pulse. the peel away sheath was still place in the intensive care unit. distal perfusion catheter placed, and extremity flow was evaluated. the patient hemodynamics improved, and pressors/inotropes were being weaned. the patient survived the event. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ h.6 revised component code as an incorrect selection was made on initial manufacturer device report number 1220648-2024-20717. additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-20717 in accordance with updated procedures. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17087 | 1220648-2024-20719 | 20494801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20494801&pc=OZD | 09/25/2024 | 11/22/2024 |  | 10/21/2024 |  |  | 09/25/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544153 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had extracorporeal membrane oxygenation cannulation and the patient experienced a brain bleed after the cannulation. heparin was discontinued and the impella was explanted shortly after. it is unknown the relationship between the impella and the patient's brain bleed. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17088 | 1220648-2024-20720 | 20495067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20495067&pc=OZD | 09/24/2024 | 11/24/2024 |  | 10/21/2024 |  |  | 09/24/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450908 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp was removed, thrombus was found on the cp side requiring a thrombectomy. the patient remained stable and survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details.
17089 | 1220648-2024-20718 | 20495164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20495164&pc=OZD | 08/29/2024 | 12/09/2024 |  | 10/21/2024 |  |  | 09/24/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025477614 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 smart assist support and during the insertion, air could not be removed from the side tube of the 23 french sheath in the graft insertion. the was no reported patient harm. | [5 empty] | The investigation for the introducer damage has been completed. data logs are not relevant to the complication and product was not returned for investigation. due to the limited clinical details and product not being returned, the root cause of the introducer damage could not be determined. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
17090 | 1220648-2024-20721 | 20495353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20495353&pc=OZD | 09/25/2024 | 11/18/2024 |  | 10/21/2024 |  |  | 09/25/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025512882 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and they had access site bleeding. anticoagulation was withheld and two units of packed red blood cells were administered. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
17091 | 1220648-2024-20722 | 20495427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20495427&pc=OZD | 09/19/2024 | 11/30/2024 |  | 10/21/2024 |  |  | 09/24/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025500186 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | Thrombosis/Thrombus | [10 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella cp support ongoing for a patient admitted in acute myocardial infarction/cardiogenic shock. after the impella cp was replaced, the arterial pressure waveform in the lower limb opposite the impella insertion disappeared, and it was determined that there was ischemia. a fogarty thrombectomy was performed and blood flow was confirmed. after returning to the intensive care unit, ischemia was again observed in both lower limbs. the potassium level also rose sharply to the 7 range, and dialysis was initiated. while preparing for surgical intervention for lower limb ischemia, the patient developed ventricular tachycardia and ventricular fibrillation. a computed tomography (ct) scan also showed lower limb ischemia and cerebral infarction, and the patient was confirmed dead shortly after the decision to do not resuscitation was made. there had been a renal infarction in the patient's 20s, so there was a thrombus in the aorta, but a review of the ct scan showed a thrombus in the impella shaft, which was likely dispersed when the thrombus was removed. patient factors and an embolism caused by the impella are considered to be the cause. | [5 empty] | The investigation for the thrombus, limb ischemia, and the stroke has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia, thrombus, stroke could not be determined as the device was not returned nor sufficient clinical details. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17092 | 1220648-2024-20723 | 20495620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20495620&pc=OZD | 06/14/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 06/14/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1337720 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced purge disc/flag issues. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. the purge disc retainer was completely cracked off and the purge flag was missing. the failure mode experienced in the field was able to be repeated. the root cause of the issue was a broken purge disc retainer and missing purge flag. this report is being filed as part of a retrospective review of historical records.
17093 | 1220648-2024-20728 | 20495812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20495812&pc=OZD | 05/31/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 05/31/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1439249 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller had a broken purge clip. no patient involvement was reported. | [5 empty] | The investigation for the reported purge disc/flag issues has been completed. the logs showed the last case in the field before the console was sent in for investigation. the product was returned. the purge disc retainer was slightly cracked and lifted off the purge assembly. it is likely this failure occurred when the purge disc was being removed at the end of the case which is why the logs do not reflect any failures. the cause of the issue was a broken purge disc retainer. this report is being filed as part of a retrospective review of historical records.
17094 | 1220648-2024-20725 | 20496223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20496223&pc=OZD | 06/20/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 06/20/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251206 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced purge disc/flag issues. the event occurred during setup. no patient was involved. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. this root cause of the issue was a broken purge disc retainer. this report is being filed as part of a retrospective review of historical records.
17095 | 1220648-2024-20743 | 20496395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20496395&pc=OZD | 10/03/2024 | 12/04/2024 |  | 10/21/2024 |  |  | 10/03/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547828 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that a patient was on impella cp support due to non-st elevated myocardial infarction (nstemi) and severe multi-vessel disease. while on support, ct results showed pseudoaneurysm at impella site. impella was explanted. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17096 | 1220648-2024-20733 | 20496701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20496701&pc=OZD | 05/05/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 05/05/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1176689 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced purge disc/flag issues. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic3517) was sent back for further investigation. aic logs confirmed the reported failure. visual inspection of the purge assembly area confirmed a ripped/ missing blue retainer. the failure mode was due to broken/ missing purge retainer. this report is being filed as part of a retrospective review of historical records.
17097 | 1220648-2024-20727 | 20496917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20496917&pc=OZD | 06/22/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 06/22/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1105809 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Audible Alarm | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced all alarms not audible. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic1681) was sent back for further investigation. during visual inspection in the interior of the aic circuitry, it was observed that the speaker connector to the self-expandable impella cp (epc) at terminal x30 was disconnected. the inaudible alarm was reproduced in the lab. when the speaker cable was unplugged, alarms were inaudible. when plugged in, the alarms were audible. there was no other visible damage to aic. the cause of the aic issue was a loose alarm connector. this report is being filed as part of a retrospective review of historical records.
17098 | 1220648-2024-20734 | 20497309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20497309&pc=OZD | 05/05/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 05/09/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1468353 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced purge disc/flag issues. no report of patient involvement. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. visual inspection of the purge assembly area confirmed a cracked blue retainer. the root cause of the issue was broken purge disk retainer. this report is being filed as part of a retrospective review of historical records.
17099 | 1220648-2024-20724 | 20497398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20497398&pc=OZD | 09/25/2024 | 11/24/2024 |  | 10/21/2024 |  |  | 09/25/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519904 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for support of a patient admitted in with cardiac history and prior coronary artery bypass graft. the patient had pre existing right bundle branch block, following a successful impella supported high risk percutaneous cardiac insertion. the patient developed complete heart block while the impella was being weaned. the impella did have to be increased for hemodynamic support and a temporary pacer was placed. following the pacer initiation, the patient's vitals were stabilized and the impella was weaned and explanted. the physician felt that the impella was across the aortic valve could have caused this arrythmia. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into vf/vt/af/at has been completed. the root cause of vt/vf cannot be determined since insufficient clinical details were provided. d9 date device returned to manufacturer added.
17100 | 1220648-2024-20726 | 20497792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20497792&pc=OZD | 06/21/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 06/21/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1345073 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller (aic) experienced purge disc/flag issues that were resolved by changing the controller. no patient harm was reported. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. the root cause of the aic issue was a defective optical bench. this report is being filed as part of a retrospective review of historical records.
17101 | 1220648-2024-20738 | 20497836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20497836&pc=OZD | 05/02/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 05/02/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1105811 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller (aic) experienced a purge disc/flag issue with no resolution specified. no patient harm was reported. | [5 empty] | The investigation has been completed. the console (ic1683) was sent back for further investigation. the console was tested after arriving. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer - see attached image). also, the inner side of purge disk was contaminated with massive glucose. the root cause of the issue was broken/ missing purge flag. this report is being filed as part of a retrospective review of historical records.
17102 | 1220648-2024-20740 | 20498065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20498065&pc=PYX | 10/03/2024 | 12/03/2024 |  | 10/21/2024 |  |  | 10/03/2024 | 10/21/2024 |  |  | 12/31/2024 |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024295630 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | SWAN-GANZ CATHETER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had placed a rp pump for support of a patient admitted in with acute myocardial infarction / cardiogenic shock and multiple cardiac and systemic issues, inclusive of dialysis. the 23 fr sheath was placed and the rp was placed, but it became enwrapped with the swan. the swan was seen within the pump's inlet cage. the team explanted and replaced with a new pump that successfully was delivered. the patient remains on impella support. | [5 empty] | The investigation for the delivery issues has been completed. clinical details state that the pump was on and running before complaint/issue occurred causing it to stop. the pump was inspected and the swan wire wrapped around the impeller blade was observed. the data logs confirm pump stop - motor current high alarms. the root cause of the pump stop was determined to be due to interaction with other device. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17103 | 1220648-2024-20739 | 20498377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20498377&pc=OZD | 04/21/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 04/21/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1400770 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced purge disc/flag issues, which was resolved by using a new console. actions taken are unknown. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic6921) was sent back for further investigation. the reported issue ¿ difficulty in insertion and removal of purge disc was confirmed. upon inspection on the purge assembly, glucose residue was found along the purge cassette insert and on purge retainer. the contamination was likely due to liquid leakage from previous usage. and the reported purge disc leaking was unable to confirm based on the investigation of console. the root cause of the failure mode was due to glucose contamination in purge area. this report is being filed as part of a retrospective review of historical records.
17104 | 1220648-2024-20736 | 20498433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20498433&pc=OZD | 05/05/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 05/05/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1380525 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a pc note detected failure that was resolved by replacing the aic. no patient harm was reported. | [5 empty] | The investigation has been completed. the console (ic6417) was sent back for further investigation. upon investigation, the purge assembly had a buildup of glucose and multiple signs of leaking that contributed to the purge failures. after cleaning, the console was able to function without failures. this report is being filed as part of a retrospective review of historical records.
17105 | 1220648-2024-20731 | 20498874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20498874&pc=OZD | 06/30/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 07/05/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1675422 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller (aic) experienced a controller error yellow alarm that was resolved by turning the pump down to p6. no patient harm was reported. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. the console was tested on an engineering lab bench and could not reproduce the issue. there was no issue with the console hardware. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. found a crack in the purge retainer (not related to the reported failure). the root cause for the controller error was not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
17106 | 1220648-2024-20735 | 20498886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20498886&pc=OZD | 06/01/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 06/01/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1461597 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller experienced a controller error alarm, which was resolved by changing the console. no patient harm was reported. | [5 empty] | The investigation has been completed. the console was returned for investigation. the logs confirmed that the reported failure mode, controller error alarms. the real time logs confirmed that this was a pattern failure which was characterized by a temporary loss of optical dataand triggering of a controller error alarm, has a low probability of reoccurrence. the cause of the transient loss of optical data was not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
17107 | 1220648-2024-20741 | 20499061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20499061&pc=OZD | 04/15/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 04/20/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1496398 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced purge disc/flag issues. actions taken are unknown. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. visual inspection of the purge assembly area confirmed a broken blue retainer on its screw stand, and purge flag was unable to be held in. the root cause of the issue was broken purge disk retainer. this report is being filed as part of a retrospective review of historical records.
17108 | 1220648-2024-20729 | 20499183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20499183&pc=OZD | 06/25/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 06/25/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1089169 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller (aic) experienced purge disc/flag issues that were resolved by changing the controller. no patient harm was reported. | [5 empty] | The investigation has been completed. the console (ic1300) was sent back for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. the root cause of the issue was a broken/missing purge flag. this report is being filed as part of a retrospective review of historical records.
17109 | 1220648-2024-20732 | 20499706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20499706&pc=OZD | 05/05/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 05/05/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1107439 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced purge disc/flag issues. the aic was replaced and the new device recognized the pump cassette with no issues. it was reported that this issue occurred during patient support and there was no harm to the patient. | [5 empty] | The investigation has been completed. the console (ic1705) was sent back for further investigation and tested. the issue with properly detecting the purge pressure disc was reproduced. the root cause of the issue was component failure, the purge flag was broken. this report is being filed as part of a retrospective review of historical records.
17110 | 1220648-2024-20737 | 20500034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20500034&pc=OZD | 05/04/2023 | 10/21/2024 |  | 10/21/2024 |  |  | 05/04/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276245 |  |  |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a purge pressure disk not being recognized alarm, which was resolved by using a new aic. there was no patient harm reported. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation and the device was tested. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the issue was broken/ missing purge flag. this report is being filed as part of a retrospective review of historical records.
17111 | 1220648-2024-20730 | 20500764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20500764&pc=OZD | 06/27/2023 | 09/26/2025 |  | 10/21/2024 |  |  | 06/28/2023 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1330939 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm. the event occurred during a training setup. no patient was involved. | The customer reported that the device exhibited an optical signal issue. |  |  |  |  | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. on boot up, the alarm was not able to be reproduced. the console was inspected for any visual malfunction. unrelated to the complaint, there was tear on the purge pressure transducer membrane. the controller error was generated due to a bad lamp. the root cause could not be determined because the malfunction could not be reproduced but was most probably an anomalous malfunction of the optical bench. this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d2, common device name. e2, health professional? h2, type of reportable event. h6, medical device problem code. h10, additional manufacturer narrative.
17112 | 1220648-2024-20753 | 20501003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20501003&pc=OZD | 11/22/2021 | 05/23/2025 |  | 10/21/2024 |  |  | 11/26/2021 | 10/21/2024 |  |  | 05/31/2023 |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094926 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding, including hematoma occurred and a transfusion was required. the patient expired due to multiple organ failure during impella support. the physician attributed no causal relationship between the device and the patient's death. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17113 | 1220648-2024-20742 | 20501057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20501057&pc=OZD | 10/03/2024 | 12/10/2024 |  | 10/21/2024 |  |  | 10/03/2024 | 10/21/2024 |  |  |  |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482030 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported the patient was on impella cp support. while on support, the patient developed a large hematoma. two units of packed red bloods cells were given, pressure was held, and the physician attempted to tighten perclose to help with oozing at the impella site. the patient remains on impella support. |  |  |  |  | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20742. b2 death and date of death added. b5 patient outcome and death rationale statement added. d6b missing. g1 reporting contact email. h1 type of reportable event. h10 health effect - impact code 1802 added. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17114 | 1220648-2024-20750 | 20501275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20501275&pc=OZD | 02/09/2023 | 10/06/2025 |  | 10/21/2024 |  |  | 02/14/2023 | 10/21/2024 |  |  | 11/30/2023 |  | 10/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151493 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A patient with a history of coronary artery bypass graft and pacemaker that was placed which resulted in infection and sepsis was placed on an impella cp device for mechanical circulatory support after the patient experienced a decline in cardiac function. it was reported that the first attempt to place the impella cp with a 14 french short sheath was unsuccessful as the pump appeared to be clogged and would not prime. the pump was removed and rinsed with distilled water. the short sheath was replaced by a long sheath and another attempt was made to place the impella cp. the second attempt was successful, and the patient was successfully support with the impella. it was reported the patient later expired while on impella cp support. it was reported the cause of death is unknown, but it was reported the patient's condition was severe even before the impella was inserted. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. difficulty inserting/removing pump through introducer: delivery issues was changed to difficulty inserting/removing pump through introducer. the root cause of the difficulty inserting/ removing pump through introducer issue was unable to be determined since the product was not returned for evaluation. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-20750 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-20750 in accordance with updated procedures.
17115 | 1220648-2024-20632 | 20489536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489536&pc=OZD | 01/29/2022 | 10/20/2024 |  | 10/20/2024 |  |  | 02/01/2022 | 10/20/2024 |  |  | 08/31/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127544 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  |  | Asian | Malfunction | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a 0.018" guidewire became caught on a diagnostic catheter. the wire was removed and then reinserted. the patient ultimately expired due to multiple organ failure, and the physician noted that there was no allegation of a device issue related to this outcome. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported product damage (guidewire damage - peripheral vessels) has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however the cause of the failure mode was use related since the guidewire got stuck in another device. this report is being filed as part of a retrospective review of historical records.
17116 | 1220648-2024-20633 | 20489792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489792&pc=OZD | 03/12/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 03/12/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1481995 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that that the automated impella controller (aic) experienced purge disc/flag issues. no patient was involved. | [5 empty] | E1: name of initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the failure mode was due to broken purge retainer. this report is being filed as part of a retrospective review of historical records.
17117 | 1220648-2024-20634 | 20489798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489798&pc=OZD | 03/07/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 03/07/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1608973 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no patient was involved. | [5 empty] | E1: name of initial reporter is unknown e2/e3: unknown the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
17118 | 1220648-2024-20635 | 20489807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489807&pc=OZD | 03/17/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 03/17/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1152883 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no patient was involved. | [5 empty] | E1: name of initial reporter is unknown. e2/e3: unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
17119 | 1220648-2024-20636 | 20489819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489819&pc=OZD | 03/21/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 03/21/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1610582 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm which was resolved by replacing the aic. no clinical details regarding patient condition were provided. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360). the logs showed a cell voltage decline in only one of battery 2¿s cells which occurred as the battery failure alarm triggered. the failure mode was reproduced. the battery 2 lcd indicator was blank (fully discharged). the battery failure was due to a defective battery 2. the root cause of the defective battery 2 was due to single cell failure. the root cause of the battery failure was due to a defective battery 2. the root cause of the defective battery 2 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
17120 | 1220648-2024-20637 | 20489831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489831&pc=OZD | 01/30/2022 | 10/20/2024 |  | 10/20/2024 |  |  | 01/31/2022 | 10/20/2024 |  |  | 07/31/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022113038 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 17 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during use, the pump stopped and was removed. the patient ultimately expired, but there was no allegation that the patient's expiration was related to the device. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the pump stop could not be determined. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17121 | 1220648-2024-20638 | 20489832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489832&pc=OZD | 02/12/2022 | 10/20/2024 |  | 10/20/2024 |  |  | 02/14/2022 | 10/20/2024 |  |  | 12/31/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164408 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 19 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. during support, the device exhibited a high purge pressure. multiple interventions were made, but these did not reportedly resolve the issue. the pump was explanted. the patient expired the following day, but there was no allegation that this was related to the impella pump. | [5 empty] | The investigation for the reported high purge pressure has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the high purge pressure was a kink in the purge line resulting in a purge blockage. this report is being filed as part of a retrospective review of historical records.
17122 | 1220648-2024-20639 | 20489833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489833&pc=OZD | 07/14/2021 | 10/20/2024 |  | 10/20/2024 |  |  | 07/14/2021 | 10/20/2024 |  |  | 03/31/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022057858 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, lower than expected flows during device use. the pump was removed, and a new catheter was connected. prior to the pump being replaced, the patient coded, and cardio-pulmonary resuscitation was initiated, and patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported product damage (cannula kink) has been completed. the product was returned. per the results of the clinical review and investigation: analysis of the returned product revealed a dried, white biological substance in the outflow and wedged between the impeller and the cannula. however, the pump was run in the lab and the low flow could not be reproduced. as data logs were not returned and low flow was not reproduced, the exact root cause of the low flow could not be determined. this report is being filed as part of a retrospective review of historical records.
17123 | 1220648-2024-20641 | 20489834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489834&pc=OZD | 08/05/2021 | 10/20/2024 |  | 10/20/2024 |  |  | 08/05/2021 | 10/20/2024 |  |  | 04/30/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED. INC. |  |  | US |  | 2022081066 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | Asian | Death | No | Yes | Pumping Stopped | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited a pump stop and could not be restarted. it was noted that the device was used past the expected life. it was reported that the patient died, reason unknown. no further information is available. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the investigation for the reported product damage (cannula kink) has been completed. the product was returned, and the pump stop was confirmed. per the results of the clinical review and investigation: the root cause of the cannula kink was determined to be wireless re-access which is not permitted with the use of impella cp. as noted in the impella ifu: impella cp with smart assist system. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
17124 | 1220648-2024-20640 | 20489835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489835&pc=OZD | 07/28/2021 | 10/09/2025 |  | 10/20/2024 |  |  | 07/28/2021 | 10/20/2024 |  |  | 06/30/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP | [5 empty] | US |  | 2022094327 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump stopped during use of the device. the physician was successful in restarting the device after reseating the white cable. it was reported that the patient died, no additional information was provided. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | Us complainant reported that the pt was placed on impella cp support for hemodynamic support. it was reported that the pump stopped during use of the device. the physician was successful in restarting the device after reseating the white cable. |  |  |  |  | B2: date of patient death is unknown. the investigation for the reported pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were sufficient to determine a root cause. the cause of the pump pause was most likely compression to the impella catheter by the surgical clamps that obstructed the pump's motor and optical cables. the root cause is excessive force. as noted in the impella ifu: impella cp with smart assist system. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records. | The product was not returned for investigation, only data logs were returned. data logs from the case showed that, about 2 hours into support, the pump stopped immediately upon the occurrence of the "impella stopped - motor current high" alarm, which was 12 minutes after surgical mode was disabled. it was also noted that the placement signal went to 3000 and the raw optical sensor went to -30000 simultaneously with the pump stop. the pump successfully restarted about 1 hour later, following reconnecting the impella and enabling surgical mode. in conclusion, after reviewing the clinical information, it was determined that the cause of the pump pause was most likely compression to the impella catheter by the surgical clamps that obstructed the pump's motor and optical cables.
17125 | 1220648-2024-20642 | 20489836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489836&pc=OZD | 12/28/2021 | 10/20/2024 |  | 10/20/2024 |  |  | 12/29/2021 | 10/20/2024 |  |  | 02/28/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067761 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced a hemorrhagic cerebrovascular accident (cva). the physician stated that this may have been related to over-anticoagulation and did not allege a relationship to the device, but the patient expired during impella support. | [5 empty] | The investigation for the reported stroke/cva has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the cerebral hemorrhage cva/stroke was determined to be unrelated to impella. no evidence or clinical imaging was provided. as noted in the impella ifu: stroke/cva: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17126 | 1220648-2024-20643 | 20489837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489837&pc=OZD | 07/04/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 07/04/2023 | 10/20/2024 |  |  | 03/31/2025 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346546 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. there was an access site bleed during impella cp support. upon local team arrival, impella was positioned fully in the aorta with repositioning sheath in place. the doctor unsuccessfully attempted to advance the impella across valve using repo sheath and attempted to use stiff guidewire to cross the aortic valve to assist in impella crossing. decision was made to remove impella entirely over stiff guidewire through repositioning sheath and use 25cm introducer for access at same site. impella was successfully placed a second time and repo sheath inserted. some steady bleeding was observed post-repo sheath placement. figure-of-eight stitch placed and manual pressure was held for 10 mins which resolved bleeding. upon move to cart, the patient went into supraventricular tachycardia which did not respond to pharmacological therapy and required synchronized cardioversion. patient was transferred to icu without incident. patient was unstable in icu with frequent arrhythmias. pt began to have unstable ventricular tachycardia which required chest compressions and defibrillation. the patient eventually went into asystole which could not be terminated. it was reported that the patient later expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17127 | 1220648-2024-20645 | 20489840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489840&pc=OZD | 01/31/2022 | 10/20/2024 |  | 10/20/2024 |  |  | 02/01/2022 | 10/20/2024 |  |  | 08/31/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119981 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. however, the physician was unable to advance the pump past the aortic arch due to calcification. the patient ultimately expired, but there was no allegation that the impella was related to this event. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was calcified vessels and a tight aortic arch. as noted in the impella ifu: delivery issue: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
17128 | 1220648-2024-20644 | 20489841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489841&pc=OZD | 10/01/2021 | 10/20/2024 |  | 10/20/2024 |  |  | 10/01/2021 | 10/20/2024 |  |  | 05/31/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022086237 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the impella pump was implanted emergently after cardiopulmonary resuscitation (cpr). it was reported that the pump showed signs of flow saturation. there is no indication of how this complication was managed. it was reported that the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported high purge pressure issues has been completed. the impella device has not been returned for evaluation. clinical details were unable to determine the root cause of the high inaccurate flow since product was not returned. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
17129 | 1220648-2024-20646 | 20489842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489842&pc=OZD | 04/15/2022 | 10/20/2024 |  | 10/20/2024 |  |  | 04/20/2022 | 10/20/2024 |  |  | 11/30/2022 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022031198 |  | 0048-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  |  | Asian | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported low pump flows during use of the device. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported low pump flows has been completed. the impella device or logs were not received from the customer. no impella device was received for evaluation. therefore, the root cause cannot be determined. this report is being filed as part of a retrospective review of historical records.
17130 | 1220648-2024-20649 | 20489843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489843&pc=OZD | 10/11/2021 | 10/20/2024 |  | 10/20/2024 |  |  | 10/11/2021 | 10/20/2024 |  |  | 07/31/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022122284 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Failure to Sense | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that placement signal issues, suction alarms, low flows, and significant bleeding form the pericardiocentesis drain occurred during support. heparin was put on hold due to bleeding. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issues has been completed. the impella device has not been returned for evaluation. a specific cause for the reported issues could not be determined through this investigation as no product was returned. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17131 | 1220648-2024-20648 | 20489844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489844&pc=OZD | 03/24/2023 | 03/07/2025 |  | 10/20/2024 |  |  | 03/24/2023 | 10/20/2024 |  |  | 09/30/2024 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023264393 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy and acute exacerbation of chronic heart failure was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced pulseless electrical activity. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information: b2 date of patient death: unknown. the investigation for the reported optical signal and death issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that no issues or abnormalities were found with the returned device, and it operated to specification within the lab. data logs were reviewed, and flows remained as expected for the duration of support. a loss in motor current and ps pulsatility is observed ~6 hours into the case, which aligns with the clinical account that the patient heart was not moving at all. therefore, the cause of the issue was patient condition. this report is being filed as part of a retrospective review of historical records.
17132 | 1220648-2024-20647 | 20489845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489845&pc=OZD | 03/20/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 03/20/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1550078 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm during a preventative maintenance (pm) check. no patient was involved. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. the failure mode was able to be reproduced during investigation. the battery 2 pack voltage was measured to be 3.5 volts rather than the expected 16 volts. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
17133 | 1220648-2024-20650 | 20489846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489846&pc=OZD | 09/12/2021 | 09/25/2025 |  | 10/20/2024 |  |  | 10/19/2021 | 10/20/2024 |  |  | 02/28/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067837 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Unknown |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella insertion site had infection. replacement of the pump was carried out, but did not resolve the issue. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-20650 g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
17134 | 1220648-2024-20651 | 20489847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489847&pc=OZD | 05/08/2021 | 10/20/2024 |  | 10/20/2024 |  |  | 10/29/2021 | 10/20/2024 |  |  | 01/31/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022038035 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. it was reported via a publication that there was left main trunk occlusion/thrombus while on impella cp optical support. the site reportedly believes that the occlusion was caused by impella-related thrombi and retrograde aortic bloodflow by ecmo (extracorporeal membrane oxygenation). additionally, subsequent tee (transesophageal echocardiography) revealed a thrombus in the impella inlet and during an exchange to impella 5.0, a migrated thrombus in the left atrium was also observed. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported thrombus issue has been completed. no device was received for evaluation. a root cause for the reported thrombus could not be determined through this investigation as no product or data logs were provided. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smartassist catheter insertion. section: axillary insertion of the impella cp with smartassist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17135 | 1220648-2024-20652 | 20489848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489848&pc=OZD | 03/25/2022 | 10/20/2024 |  | 10/20/2024 |  |  | 03/25/2022 | 10/20/2024 |  |  | 06/30/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100743 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant suspected hemolysis during use of the device, however, no intervention was reported. the patient expired the day following the device's removal. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the relationship between the use of impella and the hemolysis was unable to be determined as not product, logs, or clinical details were returned for analysis. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17136 | 1220648-2024-20653 | 20489849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489849&pc=OZD | 05/23/2022 | 10/20/2024 |  | 10/20/2024 |  |  | 05/23/2022 | 10/20/2024 |  |  | 12/31/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179234 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. during support, the device exhibited high purge pressure and heparin was discontinued. the procedural outcome is patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The investigation for the reported high purge pressure has been completed. the impella device was not received from the customer; however, the logs were returned for evaluation. after review of the logs, it was concluded the high purge pressure is most likely biomaterial buildup. this report is being filed as part of a retrospective review of historical records.
17137 | 1220648-2024-20654 | 20489850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489850&pc=OZD | 05/26/2022 | 10/20/2024 |  | 10/20/2024 |  |  | 05/26/2022 | 10/20/2024 |  |  | 01/31/2024 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183658 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, several hours after insertion the red hub started leaking blood at the red plug. the pump remained on for support for a day and was then explanted. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. no clinical details provided were sufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
17138 | 1220648-2024-20656 | 20489851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489851&pc=OZD | 06/02/2022 | 03/26/2025 |  | 10/20/2024 |  |  | 06/02/2022 | 10/20/2024 |  |  | 11/30/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169126 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited saturated flow, the resolution for this issue is unknown. additional information noted that the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. patient's peri-procedural condition was critical. | [5 empty] | This report is being filed as part of a retrospective review of historical records. corrections to initial mfr 1220648-2024-20656: b1: adverse event should not been selected. h1 type of reportable event: corrected to product malfunction h6 health effect - impact code: 1802 death should not have been reported. | The impella device and logs were returned for evaluation. upon evaluation of the device, saturated flows could not be replicated. this report is being filed as part of a retrospective review of historical records.
17139 | 1220648-2024-20655 | 20489852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489852&pc=OZD | 05/28/2022 | 10/20/2024 |  | 10/20/2024 |  |  | 05/28/2022 | 10/20/2024 |  |  | 11/30/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161267 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic |  | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. the physician was unable to place the impella, as the first look at the patients right iliac revealed the anatomy was too small to accommodate the impella. additionally, the physician checked the left femoral artery, and it was still too small. the team proceeded to administer cardio-pulmonary resuscitation and attempted percutaneous coronary intervention. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. per clinical review the cause of the delivery issues was the patients iliac was too small. the cause of the issue was patient condition related. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17140 | 1220648-2024-20657 | 20489853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489853&pc=OZD | 06/13/2022 | 10/20/2024 |  | 10/20/2024 |  |  | 06/13/2022 | 10/20/2024 |  |  | 03/31/2024 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206506 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed a brain bleed while being shifted. the patient ultimately expired, but it was not alleged that this was related to the impella. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported stroke/cva has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the hemorrhagic cva was most likely patient condition, specifically the patient's coagulopathy. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17141 | 1220648-2024-20658 | 20489854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489854&pc=OZD | 06/16/2022 | 10/20/2024 |  | 10/20/2024 |  |  | 06/16/2022 | 10/20/2024 |  |  | 11/30/2023 |  | 10/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161650 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for cardiomyopathy. the complainant reported that the patient expired of sepsis due to an infection during impella cp support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17142 | 1220648-2024-20659 | 20489900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489900&pc=OZD | 03/23/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 03/23/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1094051 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm during annual maintenance check. no patient was involved. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. logs showed that the batteries were intermittently failing to charge when ac power was connected. failure was reproduced and after replacing the power supply, the console was able to be fully charged up properly. the cause of the aic issue was a defective power supply. this report is being filed as part of a retrospective review of historical records.
17143 | 1220648-2024-20660 | 20489915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489915&pc=OZD | 03/26/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 03/26/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1613084 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no patient was involved. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) due to low pack voltage. additionally, the lt logs showed that battery 2 had a cell imbalance since cell v2 was 300 mv below the other cells. the failure mode was able to be reproduced. the battery 2 pack voltage was measured to be 3.5 volts rather than the expected 16 volts, and the battery lcd indicator was blank (fully discharged). the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
17144 | 1220648-2024-20661 | 20489927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489927&pc=OZD | 03/28/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 03/28/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401384 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller exhibited a black flag in the purge system broken message, the device was not recognizing the purge disk. no report of patient injury or harm. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. additionally, there was debris around the purge assembly and the gasket around the purge assembly was partially dislodged. therefore, the root cause of the issue was a broken/missing purge flag. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17145 | 1220648-2024-20662 | 20489969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489969&pc=OZD | 03/22/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 03/22/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1345540 |  | 0042-0000-CA |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no clinical details regarding patient condition or resolution were provided. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) and battery level low alarm (50% - event set #23) when ac not plugged in due to one of the batteries not working. the logs from the complaint date showed a cell voltage decline in only one of battery 1¿s cells which occurred as the battery failure alarm triggered. the failure mode was reproduced on an engineering bench. the battery 1 lcd indicator was blank (fully discharged). the battery failure was due to a defective battery 1. the root cause of the defective battery 1 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
17146 | 1220648-2024-20663 | 20489987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489987&pc=OZD | 03/29/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 03/30/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1586834 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no clinical details regarding patient condition or resolution were provided. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) due to low pack voltage. additionally, the lt logs showed that battery 1 had a cell imbalance since cell v1 was 400 mv below the other cells. the failure mode was able to be reproduced. the battery 1 pack voltage was measured to be 3.5 v rather than the expected 16 v, and the battery lcd indicator was blank (fully discharged). the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
17147 | 1220648-2024-20664 | 20489994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489994&pc=OZD | 03/29/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 03/29/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1336757 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no clinical details regarding patient condition or resolution were provided. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) due to low pack voltage. the failure mode was reproduced on an engineering bench. the battery 1 pack voltage was measured to be 3.5 v rather than the expected 16 v, and the battery lcd indicator was blank (fully discharged). batttest showed battery 1 had cell imbalance. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure and a depleted battery. this report is being filed as part of a retrospective review of historical records.
17148 | 1220648-2024-20665 | 20490007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490007&pc=OZD | 04/03/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 04/03/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1153140 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no clinical details regarding patient condition or resolution were provided. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360). the logs from the complaint date showed a cell voltage decline in only one of battery 2¿s cells which occurred as the battery failure alarm triggered. the failure mode was reproduced on an engineering bench. the battery 2 lcd indicator was blank (fully discharged). the battery failure was due to a defective battery 2. the root cause of the defective battery 2 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
17149 | 1220648-2024-20666 | 20490033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490033&pc=OZD | 05/03/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 05/03/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251888 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 28 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the blue purge holder broke and punctured a hole in the purge disc membrane. the device was replaced. no reported harm or injury to the patient was reported. | [5 empty] | The investigation has been completed. the console was returned for evaluation. the logs showed purge cassette removed at the beginning of the case when the blue receptacle had broken off. the blue receptacle was completely broken off that would have caused the inability to start a case. the cause of the issue was a broken blue receptacle. this report is being filed as part of a retrospective review of historical records.
17150 | 1220648-2024-20667 | 20490035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490035&pc=OZD | 03/31/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 03/31/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1094544 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no clinical details regarding patient condition or resolution were provided. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) due to low pack voltage. the failure mode was reproduced on an engineering bench. the battery 1 and 2 pack voltages were measured to be 3.5 v rather than the expected 16 v, and the battery lcd indicator was blank (fully discharged). replacing both batteries removed the failure modes so pbm was confirmed to be functional. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure and a depleted battery. this report is being filed as part of a retrospective review of historical records.
17151 | 1220648-2024-20668 | 20490040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490040&pc=OZD | 05/02/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 05/02/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210717 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller could not detect purge flow. multiple purge cassettes were used and issue could not be resolved. upon changing impella/purge cassette to a different console, the issue was resolved. no report of patient involvement. | [5 empty] | The investigation was completed. the console was returned for evaluation. the device logs showed repeated errors detecting purge fluid flow during the procedure. upon inspection of the console, the blue receptacle had a hairline fracture and was raised away from the purge assembly that would have caused the purge assembly to be unable to detect purge fluid during operation. therefore, the root cause purge disc/flag issue was due to a broken blue receptacle. this report is being filed as part of a retrospective review of historical records.
17152 | 1220648-2024-20669 | 20490053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490053&pc=OZD | 05/03/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 05/03/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1242767 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller had a damaged purge disk holder. a loaner was sent to the customer. no report of patient involvement was noted. | [5 empty] | The investigation was completed. the console was returned for evaluation. the device logs showed that the purge cassette and pump were removed in the middle of the case when the blue receptacle had broken off. upon inspection of the console, the blue receptacle was completely broken cracked and dislodge that would have caused the inability to start a case. therefore, the root cause of the purge disc/flag issue was due to a broken blue receptacle. this report is being filed as part of a retrospective review of historical records.
17153 | 1220648-2024-20670 | 20490060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490060&pc=OZD | 04/10/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 04/10/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1651013 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the nurse discovered that the purge disk area of the automated impella controller was cracked. no report of patient harm or injury. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the blue receptacle was completely broken cracked and dislodge (ic9651 broken blue receptacle) that would have caused the inability to complete the case wizard. therefore, root cause of purge disc/ flag issues was a broken blue receptacle. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17154 | 1220648-2024-20671 | 20490062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490062&pc=OZD | 05/09/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 05/11/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1286943 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 37 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller for mechanical circulatory support. the device exhibited a "purge line click-on not detected" error and the device was replaced successfully without further reported issues. when visually checked, the position sensing lever of the purge sensor unit was confirmed to have been damaged. no report of patient harm or injury. | [5 empty] | The investigation was completed. the console was returned for evaluation. the device logs revealed that the console issued the purge disc not detected alarm several times. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. the root cause of the purge disc/flag issues was a broken/missing purge flag. this report is being filed as part of a retrospective review of historical records.
17155 | 1220648-2024-20672 | 20490077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490077&pc=OZD | 05/14/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 05/16/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a pressure transmitter not detected during purge cassette change and the aic was replaced without further reported issues. the nurse stated that they were unsure which patient the aic was used on. however, there is no report of patient harm or injury. | [5 empty] | The investigation was completed. the console was returned for evaluation. the device log review from the complaint date revealed that the console issued the purge disc not detected alarm. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced and purge flag was confirmed to be missing. therefore, the root cause was a broken/missing purge flag. this report is being filed as part of a retrospective review of historical records.
17156 | 1220648-2024-20673 | 20490078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490078&pc=OZD | 04/17/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 04/17/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1327996 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller exhibited a purge disc not detected alarm. it was reported that the alarm was not associated with a bag or cassette change. the cassette was replaced but would not advance past the purge disc prompt. the console was replaced and no further issues were reported. no reported harm to the patient. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. therefore, the root cause of the issue was a broken/missing purge flag. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17157 | 1220648-2024-20674 | 20490080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490080&pc=OZD | 05/15/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 05/15/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella for mechanical circulatory support. during support, the automated impella controller (aic) displayed a "purge disc not recognized" alarm. the alarm could not be resolved, so the aic was replaced. there was no reported patient harm. | [5 empty] | The investigation was completed. the device was returned and evaluated. the logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. the console was tested. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing along with glucose buildup around the purge assembly. the cause of the issue was a broken/missing purge flag. this report is being filed as part of a retrospective review of historical records.
17158 | 1220648-2024-20675 | 20490092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490092&pc=OZD | 05/16/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 05/16/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1239546 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller displayed a "purge disc not detected" error while setting up a new pump. there was no patient involvement noted. | [5 empty] | The investigation was completed. the device was returned and evaluated. the logs from the complaint date revealed that the console issued the purge disc not detected alarm several times and was unable to complete case wizard. the console was tested. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. the cause of the issue was a defective component (broken/missing purge flag). this report is being filed as part of a retrospective review of historical records.
17159 | 1220648-2024-20676 | 20490096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490096&pc=OZD | 04/24/2023 | 11/08/2024 |  | 10/20/2024 |  |  | 04/24/2023 | 10/20/2024 |  |  | 08/11/2017 |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1295261 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Additional information noted that the care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during set up of the device, the device exhibited a "purge disc alarm". the staff tried a few times to remove/reset but message did not go away. the purge cassette kit set was replaced, and same message appeared. the aic was replaced, and no further issues were reported. no report of patient harm or injury. |  |  |  |  | Corrected information for the initial mfr 1220648-2024-20676 was provided in sections b1, b2, b5, h1, and h6. | The console (aic) was returned for evaluation. upon inspection of the console, there were no anomalies discovered, and the reported purge disc/flag issue could not be duplicated. therefore, the root cause could not be determined the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17160 | 1220648-2024-20677 | 20490102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490102&pc=OZD | 04/25/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 04/25/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431034 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller exhibited purge disk not detected alarm. the aic was replaced with no additional reported issues. no patient harm or injury reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. therefore, the root cause was due to a broken/missing purge flag. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17161 | 1220648-2024-20679 | 20490105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490105&pc=OZD | 04/21/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 04/21/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1486033 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller had a broken purge clip. a loaner was sent to the customer. no report of patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the blue receptacle was completely broken off. therefore, the root cause of purge disc/flag issues was a broken blue receptacle. (broken component) the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17162 | 1220648-2024-20678 | 20490106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490106&pc=OZD | 05/20/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 05/21/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1244243 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm when the pump was turned on. the pump was plugged in and the battry was in the on position. a different outlet was tried but the alarm presisted. no clinical details regarding resolution or patient condition were provided. | [5 empty] | The investigation for the reported automated impella controller (aic) - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. the aic logs confirmed the reported failure. there was a battery failure alarm due to low pack voltage. the failure mode was reproduced on an engineering bench. the battery liquid crystal display indicator was blank for one of the batteries of the pair. batttest showed cell 4 of battery 2 was critically depleted and battery 2 had cell imbalance. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure/imbalance causing a battery lockout. this report is being filed as part of a retrospective review of historical records.
17163 | 1220648-2024-20680 | 20490107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490107&pc=OZD | 04/25/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 04/25/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084965 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller exhibited an alarm that a purge disc is not detected even though it has one properly placed. a loaner was sent to the customer. no report of patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. therefore, the root cause of the purge disc/flag issue was a broken/missing purge flag. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17164 | 1220648-2024-20681 | 20490137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490137&pc=OZD | 04/28/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 04/28/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122397 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller was being used for an impella 5.5 implant and would not recognize the purge disc. the staff replaced the device with a new automated impella controller with no further issues. no harm or injury to patient reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. therefore, the root cause was a broken/missing purge flag. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17165 | 1220648-2024-20682 | 20490140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20490140&pc=OZD | 05/02/2023 | 10/20/2024 |  | 10/20/2024 |  |  | 05/02/2023 | 10/20/2024 |  |  |  |  | 10/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1253203 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. it was reported that an aic to aic transfer was completed but placement signal and waveforms did not appear on the aic. the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received. the complaint was alleged against the console. however, reviewing the cloud logs for the console, the complaint was not evident in the logs. thus, the complaint was most likely alleged against the console it was transferred to but that console's serial number could not be ascertained and thus could not be requested to be returned for investigation. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
17166 | 1220648-2024-20565 | 20489100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489100&pc=OZD | 11/24/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 11/24/2021 | 10/19/2024 |  |  | 03/31/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080908 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Death | Yes | No | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. post procedure, impella was inadvertently pulled into aorta and had to be removed from patient, flushed under table fluid, re-inserted successfully. patient required cpr during impella replacement. the complainant also reported a pump stopping during support, then restarting after approximately 1-2 seconds. the patient ultimately expired on support; however, there was no allegation of an impella cp related issue that caused/contributed to the death. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17167 | 1220648-2024-20566 | 20489101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489101&pc=OZD | 10/11/2022 | 05/23/2025 |  | 10/19/2024 |  |  | 10/11/2022 | 10/19/2024 |  |  | 02/29/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193757 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient died of sepsis. however, the relationship between the impella and sepsis was undetermined, and it was stated that the impella was not the direct cause of death. the patient expired after 18+ days of support, beyond the pump's indicated use. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17168 | 1220648-2024-20567 | 20489102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489102&pc=OZD | 11/08/2022 | 10/07/2025 |  | 10/19/2024 |  |  | 11/11/2022 | 10/19/2024 |  |  | 01/31/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190931 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. from the time of admission, the condition was severe, and the patient expired without recovery of cardiac function. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-20567. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-20567 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-20567 in accordance with updated procedures. h6 investigation conclusions code 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-20567.
17169 | 1220648-2024-20568 | 20489103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489103&pc=OZD | 12/05/2022 | 10/07/2025 |  | 10/19/2024 |  |  | 12/12/2022 | 10/19/2024 |  |  | 06/30/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023233645 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced a cerebral hemorrhage while on impella support. heparin had been administered for anticoagulant therapy with the impella. it is unknown what triggered the sudden cerebral hemorrhage, however the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-20568. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-20568 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-20568 in accordance with updated procedures. h6 investigation conclusions 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-20568.
17170 | 1220648-2024-20569 | 20489104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489104&pc=OZD | 11/15/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 11/15/2021 | 10/19/2024 |  |  | 08/31/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135016 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the pump had extremely low flows, and the patient never had a return of spontaneous circulation (rosc). the decision was made to remove the impella cp and place it back in the ventricle. however, the patient continued to require cardiopulmonary resuscitation (cpr). survival outcome at explant noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the low pump flow was patient condition. this report is being filed as part of a retrospective review of historical records.
17171 | 1220648-2024-20570 | 20489105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489105&pc=OZD | 11/20/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 11/20/2021 | 10/19/2024 |  |  | 06/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094986 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the pump was placed and immediately the pump would not achieve flow of more than .5 liters, with constant suction. the team suspected right heart failure, although the pump acted as if it had ingested a clot. this impella was removed and a new impella placed. additional information noted that the patient died, no further information was provided. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported low flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the low flow and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
17172 | 1220648-2024-20571 | 20489106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489106&pc=OZD | 12/05/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 12/06/2021 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Death | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient of unknown age/race/ethnicity and with an unknown primary indication was implanted with an impella cp for mechanical circulatory support. during support, the pump stopped. the pump was then explanted and replaced. the procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Date of patient death is unknown. the investigation for the reported pump stop has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the pump stop could not be determined. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17173 | 1220648-2024-20572 | 20489107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489107&pc=OZD | 12/21/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 12/21/2021 | 10/19/2024 |  |  | 08/31/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127476 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had a very short aortic root, while exchanging to a repositioning sheath a kink was observed. reportedly attempts to straighten the kink were unsuccessful, this device was replaced due to the low pump suction issues and the kinking. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown the investigation into the low pump flow issue has been completed. the impella pump was returned for investigation. the data logs showed the occurrence of the "suction" alarm several minutes after pump activation, accompanied by non-pulsatile motor current and decreased flow rates. prior to the "suction" alarm, there is a brief ventricular ps waveform. two notable kinks were found on the cannula, one was located near the outflow cage and the other was at the cannula bend. no other damage or abnormalities were found. in conclusion, it was determined that the cause of the low pump flow and suction was a kinked cannula that resulted that resulted from improper pump positioning. this report is being filed as part of a retrospective review of historical records.
17174 | 1220648-2024-20574 | 20489108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489108&pc=OZD | 07/22/2024 | 10/19/2024 |  | 10/19/2024 |  |  | 10/16/2024 | 10/19/2024 |  |  | 12/31/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163660 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support it was noted that the patient had poor systemic vascular properties. the patient's blood pressure gradually decreased, and eventually the auxiliary circulation stopped, and the patient died. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. resolved with treatment/premature device removal: an inadequate blood supply to the device arm or leg. diagnosed via cold limb or loss of pulse. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17175 | 1220648-2024-20573 | 20489109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489109&pc=OZD | 03/10/2022 | 10/07/2025 |  | 10/19/2024 |  |  | 10/16/2024 | 10/19/2024 |  |  | 11/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160738 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Thrombocytopenia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support it was noted that the patient developed hemolysis and thrombocytopenia. the patient went into respiratory, liver and renal failures with no noted intervention. additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-20573. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-20573 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-20573 in accordance with updated procedures. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2024-20573. h6 investigation conclusions code 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-20573.
17176 | 1220648-2024-20575 | 20489110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489110&pc=OZD | 04/21/2022 | 05/28/2025 |  | 10/19/2024 |  |  | 04/21/2022 | 10/19/2024 |  |  | 11/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160706 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that post impella insertion, hemolysis was confirmed. insertion site bleeding was also observed, but resolved after refixation of the indwelling sheath. mitral valve insufficiency was also noted (pulled around the chordae tendineae). it was stated that care was escalated to the impella 5.0. however, shortly after placement, the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17177 | 1220648-2024-20576 | 20489111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489111&pc=OZD | 03/03/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 10/16/2024 | 10/19/2024 |  |  | 11/30/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024299966 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the peel away was removed and the sheath was repositioned and advanced into place. no distal pulses bilaterally, and the patient had a very poor perfusion prior to the cp placement. the family withdrew care after 88.55 hours of support the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device has not been returned for evaluation. however, the root cause of the limb ischemia issue was patient condition related. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
17178 | 1220648-2024-20577 | 20489112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489112&pc=OZD | 07/02/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 10/16/2024 | 10/19/2024 |  |  | 02/28/2025 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342018 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia, reportedly the patients¿ lower extremities were cold, and they were unable to doppler the pulse. additionally, there were persistent suction alarms, the clinician discussed reducing the pressure to p1 with the physician, however the physician was aware and chose to stay at p2 with audio off. additional information was provided that noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the limb ischemia and low pump flow issues have been completed. the data logs were reviewed. the cause of the limb ischemia issue was most likely patient condition related and unknown relation to impella with ischemia observed on both legs per clinical review. the cause of the low pump flow issue was not determined due to product not being returned and clinical details provided were insufficient to determine a root cause. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17179 | 1220648-2024-20578 | 20489113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489113&pc=OZD | 08/13/2024 | 10/19/2024 |  | 10/19/2024 |  |  | 10/16/2024 | 10/19/2024 |  |  | 02/28/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022050291 |  | 0048-0034 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding including hematoma occurred. a blood transfusion was required. the complainant reported that although aggressive treatment was performed, when the flow of venoarterial extracorporeal membrane oxygenation (va-ecmo) was lowered, hemodynamics could not be maintained, and pulse pressure did not appear. after discussing with the doctor and family, the patient was do not attempt resuscitation (dnar). the patient was only to be supported if the pump turned on. the patient expired due to the underlying disease. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17180 | 1220648-2024-20579 | 20489115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489115&pc=OZD | 05/31/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 10/16/2024 | 10/19/2024 |  |  | 02/29/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023202985 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support the pump stopped during use. the team attempted to restart the pump several times without success. the pump was removed and replaced. additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was not received from the customer; however, the logs were returned for evaluation. after review of the logs it was determined the root cause of the pump stop was due to a result of high purge pressure likely due to biomaterial. as noted in the impella ifu: impella cp with smart assist system troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
17181 | 1220648-2024-20580 | 20489116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489116&pc=OZD | 06/05/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 10/16/2024 | 10/19/2024 |  |  | 06/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022101801 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited low flows. the pump was removed and replaced, but the patient subsequently expired. the doctor reported that the patient had global hypokinesis and was in pulseless electrical activity (pea). there was no allegation that the impella was related to the patient's death. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the low pump flow could not be determined. this report is being filed as part of a retrospective review of historical records.
17182 | 1220648-2024-20581 | 20489117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489117&pc=OZD | 06/09/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 06/09/2022 | 10/19/2024 |  |  | 02/29/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191786 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device came out of the left ventricle and could not be repositioned. the team elected to remove and replace the plump. in addition, the patient experienced hemolysis, though no intervention or resolution was reported regarding this issue. additional information noted care the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17183 | 1220648-2024-20582 | 20489118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489118&pc=OZD | 03/21/2023 | 05/29/2025 |  | 10/19/2024 |  |  | 07/29/2024 | 10/19/2024 |  |  | 06/30/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231994 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced cerebral hemorrhage and gastrointestinal hemorrhage. additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined
17184 | 1220648-2024-20584 | 20489119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489119&pc=OZD | 06/29/2023 | 05/23/2025 |  | 10/19/2024 |  |  | 08/12/2024 | 10/19/2024 |  |  | 11/30/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281949 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced a suspected infection. the catheter was removed, and an antibiotic treatment was started, but there was no improvement. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17185 | 1220648-2024-20585 | 20489120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489120&pc=OZD | 07/03/2023 | 05/23/2025 |  | 10/19/2024 |  |  | 07/03/2023 | 10/19/2024 |  |  | 01/31/2025 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314591 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient developed sepsis and died of multiple organ failure.medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). there was no reported impella access site infection. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17186 | 1220648-2024-20586 | 20489121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489121&pc=OZD | 12/29/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 12/29/2022 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1094057 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported by the biomedical engineer that the console(aic) had a optical sensor not supported message to occur. the console was switched out and the patient was successfully supported. the aic field software upgrade to12.1 was performed; however, there was no resolution for the optical sensor issues. there was no patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon review of the console logs, after the software upgrade the optical bench was not being initialized during console boot-up, and the pmd hw revision bit was set to ¿1¿, which prevented the optical bench from being utilized. therefore, the root cause of the aic issue was a work instruction/technical error. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17187 | 1220648-2024-20587 | 20489122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489122&pc=OZD | 01/10/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 01/10/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1418929 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a broken purge clip on this aic (automated impella controller) device. there was no adverse impact to any patient treatment reported. | [5 empty] | The console (aic) was returned for evaluation. upon visual inspection, the purge assembly area had a ripped/ missing blue retainer. the purge flag and retainer were replaced, the purge assembly was cleaned, preventive maintenance and functional check was performed on the console; before returning it to the customer. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17188 | 1220648-2024-20588 | 20489126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489126&pc=OZD | 01/17/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 01/17/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1127068 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported by the complainant that the automated impella controller had a red battery failure alarm. there was no reported patient harm. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of device, the battery failure alarm was reproduced; however, no unexpected controller shutdown was observed. batt test revealed that battery 1 shown an ¿smb¿ indicating the battery communication failure was most likely, due to low voltage. therefore, both batteries were replaced with new batteries, which were tested and proved to be in good working conditions; while maintaining a charge and a discharge. inconclusion, the root cause of the battery failure was, due to both batteries were depleted. the internal low voltage safety lockout mechanism, prevented both batteries from providing voltage and to be charged. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17189 | 1220648-2024-20589 | 20489127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489127&pc=OZD | 01/12/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 01/12/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276157 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an aic that failed the self check. the syscall error prompted the ifu to be followed. the aic was removed from service/patient use and a backup controller is being used. no patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the pbm was confirmed to have corrosion on the communication traces next to the super caps. fcn 71 had been implemented to address other pbm boards manufactured between march and november 2013 that do not currently show signs of damage, due to leaking capacitors. in conclusion, the root cause of the system self-check error was due to a defective pbm, which had corrosion that had leaking super capacitors. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17190 | 1220648-2024-20590 | 20489131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489131&pc=OZD | 01/05/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 01/05/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1662880 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure that occurred with this device. there was no adverse impact to patient support indicated. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the reported failure mode was reproduced on an engineering bench. the battery 2 lcd indicator was blinking (the blink pattern means malfunction). battery 2 was replaced which resolved the failure alarm. therefore, the root cause of the battery failure(red alarm) was due to single cell failure. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17191 | 1220648-2024-20591 | 20489132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489132&pc=OZD | 01/26/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 01/26/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1105128 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an aic (automated impella controller) alarmed with a controller error alarm. the alarm resolved itself within thirty seconds, however the decision was made to swap the console for service. the unit was removed and replaced by a new aic for ongoing patient support. there was no harm to the patient. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the logs revealed the reported issue and the controller error alarm were presented at the exact time reported, and self-resolved in 29 seconds. currently, there is a corrective action in progress for the reported issue. therefore, the root cause of the aic loss of opm could not be determined, due to the issue could not be reproduced. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17192 | 1220648-2024-20592 | 20489133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489133&pc=OZD | 01/28/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 01/28/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1262851 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had a purge disc failure on the aic (automated impella controller) for a 5.5 support. the team followed ifu recommendations and replaced the aic with the backup console. no harm or injury noted. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, testing was performed. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). the purge flag was replaced, the purge assembly was cleaned, preventive maintenance was performed, a functional check on the console and optical system conducted before the console was returned to the customer. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17193 | 1220648-2024-20593 | 20489136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489136&pc=OZD | 01/09/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 01/09/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1402264 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a false in aorta alarm following insertion for impella support. the alarm was subsequently disabled on the console and patient support continued. there was no harm to the patient. | [5 empty] | The console (aic) was not returned for evaluation. upon completion of the investigation, a supplemental report will be submitted.
17194 | 1220648-2024-20594 | 20489146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489146&pc=OZD | 02/07/2023 | 09/08/2025 |  | 10/19/2024 |  |  | 02/07/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1586601 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a placement signal not reliable alarm during use of the device. the pump was switched to a secondary console, resolving the issue. there was no harm to the patient. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, when the pump was likely physically disconnected from the console, it could not be detected due to a biased snr signal (greater than the programmed 150 limit) with no pump connected, indicating this symptom was due to an optical system-related issue. therefore, the root cause of the optical signal issue, could not be determined due the inability to be reproduce during investigation. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. when the pump was likely physically disconnected from the console, it could not be detected due to a biased snr signal (greater than the programmed 150 limit) with no pump connected, indicating this symptom was due to an optical system-related issue. to shutdown the console, the user had to perform an emergency shutdown. see ic9020 pump unplugged with snr still above 150 limit in the attachments section. the root cause of the optical signal issues was not determined due to the inability to reproduce. before sending this console back to the customer, fs was instructed to replace the optical bench (0042-1012) and front panel optical connector (0042-2736 or newer 0042-2750). fs also routinely perform a full functional check on the console and optical system (0042-7316) a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 added "2917" for optical signal issue.
17195 | 1220648-2024-20595 | 20489167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489167&pc=OZD | 02/07/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 02/07/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1128850 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Red alarm with purg disc not detected: replaced with 2 purge cassettes, these were also not recognized. replacing aic resolved the issue. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the reported issue of detecting the purge pressure disc was reproduced and the purge flag was confirmed to be broken (missing at blue disc retainer). the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17196 | 1220648-2024-20596 | 20489183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489183&pc=OZD | 02/14/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 02/14/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1155470 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Impella cp with smartassist (sn (b)(6)) stopped immediately after starting the pump in auto with ic (b)(6). restart attempts x2 were unsuccessful. aic failure alarm displayed. aic exchange to ic (b)(6) with successful pump start. no harm done to the patient. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show that after the last step of case start wizard pump was started but then stopped immadiately and console presented with alarm #13 (impella stopped due to motor current high). pump was restarted (at p-2 this time), but stopped again after 1s, with same alarm. another restart attempt (at p-4) was unsuccessful, as pump stopped after 1 with alarm #138 (pmd detecting pump speed deviation). console logs indicated that event #137 (pump speed deviation detected by cpld) might also have occurred, but there was no corresponding alarm. the next attempt to restart pump at p-4 was again unsuccessful, with pump stopping after 1s, with alarm #138. pump had been disconnected and transferred to (b)(6), where it had been started without any issues or alarms. console (b)(6) had been later that day field-tested with a demo pump, and it was able to run it, without alarms. the complaint console was tested in danvers with an engineering pump and pump simulator, but the issue was not reproduced, and each time pump was able to start. no irregularities in console operation were observed. separate pump investigation ((b)(6)) was concluded as ¿udt/nei¿. however, the results of pump testing were inconclusive: although no defects or malfunction of the returned pump were observed, any prior presence of biomaterial (that was later dislodged) could not be disproved.
17197 | 1220648-2024-20597 | 20489224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489224&pc=OZD | 09/16/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 09/16/2021 | 10/19/2024 |  |  | 06/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103727 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped shortly after being transferred to the unit. the pump was able to be restarted, however the patient¿s condition deteriorated, and cardiopulmonary resuscitation (cpr) was performed. multiple rounds of acls protocol were performed, but subsequently the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17198 | 1220648-2024-20598 | 20489233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489233&pc=OZD | 11/19/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 02/22/2022 | 10/19/2024 |  |  | 01/31/2022 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1473816 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, cerebral hemorrhage was confirmed during impella management. the device was explanted and the physician noted that the patient expired of cerebral hemorrhage the following day. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17199 | 1220648-2024-20599 | 20489239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489239&pc=OZD | 03/02/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 03/02/2022 | 10/19/2024 |  |  | 11/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169057 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump came to an abrupt stop, the motor current was reported to be within normal range prior to the pump stop. three unsuccessful attempts were made to restart the pump. the pump was pulled back to the aorta and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a large purge gap was found. the root cause of the pump stop was determined to be bearing wear. the pump stop occurred on day 14 of support which is beyond the design life of 8 days per design trace matrix 0046-9910. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17200 | 1220648-2024-20600 | 20489241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489241&pc=OZD | 04/29/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 05/09/2022 | 10/19/2024 |  |  | 02/28/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022050267 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, over 11 days the patient expired from cerebral hemorrhage due to cerebral vascular rupture cerebral bleed. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17201 | 1220648-2024-20601 | 20489243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489243&pc=OZD | 09/22/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 09/24/2021 | 10/19/2024 |  |  | 07/31/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105846 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, during the post operative process, the patient suffered a cerebral hemorrhage. the postoperative progress was good, and impella was weaned off to improve cardiac function, but the patient suffered a cerebral hemorrhage and expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17202 | 1220648-2024-20602 | 20489244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489244&pc=OZD | 09/24/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 09/28/2021 | 10/19/2024 |  |  | 07/31/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108791 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the impella device was removed after the patient¿s circulatory dynamics were disrupted due to suspected septic shock. the initial device was reinserted for the purpose of venting. the patient suffered a cerebral hemorrhage, and the procedural outcome was that the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17203 | 1220648-2024-20603 | 20489245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489245&pc=OZD | 12/16/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 12/23/2021 | 10/19/2024 |  |  | 08/31/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022122186 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Perforation of Vessels | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a rupture before impella insertion, a second rupture occurred after the pump was removed, the relationship between the rupture and impella is unknown. the patient also experienced renal failure, cardiac tamponade, ventricular perforation; all unknown relationship with no intervention reported. the patient had access site bleeding that required blood transfusion. moreover, the patient had cardiac arrest and the survival outcome at explant noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
17204 | 1220648-2024-20604 | 20489251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489251&pc=OZD | 01/28/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 02/02/2022 | 10/19/2024 |  |  | 09/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022128478 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Death | Yes | Yes | Fluid/Blood Leak | Material Split, Cut or Torn | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a purge leak, and the purge cassette was replaced. bleeding, including hematoma was also noted and a transfusion was required. the physician also reported causing damage to the sleeve of the pump. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported purge leak, sterile sheath damage issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the purge leak was not determined as the cassette was not returned for analysis. however, clinical details provided were sufficient to determine a root cause. the root cause of the sterile sheath damage was most likely use related. the patient death was determined to be unrelated to the use of impella as the cause of death was determined to be multiple organ damage. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17205 | 1220648-2024-20605 | 20489253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489253&pc=OZD | 02/04/2022 | 05/23/2025 |  | 10/19/2024 |  |  | 02/07/2022 | 10/19/2024 |  |  | 10/31/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137433 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient endured access site bleed from a 16fr sheath. the doctor stated that due to bleeding, the left ventricular filling could not be maintained, and the level of impella support could not be maintained. and ultimately the patient expired. bleeding was noted as the cause of death. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17206 | 1220648-2024-20606 | 20489264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489264&pc=OZD | 02/15/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 02/15/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1448687 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the console device failing to recognize the purge cassette. a backup console was used without further issue. there was no patient harm. | [5 empty] | The console (aic) was returned for evaluation. upon observation of the console, properly detecting the purge pressure disc was reproduced, and the purge flag was confirmed to be broken (missing at blue disc retainer). the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17207 | 1220648-2024-20607 | 20489276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489276&pc=OZD | 02/17/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 02/17/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1165236 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | High purge pressure issues. checked the purge lines and found no issues, switch patient over to tpa in case of clot. patient was then taken to ct and got a console failure red alarm. switched to a new aic and all issues were resolved. | [5 empty] | The console (aic) was returned for evaluation. the root cause of the purge pressure high alarm was highly likely due to a blocked purge line. the root cause of the controller failure red alarm was due to defective pbm. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
17208 | 1220648-2024-20608 | 20489294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489294&pc=OZD | 02/19/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 02/19/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1308008 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support due to cardiomyopathy. while on support, the automated impella controller (aic) experienced a controller error yellow alarm that was resolved by switching to a new console. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and there was no issue found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
17209 | 1220648-2024-20609 | 20489367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489367&pc=OZD | 03/01/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 03/01/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1088715 |  | 0042-0000-EU |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during pump set-up the automated impella controller (aic) did not initially recognize the catheter connected (impella cp with smart assist), configuring the set-up steps of a conventional cp (connector cable connection step). after another attempt successful configuration and case. no harm done to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. based on the review of logs, the cause of the aic (automated impella controller) issue was most likely a malfunctioning impellatronics board, combined with the failure of software mitigation mechanism. this report is being filed as part of a retrospective review of historical records.
17210 | 1220648-2024-20610 | 20489416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489416&pc=OZD | 02/28/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 02/28/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED INC |  |  | US |  | 1283668 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an automated impella controller (aic) with a syscal error after being powered up. the instructions for use (ifu) was followed and the aic was replaced with a backup and field service notified. there was no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the cause of the aic (automated impella controller) issue was contamination. this report is being filed as part of a retrospective review of historical records.
17211 | 1220648-2024-20611 | 20489428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489428&pc=OZD | 03/15/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 03/15/2022 | 10/19/2024 |  |  | 11/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157955 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a lung contusion from the time of introduction of veno-arterial extracorporeal membrane oxygenation (va ecmo) and the impella due to the effect of cardiac massage during cpr, and continued bleeding from the alveoli and thoracic cavity. while on support, additional treatments included blood transfusions, heparin dose adjustments, and chest tube drainage. however, the patient's body volume gradually decreased, and eventually va ecmo and impella's proper support could not be continued. when the above progress was notified to the patient's family, it became a policy of care, and the patient expired. the physician commented that the cause of death was myocardial infarction and not the impella. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported stroke/cva issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17212 | 1220648-2024-20612 | 20489435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489435&pc=OZD | 01/09/2022 | 05/23/2025 |  | 10/19/2024 |  |  | 03/08/2022 | 10/19/2024 |  |  | 06/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092861 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had some bleeding that required a blood transfusion. the survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17213 | 1220648-2024-20613 | 20489436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489436&pc=OZD | 03/06/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 03/06/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1398838 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  |  | Asian | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Japan complainant reported that that the patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced purge disc/flag issues that were resolved by replacing the aic. the patient ultimately expired, but there was no allegation against the aic or impella related to the patient¿s death. | [5 empty] | The investigation for the reported issue has been completed. the cause of the issue was broken/ missing purge flag. this report is being filed as part of a retrospective review of historical records.
17214 | 1220648-2024-20614 | 20489437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489437&pc=OZD | 08/02/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 08/02/2021 | 10/19/2024 |  |  | 04/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022081058 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding occurred and hemostasis was performed, and it was confirmed that the bleeding was from the site of central anastomosis of the coronary artery bypass grafting (cabg) and bleeding was not related to the impella. the survival outcome at explant noted the patient expired. there was no causal relationship between the bleed or the patient's death and the impella device. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
17215 | 1220648-2024-20615 | 20489438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489438&pc=OZD | 05/13/2022 | 10/07/2025 |  | 10/19/2024 |  |  | 05/13/2022 | 10/19/2024 |  |  | 11/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151330 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 78 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support for just under 2 weeks, the patient expired due to sepsis. the infection could not be controlled, and the patient died before being discharged. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | B2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2024-20615. g1 reporting contact email revised on manufacturer device report 1220648-2024-20615 in accordance with updated procedures. g2 report source; foreign was missing from manufacturer device report 1220648-2024-20615 h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2024-20615. h6 health effect - impact code 4648 and investigation conclusions code 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-20615.
17216 | 1220648-2024-20616 | 20489439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489439&pc=OZD | 01/03/2023 | 10/08/2025 |  | 10/19/2024 |  |  | 01/06/2023 | 10/19/2024 |  |  | 10/31/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276695 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient died of hemodynamic failure due to gastrointestinal bleeding. it was observed that there were bleeding tendencies throughout the body, and circulation could not be maintained, so care policy was adopted and the patient expired. there was no operational problems. it was indicated that it is unknown if the death is related to the device. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | G1 reporting contact email revised on manufacturer device report 1220648-2024-20616 in accordance with updated procedures. upon further review heparin was used in the purge solution the following fields were revised from the initial submission 1220648-2024-20616 in accordance with updated procedures. b2 death and date of death. h1 type of report. h6 health effect - impact code 1802.
17217 | 1220648-2024-20617 | 20489440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489440&pc=OZD | 03/19/2023 | 10/08/2025 |  | 10/19/2024 |  |  | 03/22/2023 | 10/19/2024 |  |  | 03/31/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200225 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had hemodynamic breakdown after impella insertion, and extracorporeal membrane oxygenation (ecmo) was added. the patient had bleeding from the respiratory tract, gastrointestinal tract, etc. with no bleeding at the impella insertion site and the patient expired. the cause of death was sudden cardiac arrest and with no relationship to the impella. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | G1 reporting contact email revised on manufacturer device report 1220648-2024-20617 in accordance with updated procedures. h6 investigation conclusions 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-20617. upon further review heparin was used in the purge solution the following fields were revised from the initial submission 1220648-2024-20617 in accordance with updated procedures. b2 death and date of death. h1 type of report. h6 health effect - impact code 1802.
17218 | 1220648-2024-20618 | 20489446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489446&pc=OZD | 08/06/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 08/06/2022 | 10/19/2024 |  |  | 04/30/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211942 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, gastrointestinal bleeding occurred. bleeding was observed in both the upper and lower gastrointestinal tracts. since no endoscopy was performed, it was not possible to identify the bleeding site. for activated clotting time (act), heparin was used, and no treatment such as discontinuation of anticoagulant therapy had been performed. the bleeding could not be controlled, the treatment was discontinued, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
17219 | 1220648-2024-20619 | 20489447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489447&pc=OZD | 01/28/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 01/28/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1450701 |  | 0042-0040-US |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced optical signal issues that were resolved by switching to a backup console. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the optical signal failure was an air gap from between the optical cable and the pump. the cause of the positioning issue could have been patient condition related which resolved during the console swap. this report is being filed as part of a retrospective review of historical records.
17220 | 1220648-2024-20620 | 20489449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489449&pc=OZD | 03/08/2023 | 11/08/2024 |  | 10/19/2024 |  |  | 03/08/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1169582 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported that that the patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced a battery failure red alarm that required removal of the aic. the patient ultimately expired as care was withdrawn, but there was no allegation against the aic or impella related to the patient¿s death. |  |  |  |  | Corrected information for the initial mfr 1220648-2024-20620 was provided in sections b1, b2, b5, h1, and h6. | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure was due to a defective battery 1. the root cause of the defective battery 1 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
17221 | 1220648-2024-20621 | 20489459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489459&pc=OZD | 03/08/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 03/08/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1134169 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that that the automated impella controller (aic) experienced purge disc/flag issues. no clinical details with regard to patient condition and resolution were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. root cause of the issue was due to the defective purge retainer. this report is being filed as part of a retrospective review of historical records.
17222 | 1220648-2024-20622 | 20489469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489469&pc=OZD | 03/12/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 03/12/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1489955 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Decrease in Pressure | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention and cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the console gave a controller failure alarm (red alarm) that self-resolved and then several random air in purge alarms. while on impella, it appeared that the pump had just been started and all previous data is no longer visible. customer decided to continue without console exchange. procedure successfully completed; patient stable. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was most likely use related. this report is being filed as part of a retrospective review of historical records.
17223 | 1220648-2024-20623 | 20489493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489493&pc=OZD | 03/13/2023 | 10/19/2024 |  | 10/19/2024 |  |  | 03/13/2023 | 10/19/2024 |  |  |  |  | 10/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1135097 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no clinical details regarding patient condition and resolution were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure was due to a defective battery two. the root cause of the defective battery two was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
17224 | 1220648-2024-20624 | 20489523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489523&pc=OZD | 01/09/2023 | 10/07/2025 |  | 10/19/2024 |  |  | 01/11/2023 | 10/19/2024 |  |  | 10/31/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276018 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, due to the development of multiple complications, mainly sepsis, no further treatment was given. the patient subsequently expired. it was indicated that there were no problems with the impella device. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-20624. g1 reporting contact email revised on manufacturer device report 1220648-2024-20624 in accordance with updated procedures. h6 investigation conclusions 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-20624.
17225 | 1220648-2024-20625 | 20489525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489525&pc=OZD | 02/13/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 02/14/2022 | 10/19/2024 |  |  | 07/31/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114671 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited a placement signal failure. the pump was removed and replace with an intra-aortic balloon pump. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The device was returned and evaluated, and the investigation for the reported optical signal placement issue has been completed. the pump was run in the lab and worked as intended, no issues were found. thus, the lv waveform was likely not seen as a result of the low flow and suction, likely caused by patient condition. this report is being filed as part of a retrospective review of historical records.
17226 | 1220648-2024-20626 | 20489528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489528&pc=OZD | 03/16/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 03/16/2022 | 10/19/2024 |  |  | 11/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165817 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. before using the impella, a thrombus had already been confirmed in the left chamber, but impella was inserted before coronary artery bypass grafting (cabg). after that, it was transferred to the cardiopulmonary system, and when the operation was transferred from the cardiopulmonary system to impella at the end of the operation, the impella flow rate did not increase from 0.5 l / min. echo was confirmed, but there was no problem with the volume at the position, so it was possible that a thrombus was inhaled. when it was taken out of the body, a thrombus was confirmed in the inhalation part. after that, impella was replaced with an intra-aortic balloon pump (iabp). | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the low pump flow was ingested biomaterial. the device is contraindicated for use with preexisting mural thrombus in the left ventricle per 0048-9007. this report is being filed as part of a retrospective review of historical records.
17227 | 1220648-2024-20627 | 20489529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489529&pc=OZD | 04/15/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 04/15/2022 | 10/19/2024 |  |  | 06/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100556 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a dampened placement signal. upon attempts to reposition, the pump fell out of the ventricle and could not be re-advanced. a new pump was used. the procedure outcome was noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the optical signal issue was most likely user error, due to the pump being pulled back too far into the aorta which caused bending to the device. this report is being filed as part of a retrospective review of historical records.
17228 | 1220648-2024-20628 | 20489530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489530&pc=OZD | 04/06/2022 | 05/28/2025 |  | 10/19/2024 |  |  | 10/16/2024 | 10/19/2024 |  |  | 06/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100975 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | MEDIKIT 16FR SHEATH. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during implantation, the impella was inserted from the left femoral artery using a sheath made by another company. access trouble (high calcification, stenosis) was observed at the time of initial insertion. the impella was inserted again after expansion with a percutaneous transluminal angioplasty (pta) balloon. while on support, lower limb ischemia was noted. there was a stenosis at the tip of the ecmo blood vessel, and blood removal was not sufficient. therefore, it is unknown whether the lower limb ischemia improved. the patient's condition was severe, and it was decided not to provide additional treatment. the patient died of multiple organ failure. there was no direct relationship between the death and the impella. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the delivery issues was determined to be patient condition as the patient had severe calcification and stenosis. the root cause of the leg ischemia was determined to be unrelated to impella as it happened at the ecmo access site. this report is being filed as part of a retrospective review of historical records.
17229 | 1220648-2024-20629 | 20489531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489531&pc=OZD | 12/08/2022 | 10/19/2024 |  | 10/19/2024 |  |  | 12/12/2022 | 10/19/2024 |  |  | 06/30/2024 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023233658 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, after the patient returned to the room, there was a suspicion of gastrointestinal bleeding, and the act had extended to 280 seconds at the time of insertion, so the purging liquid was used without heparin until it reached about 220 seconds. the physician was aware of the risks of managing without heparin and asked to add heparin as soon as possible, even if it is only 5u/ml. however, the patient expired during support due to multiple organ failure with no issue or causal relationship to the impella device alleged. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17230 | 1220648-2024-20630 | 20489533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489533&pc=OZD | 11/09/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 11/09/2021 | 10/19/2024 |  |  | 06/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022101015 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had hemolysis and expired due to sepsis that was unrelated to impella as the infection was uncontrollable. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17231 | 1220648-2024-20631 | 20489534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20489534&pc=OZD | 10/27/2021 | 10/19/2024 |  | 10/19/2024 |  |  | 10/17/2024 | 10/19/2024 |  |  | 06/30/2023 |  | 10/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022097305 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Appropriate Term / Code Not Available | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was lower than expected pump flow. the pump was removed and the patient remained in stable condition. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has not been returned for evaluation. the cause of the issue was not determined since product was not returned. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
17232 | 1220648-2024-20509 | 20480295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20480295&pc=OZD | 09/16/2024 | 11/29/2024 |  | 10/18/2024 |  |  | 10/01/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147694 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that a patient was on impella 5.5 support. while on support, impella stopped displayed on the automated impella controller (aic). the patient¿s outcome is unknown. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console-pump stop has been completed. the console log confirmed the reported failure mode given there were multiple instances of the impella stopped (mechanical/electrical issue) - alarm (imcgui: event set: #115 from 9 (0)) and impella stopped. retrograde flow. - alarm issued (imcgui: event set: #18 from 2 (0)) triggered during the clinical procedure. additionally, several occurrences of the pmd has detected a pump speed deviation - alarm issued (imcgui: event set: #138 from 50 (0)) alarm before the pump was unplugged from the console. real time (rt) logs showed multiple instances of motor current fluctuations, which eventually led to the pump speed deviation alarm, with the motor speed dropping to zero. see the attached image. the console was returned and upon visual inspection of the internal circuitry of the console did not reveal any anomaly. the console was run with a known good pump and purge system with no observable anomaly. the root cause for the pump stop controller error (red alarm) i.e. impella stopped and controller error (yellow alarm) i.e. pmd has detected a pump speed deviation was not related to any problems with the aic. no issues were found within the aic.
17233 | 1220648-2024-20510 | 20480315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20480315&pc=OZD | 10/01/2024 | 12/02/2024 |  | 10/18/2024 |  |  | 10/01/2024 | 10/18/2024 |  |  | 11/30/2025 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430561 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Device Embedded In Tissue or Plaque | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that upon insertion of the impella 5.5, resistance was initially met going through the graft. it was noted that the patient had multiple fresh blood clots, despite being anticoagulated. the decision was made to abort the procedure. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since patients had clots.
17234 | 1220648-2024-20511 | 20480461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20480461&pc=OZD | 09/21/2024 | 11/23/2024 |  | 10/18/2024 |  |  | 09/23/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507290 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and while on support, the impella had high purge pressure and critically low purge flows. tissue plasma activator was added to the purge and this resolved the purge flow issue. the patient was noted to be stable however outcome and explant date are unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for high purge pressure has been completed. no product was returned for evaluation. data logs were reviewed and lt log from time of alarm shows a gradual rise in purge pressure over approximately 19 hours was observed before high purge pressure alarms were triggered. no motor current spike was noted before purge rise. purge pressure remained high for about 34 hours before beginning to resolve. clinical details noted that tpa protocol was administered and was able to resolve the high purge pressure alarms. the root cause of the high purge pressure was most likely due to biomaterial buildup as it was resolved with tissue plasminogen activator (tpa). added-h6 impact code 4644.
17235 | 1220648-2024-20512 | 20480467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20480467&pc=OZD | 09/29/2024 | 11/25/2024 |  | 10/18/2024 |  |  | 09/29/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486976 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the physician was unable to reposition the impella 5.5. the patient was taken to the operation room for impella removal. the patient remained stable and survived the event. | [5 empty] | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
17236 | 1220648-2024-20513 | 20480480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20480480&pc=OZD | 10/03/2024 | 12/06/2024 |  | 10/18/2024 |  |  | 10/03/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025518513 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and there were concerns with the patient's hemolysis labs due to possible clot ingestion. the impella was explanted. the patient survived the event. | [5 empty] | The investigation into hemolysis and thrombus has been completed. the device was received for evaluation. visual inspection found biomaterial inside the pump housing & wrapped around the impeller. data logs confirmed the failure mode & clinical narrative. logs showed generally pump ran without issue during the case. multiple suction alarms were triggered near the end of support. pump then was disconnected from the console. product analysis: visual inspection found biomaterial inside the pump housing & wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of hemolysis was most likely patient condition. as the necessary information was not provided, the root cause of the thrombus was not determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20513. b2 missing. e4 should have been left blank. g1 reporting contact fax number missing. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death.". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
17237 | 1220648-2024-20514 | 20480487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20480487&pc=OZD | 10/03/2024 | 12/06/2024 |  | 10/18/2024 |  |  | 10/03/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546796 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump that failed delivery. the patient was trying to have a high risk percutaneous coronary intervention under the 5.5 support. when delivery failed the 5.5 was aborted and an impella cp was placed. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature as per the clinical description provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20514.
17238 | 1220648-2024-20515 | 20480493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20480493&pc=OZD | 02/28/2024 | 02/12/2025 |  | 10/18/2024 |  |  | 02/28/2024 | 10/18/2024 |  |  | 08/31/2025 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443592 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Renal Failure | Tachycardia | Hematuria | [8 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured ventricular tachycardia (vt) arrest with a probable relationship to the impella device used: participant transferred to cardiac care unit after catheterization lab procedure, including impella placement. participant was pale and severely diaphoretic. participant quickly decompensated with worsening hypotension and hemodynamics. participant decompensated further into pulseless vt x 3 with return of spontaneous circulation (rosc) established after each occurrence. participant started on amiodarone and impella pulled back under echocardiography guidance. infusions of epinephrine, norepinephrine, and vasopressin maximized. the patient recovered with no sequelae. additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a probable relationship to the impella device used: upon transfer to uwmc, participant found to have plasma free hemoglobin of 112 mg/dl ((b)(6) 2024 0247) and lactate dehydrogenase of 1250 u/l ((b)(6) 2024 0559). participant was noted to have black urine and pigment nephropathy from on-going hemolysis in the setting of impella cp. additional values: plasma free hemoglobin (pfh): 90 mg/dl at 0559; 31 mg/dl at 0839 lactate dehydrogenase (ldh): 2475 u/l at 0839. sae unresolved at time of participant expiration. it was reported that the patient/event has not recovered/not resolved. the patient expired after 20.78 hours of patient support. the relationship between the cause of death and the impella is unknown. medical safety reviewed the case and states there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issues has been completed. device history record review confirmed the pump passed all post sterile inspection checks. no pump was received for evaluation. data logs show suction alarms and positioning alarms. no motor current spikes outside p level changes. clinical has no mentions of pump positioning status. the cause of the reported issues was not determined since product was not returned and insufficient clinical information was received. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
17239 | 1220648-2024-20516 | 20480503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20480503&pc=OZD | 08/07/2024 | 10/22/2024 |  | 10/18/2024 |  |  | 08/07/2024 | 10/18/2024 |  |  | 11/30/2024 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024304201 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support due to anterior st elevated myocardial infarction (stemi). while on support, heparin was paused due to blood in orogastric (og) tube. no anticoagulation was given due to sub-arachnoid bleed. additionally, it was noted that the device was in the wrong position. care was withdrawn due to poor prognosis. patient expired. medical safety has reviewed the event and has stated that use of the device cannot conclusively be associated with the outcome. | [5 empty] | As noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b.2 other serious was entered incorrectly on the initial manufacturer device report number 1220648-2024-20516 and the correct selections were selected on this report. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-20516 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-20516 was submitted. g.3 new information was received that the date should be 20-aug-2024 and not 07-aug-2024 as reported on the initial manufacturer device report number 1220648-2024-20516. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-20516 and were added. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. as this clinical information was not provided, the true root cause of the stroke / cerebrovascular accident could not be determined. the root cause of the positioning issues cannot be determined due to insufficient information provided. due to limited clinical details, the root cause of the vascular damage - peripheral vessels could not be determined.
17240 | 1220648-2024-20517 | 20480517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20480517&pc=PYX | 09/23/2024 | 11/18/2024 |  | 10/18/2024 |  |  | 09/24/2024 | 10/18/2024 |  |  | 11/30/2025 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024428057 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed a pericardial effusion while on impella rp support. the team is unsure of origination or effusion. the patient was taken for emergent window due to tamponade. rp flex removed during this procedure, but not due to complication. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation vascular damage - great vessel has been completed. data logs are not relevant to the investigation. returned product was investigated and there was no damage to the cannula, pigtail, or inflow/outflow cages. there was a slight deformation to the tip of the repositioning unit, but it cannot be determined how or when it was damaged, as it was not reported. the root cause of the vascular damage could not be determined due to lack of clinical detail. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20517. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
17241 | 1220648-2024-20518 | 20480971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20480971&pc=OZD | 09/22/2024 | 11/21/2024 |  | 10/18/2024 |  |  | 09/22/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025507479 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and while on support, the patient had an ischemic stroke. heparin was started and anticoagulation was monitored. the patient was weaned successfully and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17242 | 1220648-2024-20519 | 20481011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20481011&pc=OZD | 06/07/2024 | 03/05/2025 |  | 10/18/2024 |  |  | 06/07/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482077 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the patient had an attempted impella 5.5 implant and the pump was able to fully enter the vessel but could not make the last turn into the ascending. the pump was removed from the body and was discarded after a kink was observed in the cannula close to the outlet. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because insertion was aborted due to patient anatomy. g1 reporting contact email revised on manufacturer device report 1220648-2024-20519 in accordance with updated procedures.
17243 | 1220648-2024-20520 | 20481125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20481125&pc=OZD | 08/14/2024 | 02/14/2025 |  | 10/18/2024 |  |  | 08/14/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025452631 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during the implant, the patient went into ventricular fibrillation (vf). additionally, the patient was noted to not have distal pulses in their feet while on support. the cp was explanted and replaced with an impella 5.5. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia and arrhythmia has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and arrhythmia could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email.
17244 | 1220648-2024-20521 | 20481127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20481127&pc=OZD | 09/22/2024 | 11/16/2024 |  | 10/18/2024 |  |  | 09/22/2024 | 10/18/2024 |  |  | 12/31/2024 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024301577 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump support ongoing with a patient stable enough to ambulated while on support. the axillary site had a bleed that prompted the team to stop heparin. the systemic heparin was stopped and then restarted. the patient was noted to be stable. however, outcome and explant date are unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was determined to be anticoagulation management because the hemostasis was achieved by stopping the heparin.
17245 | 1220648-2024-20522 | 20481350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20481350&pc=OZD | 09/13/2024 | 11/20/2024 |  | 10/18/2024 |  |  | 09/23/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442929 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support ongoing for a patient awaiting transplant. the pump was explanted after 20 days for the heart transplant. the patient had a seizure upon waking and had an ischemic stroke. the causal relationship of the pump to the stroke was not known to the team. the patient survived the procedure. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17246 | 1220648-2024-20523 | 20481721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20481721&pc=OZD | 09/18/2024 | 07/18/2025 |  | 10/18/2024 |  |  | 09/18/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509350 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 74-year-old male patient with cardiomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the patient arrived in the intensive care unit with an impella 5.5 inserted. hemoglobin dropped and the patient developed bleeding from the impella access site. manual compression, adjustment of partial thromboplastin, and activated clotting time protamine application were done. heparin was paused. six red blood cells units and four fresh frozen plasma units were applied. bleeding was resolved. additionally, the impella was constantly bumping against the ventricular wall, which led to constant ventricular fibrillation due to the stenosis. the impella could not really be positioned well. care was ultimately withdrawn. medical safety review of the event noted there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | Access site bleeding - major: after transfer to icu, bleeding was noted coming from pt's drains. the pt was noted to have a act of 210. the pump was not returned. the cause of the access site bleeding - major was most likely high patient act values. positioning issues: after 4 days on support, it was noted the pump is bumping against the ventricular wall but pump cannot be repositioned well due to stenosis. the pump was not returned. data logs were not recovered. the cause of the positioning issue was most likely patient condition related due to the noted stenosis in the vessels. renal failure: after 5 days on support, pt was noted to have renal failure. troubleshooting methods were not stated and it is unknown if the issue resolved. the cause of the renal failure was not determined due to insufficient clinical details. low pump flow: after device placement in the or, suction and reduced flow rates due to va ecmo and poor rv function / volume status. volume given and flow rates returned to 2.8-3.0 l/min which is nominal at p-5 (2.8 - 3.7l/min). the pump was not returned. data logs were not recovered. the cause of the low pump flow was most likely patient condition related due to lack of blood volume with pt also on va ecmo and poor rv function as stated in clinical. flow rates stated to be nominal after blood volume given. vf/vt/af/at: after 1 and 8 days on support, pt noted to have vt. after 7 days on support, pt noted to have cardiac arrhythmias. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities response to any antiarrhythmic treatments or interventions during the event any documented device-related issues or complications during impella support evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
17247 | 1220648-2024-20524 | 20481821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20481821&pc=OZD | 09/19/2024 | 11/21/2024 |  | 10/18/2024 |  |  | 09/19/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438136 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 71-year-old female patient was on impella 5.5 support and the patient had hemolysis. the patient was already on continuous renal replacement therapy. one unit of blood products were given to the patient and support was continued. the patient was on impella support for 329.88 hours before the patient expired. the relationship between the impella and cause of death is unknown. medical safety review of the event noted there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. no product was returned for evaluation. data logs were reviewed and no persistent positioning alarms or suction alarms were noted. no motor current spikes noted outside of p-level changes. placement signal is slightly variable throughout case. clinical details noted that hemolysis improved after blood volume was given. the root cause of the hemolysis was most likely patient condition related as hemolysis improved after blood volume was given.
17248 | 1220648-2024-20525 | 20481824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20481824&pc=OZD | 09/23/2024 | 11/22/2024 |  | 10/18/2024 |  |  | 09/23/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440369 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and there were low pump flows. the pump was repositioned and this did not resolve the pump flow issue. the doctor decided to replace the pump. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the low pump flow has been completed. the impella device nor the data logs were returned for evaluation. therefore, without sufficient clinical details, the root cause of the low pump flow could not be determined. added-h6 impact code 4628. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
17249 | 1220648-2024-20526 | 20481974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20481974&pc=PYX | 09/19/2024 | 11/21/2024 |  | 10/18/2024 |  |  | 09/19/2024 | 10/18/2024 |  |  | 06/30/2025 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024359000 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during an impella rp implant procedure, the physician could not successfully get the 0.027 wire through the balloon-tipped catheter, nor backload the impella on the 0.027 wire. the entire process was restarted again with a new impella rp kit and still was unsuccessful. the procedure was aborted and the patient expired in the procedural area. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome (patient's demise). patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the pump was not able to be placed due to cardiac anatomy. added- h6 impact code 4629, 4627 | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella rp smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
17250 | 1220648-2024-20527 | 20482204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20482204&pc=OZD | 09/24/2024 | 11/21/2024 |  | 10/18/2024 |  |  | 09/24/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024436433 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump being delivered to the 55-year-old female patient admitted in for surgery/cabg. the pump failed to deliver due to anatomic limitations and so the team instead placed a cp. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because unable to advance impella 5.5 due to small vessel size. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
17251 | 1220648-2024-20528 | 20482374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20482374&pc=OZD | 09/20/2024 | 11/16/2024 |  | 10/18/2024 |  |  | 09/20/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486958 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, heparin was on hold due to femoral hematoma post intra-aortic balloon removal. additionally, the developed had a gastrointestinal (gi) bleed. systemic anticoagulation was on hold. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the gastrointestinal bleed (gi bleed) has been completed. pump was not returned for investigation. clinical information provided is not sufficient and gives no information about the reason for the damage. therefore, the root cause of the gi bleed was not determined since product was not returned and insufficient information provided. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
17252 | 1220648-2024-20529 | 20482384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20482384&pc=OZD | 09/19/2024 | 11/13/2024 |  | 10/18/2024 |  |  | 09/19/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546185 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to ejection fraction of 30%, mitral regurgitation, and right ventricle regurge. while on support, some oozing was noted at the groin site. femstop was placed and the patient was given four units fresh frozen plasma. suspected sepsis was also noted. it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported sepsis and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of sepsis and access site bleed could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17253 | 1220648-2024-20530 | 20482545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20482545&pc=OZD | 09/24/2024 | 11/22/2024 |  | 10/18/2024 |  |  | 09/24/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025458128 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp support for a 46-year-old male patient admitted in for a high risk surgery of the mitral valve and coronary artery bypass graft (cabg). the pump came out of position and could not be redelivered. the pump was explanted after just under 7 hours of support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely attempt to recross the valve wirelessly.
17254 | 1220648-2024-20531 | 20482567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20482567&pc=PYX | 09/24/2024 | 10/18/2024 |  | 10/18/2024 |  |  | 09/24/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025529726 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a 52-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella rp device for mechanical circulatory support. while on support, the patient developed hemolysis. the pump was explanted. | [5 empty] | The investigation into the reported "hemolysis" has been completed since the original report was filed. the product was returned and evaluated. there was biomaterial found on the impeller. no other abnormalities were found. during data analysis, the log review showed a motor current spike on the first day of support along with suction alarms. on the 8th day of support, there was a small shift up in motor current and placement signal but no change in p-level. the pump (sn (b)(6)) passed all post-sterile inspection. the root cause of the hemolysis was most likely biomaterial.
17255 | 1220648-2024-20532 | 20482594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20482594&pc=OZD | 09/24/2024 | 12/10/2024 |  | 10/18/2024 |  |  | 09/24/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506599 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 53-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. while on support, hemolysis developed. origin is unclear. the impella position was deemed correct. impella was explanted. | [5 empty] | The investigation for the hemolysis has been completed. neither product nor data logs were returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
17256 | 1220648-2024-20533 | 20482734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20482734&pc=OZD | 09/24/2024 | 11/22/2024 |  | 10/18/2024 |  |  | 09/24/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519874 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and they had a large mural thrombus in the left ventricle. the pump was explanted due to the thrombus. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
17257 | 1220648-2024-20534 | 20482760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20482760&pc=OZD | 09/24/2024 | 12/06/2024 |  | 10/18/2024 |  |  | 09/24/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485528 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 placement, the impella 5.5 passed into the auxiliary artery without problems but was not able to pass into the truncus and aorta. the impella insertion was aborted due to the patient¿s anatomy. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant was aborted due to patient's anatomy as plaque had built up.
17258 | 1220648-2024-20535 | 20483040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20483040&pc=OZD | 09/24/2024 | 12/04/2024 |  | 10/18/2024 |  |  | 09/24/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025514118 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Disability | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. while the patient was being supported by impella cp for 10 days, the impella stopped. attempts were made to restart the impella cp, but it was unsuccessful. additionally, thrombus was found in the common iliac artery. the impella was removed and ecmo was placed. | [5 empty] | The investigation into pump stop and thrombus has been completed. as the necessary information was not provided, the root cause of the thrombus was not determined. the data logs show a spike in purge pressure and a motor current spike before pump stop and confirm failed restart attempts. the pump was unable to be run in the lab due to severed state. the root cause of the pump stop was most likely due to bearing wear. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20535 a5 ethnicity and race missing b2 disability or permanent damage inadvertently selected. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
17259 | 1220648-2024-20536 | 20483223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20483223&pc=OZD | 09/23/2024 | 12/09/2024 |  | 10/18/2024 |  |  | 09/24/2024 | 10/18/2024 |  |  | 04/30/2025 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347897 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support ongoing for a patient admitted in with acute myocardial infarction and cardiogenic shock. it was reported that suction/low flows during the support. volume was given and a day later suction alarms were noted to be resolved. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into low pump flow has been completed. the impella device was not returned for evaluation. the root cause of the low flow was most likely a patient condition as suction alarms were resolved under volume management. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20536. g1 reporting contact email. h6 code 4644 was reported incorrectly. new code was added to health effect - impact code. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
17260 | 1220648-2024-20537 | 20483269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20483269&pc=OZD | 07/11/2024 | 11/16/2024 |  | 10/18/2024 |  |  | 09/24/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447844 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute kidney injury with a "possible" relationship to the impella cp device used: ¿acute kidney injury requiring continuous renal replacement therapy for volume management.¿ it was reported that the patient/event has not recovered/not resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the renal failure has been completed. pump was not returned for investigation. clinical information provided says that the patient was diagnosed with tukotsubo syndrome and had multi organ failure. data logs provided have suction alarms but no motor current spikes or positioning alarms. as per clinical information provided, the root cause of the renal failure issue was most likely patient condition related.
17261 | 1220648-2024-20538 | 20483282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20483282&pc=OZD | 09/25/2024 | 12/10/2024 |  | 10/18/2024 |  |  | 09/25/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527019 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed hemolysis and bleeding at the impella access site. the dressing was changed and one unit of packed red blood cells was given. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the hemolysis and access site bleed has been completed. the impella device was not returned for evaluation. data logs showed suction alarms at starting that resolved. no motor current (mc) spikes outside of p-levels changes. no positioning issues mentioned. the cause of the access site bleeding - major most likely was patient movement related as it occurred after transferring the patient to icu. the cause of the hemolysis issue cannot be determined since the complaint device not returned for investigation and lack of clinical details. corrections to the initial mfr report: g1-mdr reporting contact email g1-mfr site name and address h6-clinical code added 1886. h6 impact code 4627. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17262 | 1220648-2024-20539 | 20483286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20483286&pc=OZD | 09/25/2024 | 11/17/2024 |  | 10/18/2024 |  |  | 09/25/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1329519 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller displayed red controller failure mode on startup. issue was found during training, not in clinical use therefore no patient involvement. | [5 empty] | The investigation for the console (aic) controller failure red alarm has been completed. console logs from the reported day of event are consistent with complaint and show persistent controller failure alarm due to epm software corruption (#1029), immediately after boot-up, on each consecutive boot-up. the console was returned for evaluation. upon evaluation, the issue was reproduced (controller failure alarm). after extracting epm firmware from microprocessor in impellatronic board and comparing to to known good code, we were able to confirm the corruption. after re-flashing the firmware, the issue was resolved and there were no alarms upon console boot-up. the corruption did not reoccur after several power-cycles of the console. the cause of the corruption was unable to be determined. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
17263 | 1220648-2024-20540 | 20483291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20483291&pc=PYX | 09/25/2024 | 11/19/2024 |  | 10/18/2024 |  |  | 09/25/2024 | 10/18/2024 |  |  | 10/31/2025 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024413721 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support and there was a clot ingested into the impella rp. flows decreased with high motor current. the pump stopped and the pump was removed and replaced with another pump. there is no reported harm to the patient. | [5 empty] | The investigation for the thrombus, low pump flow, and saturated flow has been completed. the pump was returned for investigation. the returned pump had some biomaterial and blue fibers ingested into the inflow cage. no other physical abnormalities were observed. the pump passed all optical and electrical testing. no damage was noted on the dp sensor. pump flow appeared to be saturated during testing in a bucket of water. the blue fibers could not be flushed out after testing. data logs show a motor current spike after case start while running at a steady p-level, with a drop in snr and an elevated placement signal, followed by low pump flow. the pump was stopped and restarted after a couple of minutes, and only the motor current trend was observed, with the placement signal continuing to trend. as per the clinical report, low pump flow was observed, followed by a motor current spike. additionally, some biomaterial was observed on the outflow cage. the data log and clinical reports suggest that some ingestion occurred during the case run, causing the low pump flow. the time of injected biomaterial or blue fibers entraining the impeller cannot be verified without sufficient clinical details. clinical report and data logs review does not support the pump stop issue hence it was removed. also, saturated failure mode was changed to low pump flow as investigated after clinical and data log review. the cause of the low pump flow issue was most likely biomaterial and blue fiber ingestion based on data log review and returned product investigation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17264 | 1220648-2024-20541 | 20483303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20483303&pc=OZD | 10/01/2024 | 12/11/2024 |  | 10/18/2024 |  |  | 10/01/2024 | 10/18/2024 |  |  | 09/30/2025 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407640 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a persistent pump in aorta alarm was triggered during impella cp support. the physician declined troubleshooting / repositioning attempts as explant was planned for the following day. it is unknown if the positioning of the pump was compromised. there was no patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was use related due to improper technique as the physician wanted to leave the pump in that position because he had plan to explant it next day. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
17265 | 1220648-2024-20542 | 20483314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20483314&pc=OZD | 10/04/2024 | 11/29/2024 |  | 10/18/2024 |  |  | 10/04/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544199 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that after removing the impella cp when doing the perclose closure, the perclose failed and arterial access was lost. femoral cutdown for arterial repair, pressure held, and site sutured close was done for hemostasis. three units of blood were given. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the reported issue was most likely use issue since the perclose failed. the following have been reported incorrectly or missing from manufacturer device report. e4 should have been left blank g1 reporting contact fax number missing.
17266 | 1220648-2024-20543 | 20483423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20483423&pc=OZD | 07/13/2024 | 11/21/2024 |  | 10/18/2024 |  |  | 09/24/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486499 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Respiratory Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute hypoxic respiratory failure with an alleged unknown (unrelated) relationship to the impella device used. the patient presented with shortness of breath and cough ongoing for 4 days. finding of hypoxemia for which they were placed on bilevel positive airway pressure (bipap). the patient recovered with no sequelae. | [5 empty] | The investigation has been completed and the root cause of the reported issue was unable to be determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
17267 | 1220648-2024-20544 | 20483528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20483528&pc=OZD | 09/25/2024 | 11/21/2024 |  | 10/18/2024 |  |  | 09/25/2024 | 10/18/2024 |  |  | 10/31/2025 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421510 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support needing mechanical circulatory support. the patient had runs of ventricular tachycardia and coded. the medical doctor made the choice to explant the pump and treat medically. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
17268 | 1220648-2024-20545 | 20485306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20485306&pc=OZD | 09/27/2024 | 11/25/2024 |  | 10/18/2024 |  |  | 09/27/2024 | 10/18/2024 |  |  | 07/31/2025 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442217 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. a clot was found encasing the impella in descending aorta. additionally, the impella could not maintain good flows. the decision was made to exchange for new cp. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". | The investigation for the low pump flow and thrombus has been completed. data logs confirmed the failure mode and clinical narrative. logs showed after pump started and run for about 30 minutes, low flow occurred that triggered auto suction response and suction alarms. this event remained to the rest of the case. product was not returned for review. as the patient had rv failure, the root cause of low flow was most likely patient condition.
17269 | 1220648-2024-20546 | 20485446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20485446&pc=OZD | 09/28/2024 | 11/20/2024 |  | 10/18/2024 |  |  | 09/30/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117656 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | Application Program Problem | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The investigation was completed with the findings from the log reports that the automated impella controller presented with a placement signal not reliable alarm, and two minutes later with a controller error alarm. both alarms kept self-resolving intermittently for the next four hours until the pump was transferred to a backup console where the reported issue was resolved. the patient is also on extracorporeal membrane oxygenation. | Us complainant reported the automated impella controller's (aic) impella connect platform was having connectivity issues. the screen was also going in and out. no known patient harm or injury. this issue occurred during patient use. the aic was replaced. |  |  |  |  | The investigation has been completed. console logs from the reported day of event are consistent with complaint and showed that 12 days after pump was started, the console presented with placement signal not reliable alarm and two minutes later with a controller error alarm. both alarms kept self-resolving intermittently for the next four hours, until the pump was transferred to a backup console, where the reported issue was resolved. long term log and real time log data showed signal patterns consistent with what appears to be two failure modes: loss of light from optical bench lamp resulting in very low signal-to noise ratio (snr) and oscillating contrast values and optical bench communication issues. although the occurrences of both failure modes were almost simultaneous, suggesting their correlation, we found no direct evidence to confirm it. console logs also showed several interruptions of data streaming and repeated attempts to reestablish handshake. rlm journals from the same time showed multiple reboots of the rlm, most of which coincided with interruption of data connection from automated impella controller (aic) to rlm (¿bad file descriptor¿ or ¿icm not ready for file receive¿ entries in aic logs). the console was extensively tested and peak signal-to-noise ratio (psnr) issue was reproduced. it has been observed that the lamp assembly on the optical bench would intermittently fail, resulting in loss of light in the optical sensor. while psnr condition was on-going, the optical bench maintained communication with the console, and no crc errors were present in its data. analysis of rlm journal indicated multiple unexpected reboots of the rlm, which were not observed during testing. rlm maintained stable connection to a test wifi network, with no interruptions of data streaming or video. majority of rlm reboots in the field were triggered by the rlm software detecting loss of received data (¿90 seconds of inactive aic connection - resetting the usb stack.¿). however, the cause for the preceding data interruption could not be identified. several components of the impella system (spanning aic and rlm) participate in usb data transfer and could have contributed to communication issue: 4-in-1 carrier, aic i/o panel and its harness, rlm backstrap cable and the rlm. the cause of the issue was a defective component. the cause of the aic issue was most likely an intermittent malfunction of optical bench. b.5 added new information from investigation results since initial manufacturer device report number 1220648-2024-20546 was submitted. h.6 due to investigation results, medical device problem code 1184 and component code 416 were incorrect to report on initial manufacturer device report number 1220648-2024-20546. new codes were added to health effect -impact code, medical device problem and component code sections based of the investigation results. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
17270 | 1220648-2024-20547 | 20485510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20485510&pc=OZD | 09/30/2024 | 11/18/2024 |  | 10/18/2024 |  |  | 09/30/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551931 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Obstruction/Occlusion | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an impella cp implant and during the implant, there were delivery issues. the initial access in the right femoral was obtained. two perclose were deployed, but were unable to take. the staff changed the access to the left femoral with four perclose and two eventually took and the impella was able to be placed. due to severe plaque of both left and right arteries, it would be very dangerous to place the impella with no closure device per the physician. after the procedure was completed, the perclose did not work and 14fr cook sheath placed in place of peel away. the patient was sent to the intensive care unit to await lower activated clotting time and be able to place fem stop as well as place manual pressure to obtain hemostasis in both groins. the patient survived the event. | [5 empty] | The investigation for the reported issue has been completed. the impella device was notreceived from the customer. the root cause of delivery issue is determined to be patient condition because of severe plaque of both left and right arteries and decided it would be very dangerous to place impella. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17271 | 1220648-2024-20548 | 20485815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20485815&pc=OZD | 11/10/2022 | 12/10/2024 |  | 10/18/2024 |  |  | 10/01/2024 | 10/18/2024 |  |  | 03/31/2024 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200230 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 66 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Information obtained from literature title: case report of right-to-left shunt caused by insertion of impella for ventricular septal perforation extracorporeal circulation technique 2023; vol.50. no4,466-467 since the exact date of hemolysis was unknown, the date of replacement of cardiac phase space analysis (cpsa) with 5.5 (date of cpsa removal) was entered. case 1 a woman in her 60s was rushed to her previous hospital because of chest pain and was then transferred to our hospital with cardiogenic shock due to acute myocardial infarction. after the transfer, echocardiography revealed a ventricular septal perforation (vsp) (2x 1cm) on the apical side. due to shock vitals, an impella cp was urgently inserted, and the patient was admitted to the intensive care unit (icu). due to severe hemolysis and a long time until surgery, the impella cp was replaced with an impella 5.5 on the fourth day of admission to the icu. the next day, the movement of impella's blood flowing out, causing a right-to-left shunt from the vsp, and a sudden drop in spo2 was observed. volume management was performed to prevent the left ventricle from collapsing, and although it was difficult, the patient was managed with impella alone until surgery. the patient survived the event | [5 empty] | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
17272 | 1220648-2024-20549 | 20487147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20487147&pc=OZD | 09/20/2024 | 11/21/2024 |  | 10/18/2024 |  |  | 09/20/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530640 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  |  | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | Ischemia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, there were signs of hemolysis/hematuria, and the leg was noted to be cool and doppler signals were absent. the 14 french sheath was removed and the repositioning sheath was inserted. a repeat doppler then had positive signals with the patient's foot increasing in temperature. additionally, the impella screen was stuck during the on the priming loading screen after attempting to de-air purge bag. this purge system was hooked to a pressure bag per the team's guidance as they were having prior purge pressure issues. troubleshooting efforts were unsuccessful. the patient ultimately expired as they were coding while the purge issues were ongoing. there was no resolution to purge pressure low. medical safety has reviewed the case and has stated that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the low pump flow, hemolysis, and limb ischemia has been completed. only the purge cassette was returned for evaluation. the field data logs were reviewed and the complaint purge cassette was not able to prime as the purge ticks were not seen and there was low purge pressure. the imc logs showed it stuck in priming. the returned purge cassette was connected to an aic to prime and upon priming it was stuck on the de-airing screen. the piston could not be heard driving. it was opened and attempted to prime again but the piston would not move. the gear was then manually turned by hand and resistance was felt at the sticking point. the root cause of the priming issue was due to a damaged driveshaft resulting in piston failure. data logs also found that there was a spike in motor current, placement signal and purge pressure followed by a decrease in purge flow. the root cause of the hemolysis is most likely due to biomaterial. the root cause of the limb ischemia could not be determined without additional information. d9 device return date. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. the purge cassette and data stick, however, have been returned. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17273 | 1220648-2024-20550 | 20487904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20487904&pc=OZD | 03/29/2022 | 05/23/2025 |  | 10/18/2024 |  |  | 10/14/2024 | 10/18/2024 |  |  | 10/31/2022 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021014949 |  | 0048-0034 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Stenosis | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after six days the patient was successfully weaned from the pump and the device explanted. it was later noted that the patient had aortic insufficiency, hemolysis and bleeding that required blood transfusions. additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm wheth er a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17274 | 1220648-2024-20553 | 20487959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20487959&pc=OZD | 08/27/2021 | 05/02/2025 |  | 10/18/2024 |  |  | 08/27/2021 | 10/18/2024 |  |  | 03/31/2023 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022068411 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support it was confirmed that there was a liquid leak from the purge system. the purge system was replaced, and the complication was managed. however, traces of water accumulating in the place where the purge cassette was inserted were confirmed. additional information was provided that noted care was withdrawn and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. neither the product nor the data logs were returned for investigation. after reviewing the clinical information, the cause of the purge leak was sidearm filter damage, most likely caused by ipa interaction.
17275 | 1220648-2024-20551 | 20487966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20487966&pc=OZD | 07/26/2021 | 10/18/2024 |  | 10/18/2024 |  |  | 10/16/2024 | 10/18/2024 |  |  | 05/31/2023 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022085728 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was a pump stop alarm, purge flows high and with low purge flows. the motor current climbed and eventually shut off. the team was unable to restart the pump after several attempts. decision was made to remove the pump and withdraw care. after the pump was removed, the patient extubated and died shortly after. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause could not be determined as product was not returned. this report is being filed as part of a retrospective review of historical records.
17276 | 1220648-2024-20554 | 20487979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20487979&pc=OZD | 01/21/2022 | 10/18/2024 |  | 10/18/2024 |  |  | 10/14/2024 | 10/18/2024 |  |  | 07/31/2023 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022122757 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, a broken sidearm was found on the device. a sidearm bypass was applied and the device was removed, revealing a catheter kink proximal to the pump handle. the patient ultimately expired, but no direct allegation was made relating the impella to this event. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The device was returned and evaluated, and the investigation for the reported purge leak has been completed. a catheter kink was found proximal to the pump handle. after reviewing the clinical information, it was determined that the cause of the purge leak was most likely sidearm bond damage due to excessive force on the sidearm during use. as noted in the impella ifu: purge leak: pump impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." this report is being filed as part of a retrospective review of historical records.
17277 | 1220648-2024-20555 | 20487983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20487983&pc=OZD | 05/16/2022 | 10/18/2024 |  | 10/18/2024 |  |  | 10/15/2024 | 10/18/2024 |  |  | 02/29/2024 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198859 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support it was reported that the device became kinked during use with low pump flows. the cannula kinked on itself inside the patient and there were not optimal flows. the pump was removed, and a new pump placed. additional information was provided that noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported low pump flow and cannula kink has been completed. the impella device was not received from the customer. however, details were sufficient to determine the root cause to be a kinked cannula. this report is being filed as part of a retrospective review of historical records.
17278 | 1220648-2024-20552 | 20487984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20487984&pc=OZD | 09/18/2024 | 04/28/2025 |  | 10/18/2024 |  |  | 09/24/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025537470 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 39-year-old female patient with caridomyopathy and other systemic comorbidities had an impella 5.5 pump placed for care escalation from the impella cp and ecmo. the impella 5.5 was replaced due to placement signal issues after just 2 days of support. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The returned pump was inspected and there were no physical abnormalities. there was no biomaterial on the impeller. the 5s parameters were within specification and the pump passed the light continuity test. the pump was run in the lab and neither the low pump flows nor the saturated flows reproduced. optical signal issue: the data logs from the case show intermittent 'impella position in aorta' alarm after observed ingestion events. the ingestion events show motor current spikes with speed deviations and loss of motor current pulsatility. there were multiple motor current spikes possibly as a result of repeated contact with the impeller. the pump flows also became saturated. it is likely that the false 'impella position in aorta' alarm was as a result of the ingestion events. the other alarm occurrence was likely when the pump was being weaned for explant at p2. the root cause of the optical signal issue was unable to be definitively determined as false position alarm could be as a result of biomaterial ingestion, but limited clinical information was provided related to the failure. saturated flow: the data logs show ingestion pattern with motor current spikes with speed deviations and loss of motor current pulsatility. there were multiple motor current spikes possibly as a result of repeated contact with the impeller. the pump flows became saturated as a result of this. the returned pump was run in the lab and saturated flows did not reproduce. the root cause of the saturated flow was determined to be due to biomaterial as evidenced by the ingestion pattern observed in the data logs. low pump flow: the data logs confirm #14 suction alarms which triggered for about 30 seconds while the pump was at p4. there were no other suction alarms throughout the case. there was low flow at p4 flowing below the expected range. the returned pump was run in the lab and low pump flows did not reproduce. the root cause of the low pump flow issue was not determined as limited clinical information related to failure was provided and no data log abnormalities were observed around occurrence. b3 date of event was erroneously entered on manufacturer device report 1220648-2024-20552 g1 reporting contact email revised on manufacturer device report 1220648-2024-20552 in accordance with updated procedures.
17279 | 1220648-2024-20556 | 20487992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20487992&pc=OZD | 05/18/2022 | 10/18/2024 |  | 10/18/2024 |  |  | 10/14/2024 | 10/18/2024 |  |  | 02/29/2024 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195666 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was a leak due to a purge line break at the filter. the team had initially tried to remedy the issue by replacing the purge cassette, however, this did not work. a new impella was opened and placed successfully. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was returned for evaluation. upon visual inspection, the side arm was broken off in-between the filter and the reservoir. in conclusion, the root cause is due to a sidearm bond break as a result of excessive force. no other damage was found to the sidearm. this report is being filed as part of a retrospective review of historical records.
17280 | 1220648-2024-20557 | 20487994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20487994&pc=OZD | 02/20/2023 | 10/13/2025 |  | 10/18/2024 |  |  | 10/15/2024 | 10/18/2024 |  |  | 03/31/2024 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213029 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Defective Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella cp device for mechanical circulatory support. the complainant reported that the peel away sheath cracked in the arteriotomy prior to the impella cp being advanced across the valve. delivery was still managed successfully and the peel away sheath was removed without complication. following the delivery, oozing was noted at the access site prompting the utilization of femostop and manual pressure to achieve hemostasis. the impella was re-inserted without issue, however, the patient experienced a ventricular tachycardia and was defibrillated. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
17281 | 1220648-2024-20558 | 20487999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20487999&pc=OZD | 01/20/2022 | 10/18/2024 |  | 10/18/2024 |  |  | 01/25/2022 | 10/18/2024 |  |  | 07/31/2023 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022123953 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced significant bleeding despite multiple sutures. it was suspected that the device had become dislodged or malpositioned, though this was not confirmed at the time. the patient expired while on impella support, but there was no allegation that the death was related to the impella. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17282 | 1220648-2024-20559 | 20488002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20488002&pc=OZD | 03/22/2024 | 10/18/2024 |  | 10/18/2024 |  |  | 07/22/2024 | 10/18/2024 |  |  | 11/30/2023 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151341 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced limb ischemia during support. actions taken to resolve the issue are unknown. additional information noted the patient expired and it was noted that the patient's death was due to multiple organ failure. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported issue of limb ischemia has been completed. the impella device was not received from the customer. the root cause of the limb ischemia cannot be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17283 | 1220648-2024-20560 | 20488011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20488011&pc=OZD | 04/04/2023 | 10/18/2024 |  | 10/18/2024 |  |  | 04/04/2023 | 10/18/2024 |  |  | 12/31/2024 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303761 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced perforation/left ventricle rupture. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
17284 | 1220648-2024-20561 | 20488013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20488013&pc=OZD | 06/30/2023 | 05/19/2025 |  | 10/18/2024 |  |  | 06/30/2023 | 10/18/2024 |  |  | 02/28/2025 |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321948 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. pump stop occurred and prompted pump removal to allow the patient to expire per the do not resuscitate plan. the pump had been used well beyond indication. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. the cpsa pump had been used over a month which is beyond the design life of this pump of 8 days. the cause of the pump stop cannot be determined since the complaint device was not returned for investigation and no data logs returned to review. the patient death was unrelated to the impella pump, resulting in death due to systemic circulatory failure. not affiliated with impella. this report is being filed as part of a retrospective review of historical records.
17285 | 1220648-2024-20562 | 20488838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20488838&pc=OZD | 09/20/2024 | 10/24/2024 |  | 10/18/2024 |  |  | 09/20/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515997 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 57-year-old male patient was on impella cp support and after the implant, the patient did not have flows past the impella repositioning sheath. the doctor placed a reperfusion sheath to provide flow to the distal limb. flow was restored with the reperfusion sheath. additionally, the patient had access site bleeding while on support. purse string sutures were applied and the doctor ordered blood transfusions to increase volume. care was ultimately withdrawn and the patient expired. medical safety review of the event noted there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | H.6 code was inadvertently omitted from the initial manufacturer device report number 1220648-2024-20562 and was added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17286 | 1220648-2024-20563 | 20488842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20488842&pc=OZD | 09/20/2024 | 11/21/2024 |  | 10/18/2024 |  |  | 09/20/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484423 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, a 75-year-old male patient had hemolysis. the patient experienced mean arterial pressure of 110-115 mmhg throughout support which was ultimately brought down to 86 mmhg with weaning of inotropic support. the patient was also given two units of blood products and the hemolysis improved. the patient was placed on comfort care and ultimately expired. medical safety review noted that there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. the impella was returned for evaluation and visually inspected. no biomaterial observed. no broken blade or sharp edge found. cannula shape and pigtail deformed. data log reviewed: many suction alarms occurred. no motor current (mc) spike observed. no positioning issue alarms occurred and the logs show declining placement signal (ps) trend at end with suction alarms. hemolysis began to resolve with decreased afterload. the cause of the hemolysis issue most likely patient condition related.
17287 | 1220648-2024-20564 | 20488862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20488862&pc=OZD | 09/21/2024 | 11/16/2024 |  | 10/18/2024 |  |  | 09/21/2024 | 10/18/2024 |  |  |  |  | 10/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521537 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and there were purge pressure high alarms. no kinks were found and the purge cassette was replaced with no resolution. tissue plasminogen activator was implemented in the purge. staff monitored. the patient was on impella support for 34.61 hours before the patient expired. medical safety review noted that there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. the root cause of the high purge pressure cannot be determined as no logs were returned for evaluation and it was not clear if tissue plasminogen activator (tpa) resolved the high purge pressure alarms. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
17288 | 1220648-2024-20472 | 20468405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20468405&pc=OZD | 06/27/2021 | 10/17/2024 |  | 10/17/2024 |  |  | 07/07/2021 | 10/17/2024 |  |  | 01/31/2023 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022045951 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | Renal Failure | Sepsis | Hemorrhagic Stroke | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed sepsis, hemolysis and renal failure. it was also reported that there had been a pump stop. the pump was stopped in an emergency. an attempt was made to restart the device. hemodynamics could be maintained to some extent, and it was decided to explant the device. additionally, the patient suffered cerebral bleeding and expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop, renal failure, infection/sepsis, hemolysis, stroke/cva issues has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was confirmed to have bearing wear. the cause of the pump stop was bearing wear. the pump stop occurred beyond the 8 day impella cp indication for use per design trace matrix 0046-9910. the impella device was confirmed to have been sterilized under validated conditions and there was no evidence to indicate that the sepsis was related to impella. there was no evidence or clinical imaging to indicate that the stroke was caused by impella. the cause of the hemolysis was not determined. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17289 | 1220648-2024-20473 | 20468412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20468412&pc=OZD | 06/30/2021 | 10/17/2024 |  | 10/17/2024 |  |  | 07/20/2021 | 10/17/2024 |  |  | 12/31/2022 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022039835 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombosis/Thrombus | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed sepsis. additionally, upon removal of the impella a thrombus was adhered to the aortic valve, no events were related to the thrombus. the patient expired 2 days after the device was removed. the cause of death was reported to be due to the sepsis. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported thrombus and infection/sepsis issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smartassist instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17290 | 1220648-2024-20474 | 20468441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20468441&pc=OZD | 08/11/2021 | 01/03/2025 |  | 10/17/2024 |  |  | 08/17/2021 | 10/17/2024 |  |  | 05/31/2023 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089427 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was additionally reported that the patient experienced limb ischemia during impella cp support. upon review of the information for the case, with additional information provided to abiomed that the impella may have been a factor in the patient's pericardial bleed, the use of the device cannot be excluded from the patient's outcome. | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, pericardial fluid accumulation was noted from the beginning, and drainage was performed. the bleeding was controlled, and the patient was taken off ecmo to improve hemodynamics with impella. immediately after being taken off ecmo, massive bleeding was observed again from the pericardium, and it was determined that there was bleeding in the anterior wall of the right atrium due to chest compression. surgical repair was considered, but the family did not request any more aggressive treatment, and the patient expired after removal. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | Additional information for the initial mfr 1220648-2024-20474 was provided in sections b1, b2, b5, h1, and h6 to reflect new information provided to abiomed regarding limb ischemia and the possible relationship between the patient's outcome and the use of the impella device. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17291 | 1220648-2024-20475 | 20468472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20468472&pc=OZD | 09/09/2021 | 10/17/2024 |  | 10/17/2024 |  |  | 09/09/2021 | 10/17/2024 |  |  | 06/30/2023 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108284 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were multiple cp pump stops and ongoing ventricular tachycardia (vt). shortly after the patient was made do not resuscitate (dnr) and after other care had been withdrawn, the pump failed one last time. several unsuccessful attempts to restart the pump were completed. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk cpi instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17292 | 1220648-2024-20476 | 20470329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20470329&pc=PYX | 09/29/2024 | 12/03/2024 |  | 10/17/2024 |  |  | 09/29/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025453773 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp flex support and shortly after the implant, the patient was being cleaned and prepped for transfer and the pump became malpositioned. the doctor had to stop the rp flex, pull it back through the sheath until the inlet cage appeared, then rewire through the cage then remove the flex completely so that they could replace the pump. the pump was replaced. the patient survived the event. | [5 empty] | The investigation for the reported positioning issue has been completed. product was not returned for review. the pump had not been sutured in place yet at the time of the dislodgement. the root cause of positioning issue was most likely use-related. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
17293 | 1220648-2024-20477 | 20470559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20470559&pc=OZD | 09/29/2024 | 11/29/2024 |  | 10/17/2024 |  |  | 09/29/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025487179 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and the patient was having access site bleeding. heparin was withheld and support continued. additionally, while on support the patient was having ventricular tachycardia (vt). the impella was pulled back 1 cm. support continued, and the patient survived the event. | [5 empty] | The investigation for the reported access site bleed and ventricular tachycardia (vt) has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient act values, heparin was stopped to achieve hemostasis as per the clinical description. the root cause of the vt could not be determined without additional information. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17294 | 1220648-2024-20478 | 20470597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20470597&pc=OZD | 09/30/2024 | 11/17/2024 |  | 10/17/2024 |  |  | 09/30/2024 | 10/17/2024 |  |  | 09/30/2025 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403936 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was inserted, and, after percutaneous coronary intervention was performed, the patient¿s groin started oozing. the physician elected to suture around the insertion site to see if it would help with hemostasis. in the intensive care unit, the groin continued oozing despite angle matching, forward tension sutures, tightening perclose, no heparin, and knee immobilizer. pressure was held on artery above impella and on site with no improvement. in total 4 units of red blood cells were given. the impella was removed with a surgical cutdown. the patient remained stable and survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since the bleeding was managed by forward tension sutures, angle matching, tightening perclose, knee immobilizer, act in range.
17295 | 1220648-2024-20479 | 20470943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20470943&pc=OZD | 09/13/2024 | 12/04/2024 |  | 10/17/2024 |  |  | 09/23/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531231 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Cardiogenic Shock | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "definite" relationship to the impella cp device used. the patient recovered with no sequelae. | Additional information received from via abiomed¿s long-term outcome and quality indicator (loqi) that while on support, the patient endured cardiogenic shock after impella dislodged with a definite relationship to the impella device used. impella cp initially at p9 in catheriterization lab with adequate flows. upon being reversed from paralytics, patient coughing / sitting up in bed and noted uncoupling of waveforms, with repeated suction events at p6 and p5. the impella was repositioned and the patient recovered with no sequelae |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | B3 updated to reflect date of event of the new information received. b5 updated with additional information received. h6 health effect - clinical code and health effect - impact code added to reflect new information. e4 was missing on manufacturer device report 1220648-2024-20479. | The investigation for the hemolysis and cardiogenic shock has been completed. the pump was not returned for investigation. according to the clinical data, hemolysis during impella support and improved the following day. uncoupling placement waveforms was noted while the patient was coughing and sitting up in bed, accompanied by a suction alarm. pump was dislodged, and the impella was repositioned after cardiogenic shock. data logs show several suction alarms followed by an impella position in aorta alarm late. the pump was unable to achieve flow at higher p-levels. the cause of the hemolysis issue was improper positioning of the device based on data log review and given clinical details. the following have been reported incorrectly on manufacturer device report 1220648-2024-20479.
17296 | 1220648-2024-20480 | 20471272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20471272&pc=OZD | 09/14/2024 | 12/11/2024 |  | 10/17/2024 |  |  | 09/24/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547858 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump being placed to support a patient admitted in acute coronary syndromes. it was reported that resistance was felt in the brachiocephalic artery and the impella could not be advanced any further, the shaft was found to be kinked when the impella was removed. insertion was abandoned and the patient was managed solely with the extracorporeal membrane oxygenation that had already been installed. no patient harm was reported. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of the calcified vessel condition. the shaft was kinked during insertion and insertion was aborted. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-20480. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
17297 | 1220648-2024-20481 | 20471728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20471728&pc=OZD | 09/18/2024 | 09/17/2025 |  | 10/17/2024 |  |  | 09/20/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025517777 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | Additional information received on 09/25/2024 via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding the patient that endured cva with a "unknown (undetermined)" relationship to the impella 5.5 device used: "patient not moving the left side of his body, ct imaging completed showing large right middle cerebral artery infarct with second marginal coronary artery branch occlusion. worsening midline shift of 8mm. decompressive craniectomy performed.¿. the patient recovered with sequelae. | While on support, placement signal (ps) pressure readings elevated in 200s both aorta and left ventricle. it was reported that the issue was resolved. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of stroke/cva and thrombus has been completed. as the necessary information was not provided, the root cause of the thrombus and stroke/va was not determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-20481 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-20481 in accordance with updated procedures. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. impella 5.5 was returned for evaluation. optical signal issue: as the failure mode of "placement signal issue" is for investigating failures of the differential pressure sensor and the impella 5.5 does not have a dp sensor, the reported failure mode of "placement signal issue" was investigated under the failure mode of "optical signal issue". no visual abnormalities were noted on returned pump. no biomaterial noted and pump passed laser continuity test. pump's optical parameters did not show any hard failures of the optical sensor. data logs were reviewed and no hard failures of the optical sensor were noted. snr throughout case was observed to be continuously low, with contrast also periodically low. the data logs from the previous pump run and the pump run after the impella 5.5 sn (b)(6) on ic1675 were reviewed and low snr was observed on both pumps. clinical details noted that the ao and lv signal were elevated in the 200s. volume was given in attempt to resolve issue and it was not noted if placement signal issue resolved. upon review of the data logs and evaluation of the returned pump, it is apparent that the console is the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. no problem is found with the pump during the case. thrombus: the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. stroke/cva: strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation. timing of the stroke in relation to impella support (during or post-implant). detailed description of the symptoms experienced by the patient. neurological examination findings and any focal deficits observed. diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic). laboratory test results, including coagulation profiles and cardiac markers. any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. response to any previous anticoagulation or antiplatelet therapies. evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
17298 | 1220648-2024-20482 | 20471742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20471742&pc=OZD | 09/23/2024 | 11/17/2024 |  | 10/17/2024 |  |  | 09/23/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474085 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  |  | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support to assist in coming off cardiopulmonary bypass. while on support, there was decreasing flows on the impella so blood products were given. additionally, the doctor wanted to perform hemi sternotomy for arterial repair of innominate artery after removal of pump. the doctor opened the pocket and dissected tissue down to graft attachment on innominate artery. pump removed, but a tear on innominate was noted. the doctor placed their hand in the pocket and pinched the artery. decision was made not to canulate for venoarterial extracorporeal membrane oxygenation or attempt new impella support based on prognosis with advanced therapies. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation for the low pump flow and vascular damage has been completed. product was not returned. clinical details suggest that leading up to explant, the physician's plan was to perform a hemi sternotomy to repair an artery. when explanting the pump, they did find vascular damage to the innominate artery. there is no further detail provided as to why there was suspicion of vascular damage or how it could have occurred. product was not returned for investigation. the root cause of the vascular damage could not be determined due to lack of clinical details. data logs were investigated, and the low pump flows were confirmed. particularly early on in the case, pump flows are below specification for the corresponding p-level. additionally, there were 21 suction alarms that triggered in the first 6 hours of support. for the remainder of the case, pump flows continued to be low or were variable/unstable. there were also "impella position in aorta" alarms that began to trigger at 16:11. however, this is after the report of low pump flows, and these alarms all resolved quickly. product was not returned for investigation. based on clinical details and data log review, the root cause of the low pump flows was determined to most likely be patient condition.
17299 | 1220648-2024-20483 | 20471752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20471752&pc=OZD | 09/24/2024 | 12/09/2024 |  | 10/17/2024 |  |  | 09/24/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025484670 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported that impella 5.5 insertion was aborted due to vessel anatomy complications. no reported patient harm. | [5 empty] | The investigation into delivery issues has been completed. the root cause of the delivery issue was most likely patient condition related, impella insertion aborted due to vessel anatomy complications as per the clinical description provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20483 e4 should have been left blank g1 reporting contact email g2 report source foreign missing h6 code 4627 was reported incorrectly. new code was added to health effect - impact code | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
17300 | 1220648-2024-20485 | 20471770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20471770&pc=PYX | 09/24/2024 | 11/11/2024 |  | 10/17/2024 |  |  | 09/24/2024 | 10/17/2024 |  |  | 03/31/2025 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024394966 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a patient suffering from acute myocardial infarction and cardiogenic shock placed on impella cp and impella rp flex support. the rp came out of position after the cp was repo sheath advanced. the rp was explanted due to the malpositioning. the patient expired with care withdrawal. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. based on clinical description, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
17301 | 1220648-2024-20486 | 20471782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20471782&pc=OZD | 09/20/2024 | 11/20/2024 |  | 10/17/2024 |  |  | 09/20/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025543837 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Failure to Sense | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, trans-esophogeal echocardiography showed good placement across the valve. however, the console was alerting impella in aorta alarms and the left ventricle waveforms were very high. the decision was made to explant the impella 5.5. the patient survived the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation has been completed. the returned pump was inspected and biomaterial was found on the impeller. the 5s parameters were within specification and the pump passed the light continuity test. the real time logs from the case showed 'impella position in aorta' alarm before placement signal monitoring was disabled. according to the clinical details, the position was verified via transesophageal echocardiogram. the root cause of the placement signal issue was most likely patient condition related. h.6 medical device problem codes 1158 and 1559 were reported in error on the initial manufacturer device report 1220648-2024-20486 and a new code was added. according to the investigation results, the failure mode 'positioning failure' was removed because the issue reported was related to placement signal monitoring of the optical sensor and not physical pump positioning issues. the failure mode 'placement signal issue' was changed to optical signal issue because the issue is related to optical sensor.
17302 | 1220648-2024-20488 | 20471960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20471960&pc=OZD | 10/02/2024 | 11/29/2024 |  | 10/17/2024 |  |  | 10/02/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548079 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp was placed, percutaneous coronary intervention was performed. when performing the sheath exchange the peel away sheath was pulled back but there was difficulty breaking away and impella was pulled completely into the aorta. after trying to get impella back across valve without success, a decision was made to explant cp and place a new one. the patient remained stable and survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issues was most likely a use related issue. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20488. e4 should have been left blank. g1 reporting contact fax number missing.
17303 | 1220648-2024-20489 | 20471967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20471967&pc=OZD | 10/03/2024 | 12/06/2024 |  | 10/17/2024 |  |  | 10/03/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531728 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that hematoma noted, and manual pressure held. once patient was in the intensive care unit, hematoma continued. the physician decided to remove the impella cp at bedside. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20489.
17304 | 1220648-2024-20484 | 20471993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20471993&pc=OZD | 09/24/2024 | 11/17/2024 |  | 10/17/2024 |  |  | 09/24/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516610 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient was scheduled for impella 5.5 implant for mechanical circulatory support. during attempted delivery, it was noted that the pump was unable to advance into position and the implant was aborted. support continued for the patient via an already in place impella cp device. a few days following explant, it was reported that bleeding had developed at the explant site of the impella 5.5 pump. systemic heparin was stopped to manage the condition. | [5 empty] | The investigation for the access site bleed and the delivery issue has been completed. the pump was not returned for evaluation. however, clinical details was sufficient to determine root cause. the root cause of delivery issue is determined to be patient condition because despite the multiple attempts with two pumps, both the pumps are unable to cross the subclavian artery and first pump was kinked during the attempt. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
17305 | 1220648-2024-20487 | 20471998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20471998&pc=OZD | 06/11/2024 | 11/16/2024 |  | 10/17/2024 |  |  | 09/25/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025473970 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump support ongoing in (b)(6) 2024 for a patient admitted in cardiogenic shock / cardiomyopathy. the support was successful for over 4 days and then explanted for a left ventricular assist device. later in (b)(6) 2024 abiomed¿s long-term outcome and quality indicator (loqi) impella registry database reported upon thrombosis and a stroke during the support. the allegation was an unknown relationship to the use of the impella. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported stroke and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of stroke and thrombus could not be determined as the device was not returned nor sufficient clinical details. added h5 yes, h8 initial use of device.
17306 | 1220648-2024-20490 | 20472638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20472638&pc=OZD | 07/26/2024 | 10/17/2024 |  | 10/17/2024 |  |  | 09/20/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1211522 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that after the automated impella controller (aic) was removed from service, an error message of the usb flash drive was not detected or supported. troubleshooting was done unsuccessfully. there was no reported patient harm. investigation results determined this issue was reportable. | [5 empty] | The investigation into the reported event has been completed. the console logs from the reported day of event are consistent with the complaint and showed four failed attempts to download console logs on a usb stick, each time resulting in message ¿usb flash drive is not detected or supported¿. the console logs did not show interruptions of data streaming, indicating that any connectivity issues are unrelated to data transfer. journals for the reported day of event were unavailable but journals three days later showed frequent disconnections as reconnections to site network, and inability to reach ¿endpoints¿ (external servers necessary for impella connect operation). the console was tested and the issue of log downloading was not reproduced - with a known-good usb stick. the console was able to upload logs onto it. one of the usb sticks in the field were received for analysis, and was confirmed to have been incompatible with consoles, as it did not work with a known-good engineering console. testing showed that the unit was connecting to abiomed test wifi, connection was stable, and data and video were streamed. the unit operated within specification.
17307 | 1220648-2024-20492 | 20472751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20472751&pc=OZD | 09/30/2024 | 11/17/2024 |  | 10/17/2024 |  |  | 09/30/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025543806 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and they had bleeding at the insertion site. the patient lost a lot of blood. coagulation products were given to the patient and heparin was withheld. the patient remains on impella support. | [5 empty] | The investigation access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely anticoagulation management since the bleed was managed by giving coagulation products to the patient and heparin was withheld. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17308 | 1220648-2024-20491 | 20472755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20472755&pc=OZD | 09/29/2024 | 11/27/2024 |  | 10/17/2024 |  |  | 09/29/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516613 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, post-operation bleeding was noted. heparin was withheld, femoral artery required repair. additionally, impella 5.5 experienced an impella undetectable error. troubleshooting was done and the decision was made to explant pump. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the pump not detected issue has been completed. the pump was returned and evaluated. the "impella defective" alarm reproduced when connected to a lab console and the pump could not be detected. there was blood found in the red handle and a hole in the catheter between the 25 and 30 cm mark. the data logs did not have any abnormalities before the aic transfer. the root cause of the pump not being detected was blood ingress due to a hole in the catheter resulting from improper use.
17309 | 1220648-2024-20495 | 20473061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20473061&pc=OZD | 09/20/2024 | 12/06/2024 |  | 10/17/2024 |  |  | 09/20/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556209 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 66-year-old female patient with caridology and other systemic comorbidities was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient developed bleeding from the impella access site which caused a reduction in hemoglobin. two units of red blood cells were administered. the patient was successfully weaned at the end of unit support. the patient survived the explant. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17310 | 1220648-2024-20493 | 20473145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20473145&pc=OZD | 09/20/2024 | 10/24/2024 |  | 10/17/2024 |  |  | 09/20/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1173720 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a battery failure. there was no patient harm reported. | [5 empty] | The investigation for the console (aic) battery failure (red alarm) has been completed. data logs were reviewed which found there were no battery failures observed. console was returned for evaluation. the reported battery failure could not be reproduced. results of running a batt test utility on this console showed that the health of the batteries were normal. after further inspection of the aic, it was discovered that there was damage on the inside of the ac power cord connector and the aic could not detect ac power when connected to a wall outlet. it was evident that this issue with ac power detection and inability to charge is what the clinical staff meant to report. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
17311 | 1220648-2024-20494 | 20473522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20473522&pc=OZD | 09/30/2024 | 11/25/2024 |  | 10/17/2024 |  |  | 09/30/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486424 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had placed a 5.5 pump to support a patient admitted in cardiogenic shock and suffering from cardiomyopathy. the pump was placed and patient was sent to the intensive care unit. when placed, the pump position had to be adjusted as the pump was deep and ectopy was occurring. once pump repositioned, the ectopy resolved. the pump remains on for support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported arrhythmia has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of arrhythmia could not be determined as the device was not returned nor sufficient clinical details.
17312 | 1220648-2024-20496 | 20473789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20473789&pc=OZD | 09/20/2024 | 12/09/2024 |  | 10/17/2024 |  |  | 09/20/2024 | 10/17/2024 |  |  |  |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025538910 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 60-year-old female patient had severe bleeding in the right groin following impella cp insertion. the bleeding was stopped by a sharper angle and a pressure bandage but a large hematoma still formed. vascular surgeon will be evaluating the hematoma. the patient required two transfusions. the pump was successfully explanted after four days of support. the patient survived the explant. | [5 empty] | The investigation for the reported access site bleed has been completed. no issues were found in the returned product. the root cause of the access site bleeding cannot be determined since insufficient clinical details were provided. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17313 | 1220648-2024-20497 | 20474254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20474254&pc=OZD | 09/20/2024 | 11/21/2024 |  | 10/17/2024 |  |  | 09/20/2024 | 10/17/2024 |  |  | 11/30/2025 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430106 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Detachment of Device or Device Component | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and while on support, the patient had an open chest. four washouts were performed, and multiple blood products were given to the patient. heparin was in the purge. support was continued. additionally, upon removal of the pump, the pump was pulled out of the left ventricle. the nurse met significant resistance when pulling the pump out of the graft. the nurse was advised to keep pulling and ended up tearing the impella catheter off the cannula. blood loss was noted out of graft, and it was eventually clamped. the cannula was then removed. the surgeon noticed that the sutures had not been cut on the graft which is what caused the resistance. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the product damage, vascular damage, and removal issues has been completed. the product was not returned. based on the clinical details, the surgeon sewed the graft directly to the aorta, and had the nurse pull the pump even though there was significant resistance met. the significant resistance was caused by the sutures not being removed for the graft. the root cause of the removal issue is due to use. due to the sutures not being removed, the nurse tore the graft causing significant bleeding. the root cause of the vascular damage peripheral vessel is due to use. even though there was significant resistance met, the physician instructed the nurse to keep pulling. the pump was pulled out of the lv but the cannula became detached in the graft. the root cause of the product damage (detached inflow/outflow) peripheral vessel is due to use. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
17314 | 1220648-2024-20501 | 20477648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20477648&pc=OZD | 09/14/2022 | 10/29/2024 |  | 10/17/2024 |  |  | 09/14/2022 | 10/17/2024 |  |  | 03/31/2024 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202203 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced uncontrolled gastrointestinal bleeding and multiple organ failure. additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Medical device report # 1220648-2024-20501 was identified as a duplicate report of medical device report # 1220648-2024-20754. follow up reporting for the reported event will be provided under mdr #1220648-2024-20754.
17315 | 1220648-2024-20503 | 20477656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20477656&pc=OZD | 10/15/2021 | 10/17/2024 |  | 10/17/2024 |  |  | 10/15/2021 | 10/17/2024 |  |  | 05/31/2023 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089245 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella pump stopped suddenly. several attempts to restart the device were unsuccessful. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17316 | 1220648-2024-20504 | 20477662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20477662&pc=OZD | 06/10/2022 | 03/27/2025 |  | 10/17/2024 |  |  | 06/02/2022 | 10/17/2024 |  |  | 10/31/2023 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155859 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed occasional instances of ventricular tachycardia. there was also a pump stop, and multiple attempts were made to restart the device. the team then replaced the console, which did not resolve the issue. the device was explanted. additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop and vf/vt/af/at issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the device passed all post-sterile inspection checks. the evaluation of the device noted that due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was returned and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17317 | 1220648-2024-20506 | 20477720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20477720&pc=OZD | 08/10/2021 | 10/17/2024 |  | 10/17/2024 |  |  | 08/10/2021 | 10/17/2024 |  |  | 01/31/2023 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022046024 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis and later went into renal failure. it was also noted was the patient experienced bleeding/hematoma, which required a blood transfusion. additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17318 | 1220648-2024-20507 | 20477728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20477728&pc=OZD | 06/08/2022 | 10/17/2024 |  | 10/17/2024 |  |  | 06/08/2022 | 10/17/2024 |  |  | 02/29/2024 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192054 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. there was a slight tendency for ischemia in both lower limbs. the patient was observed without additional treatment per the family¿s decision. it was later noted that the patient expired due to multi-organ failure due to lack of additional treatment. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17319 | 1220648-2024-20502 | 20477745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20477745&pc=OZD | 04/12/2022 | 10/17/2024 |  | 10/17/2024 |  |  | 04/12/2022 | 10/17/2024 |  |  | 11/30/2023 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157951 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. after the patient expired and the impella was explanted a blood clot was found in the pig tail. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause was most likely patient condition related, and the thrombus developed due to the lack of heparin within the purge. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17320 | 1220648-2024-20505 | 20477764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20477764&pc=OZD | 06/23/2021 | 10/17/2024 |  | 10/17/2024 |  |  | 06/23/2021 | 10/17/2024 |  |  | 08/31/2022 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1501563 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. it was reported that the patient was placed on amiodarone for wide ventricular tachycardia, which subsided. it was reported that there was a purge system blocked alarm. tissue plasminogen activator (tpa) protocol was being utilized for high purge pressure/low purge flow. it was also noted that the pump stopped. a new pump was implanted. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17321 | 1220648-2024-20499 | 20477772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20477772&pc=OZD | 03/20/2022 | 10/17/2024 |  | 10/17/2024 |  |  | 03/20/2022 | 10/17/2024 |  |  | 12/31/2023 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172103 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. temporary pump stop was noted. the pump was restarted then stopped again. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17322 | 1220648-2024-20500 | 20477774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20477774&pc=OZD | 07/12/2022 | 05/22/2025 |  | 10/17/2024 |  |  | 07/12/2022 | 10/17/2024 |  |  | 02/29/2024 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192070 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  |  | Asian | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the impella was unable to be delivered initially. after ballooning of the lower limb, impella was inserted. subsequently, the patient expired from the drop in hemoglobin and an inability to maintain hemodynamics. it was noted causality is unknown because it is undeniable that the patient was on anticoagulation with impella, which may have contributed to the hemorrhage. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17323 | 1220648-2024-20508 | 20477789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20477789&pc=OZD | 06/16/2022 | 10/17/2024 |  | 10/17/2024 |  |  | 06/16/2022 | 10/17/2024 |  |  | 02/29/2024 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202265 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. impella cp had a pump stop alarm. the impella was removed and the doctor did not insert a new impella device. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the pump stop has been completed. no product data or product was returned, and the clinical details provided were insufficient to determine a root cause. the cause of the issue was not determined since the product was not returned. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17324 | 1220648-2024-20498 | 20477915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20477915&pc=OZD | 03/06/2022 | 10/17/2024 |  | 10/17/2024 |  |  | 03/09/2022 | 10/17/2024 |  |  | 09/30/2023 |  | 10/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129776 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient's hemoglobin dropped sharply, and ct examination revealed hemothorax and cerebral hemorrhage. it was judged that neurological recovery was difficult due to the degree of cerebral hemorrhage. the patient was moved to palliative treatment and expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17325 | 20461125 | 20461125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20461125&pc=QOC | 06/01/2024 | 07/09/2024 | 07/09/2024 | 10/16/2024 | 07/09/2024 | 10/16/2024 |  | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | QOC | [5 empty] | ABIOMED INC |  |  | US | [5 empty] | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 1 MO |  | Female |  |  |  | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | We have had several instances recently of impella 5.5 catheter covers being torn/damaged. i discussed the issue with the abiomed representatives, and it appears that the proximal and distal ends of the catheter's cover have white plastic rings that hold the clear plastic cover in place. these plastic rings are being dislodged leading to the plastic cover to be compromised. abiomed has elevated the issue to their uplines to see if they can make adjustments to the adhesives. in the meantime, they recommend that we secure the plastic tips with tape/ tegaderm dressings.
17326 | 1220648-2024-20428 | 20461421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20461421&pc=OZD | 11/10/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 11/22/2021 | 10/16/2024 |  |  | 07/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105859 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, hemolysis was observed. the device was explanted, and the procedural outcome was the patient expired on support. it was also reported that there was no causal relationship with the impella and the patient expiring. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17327 | 1220648-2024-20427 | 20461432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20461432&pc=OZD | 10/07/2021 | 05/23/2025 |  | 10/16/2024 |  |  | 11/25/2021 | 10/16/2024 |  |  | 07/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110749 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had bleeding/hematoma that required a blood transfusion. the patient ultimately expired, but it was not indicated that it was not related to the impella. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17328 | 1220648-2024-20429 | 20461537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20461537&pc=OZD | 11/22/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 11/26/2021 | 10/16/2024 |  |  | 05/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092768 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had thrombocytopenia. it was also reported that the patient expired while on support due to the causative disease with no allegations made toward the impella as the cause of death. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17329 | 1220648-2024-20431 | 20461598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20461598&pc=OZD | 08/20/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 08/20/2021 | 10/16/2024 |  |  | 06/30/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100688 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a hematoma and developed limb ischemia. the hematoma formation site was on the inguinal side of the insertion site of impella, and the bleeding could be visually confirmed. the limb ischemia was reported due to the impella was inserted into the lower limb on the impella insertion side, the relevance to the ischemia could not be ruled out. while on support, the expired, the cause of death was circulatory failure, or rather chronic heart failure, and there was no causal relationship with impella or any trouble during operation. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported access site bleeding and limb ischemia issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding was high patient act values. the root cause of the limb ischemia was most likely the condition/size of the patient vessel. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17330 | 1220648-2024-20430 | 20461765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20461765&pc=OZD | 10/07/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 10/07/2021 | 10/16/2024 |  |  | 05/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083755 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was saturated flows and hemolysis. other than reducing the device's performance level, no intervention was required in response to the saturated flows/hemolysis. the patient became confused as time went on and the family made the decision to withdraw care. survival outcome at explant noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported inaccurate high flow and hemolysis issues has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned product was visually inspected, and biomaterial was observed in the outflow and in the purge gap. the impeller was difficult to rotate by hand. delaminated epotek was observed. the cause of the inaccurate high flow was most likely ingested biomaterial applying resistance to the pump rotor. the cause of the hemolysis was most likely delaminated epo-tek. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17331 | 1220648-2024-20433 | 20461843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20461843&pc=OZD | 06/28/2023 | 10/16/2024 |  | 10/16/2024 |  |  | 06/28/2023 | 10/16/2024 |  |  | 03/31/2025 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349504 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during preparation the pump was unable to prime for implant. troubleshooting was attempted, however the issue remained. the procedure was aborted due to the priming failure. there was no patient involvement at the time of the failure. survival outcome at explant noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was discarded by the customer and no clinical information was provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
17332 | 1220648-2024-20434 | 20462042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462042&pc=OZD | 12/10/2021 | 05/23/2025 |  | 10/16/2024 |  |  | 12/29/2021 | 10/16/2024 |  |  | 07/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114359 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was bleeding and hematoma that required transfusion. the patient had expired due to multiple organ failure during support and was not related to the impella. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17333 | 1220648-2024-20432 | 20462061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462061&pc=OZD | 06/19/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 08/03/2021 | 10/16/2024 |  |  | 02/28/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022050207 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed aortic regurgitation after inserting the impella device. it is unknown how the reported issue was resolved. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported aortic regurgitation issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion. ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve.¿ section: use of echocardiography for positioning of the impella catheter. impella catheter inlet to the aorta. ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ this report is being filed as part of a retrospective review of historical records.
17334 | 1220648-2024-20437 | 20462096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462096&pc=OZD | 09/23/2024 | 11/23/2024 |  | 10/16/2024 |  |  | 09/23/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025531933 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 placement was aborted due to the patient¿s anatomy. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation delivery issues has been completed. the returned product was visually inspected and no abnormalities were found. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature as per the clinical description provided.
17335 | 1220648-2024-20435 | 20462104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462104&pc=OZD | 09/10/2024 | 06/12/2025 |  | 10/16/2024 |  |  | 09/20/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025517831 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Tachycardia | [10 empty] | Other | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A notification was received on (b)(6)2024 via abiomed¿s long-term outcome and quality (loqi) indicator impella registry indicating an allegation of supraventricular tachycardia: synchronized cardioversion at bedside, returned to sinus rhythm. the patient recovered with no sequelae. the patient remains on impella support. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | A notification was received on 5/29/2025 via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that starting on (b)(6) 2024 endured hemolysis with a "definite" relationship to the impella device used: ¿ldh on admit from osh 2467. dropped throughout admission. ldh (b)(6) 2024 417" |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported supraventricular tachycardia (svt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the svt could not be determined. corrections to the initial mfr report: b2-selected death and date of death. b5 mso review. h1-type of reportable event changed to death. h6 impact code added 1802. | B5 revised with updated information received from the clinical site. h6 revised to reflect the additional failure mode. h10 instructions for use. data review: data logs showed pump ran without issue during support. there were suction alarms triggered in early of the case but resolved. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. should any new information or investigation results for hemolysis are received, a supplemental report will be submitted. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ g1 reporting contact email revised on manufacturer device report 1220648-2024-20435 in accordance with updated procedures.
17336 | 1220648-2024-20438 | 20462407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462407&pc=OZD | 09/11/2024 | 02/12/2025 |  | 10/16/2024 |  |  | 09/20/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531358 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team was forwarded a protect iv study that found the patient had hypotension (blood pressure dropped to 78/53 very breifly) and this was probably related to the impella and impella procedure. medical safety determined that this was likely due to the impella supporting the patient inadequately. normal saline bolus was give and 200 micrograms of neo-synephrine was given. the patient's blood pressure quickly rose back to 112/62. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into the reported issue was completed. data logs showed pump ran without issue during support. device history record review confirmed that the pump passed all post sterile inspection checks. product was not received for evaluation. the root cause of low flow was unable to be determined as limited clinical details were provided and no device was received for evaluation. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
17337 | 1220648-2024-20440 | 20462553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462553&pc=OZD | 10/01/2024 | 12/04/2024 |  | 10/16/2024 |  |  | 10/01/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547570 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | Infusion or Flow Problem |  |  |  |  | Thrombosis/Thrombus | Cardiovascular Dysfunction/Insufficiency | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during support, flows were below 1 l at p4 with ongoing suction alarms. patient¿s creatine and liver function test (lft) were on the rise. motor current was also spiking every other beat and flow saturation waveforms were intermittent with equalization of flows at 4.5l at p4. staff were alerted that there was a high potential for a pump stop. blood products were given to the patient. the pump was alarming impella in aorta when clearly on echocardiography, impella was transvalvular during suction and flow saturation. the pump was removed and the patient's lft normalized. additionally, biomaterial was observed. the patient survived the event. | [5 empty] | The investigation into hemolysis, optical signal issue, saturated flow, low pump flow, and thrombus has been completed. the product was returned and evaluated. there was significant biomaterial wrapped around the impeller. no other abnormalities were found. log review showed pump flow begin to significantly decrease during the third day of support. the flows were intermittently below 1l/min at p4. on the last day of support, the pump was turned to p-2 and pump flow increased to above 4.0l/min. there were multiple large motor current spikes in the last two days of support. the thrombus failure mode (fm) was removed since it was incorporated into the root cause of the other fms. log review shows a high number of suction alarms and multiple motor current spikes before the reported hemolysis. the root cause of hemolysis was most likely biomaterial. the root cause of the optical signal issue was most likely biomaterial. removed saturated flow as failure mode because it was subsequent and correlated to the low pump flow. duplicate root causes. log review showed pump flow begin to decrease while at p-4, just prior to all the suction alarms. the large motor current spikes occurred later. the root cause of the low pump flow was biomaterial. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20440 e4 should have been left blank. g1 reporting contact fax number missing. h6 4445, 4614, 4440, and 2964 were reported incorrectly. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.
17338 | 1220648-2024-20439 | 20462577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462577&pc=OZD | 10/01/2024 | 11/29/2024 |  | 10/16/2024 |  |  | 10/01/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1849119 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported that during prep, the staff went to unplug the automated impella controller (aic), and the 3-prong portion of the plug separated from the cord. no harm has been reported. | [5 empty] | The investigation for the console (aic) damage has been completed. the console nor the data logs were returned for evaluation. clinical details states that the power cord was shipped to this hospital account and this aic was repaired on site. the root cause of this issue was likely a defective power cord. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
17339 | 1220648-2024-20441 | 20462586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462586&pc=OZD | 07/16/2024 | 11/18/2024 |  | 10/16/2024 |  |  | 09/23/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025487634 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hypovolemic Shock | Cardiogenic Shock | Respiratory Failure | Multiple Organ Dysfunction Syndrome | Liver Failure | [7 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received on 09/23/2024 via abiomed¿s long-term outcome and quality (loqi) indicator impella registry indicating an allegation of respiratory failure: persistent multiorgan failure, liver failure, respiratory failure, cardiogenic shock, distributive shock. the patient outcome was fatal. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the respiratory failure, multiorgan failure, cardiogenic shock, hypovolemic shock (distributive), and liver failure has been completed. the device was not returned for evaluation. additionally, clinical details were not sufficient to determine root cause. therefore the root cause of the reported issues could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17340 | 1220648-2024-20436 | 20462657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462657&pc=OZD | 09/21/2024 | 11/20/2024 |  | 10/16/2024 |  |  | 09/23/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025531986 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and while on support, the sterile sleeve became disconnected from the plastic touhy lock exposing the 9f catheter and breaking the sterility of the device. support continued until successful bridge to transplant. | [5 empty] | The investigation of product damage (sterile sleeve damage) has been completed. the root cause of the product damage (sterile sleeve damage) was not determined as no product or images were returned for analysis. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
17341 | 1220648-2024-20442 | 20462660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462660&pc=OZD | 09/20/2024 | 12/09/2024 |  | 10/16/2024 |  |  | 09/20/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025524545 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2024 |  |  |  |  | The complainant reported that an impella 5.5 was successfully supporting a patient for several days while waiting for the transplant but air in the flushing system alarm appeared. the air-release procedure was performed according to the instructions. the alarm did not disappear. after repeating the procedure several times, it was decided to remove the impella 5.5 due to concerns about the patient's health. a second impella 5.5 was implanted. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into purge leak ¿ pump has been completed. the impella device was not returned for evaluation. the cause of the purge leak was most likely a leak from the sidearm but the exact location of the leak was unknown. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20442.
17342 | 1220648-2024-20445 | 20462702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462702&pc=OZD | 06/22/2023 | 05/29/2025 |  | 10/16/2024 |  |  | 06/22/2023 | 10/16/2024 |  |  | 02/28/2024 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321947 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was supported by extracorporeal membrane oxygenation and impella cp device for mechanical circulatory support. while on support, the patient died of hemorrhagic shock and extensive cerebral infarction after cardiac massage. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined
17343 | 1220648-2024-20443 | 20462763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462763&pc=OZD | 09/23/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 11/16/2021 | 10/16/2024 |  |  | 02/28/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022050191 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the patient had a cva/stroke. actions taken are unknown. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported cva/stroke issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the cva/stroke was determined to be unrelated to the impella cp since the event occurred 5 days after impella removal. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17344 | 1220648-2024-20446 | 20462935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20462935&pc=OZD | 09/27/2024 | 11/21/2024 |  | 10/16/2024 |  |  | 09/27/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1249511 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the purge cassette was being replaced on the automated impella controller (aic) and the purge disk holder cracked. the aic was replaced and support continued. this event occurred during patient support, and no harm has been reported. | [5 empty] | The device was returned for evaluation. upon completion of the investigation, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to this issue. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. added- h6 impact code 4629.
17345 | 1220648-2024-20444 | 20463183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20463183&pc=OZD | 12/03/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 12/03/2021 | 10/16/2024 |  |  | 07/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022108763 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis and renal failure. actions taken are unknown. it was stated that the patient experienced bradycardia and imaging revealed hypoxic encephalopathy. the patient¿s status was do not attempt resuscitation. it was stated that hypoxic encephalopathy was the main cause of the patient¿s death. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ report is being filed as part of a retrospective review of historical records.
17346 | 1220648-2024-20450 | 20463225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20463225&pc=OZD | 01/15/2022 | 10/16/2024 |  | 10/16/2024 |  |  | 01/18/2022 | 10/16/2024 |  |  | 10/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022148525 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the cp was followed by care escalation to the impella 5.0. the team observed a cerebral infarction. the infarction cause was not determined, and the team questions the travel from a cp thrombus to the brain. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported stroke issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17347 | 1220648-2024-20447 | 20463253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20463253&pc=OZD | 02/04/2022 | 10/16/2024 |  | 10/16/2024 |  |  | 02/04/2022 | 10/16/2024 |  |  | 10/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022148643 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device along with extracorporeal membrane oxygenation (ecmo) support for mechanical circulatory support. while on support the patient suffered a hemorrhagic neurological event. the event was attributed to a long ecmo run and anticoagulation. additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17348 | 1220648-2024-20448 | 20463396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20463396&pc=OZD | 09/26/2024 | 11/20/2024 |  | 10/16/2024 |  |  | 09/26/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1231461 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during prepping the automated impella controller (aic) a controller error alarm occurred. the aic was removed. no patient involvement has been reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of the automated impella controller (aic) issue has been completed. the data logs from the console confirmed that the system self check failure alarm was issued. this system self check failure was preceded by numerous powermod1 errors. the console was returned for evaluation. upon boot up, there was a system-self check failure. inspection of the power battery manager (pbm) under a digital microscope confirmed that the corrosion of copper traces in vicinity of super capacitors c69 and c70. the root cause of the system self check failure was contamination of pbm through supercapacitor electrolyte leakage leading to pbm corrosion.
17349 | 1220648-2024-20449 | 20463607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20463607&pc=OZD | 09/27/2024 | 12/04/2024 |  | 10/16/2024 |  |  | 09/27/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525911 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp had high purge pressure system blocked alarm. the purge cassette was exchanged but the same issue continued. percutaneous coronary intervention procedure completed under constant visible alarm. the impella was explanted after three hours of support. there was no harm to the patient. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the high purge pressure has been completed. the pump was inspected and biomaterial was found in the purge gap. the pump was soaked and run and the high purge pressure did not reproduce. the data logs confirm high purge pressure alarms. the logs show a motor current spike with speed deviation and drop in purge flow. after the spike there was a gradual rise in purge pressure before alarm was triggered. tissue plasminogen activator (tpa) was not added to the purge. the high purge pressure did not resolve during the case. the root cause of high purge pressure was determined to be biomaterial. the instructions for use for the impella cp with smartassist for use during cardiogenic shock states the following: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged. added-d9 device return date, h6 impact code 4627.
17350 | 1220648-2024-20452 | 20463619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20463619&pc=OZD | 07/06/2024 | 11/24/2024 |  | 10/16/2024 |  |  | 09/24/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446121 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Thrombocytopenia | [11 empty] | Required Intervention | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support for six days. the patient was admitted in with cardiogenic shock and cardiomyopathy. after the six days, the pump was explanted and venous-venous extra corporeal membrane oxygenation was placed. the support was in (b)(6) 2024. later in (b)(6) 2024, abiomed¿s long-term outcome and quality indicator (loqi) impella registry database reported upon two adverse events during the support. the thrombocytopenia and renal failure both had possible relationship to the use of impella. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation into thrombocytopenia and renal failure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b was entered incorrectly on manufacturer device report: 1220648-2024-20452.
17351 | 1220648-2024-20451 | 20464126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20464126&pc=OZD | 09/27/2024 | 12/11/2024 |  | 10/16/2024 |  |  | 09/27/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524000 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in india had a cp support initiated for a patient in acute myocardial infarction / cardiogenic shock. the pump was supporting to allow for the percutaneous coronary intervention of the coronary artery. the team found post-explant that there was an arterial dissection that prompted bleeding and intervention in the form of two units of red blood cells. the patient survived the issue, but the team made the decision to withdraw care and patient expired. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation into access site bleeding and vascular damage - peripheral vessel has been completed. as vascular damage to the femoral artery (the vessel used for access) was explicitly mentioned through the voice of the customer, the investigation into the vascular damage failure mode will cover both reported complications. vascular damage: data logs are not relevant and product was not returned for investigation. clinical details do not report any information that would explain how or when the damage to the femoral artery occurred. due to limited clinical details and product not being returned, the root cause of the vascular damage could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-20451
17352 | 1220648-2024-20457 | 20464344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20464344&pc=OZD | 09/17/2024 | 11/30/2024 |  | 10/16/2024 |  |  | 09/17/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502992 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. it was reported that the impella cp was removed due to signs of limb ischemia on the impella leg side. fem-fem crossover bypass was not an option. care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: d4-udi number.
17353 | 1220648-2024-20453 | 20464346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20464346&pc=OZD | 09/16/2024 | 11/15/2024 |  | 10/16/2024 |  |  | 09/16/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025519940 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump placed to escalate care for a patient with syncope. it was reported that there were signs of hemolysis. the patient had blood in urine. the impella was too deep and was pulled back. it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation has been completed. the device was not returned for analysis. the data logs were reviewed and there were positioning alarms seen but no motor current spikes or a trend in the placement signal. the clinical details stated that the pump was mal-positioned and upon repositioning, the blood in the urine began to clear up. the root cause of the hemolysis is most likely due to improper positioning. the failure mode positioning issue was added as it was stated that there were placement signal low alarms at the beginning and the pump was not in a good position. the data logs showed alarms but then they resolve with re-positioning the pump. the root cause of the positioning issue is most likely due to use. b.1 revised to add product problem per investigation results since the initial manufacturer device report number 1220648-2024-20453 was submitted. h.6 medical device problem code was added to reflect the addition of the failure mode positioning issue since the initial manufacturer device report number 1220648-2024-20453 was submitted.
17354 | 1220648-2024-20454 | 20464855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20464855&pc=OZD | 09/17/2024 | 11/14/2024 |  | 10/16/2024 |  |  | 09/17/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025520116 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and while on support the patient had ventricular tachycardia. the patient was shocked and the pump was repositioned. after repositioning, the patient's ectopy resolved. support was continued. the patient eventually expired after care was withdrawn. upon review by abiomed's medical safety officer the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt was not determined.
17355 | 1220648-2024-20461 | 20465874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20465874&pc=OZD | 10/04/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 10/04/2021 | 10/16/2024 |  |  | 05/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083716 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Perforation of Vessels | [11 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the complainant reported the patient expired due to acute myocardial infarction. additional information noted that the patient also experienced cardiac tamponade and hemolysis, actions taken to resolve the issues are unknown. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17356 | 1220648-2024-20460 | 20466595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20466595&pc=OZD | 01/05/2022 | 10/16/2024 |  | 10/16/2024 |  |  | 01/24/2022 | 10/16/2024 |  |  | 02/28/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022050840 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon initiating the device, the pump showed no flows and had a flat motor current and placement signal. the patient expired during impella support, however there was no allegation that this was related to the impella. no further information was provided. | [5 empty] | The investigation for the low pump flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the low pump flow and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
17357 | 1220648-2024-20458 | 20466728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20466728&pc=PYX | 09/16/2024 | 11/15/2024 |  | 10/16/2024 |  |  | 09/16/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025478615 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support and during the implant, mattress sutures were placed prior to additional dilations. the sutures snapped when attempting to tighten. additional tires were placed, but the patient continued to bleed from the rp access site. the doctor placed additional sutures. oozing continued throughout support. it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be use issue because the bleeding was controlled by placing the additional sutures at access site.
17358 | 1220648-2024-20462 | 20466882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20466882&pc=OZD | 07/22/2024 | 10/16/2024 |  | 10/16/2024 |  |  | 07/22/2024 | 10/16/2024 |  |  | 10/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151369 |  | 0048-0024-JP |  | No |  |  | No |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after normal explant of the impella cp, the percutaneous cardiopulmonary support (pcps) was difficult to insert, the patient experienced a stroke and died with insertion. no further information is available. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. ischemic stroke: without imaging evidence to confirm the source of the stroke was due to a clot that passed through the impella device, the relation to impella is unknown. hemorrhagic stroke: without a pre-existing ischemic stroke, if heparin was in the purge, it is unknown if the cause of the stroke was due to heparin. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17359 | 1220648-2024-20464 | 20466886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20466886&pc=OZD | 07/23/2024 | 10/16/2024 |  | 10/16/2024 |  |  | 07/23/2024 | 10/16/2024 |  |  | 12/31/2022 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022032117 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  | Asian | Death | No | Yes | Positioning Failure | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for cardiomyopathy. the complainant reported that they encountered a positioning issue. it was found that the impella had completely come out of the left ventricle. the pump was removed again in the catheter laboratory and reinserted. it was noted that the patient died seventeen days post explant. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. positioning issues that occur after initial delivery of device into the ventricle. excludes placement signal issues. the root cause of positioning issue was most likely patient condition. this report is being filed as part of a retrospective review of historical records.
17360 | 1220648-2024-20459 | 20466894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20466894&pc=OZD | 05/16/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 03/08/2022 | 10/16/2024 |  |  | 02/28/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022052321 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced a cardiac tamponade with unknown relationship to the impella. the patient ultimately expired. actions taken are unknown. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported issue has been completed. no device was returned for evaluation. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
17361 | 1220648-2024-20468 | 20466937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20466937&pc=OZD | 07/21/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 04/05/2022 | 10/16/2024 |  |  | 04/30/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022075933 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | A |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis during use of the device. no intervention nor premature removal was reported. additional information noted that the patient expired on support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17362 | 1220648-2024-20456 | 20467030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20467030&pc=OZD | 09/15/2024 | 11/20/2024 |  | 10/16/2024 |  |  | 09/16/2024 | 10/16/2024 |  |  | 09/30/2025 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024405076 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 75-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient experienced hemolysis. two plasma free hemoglobin draws were conducted and both were over 40. care was ultimately withdrawn and the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported hemolysis has been completed. the device, data logs or sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
17363 | 1220648-2024-20465 | 20467219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20467219&pc=OZD | 10/14/2024 | 10/29/2024 |  | 10/16/2024 |  |  | 10/14/2024 | 10/16/2024 |  |  | 05/31/2024 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED. INC. |  |  | US |  | 2023225194 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported that the patient died from a cerebral hemorrhage. no further information is available. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | Medical device report # 1220648-2024-20465 was identified as a duplicate report of medical device report #1220648-2024-20755 . follow up reporting for the reported event will be provided under mdr #1220648-2024-20755.
17364 | 1220648-2024-20463 | 20467396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20467396&pc=OZD | 07/23/2024 | 10/16/2024 |  | 10/16/2024 |  |  | 07/23/2024 | 10/16/2024 |  |  | 11/30/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023165832 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 24 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the patient died from a cerebral hemorrhage. it was also reported that the impella cp worked as expected. no further information available. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the investigation for the reported stroke issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the stroke and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17365 | 1220648-2024-20467 | 20467407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20467407&pc=OZD | 10/20/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 03/25/2022 | 10/16/2024 |  |  | 07/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103850 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, hemolysis was observed in the patient. actions taken to resolve is unknown. additional information noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17366 | 1220648-2024-20455 | 20467411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20467411&pc=OZD | 09/16/2024 | 11/18/2024 |  | 10/16/2024 |  |  | 09/16/2024 | 10/16/2024 |  |  |  |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494801 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 84-year-old female noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impella support, the pump had signs of hemolysis and ectopy during the 2 day support. the patient had care withdrawn and expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the hemolysis and ectopy has been completed. product and data logs were not returned for review. the root cause of hemolysis and ectopy was unable to be determined as limited clinical details were provided and no product was returned for review. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
17367 | 1220648-2024-20470 | 20467519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20467519&pc=OZD | 12/29/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 12/29/2021 | 10/16/2024 |  |  | 07/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144643 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Atrial Fibrillation | [12 empty] | Other | Death | [23 empty] | A |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced runs of atrial fibrillation. there was also a pump stop, flows were saturated, and the mc was very elevated, even on p3. actions taken to resolve the issues are unknown. the impella was unable to be successfully restarted. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17368 | 1220648-2024-20471 | 20467534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20467534&pc=OZD | 12/23/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 03/01/2022 | 10/16/2024 |  |  | 08/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022118811 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support the patient suffered a cerebral infarction resulting in the impella being removed. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported issue of "stroke/cva" has been completed. the impella device was not received from the customer, therefore, investigation of the device was not possible. the cause of the stroke/cva and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: as noted in the impella ifu "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17369 | 1220648-2024-20466 | 20467579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20467579&pc=PYX | 09/16/2024 | 11/20/2024 |  | 10/16/2024 |  |  | 09/16/2024 | 10/16/2024 |  |  | 10/31/2025 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024413658 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the tip of 23 french dilator of impella rp flex split. they were unable to advance into the patient for imepella placement. a new sheath was used without any issue and the impella was placed. it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the introducer damage has been completed. no product or images were returned for analysis. the root cause of the introducer damage - mechanical was not determined as no product or images were returned for analysis.
17370 | 1220648-2024-20469 | 20467690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20467690&pc=OZD | 11/26/2021 | 10/16/2024 |  | 10/16/2024 |  |  | 04/07/2022 | 10/16/2024 |  |  | 07/31/2023 |  | 10/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108929 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Sepsis | Ventricular Fibrillation | [9 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 04/07/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis, renal failure, ventricular tachycardia (vt), and sepsis while on impella support. actions taken to resolve the issues is unknown. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported issues has been completed. no device was returned for evaluation. the cause of the issue was not determined since product was not returned. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17371 | 1220648-2024-20388 | 20447985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447985&pc=OZD | 03/16/2022 | 10/15/2024 |  | 10/15/2024 |  |  | 03/16/2022 | 10/15/2024 |  |  | 11/30/2023 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162693 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that an impella cp pump lost flow during support and alarmed for positioning in the aorta when it was not in the aorta. the team replaced the pump and support was continued with the new pump. there was no reported patient harm. | [5 empty] | The investigation for the reported low pump flows and positioning issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. data log review shows a sudden spike in motor current, followed by an abrupt drop in flows and loss of motor current pulsatility. the placement signal was still aortic, causing an impella in aorta alarm. based on the log review, the cause of the low flow is most likely ingested biomaterial. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ this report is being filed as part of a retrospective review of historical records.
17372 | 1220648-2024-20389 | 20448034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20448034&pc=OZD | 02/04/2023 | 10/15/2024 |  | 10/15/2024 |  |  | 02/05/2023 | 10/15/2024 |  |  | 06/30/2024 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246494 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the clinical team observed an "impella stopped" alarm during support, which resolved itself almost immediately after triggering. the patient's condition was not affected. the decision was made to proactively change the aic (automated impella controller) for another one. no other issues occurred for the remainder of support. | [5 empty] | The investigation for the reported pump stop has been completed. data logs show that an "impella stopped (rpm deviations" alarm triggered and lasted for 2 seconds before resolving and all performance metrics went to 0. the alarm resolved on its own and the pump restarted and performed with no issues. however, the cause of the pump stop was not determined since the product was not returned. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17373 | 1220648-2024-20390 | 20450077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20450077&pc=OZD | 09/20/2024 | 11/19/2024 |  | 10/15/2024 |  |  | 09/20/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025496212 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Pump | Pumping Stopped | Priming Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp had a motor stop alarm prior to insertion. the device was replaced. | [5 empty] | The investigation of pump did not start and priming issue has been completed. the pump was returned and upon inspection, the full red lumen was found in the pump. the red lumen was removed and the pump was run. the pump stop did not reproduce and the device performed to specification with no priming issue or motor current pump stop. no data logs or imc logs were provided. the root cause of the pump did not start was determined to be due to a red lumen issue. the root cause of the priming issue was not determined as no data logs were provided and insufficient information related to issue was provided the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20390. the failure mode was changed from 'pump stop' to 'pump did not start' because pump was unable to be start. a 'priming issue' failure mode was added for the 'purge issues' mentioned in the complaint intake description which occurred prior to insertion. h6 code 1503 was reported incorrectly. new codes were added to medical device problem code. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
17374 | 1220648-2024-20281 | 20450107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20450107&pc=OZD | 03/16/2022 | 10/15/2024 |  | 10/15/2024 |  |  | 03/19/2022 | 10/15/2024 |  |  | 08/31/2023 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022130771 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Sepsis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the patient became febrile. there was no information provided on intervention performed, if any, for the fever of unknown origin. it was also reported there was a high motor current pump stop that occurred on day 10 of impella cp support. the staff tried to reset the device without success, and it was noted that a clot formed in the outlet of the cannula. the decision was made to explant the device. the patient was hemodynamically stable after impella removal. | [5 empty] | The data logs were returned and confirm a high motor current pump stop on day 10 of support. the product as returned and biomaterial was found in the outflow and under the impeller. the root cause of the pump stop was determined to be ingested biomaterial. the pump stop occurred on day 10 of support which is beyond the 8 day design life per design trace matrix 0046-9717. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17375 | 1220648-2024-20391 | 20450392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20450392&pc=OZD | 09/23/2024 | 12/20/2024 |  | 10/15/2024 |  |  | 09/23/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025526019 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional treatment information was received that the patient went into operating room for evacuation of hematoma as well as irrigation and debridement of wound. patient also had jackson pratt (jp) drain placed post operative at the site. the patient recovered with no sequelae. | The complainant reported that the patient on impella 5.5 support developed a hematoma at the axillary insertion site. after the patient went to the bathroom, the patient complaint of pain in that area. bivalirudin was temporarily paused. the patient is scheduled to go for a hematoma evacuation. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. |  |  |  |  | B5 updated with additional treatment information received. h6 health effect impact code updated to reflect treatment. investigation for access site bleeding is pending should any new information be received, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was determined to be patient movement because patient complained pain and discomfort at the access site after going to the bathroom and resulted in hematoma. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-20391.
17376 | 1220648-2024-20392 | 20450414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20450414&pc=OZD | 06/05/2024 | 11/20/2024 |  | 10/15/2024 |  |  | 09/23/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1629362 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after the patient was weaned from the automated impella controller, the physician wanted to remove the purge cassette. the purge disk holder broke when the purge disk was removed. there was no reported patient harm. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. during the device analysis the purge disc retainer was found to be broken (retainer completely broken off). d9 date device returned to manufacturer added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
17377 | 1220648-2024-20393 | 20450718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20450718&pc=OZD | 09/23/2024 | 02/26/2025 |  | 10/15/2024 |  |  | 09/23/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534073 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient Device Interaction Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was inserted over the 0.18 guidewire. upon attempting to remove the 0.18 guidewire, the medical doctor felt resistance thereofre the impella and 0.18 wire were removed. it was noted that the 0.18 wire was sheared. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the complaint intake product is being updated from the impella to the guidewire to match the voice of the customer. the returned product was inspected and guidewire damage was confirmed. as evidenced by clinical details, the root cause of the delivery issue was most likely use related due to excessive force. g1 reporting contact email revised on manufacturer device report 1220648-2024-20393 in accordance with updated procedures.
17378 | 1220648-2024-20394 | 20451010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451010&pc=OZD | 09/23/2024 | 12/09/2024 |  | 10/15/2024 |  |  | 09/23/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559544 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a (b)(6) year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient developed bleeding from the impella access site. compresses applied, heparin paused, and 2 units of red blood cells were given. | It was further reported that the pump was successfully weaned and explanted. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added b5 patient outcome. corrections to the initial mfr report: g1 mdr reporting contact email.
17379 | 1220648-2024-20395 | 20451293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451293&pc=OZD | 09/19/2024 | 12/06/2024 |  | 10/15/2024 |  |  | 09/19/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025518755 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, the patient developed hemolysis. the impella was located in the papillary muscle and there were attempts to reposition the impella but were unsuccessful. there were no alarms. the patient outcome and explant date are unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. product and data logs were not returned for review. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position.
17380 | 1220648-2024-20401 | 20451305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451305&pc=OZD | 09/23/2024 | 11/19/2024 |  | 10/15/2024 |  |  | 09/23/2024 | 10/15/2024 |  |  | 10/31/2025 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024419245 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding, the complainant had placed an impella 5.5 pump to support a 50-year-old male patient admitted in covid+ and with cardiogenic shock (cgs) and cardiomyopathy. the impella 5.5 pump was placed while team awaited planned lvad placement. after 7 days the pump was explanted and hm3 was placed. in (b)(6) 2024, the patient experienced renal failure with "possible" relationship to the impella. the patient was started on continuous renal replacement therapy (crrt). | [5 empty] | The investigation for the renal failure has been completed. pump was not returned for investigation. data logs were returned and no abnormalities were observed. the pump was explanted several weeks before the renal failure was reported. without pump back the complaint cannot be further investigated. therefore, the root cause of the renal failure issue was not determined since product was not returned. added h6 impact code 4641. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17381 | 1220648-2024-20396 | 20451367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451367&pc=OZD | 09/22/2024 | 12/05/2024 |  | 10/15/2024 |  |  | 09/22/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481553 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported that a patient was on impella cp support. while on support, the nurse reported a pump stop with the error message ¿pump stop mc high¿. the probability of a thrombus was confirmed. the impella failed to start, and the decision was made to replace the impella. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into pump stop and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-20396. e4 should have been left blank . g1 reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
17382 | 1220648-2024-20397 | 20451372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451372&pc=OZD | 09/22/2024 | 11/16/2024 |  | 10/15/2024 |  |  | 09/22/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533960 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp being placed via a 14fr sheath to support a patient admitted in with acute myocardial infarction and cardiogenic shock. it was reported that the circle lock on long cp dilator broke off and was attached still to the peelaway sheath. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of introducer damage ¿ mechanical has been completed. no product or images were returned for analysis. the root cause of the introducer damage - mechanical was not determined as no product was returned for analysis. e4 should have been left blank on manufacturer device report 1220648-2024-20397. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
17383 | 1220648-2024-20400 | 20451386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451386&pc=OZD | 05/27/2024 | 11/15/2024 |  | 10/15/2024 |  |  | 09/23/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025444280 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding, a 63-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient endured recurrent ventricular tachycardia (vt) requiring defibrillations and cardio-pulmonary resuscitation (cpr) with a "possible" relationship to the impella device used. epinephrine and levophed increased. vaso also increased. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
17384 | 1220648-2024-20405 | 20451451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451451&pc=OZD | 09/30/2024 | 11/18/2024 |  | 10/15/2024 |  |  | 09/30/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523984 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and the patient had cold extremity and no pulses were dopplerable. there were limited flows in the vessels. a distal perfusion catheter was inserted which was successful. additionally, the patient had ventricular tachycardia (vt) during the impella insertion. the patient was shocked at 200j and went back to normal sinus rhythm. the patient had arrhythmia prior to the impella support, however it is unknown if the impella implant contributed to the vt. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported issue has been completed. the impella device was not received from the customer and as the necessary information was not provided, the root cause of the vt/vf and limb ischemia was not determined."
17385 | 1220648-2024-20399 | 20451503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451503&pc=OZD | 09/22/2024 | 11/17/2024 |  | 10/15/2024 |  |  | 09/22/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448610 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had hematoma and significant oozing to the access site. a femstop was placed and heparin was held. additionally, the patient urine output was red. the hemoglobin, hematocrit, and platelets were dropping significantly. one unit of red blood cells was given and the impella was removed. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation access site bleeding and hemolysis has been completed. the data logs showed pump ran without issue during support. there were suction alarms triggered during support but resolved quickly. the pump was returned for analysis. visual inspection found no issues on the pump. the root cause of hemolysis was most likely patient condition. the root cause of access site bleeding was most likely patient movement as it happened during transfer. e4 should have been left blank on manufacturer device report.
17386 | 1220648-2024-20398 | 20451542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451542&pc=OZD | 09/22/2024 | 09/22/2025 |  | 10/15/2024 |  |  | 09/22/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544165 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant had a impella cp support placed to support a patient with known coronary artery disease and getting ardiopulmonary resuscitation and defibrillation shocks. it was reported that after impella placement there were signs of hemolysis and limb ischemia. the hemolysis was seen by urine color change and were given p level adjustment and fluid bolus. the limb was given a distal perfusion cannula. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and limb ischemia has been completed. the impella was not returned for evaluation. log review did not show any motor current spikes or suction alarms before the reported hemolysis that would indicate biomaterial ingestion. the placement signal began trending down shortly after the reported hemolysis. the root cause of the hemolysis was not determined since product was not returned. the root cause of the limb ischemia could not be determined without additional information. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned. log review showed the placement signal trending down in the beginning of the case. the "placement signal low" alarms occurred almost 10 hours after the placement signal began to trend down. the 5s parameters did not have any abnormalities. the pump was confirmed to be correctly positioned, but the patient's lv was small. the root cause of the optical signal issue was most likely patient-condition related. the fm "thrombus" was added and investigated since there was a thrombus found in the stomach and removed. the root cause of the thrombus was not determined since insufficient clinical details were provided. the fm "vascular damage - peripheral vessel" was added and investigated since there was a gi bleed. the root cause of the vascular damage was not determined since insufficient clinical details were provided. the root cause of the vascular damage was not determined since insufficient clinical details were provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem b5 updated with additional information from the initial manufacturer device report number 1220648-2024-20398. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures from the initial manufacturer device report number 1220648-2024-20398. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-20398.
17387 | 1220648-2024-20403 | 20451549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451549&pc=OZD | 09/28/2024 | 11/26/2024 |  | 10/15/2024 |  |  | 09/28/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025531924 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Intraoperative Cardiac Injury | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported that impella 5.5 insertion via axillary method was attempted in right axillary artery. the physician was unable to successfully pass the device through axillary artery - vessel size was likely too small. the physician attempted to balloon angioplasty the artery and was still unable to get the outlet to advance much further than mid-clavicular line. during de-airing, pocket of air noted at apex of left ventricle (lv). the physician realized that with the heart manipulation that the impella catheter had perforated the lv. insertion of impella 5.5 was aborted and cp was inserted without issue. | On (b)(6) 2024 the impella 5.5 was primed and placed over 0.018 wire. upon attempt to advance through axillary artery, the physician was unable to get outlet to advance past mid-clavicular line. the physician made multiple attempts to advance device as well as balloon angioplasty and arterial nitroglycerin administration, and was unsuccessful. the impella 5.5 removed from axillary artery and it was decided to proceed with placement of impella cp. the patient was rested on impella cp and vent and on (b)(6) 2024, the patient was brought to the cardiovascular operating room to undergo impella cp removal, coronary artery bypass grafting x4/mitral valve repair. the decision was made the physician would attempt to do a direct implant of the second impella 5.5 the impella 5.5 device was wirelessly and under direct visualization through the open end of the graft anastomosis and the impella was advanced across the aortic valve. after placement, the physician continued with completion of the anastomosis of the platinum graft to the aorta. once anastomosis completed, cross clamp removed and the impella device was briefly turned on to de-air. during de-airing, pocket of air noted at apex of the left ventricle and the physician attempted to assist with air removal by manipulating heart. there was a sudden influx of air noted on transesophageal echocardiogram. upon further examination, the physician realized that with the heart manipulation that the impella catheter had perforated the left ventricle. impella device was pulled back and hole was sewn using a patch and pledgetted stitches. heart was de-aired for a second time and device position was optimized. cpb flows were then weaned as impella flows were increased to p-7 with flows of 4.1 l/min. patient remained hypotensive despite 4.1 l/min of impella flow. the decision was made to withdraw care and the patient expired. per medical safety officer review it is unknown if the delivery issues of the first impella 5.5 was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature due to patient's anatomy. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20403. b1 updated to adverse event. b2 death and date of death added. b5 updated with patient outcome and mso statement. e4 should have been left blank. g1 reporting contact fax number missing. h1 type of reportable event. h6 death added to health effect - impact code. manufacturer device report 1220648-2024-23295 reported death for the second impella device used in this case.
17388 | 1220648-2024-20402 | 20451564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451564&pc=OZD | 09/26/2024 | 12/10/2024 |  | 10/15/2024 |  |  | 09/26/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532406 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported that the impella cp was explanted due to the patient developing bleeding issue. during the implantation, the patient was confused and tended to buckle the leg. the patient is stable and survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report:g1 mdr reporting contact email.
17389 | 1220648-2024-20404 | 20451583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451583&pc=OZD | 09/29/2024 | 11/27/2024 |  | 10/15/2024 |  |  | 09/29/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523834 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Defective Device | Infusion or Flow Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient had impella cp support and during support, the long peel away sheath had oozing around the orange rim. no blood products were given. staff applied a clamp to the sheath and the issue resolved. additionally, the pump flow became saturated. the mean motor current was above 1000, flow max 3.7, min 3.7, average 3.7, p level 8. the pump was exchanged. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into introducer damage ¿ mechanical and saturated flow has been completed. the impella device was not returned for evaluation. the root cause of the saturated flow was most likely due to biomaterial. since the device was not received for investigation root cause for the introducer damage ¿ mechanical is unknown. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20404. e4 should have been left blank g1 reporting contact fax number missing.
17390 | 1220648-2024-20408 | 20451702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451702&pc=OZD | 09/21/2024 | 11/08/2024 |  | 10/15/2024 |  |  | 09/21/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545580 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 37-year-old male patient on impella cp support had oozing from the impella insertion site. heparin was held for two hours and manual pressure was used. the pump was explanted after three days of support. the patient was noted to be stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17391 | 1220648-2024-20406 | 20451716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451716&pc=OZD | 09/19/2024 | 02/13/2025 |  | 10/15/2024 |  |  | 09/24/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025513859 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 53-year-old- male patient had a replacement impella cp smart assist that was placed in the patient's right axillary. after the implant, the patient was on impella support for 5.67 hours before the patient expired. the cause of death was determined to vt/vf. the relationship between the cause of death and the impella is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-20406 in accordance with updated procedures.
17392 | 1220648-2024-20407 | 20451793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20451793&pc=OZD | 09/20/2024 | 12/03/2025 |  | 10/15/2024 |  |  | 09/20/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502288 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a vga connector failure on an automated impella controller. there was not any patient use noted at the time of the malfunction. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Data analysis: there was no log evidence available related to the reported failure mode. device analysis: this console was tested after arriving in fs and reproduced the reported failure mode. there was no output from the rlm vga connector. upon replacing the original backstrap cable with a known-good cable, the rlm vga connector still did not output. root cause: the root cause for no output from the rlm vga connector was most likely a defective rlm board. there was no display issue found with the aic during the case. g1: corrected the manufacturer email address. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation.
17393 | 1220648-2024-20409 | 20452072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20452072&pc=OZD | 09/26/2024 | 11/24/2024 |  | 10/15/2024 |  |  | 09/26/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540275 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and there was low pump flow with the motor current off as well. additionally, there was a possible interaction with the valve. the pigtail was stuck and unable to be free. the pump was explanted due to low pump flow. the patient survived the event. | [5 empty] | The investigation for the low pump flow and the positioning issues has been completed. the pump was returned for investigation. no physical abnormalities were observed on the returned product. the pump ran as expected during testing in a bucket of water. the data log shows that the pump was started on p9, with reported low pump flow and several suction alarms. a placement signal trend was observed, which correlated with the motor current and a drop in pump flow. later, the pump was set to a lower p-level, which also resulted in subsequent suction alarms. no abnormalities were reported with the positioning issue or optical signal parameters. the pump was unable to achieve good flow at higher p-levels. according to the clinical team, low pump flow and suction issues were reported throughout the case, even after the pump was repositioned and blood volume was adjusted. prior to the explanation, the pigtail was found to be stuck and could not be freed. the cause of the low pump flow issue was most likely patient condition related, based on return product investigation and data log review. the cause of the positioning issue was not determined since sufficient clinical information was not provided. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
17394 | 1220648-2024-20410 | 20452424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20452424&pc=PYX | 09/26/2024 | 12/09/2024 |  | 10/15/2024 |  |  | 09/26/2024 | 10/15/2024 |  |  | 10/31/2024 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023273326 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | HEARTMATE3 | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella rp was implanted for right ventricle failure following heartmate3 implantation. after 20 minutes of support the purge pressure high alarm triggered. the impella was running at p8 with outflow of 2.9 l/min. changing the purge cassette and lowering dextrose concentration did not solve the problem. the impella was removed, and extra corporeal membrane oxygenation was used for right support. the patient survived the event. | [5 empty] | The investigation into high purge pressure has been completed. returned product was investigated and biomaterial was found to be in the cannula and purge gap. the pump was connected to a console in the lab and high purge pressure reproduced. then, a window was cut in the catheter just proximal to the motor housing and there was no blood ingress. finally, a cut was made through the catheter at that same location, and the high purge pressure immediately resolved. from returned product it was determined that the root cause of the high purge pressure was most likely be biomaterial. d9 date device returned to manufacturer added. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20410. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
17395 | 1220648-2024-20411 | 20452516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20452516&pc=PYX | 09/22/2024 | 11/22/2024 |  | 10/15/2024 |  |  | 09/22/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025547385 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella rp flex to support a patient admitted in with right sided heart failure. the patient has multiple cardiac and systemic risk factors. the pump was placed but there were concerns of hemolysis and possible clot ingestion. the pump was explanted after two days. the patient survived the explant. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the thrombus and hemolysis has been completed. the returned pump was inspected and biomaterial was found wrapped around the impeller. there were no other visual abnormalities with the pump. the clinical details also mention patient's condition was complicated by bacteremia, covid pneumonia, and acute on chronic systolic heart failure requiring increasing doses of dobutamine. the data logs show no motor current spikes and no positioning issues. there were no suction alarms and placement signal was trending downwards. the pump also passed the hemolysis test. the root cause of the hemolysis was most likely due to patient condition. the root cause of the thrombus could not be determined without additional clinical details.
17396 | 1220648-2024-20413 | 20452783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20452783&pc=OZD | 09/16/2024 | 11/15/2024 |  | 10/15/2024 |  |  | 09/23/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524255 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Failure to Pump | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant was on impella cp support that was weaned and removed after percutaneous coronary intervention. the team then found after explant that the patient was in need of hemodynamic support and the team replaced with a new impella cp. the impella cp failed to start and a third impella cp was placed. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the pump did not start has been completed. data logs were reviewed, and there were no issues during the case. there were no abnormal alarms throughout the pci support. it can be seen in the logs that the pump is unplugged and the console is shut down. then, roughly 30 minutes later, the console is restarted and the original pump is plugged back in. the console skips to case start step 6, which is normal when a known pump is plugged back in, and then the "impella stopped. retrograde flow" alarm is issued. 40 seconds later, the user requests to shut down the console and the pump is unplugged. they then restart the console and plug in the replacement pump. the user likely thought the "impella stopped" alarm meant they could not use the pump again, however, it is really just a warning that if the pump is left not running in the body for too long, retrograde flow can occur. product was not returned for investigation. based on clinical details and data log review, the root cause of the pump stop was determined to be no problem found. impella cp with smartassist for use during cardiogenic shock; section: troubleshooting the purge system; ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ corrections to the initial mfr report: b1-incorrectly selected adverse event.
17397 | 1220648-2024-20412 | 20452854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20452854&pc=OZD | 09/26/2024 | 11/17/2024 |  | 10/15/2024 |  |  | 09/26/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519887 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and the patient had severe disease and stenosis in their femoral arteries prior to the impella implant. once the peel away sheath was placed in the left femoral, distal pulses were lost. staff placed a distal perfusion catheter, and the limb ischemia resolved. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
17398 | 1220648-2024-20414 | 20455888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20455888&pc=OZD | 09/20/2024 | 11/20/2024 |  | 10/15/2024 |  |  | 09/20/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524145 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | Life Threatening | Required Intervention | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a patient admitted in with angina. the impella cp was placed and supported for five days. the patient had multiple runs of ventricular tachycardia and was shocked. the patient did have underlying electrophysiological issues. the patient remained stable until the successful wean and explant. | It was reported that lidocaine drip was initiated. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary information was not provided, the root cause of the vt was not determined. b.5 added information that was inadvertently omitted from initial manufacturer device report number: 1220648-2024-20414. h.6 code 4644 was added to health effect - impact code due to information added in b.5 since initial manufacturer device report number: 1220648-2024-20414 was submitted.
17399 | 1220648-2024-20415 | 20455929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20455929&pc=OZD | 09/21/2024 | 11/08/2024 |  | 10/15/2024 |  |  | 09/21/2024 | 10/15/2024 |  |  |  |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539754 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and the patient had access site bleeding. one unit of blood products were transfused, heparin was discontinued and a sandbag was placed. additionally, the patient had several runs of sustained ventricular tachycardia requiring defibrillation. support continued and the patient remained stable. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported arrhythmia and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of arrhythmia and access site bleed could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
17400 | 1220648-2024-20416 | 20457609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20457609&pc=OZD | 05/23/2023 | 10/15/2024 |  | 10/15/2024 |  |  | 05/23/2023 | 10/15/2024 |  |  | 12/31/2024 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318620 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were suction issues resulting in low pump flow. repositioning attempts were made to resolve the issue; however, the patient was withdrawn from care the following day. it was noted that suction issues were related to patient's afterload and mean arterial pressure. the patient expired after 24.18 hours of support. medical safety review of the event note use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | Pump was not returned for investigation. data log shows low flow and several suction alarm. the cause of the low flow issue was not determined since product was not returned and sufficient clinical information was not provided. this report is being filed as part of a retrospective review of historical records.
17401 | 1220648-2024-20420 | 20457637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20457637&pc=OZD | 02/25/2023 | 10/15/2024 |  | 10/15/2024 |  |  | 02/25/2023 | 10/15/2024 |  |  | 08/31/2024 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259606 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced an absence of pulses in the leg. actions taken to resolve the issue are unknown. additional information noted care was withdrawn and the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation into the reported issue of limb ischemia has been recompleted. the impella device was not received from the customer, therefore, investigation of the device was not possible. no clinical details provided were sufficient to determine a root cause. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17402 | 1220648-2024-20422 | 20457643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20457643&pc=OZD | 12/06/2022 | 05/22/2025 |  | 10/15/2024 |  |  | 12/06/2022 | 10/15/2024 |  |  | 08/31/2023 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022118868 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis and bleeding was noted while on support. transfusion was required. additional information noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17403 | 1220648-2024-20418 | 20457679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20457679&pc=OZD | 08/12/2024 | 02/13/2025 |  | 10/15/2024 |  |  | 08/12/2024 | 10/15/2024 |  |  | 01/31/2023 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022060036 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced thrombocytopenia during support. actions taken to resolve the issue are unknown. additional information noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation of the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. correction section h6: the health impact code was inadvertently submitted incorrect, as the device cannot conclusively be associated with the outcome (patient's demise). the correct code has been added to this report .
17404 | 1220648-2024-20425 | 20457693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20457693&pc=OZD | 06/07/2022 | 05/23/2025 |  | 10/15/2024 |  |  | 06/07/2022 | 10/15/2024 |  |  | 10/31/2023 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147719 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced a hematoma and bleeding at the access site. the sheath was repositioned to resolve this issue. additional information noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the access site bleeding was most likely use related as the bleed was stopped by progressing the repo unit further into the vessel. the root cause of the patient death was determined to be patient condition and unrelated to the use of impella as no details were provided to suggest a causal relationship. this report is being filed as part of a retrospective review of historical records.
17405 | 1220648-2024-20421 | 20457700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20457700&pc=OZD | 06/30/2023 | 10/15/2024 |  | 10/15/2024 |  |  | 06/30/2023 | 10/15/2024 |  |  | 12/31/2024 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318173 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced bilateral limb ischemia. actions taken to resolve the issue are unknown. additional information noted care was withdrawn and the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation into the reported issue of "limb ischemia" was completed. no device was received for evaluation. the cause of the limb ischemia issue was most likely patient condition related as both limbs were noted to be ischemic and unrelated to the impella device. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17406 | 1220648-2024-20423 | 20457702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20457702&pc=OZD | 09/16/2021 | 10/15/2024 |  | 10/15/2024 |  |  | 03/24/2022 | 10/15/2024 |  |  | 06/30/2023 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100715 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis. actions taken to resolve the issue are unknown. additional information noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation into the hemolysis has been completed. the impella device was not received from the customer, therefore, investigation of the device was not possible. no data logs were provided for review. the clinical details provided were insufficient to determine a root cause. the cause of the hemolysis was not determined due to insufficient information, no data logs, and the product not being returned. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17407 | 1220648-2024-20424 | 20457726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20457726&pc=OZD | 05/24/2022 | 10/07/2025 |  | 10/15/2024 |  |  | 05/24/2022 | 10/15/2024 |  |  | 11/30/2023 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157923 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced access site bleeding requiring blood transfusion. the patient continued on support until the device was successfully weaned. additional information noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | D4 model number was erroneously entered on the initial manufacturer device report number 1220648-2024-20424. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-20424 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-20424. in accordance with updated procedures. h6 health effect - impact code 1802 and investigation conclusions code 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-20424.
17408 | 1220648-2024-20419 | 20457738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20457738&pc=OZD | 06/06/2023 | 10/15/2024 |  | 10/15/2024 |  |  | 06/06/2023 | 10/15/2024 |  |  | 03/31/2025 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342377 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced bloody urinary output and absent pulses in the foot on the impella side. actions taken to resolve the issue are unknown. additional information noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation into the reported issue of limb ischemia - reversible was completed. no device was received for evaluation. clinical information was sufficient to determine the cause of the reported issue. the cause of limb ischemia issue was most likely patient condition related based on the clinical information. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17409 | 1220648-2024-20426 | 20457997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20457997&pc=OZD | 06/11/2022 | 10/07/2025 |  | 10/15/2024 |  |  | 06/16/2022 | 10/15/2024 |  |  | 11/30/2023 |  | 10/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161718 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient developed hemolysis while on support. actions taken to resolve the issue are unknown. additional information noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-20426. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-20426 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-20426 in accordance with updated procedures. h6 health effect - impact code 1802 was erroneously entered on the initial manufacturer device report number 1220648-2024-20426. h6 investigation conclusions 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-20426.
17410 | 1220648-2024-20326 | 20436632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436632&pc=OZD | 04/06/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 04/06/2023 | 10/14/2024 |  |  | 05/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023228499 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient had developed a clot in the left leg. pedal pulses were lost on the leg, patient was taken to ccl where the superficial femoral artery and popliteal arteries were found blocked. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
17411 | 1220648-2024-20327 | 20436639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436639&pc=OZD | 02/12/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 02/12/2023 | 10/14/2024 |  |  | 09/30/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281365 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. while on support there were reports of saturated flows. in addition, high motor current alarms were noted for which troubleshooting was attempted. the pump was replaced with another impella cp device. | [5 empty] | The investigation for the reported saturated has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. this report is being filed as part of a retrospective review of historical records.
17412 | 1220648-2024-20328 | 20436643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436643&pc=OZD | 02/13/2023 | 09/30/2025 |  | 10/14/2024 |  |  | 02/13/2023 | 10/14/2024 |  |  | 10/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024286290 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Difficult to Insert | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella cp device for mechanical circulatory support. while on support there were positioning and placement signal issues. it was also noted that there was an unsuccessful attempt to place the impella utilizing a 14 french dryseal. the impella was explanted and a 19 french transcatheter aortic valve replacement (tavr) sheath was placed using the impella site. the tavr was placed without patient harm. | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. there were positioning and placement signal issues. the pump remained on for support despite the malfunctions. there was no reported harm or injury to the patient. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the physician stated that the patient had aortic insufficiency. the root cause for the positioning and optical signal issue is due to patient condition. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
17413 | 1220648-2024-20329 | 20436650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436650&pc=OZD | 03/16/2023 | 03/10/2025 |  | 10/14/2024 |  |  | 03/16/2023 | 10/14/2024 |  |  | 04/30/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226486 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, abnormal flows for p level and slightly elevated motor current, indicating possible flow saturation. a pump stop was also noted. no further information was provided. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump stop and saturated flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the device passed all post-sterile inspection checks. the evaluation of the device noted that the data logs showed that the pump stopped immediately upon a sudden motor current spike accompanied by the "impella stopped - motor current high" alarm; the pump was unable to be restarted. about 3 days prior to the pump stop, the motor current began to gradually rise, and the pump flow gradually became more elevated and saturated with no change in p-level; the increases were preceded by a sudden spike in purge pressure. visual inspection of the returned pump revealed the presence of biomaterial on the impeller within the cannula. no other damage or abnormalities. in conclusion, it was determined that the cause of the pump stop was ingested biomaterial. the root cause of the saturated flow could not be determined as insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
17414 | 1220648-2024-20330 | 20436657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436657&pc=OZD | 03/27/2023 | 10/29/2025 |  | 10/14/2024 |  |  | 03/27/2023 | 10/14/2024 |  |  | 12/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303740 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were low pump flows and a pump stop. the stopped pump was able to be restarted x2 and the team replaced the aic. the pump was later explanted as the team had the patient additionally supported by extracorporeal membrane oxygenation. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | G1 reporting contact email revised on manufacturer device report 1220648-2024-20330 in accordance with updated procedures. h6 medical device problem code 1248 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-20330.
17415 | 1220648-2024-20331 | 20436729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436729&pc=OZD | 03/29/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 03/29/2023 | 10/14/2024 |  |  | 11/30/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024310525 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support there was a low pump flow coinciding with a continuous suction issue. upon further examination, it was discovered that there was a thrombus in the inlet of the pump. the pump was replaced with another impella cp pump for ongoing support. | [5 empty] | The investigation for the reported low pump flow and thrombus has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the low pump flow was most likely patient due to biomaterial ingestion. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
17416 | 1220648-2024-20332 | 20436740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436740&pc=OZD | 04/03/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 04/03/2023 | 10/14/2024 |  |  | 10/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024296958 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient had palpable hematoma and bleeding at the left groin access site. the patient¿s urine was dark red in color. it was reported that manual pressure was applied to resolve the hematoma, and the device was repositioned to manage hemolysis. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17417 | 1220648-2024-20333 | 20438461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20438461&pc=OZD | 09/19/2024 | 11/21/2024 |  | 10/14/2024 |  |  | 09/19/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563294 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 68-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient developed bleeding at the access site. several stiches were applied and bleeding was resolved. it was further reported that the patient also experienced multiple suction alarms, having to drop down to p3. the staff will be checking the impella placement and the physician will be assessing flows to ensure the patient is receiving the flow needed. the pump was explanted the following day. the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and low pump flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
17418 | 1220648-2024-20334 | 20438579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20438579&pc=OZD | 09/19/2024 | 12/16/2024 |  | 10/14/2024 |  |  | 09/19/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1457787 |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had a red alarm and the tile was not showing in the impella connect. this error was connected to a clinical case. while looking at the case, it was noted that there was a data streaming issue. no patient harm has been reported. | [5 empty] | The investigation into automated impella controller (aic) - controller failure (red alarm) has been completed. the cause of the issue was not determined since the product was not returned. d.2 common device name was upadted since the initial manufacturer device report number 1220648-2024-20334 was submitted in accordance with update procedures. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-20334 was submitted. | A.1, a.4 and a5 are unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
17419 | 1220648-2024-20335 | 20438615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20438615&pc=OZD | 09/24/2024 | 11/17/2024 |  | 10/14/2024 |  |  | 09/24/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1684985 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us based hospital biomed department observed a newly unboxed automated impella controller to have damage. damage was on the cover and housing. there was no patient involvement. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) damage before therapy has been completed. console logs from the device do not contain any relevant data. the console was returned for evaluation. upon visual inspection, there was misaligned (rotated) pump connector cover, misaligned ¿power button¿ label, scratches on the rotating knob, chips and scratches on the front housing: upper right and bottom right corner. console testing also revealed compromised optical pump connector causing failing ¿5s¿ optical parameter test (snr and contrast out-of-spec). the root cause of the console damage was mishandling of the console.
17420 | 1220648-2024-20336 | 20438699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20438699&pc=OZD | 09/19/2024 | 11/19/2025 |  | 10/14/2024 |  |  | 09/19/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1750264 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The biomedical engineer had an automated impella controller which failed to purge the cassette. no harm or injury. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
17421 | 1220648-2024-20338 | 20438728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20438728&pc=OZD | 04/06/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 04/06/2023 | 10/14/2024 |  |  | 12/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318069 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited a low purge with no leaks. while the purge cassette was being replaced, the device produced a "placement signal unreliable" alarm. the team chose to remove and replace the pump, and no patient harm was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported purge leak and optical signal issue has been completed. data showed low purge pressure (pp) & placement signal (ps) spiked to 3000 mmhg during the case start that triggered alarms pp low & ps not reliable. pump analysis: visual inspection found no issue in the purge system of the pump. pump plug was opened to check if any damage on the components & a broken optical fiber was found during visual inspection. pump tests: 5s parameters were out of specification on the console & optical laser continuity test was failed which confirmed the ps unreliable failure mode. the cause of the purge leak was most likely due to use, and the cause of the optical signal issue was a damaged optical fiber line. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
17422 | 1220648-2024-20337 | 20438729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20438729&pc=OZD | 09/23/2024 | 11/25/2024 |  | 10/14/2024 |  |  | 09/24/2024 | 10/14/2024 |  |  | 04/30/2025 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351279 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported there was consistent suction. the pump was removed and a clot was found. the patient¿s outcome is unknown. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional. interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into low pump flow has been completed. the impella device was not returned for evaluation. complaint device not returned for investigation. data log reviewed: no motor current (mc) spike observed. suction alarms are confirmed. drop in mc observed without changing p-level. based on the clinical data clot found upon pump removal for low pump flow issue. the cause of the low pump flow issue most likely was biomaterial ingestion. the failure mode (fm) thrombus removed as the root cause of the low pump flow issue fm cover this fm. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20337. b1 should have been product problem only. h6 code 4440, 4614, and 4627 were reported incorrectly. new codes were added to health effect - clinical code and health effect - impact code.
17423 | 1220648-2024-20340 | 20438809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20438809&pc=OZD | 06/09/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 06/09/2023 | 10/14/2024 |  |  | 05/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229253 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Fluid/Blood Leak | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock was needing an impella cp for mechanical circulatory support. immediately after implantation there were purge alarms, suction alarms, and placement signal alarms. as troubleshooting was unable to resolve the issues, the pump was replaced with another impella cp device. there was no reported harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
17424 | 1220648-2024-20339 | 20438820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20438820&pc=OZD | 04/08/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 04/08/2023 | 10/14/2024 |  |  | 12/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303704 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump produced suction alarms regardless of the pump speed level. an echocardiogram then revealed that there was a clot on the pump's inlet valve. the pump was explanted, and no additional thrombus was found. the patient survived the procedure. | [5 empty] | The investigation for the reported low pump flow and thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. data log shows the low flow with reported suction alarm. doctor reported clot found on inlet cage of pump. the cause of the low flow issue was biomaterial found on inlet cage based on provided clinical information. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17425 | 1220648-2024-20341 | 20438895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20438895&pc=OZD | 09/24/2024 | 11/23/2024 |  | 10/14/2024 |  |  | 09/24/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025505346 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and they had a hematoma at the access site. pressure was held, but the bleeding worsened significantly after. products were given to the patient and the impella was removed. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation access site bleeding has been completed. the pump was returned for inspection and not abnormalities were observed. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
17426 | 1220648-2024-20342 | 20438966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20438966&pc=OZD | 04/27/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 04/28/2023 | 10/14/2024 |  |  | 07/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255739 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device reportedly stopped functioning and produced a "pump stop" alarm. the device was explanted and replaced, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the pump stop could not be determined. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17427 | 1220648-2024-20343 | 20438980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20438980&pc=OZD | 10/20/2021 | 10/14/2024 |  | 10/14/2024 |  |  | 10/20/2021 | 10/14/2024 |  |  | 06/30/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100469 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the team observed a pump stop that they attributed to use of the wrong clamp at the pump motor. the event caused no reported harm or injury. impella was removed and an intra-aortic balloon pump placed for postoperative care. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the pump stop and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17428 | 1220648-2024-20344 | 20438997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20438997&pc=OZD | 06/14/2022 | 05/02/2025 |  | 10/14/2024 |  |  | 06/14/2022 | 10/14/2024 |  |  | 12/31/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171657 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported on the 10th day of impella use, the pumped abruptly stopped. the pump was explanted. there was no patient harm reported. | Additional information noted that there was a liquid leakage occurred from housing. no action wa staken to address the issue. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the issue was not determined since product was not returned. correction section : b5, h6. this information was inadvertently left out in the initial report which has been added to this report.
17429 | 1220648-2024-20345 | 20439125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439125&pc=OZD | 09/24/2024 | 11/21/2024 |  | 10/14/2024 |  |  | 09/24/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510643 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella 5.5 implant and the pump would not cross the subclavian artery. the pump was removed and a kink was noted in the cannula. the implant was aborted. the patient survived the event. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, unable to pass through the point of the subclavian artery as per the clinical description provided. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17430 | 1220648-2024-20346 | 20439156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439156&pc=OZD | 10/13/2022 | 10/14/2024 |  | 10/14/2024 |  |  | 10/13/2022 | 10/14/2024 |  |  | 06/30/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240352 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted with post cardiotomy cardiogenic shock¿and (pccs)/low cardiac output syndrome (lcos) was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp was pushed too deep into the ventricle when the repositioning sheath was placed. as the pump was pulled back out it experienced a high motor current pump stop. after multiple unsuccessful attempts to restart the pump it was restarted but had saturated flows. due to the saturated flows, a new impella was placed. no harm to the patient resulted from the use of this device. | [5 empty] | The investigation for the reported issue has been completed. the data logs confirm a motor current spike right before the high motor current pump stop. this is indicative of biomaterial ingestion. the product was returned and biomaterial was found wrapped around the base of the impeller. the root cause of the saturated flow was determined to be biomaterial. as noted in the impella ifu: impella cp with smart assist system section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17431 | 1220648-2024-20347 | 20439278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439278&pc=OZD | 05/02/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 05/02/2023 | 10/14/2024 |  |  | 02/28/2025 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324719 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was successfully placed, however, when the fellow pulled the impella back into the aorta and they were unable to recross aortic valve. the decision was made to remove this impella due to the positioning issue and replace it with a new impella. there was no patient harm reported, and this device was replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ this report is being filed as part of a retrospective review of historical records.
17432 | 1220648-2024-20348 | 20439323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439323&pc=PYX | 09/24/2024 | 11/26/2024 |  | 10/14/2024 |  |  | 09/24/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025476526 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a rp flex support placed for a patient who had an amulet left atrial appendage device placed after atrial perforation. the rp flex was placed for the hemodynamic support. the patient was in disseminated intravascular coagulation and had surgical bleeding. the pump had low flows and was explanted after 21 hours. the pump at explant was seen to have thrombosed. the patient was lacking anticoagulation due to surgery. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | Pump was not returned for investigation. data logs were investigated and high purge pressure alarms were observed followed by low flows. therefore, failure mode low flow was changed to high purge pressure since low flow is a result of high purge pressure. there was a motor current spike observed right before the purge pressure rise. these characteristics are of sudden biomaterial ingestion. therefore, the root cause of the high purge pressure issue was most likely biomaterial. the failure mode of ¿thrombus¿ was deleted since it is covered in the high purge pressure. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20348. h6 code 4440 and 4627 were reported incorrectly. new codes were added to health effect - clinical code and health effect - impact code.
17433 | 1220648-2024-20349 | 20439338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439338&pc=OZD | 09/18/2024 | 12/03/2025 |  | 10/14/2024 |  |  | 09/18/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1768515 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The biomedical engineer reported that the impella automated impella controller had a black screen and nothing else was showing. there was no patient involvement. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. d2 common device name revised on manufacturer device report in accordance with updated procedures. g1. reporting contact email revised on manufacturer device report in accordance with updated procedures. h6. investigation conclusions revised on manufacturer device report in accordance with updated procedures. h10. additional manufacture narrative revised on manufacturer device report in accordance with updated procedures. | The investigation has been completed. data review: console logs were available via rlfs. no evidence to support the reported black screen was observed after reviewing the imc logs. device analysis: testing and analysis of the console was not available because the product was not returned. conclusion: the root cause of the reported black screen issue could not be determined because the product was not returned. d4: corrected the catalog number. g1: corrected the manufacturer email address h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation.
17434 | 1220648-2024-20350 | 20439410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439410&pc=OZD | 09/17/2024 | 07/07/2025 |  | 10/14/2024 |  |  | 09/17/2024 | 10/14/2024 |  |  | 12/31/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023164548 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Female |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp escalated to an impella 5.0 for a patient admitted in with cardiogenic shock and cardiomyopathy. the impella 5.0 was placed but the patient expired after just two hours. the impella 5.0 had low flows and there was an observation made of a left ventricle clot. of note, the facility used an expired pump. the expiration date on the packaging was 12/2023 but the pump was used in (b)(6) 2024. upon review by abiomed's medical safety officer, there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | Low pump flow: after pump implant, several alarms with impella outlet blocked alarm. one day prior to pump implant, thrombus was noted on the tee. the pump was not returned. data logs were not recovered. the cause of the low pump flow was not determined since product was not returned, data logs were not recovered, and insufficient clinical details. thrombus: 1 day prior to pump implant, thrombus was noted on the tee. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
17435 | 1220648-2024-20351 | 20439562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439562&pc=OZD | 09/18/2024 | 11/21/2024 |  | 10/14/2024 |  |  | 09/18/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531523 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump delivered to support a patient admitted in for a high risk surgery patient. the impella cp pump was cross clamped and when clamping was released the pump had low flows. the flows prompted the team to remove and replace the pump. no patient harm was reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. complaint device not returned for investigation. data log were reviewed which showed many suction alarms occurred during the case. a few short positioning alarms observed that resolved with trouble shooting. flow was in optimal range until the clinical states that the pump was clamped. flows were seen to be low after. however due to the device not been returned and unclear evidence provided the low pump flow cannot be determined.
17436 | 1220648-2024-20352 | 20439571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439571&pc=OZD | 09/24/2024 | 11/21/2024 |  | 10/14/2024 |  |  | 09/24/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532930 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed a large groin hematoma. the cp was removed surgically. the patient survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17437 | 1220648-2024-20353 | 20439653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439653&pc=OZD | 09/17/2024 | 11/14/2024 |  | 10/14/2024 |  |  | 09/17/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025504028 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had impella 5.5 support and upon explanting the impella 5.5 the radial arterial line and pulse oxygen was lost on the right side. the arm was warm and in good color but a clot was found in the proximal radial and distal brachial arteries. the decision was made to do endarterectomy. a small incision was made and a clot was removed. the patient regained flow to the distal right arm. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and thrombus were not determined.
17438 | 1220648-2024-20355 | 20439747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439747&pc=OZD | 11/06/2022 | 10/14/2024 |  | 10/14/2024 |  |  | 11/06/2022 | 10/14/2024 |  |  | 07/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255090 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient was coding while the staff was prepping the pump. after shocking the patient, the patient became unruly and kicked the pump off the sterile field. the physician used another pump to stabilize the patient. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the product damage was use related. this report is being filed as part of a retrospective review of historical records.
17439 | 1220648-2024-20354 | 20439749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439749&pc=OZD | 09/25/2024 | 11/27/2024 |  | 10/14/2024 |  |  | 09/25/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1814768 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella support when the automated impella controller (aic) had damage to the cord with slightly exposed wires. the console was running on battery power despite switching outlets. the console was replaced. no harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) battery issue and cabling issue has been completed. data logs nor the device was returned for evaluation. however, clinical details were provided which were sufficient to determine the root cause. the reported complaint was most likely due to the a failure of the wire strain relief (defective ac cord set).
17440 | 1220648-2024-20356 | 20439753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439753&pc=OZD | 09/17/2024 | 12/04/2025 |  | 10/14/2024 |  |  | 09/17/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1213653 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had been setting up the impella 5.5 and automated impella controller (aic) for support when the console failed to detect the pump and the team replaced the pump but the second pump failed also. the team then replaced the aic and this remedied the issue. no patient harm has been reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
17441 | 1220648-2024-20358 | 20439899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439899&pc=OZD | 09/24/2024 | 11/22/2024 |  | 10/14/2024 |  |  | 09/25/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516993 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us medical center placed a 5.5 pump to support a patient admitted in for a high-risk percutaneous coronary intervention due to cancer diagnosis precluding the coronary artery bypass graft surgery. the pump was placed but, there was an access site bleed on the day after implant. the team treated the hematoma by dressing changes, three units of red blood cells, evacuation of the hematoma, and heparin holding. the pump remained on for support, and the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient movement, patient became extremely agitated and began pulling at lines and overly moving his right arm as per the clinical description provided.
17442 | 1220648-2024-20357 | 20439934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20439934&pc=PYX | 09/17/2024 | 11/21/2024 |  | 10/14/2024 |  |  | 09/17/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025502660 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp flex support and they had a hematoma on their neck. pressure dressing was applied, heparin was paused for six hours and the hematoma improved. support continued until the patient was successfully weaned off the device. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be anticoagulation management because the bleeding was resolved after stopping the heparin.
17443 | 1220648-2024-20359 | 20440023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20440023&pc=PYX | 09/18/2024 | 11/21/2024 |  | 10/14/2024 |  |  | 09/18/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529020 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Sense | [6 empty] | Sepsis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 59-year-old female patient was on impella rp flex support due to st elevated myocardial infarction. while on support, the patient developed aspiration pneumonia/sepsis and was treated with an antibiotic. it was also noted that the map readings of the placement signal were inconsistent. the patient had some supraventricular tachycardia (svt) overnight. the patient was found to be hypokalemic. potassium was given and the svt was resolved. the waveform on impella rp flex was becoming more erratic and looking ventricular in nature. the position of the device was appropriate, but the waveform had become worse with negative numbers on systolic and diastolic. at this time, there was also a placement signal unreliable alarm. motor current was still present and pulsatile. according to the representative, it is likely that the differential pressure sensor has gone bad causing erratic waveforms and flows. the pump was successfully explanted after nine days of support. the patient is stable and survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the placement signal, sepsis, and ventricular tachycardia (vt) has been completed. the pump was returned for investigation and passed all electrical testing. no biomaterial was found on the impeller or in the purge gap. additionally, no issues were noted with the optical signal parameters or external damage to the dp sensor. pump flow appeared to be saturated during testing in a bucket of water. the pump was sent for 24 hour flow loop test to evaluate the dp sensor issue, which is primarily used to calculate the placement signal and pump flow. no issue was reported on dp sensor during the flow loop test. the data log showed a placement signal not reliable alarm after three hours of the case run, with low signal-to-noise ration (snr) and contrast was observed in the optical signal parameters. later during the case, several drop in snr and contrast was reported. cvp was negative throughout the entire case run. there was a loss of flow, and the placement signal spiked to 3000. cvp also spiked to 300 during this time. the same trend was observed several times during the case run, with an active alarm. a motor current spike was reported 12 hours before the case end. the cause of the placement signal issue was most likely patient condition related, based on returned product investigation and data log review. placement signal failure mode severity was changed from 3 to 1 since the pump was not removed due to placement signal issue. as the necessary information was not provided, the root cause of the vt was not determined.
17444 | 1220648-2024-20360 | 20440257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20440257&pc=OZD | 09/18/2024 | 11/20/2024 |  | 10/14/2024 |  |  | 09/18/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544930 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | Adverse Event Without Identified Device or Use Problem | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient admitted in with a cardiac history and known endovascular aneurysm repair (evar). the team placed the 14 french sheath with difficulty due to presence of the evar and native anatomy. the team observed a vessel dissection of the common iliac. this procedure was aborted due to safety concerns. the dissection was treated with balloon angioplasty from the contralateral access. the dissection resolved after treatment. this patient will be brought back later for axillary access for impella support for treatment of coronary arteries. | It was further reported that an angiogram confirmed there was a kink in the introducer when placing the sheath. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation has been completed. no product was returned for analysis. the clinical details stated that there was difficulty placing the sheath due to tortuous vessels as well as an endovascular aneurysm repair graft in place. this resulted in a kink in the sheath. the root cause of the introducer damage - mechanical is most likely due to the patients condition. the root cause of the vascular damage - peripheral vessels is most likely due to the patients condition as they had diseased vessels. the root cause of the introducer damage - mechanical is most likely due to the patients condition. the kink in the introducer was not addressed in initial manufacturer device report number: 1220648-2024-20360. therefore, the following sections were revised: b.2 added product problem, b.5 added information, h.6 medical device problem code 2889 was added.
17445 | 1220648-2024-20361 | 20440586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20440586&pc=OZD | 09/25/2024 | 11/17/2024 |  | 10/14/2024 |  |  | 09/25/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532336 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Unspecified Vascular Problem | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and after the impella explant, the perclose were only achieving partial hemostasis. the doctor elected to obtain left femoral artery access to go up and over for right femoral artery tamponade however, the sheath and balloon could not be advanced up and over. patient's blood pressure dropped to the 68/51/56 at which time the doctor moved to left radial access to perform balloon tamponade. after twenty minutes of tamponade, hemostasis was obtained. the patient survived the event. | [5 empty] | The investigation for the vascular damage has been completed. pump was not returned for investigation. as per the clinical details provided, after removal of the pump, hemostasis was not achieved even after 55 minutes of manual pressure. finally balloon tamponade was used and hemostasis was successful. the reason for the vascular damage is unknown. therefore, the root cause of vascular damage - peripheral vessel is unknown since insufficient clinical information was provided. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
17446 | 1220648-2024-20362 | 20440757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20440757&pc=OZD | 09/25/2024 | 11/21/2024 |  | 10/14/2024 |  |  | 09/25/2024 | 10/14/2024 |  |  |  |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533750 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had peripheral vascular disease before placement of the impella cp. the physician was able to access the artery with the 14f and used the repositioning sheath before bringing patient to the intensive care unit (icu). the patient was unstable before placement, but impella was mostly used for high risk percutaneous coronary intervention. when the patient arrived in icu the doppler pulse could not be dopplered. venous and arterial doppler revealed zero flow down the leg. it was decided that the pump should be pulled since patient had improved and was off pressor support. the physician pulled pump at bedside. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17447 | 1220648-2024-20363 | 20440970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20440970&pc=OZD | 07/05/2021 | 05/23/2025 |  | 10/14/2024 |  |  | 04/05/2022 | 10/14/2024 |  |  | 02/28/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022064820 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis and renal failure. no additional information on intervention nor premature removal was reported. the patient remained on extracorporeal membrane oxygenation (ecmo) support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17448 | 1220648-2024-20364 | 20441008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20441008&pc=OZD | 07/17/2021 | 10/14/2024 |  | 10/14/2024 |  |  | 07/27/2021 | 10/14/2024 |  |  | 06/30/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092261 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support for 11 days, the pump stopped. the pump was removed and replaced. it was also noted that during this time, that the device was repositioned several times due to hemolysis. | [5 empty] | On day 10 of impella support, a pump stop occurred after which the pump could not be restarted. the field lt and rt logs were reviewed which confirmed multiple impella stopped ¿ motor current high alarms on 27jul2021. visual inspection of the returned pump revealed a large purge gap. therefore, the root cause of the pump stop was determined to be bearing wear. this is due to the pump being used for 10 days, which is beyond the intended design life of 8 days (design trace matrix 0046-9910). impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17449 | 1220648-2024-20365 | 20441059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20441059&pc=OZD | 09/06/2021 | 10/14/2024 |  | 10/14/2024 |  |  | 09/06/2021 | 10/14/2024 |  |  | 07/31/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110296 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Pumping Stopped | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped for approximately 3-4 minutes and was able to be restarted. there were also runs of ventricular tachycardia noted. a new pump was inserted prior to transfer to the intensive care unit. | [5 empty] | The site reported that the pump stopped shortly after initiating support on (b)(6) 2021. the field lt and rt data logs were reviewed, and a impella stopped ¿ motor current high alarm was confirmed. the pump was able to be restarted and functioned normally for the remainder of the 3-hour case. visual inspection of the returned pump did not reveal any abnormalities that would have contributed to a pump stop. the pump was run in a water basin at various p levels with no pump stop reproduced. however, it was noted that motor current was high throughout the lab run (greater than 800 ma) and the pump was naturally purging more quickly than normal. the root cause for the pump pause could not be determined through this evaluation as the issue was unable to be reproduced. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17450 | 1220648-2024-20366 | 20441152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20441152&pc=OZD | 10/28/2021 | 10/14/2024 |  | 10/14/2024 |  |  | 10/28/2021 | 10/14/2024 |  |  | 07/31/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114418 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient showed signs of hemolysis, which prompted pump repositioning. the patient had red urine/hematuria. when attempted repositioning failed, the pump was pulled to the aorta and was replaced when re-crossing failed. the patient was medically managed when the pump was out of position. | [5 empty] | No product was returned for evaluation. evaluation of the provided field rt data logs revealed that impella position in ventricle, impella position unknown, and suction alarms occurred within one hour of pump insertion. additionally, numerous impella position in aorta alarms were triggered between 15:08 and 15:35 which coincides with the reported events. based on the provided information that the pump outlet had been at the aortic valve and the pump was measuring 7cm, as well as the information that the patient¿s urine cleared following repositioning, the root cause of the reported hemolysis was most likely improper positioning. refer to es2020-079 for further information. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17451 | 1220648-2024-20367 | 20441321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20441321&pc=OZD | 12/30/2021 | 10/14/2024 |  | 10/14/2024 |  |  | 01/07/2022 | 10/14/2024 |  |  | 06/30/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100653 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | Thrombosis/Thrombus | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis and a thrombus in the left ventricle, however, no adverse impact was noted, and no intervention was required beyond maintaining sufficient anticoagulation and waiting for the thrombus to shrink. also, sepsis of unknown origin was noted, which was treated with antibiotics. | [5 empty] | The impella device was not received from the customer. the site reported that there was a 20mm thrombus within the left ventricle approximately 4 days into impella support. although a specific cause for the reported thrombus could not be determined through this investigation, it was most likely related to patient condition. furthermore, the root cause of the reported hemolysis was most likely due to improper pump positioning as it improved with repositioning. refer to (b)(6) for further information. no product or data logs were available. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17452 | 1220648-2024-20368 | 20441384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20441384&pc=OZD | 03/14/2022 | 10/14/2024 |  | 10/14/2024 |  |  | 03/14/2022 | 10/14/2024 |  |  | 07/31/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108822 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp had a kink after insertion. the kink had suction alarms that prompted the team to troubleshoot. no lasting harm or injury from the kink. the team also reported upon aortic regurgitation (ar) during the support. the ar was attributed to the pump movement and heart shape. the ar was present even after explant and placement of the intra-aortic balloon pump (iabp) and pcps. | [5 empty] | The investigation for the reported cannula kink and heart valve damage issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the suction and low flow was a kink due to improper positioning. the aortic regurgitation was determined to be unrelated to the use of impella as the condition existed prior to the insertion of impella. impella cp with smart assist system instructions for use for the related event are as follows: heart valve damage section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve.¿ this report is being filed as part of a retrospective review of historical records.
17453 | 1220648-2024-20369 | 20441400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20441400&pc=OZD | 06/26/2021 | 10/14/2024 |  | 10/14/2024 |  |  | 04/05/2022 | 10/14/2024 |  |  | 01/31/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022060031 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis and renal failure. the team did not treat the event and there was no lasting harm or injury. in addition, it was noted the patient required blood transfusion due to bleeding. | [5 empty] | The investigation for the reported hemolysis, renal failure and access site bleeding issues has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17454 | 1220648-2024-20370 | 20441454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20441454&pc=OZD | 03/07/2022 | 10/14/2024 |  | 10/14/2024 |  |  | 03/07/2022 | 10/14/2024 |  |  | 11/30/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151790 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | Asian | Injury | Yes | No | Fluid/Blood Leak | Pumping Stopped | [5 empty] | Sepsis | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the impella cp had a pump stop that occurred after 14 days of support. the device was not restarted and support was continued with extracorporeal membrane oxygenation (ecmo). it was also noted there was sepsis from insertion to removal. it is unknown if impella caused or contributed to the reported sepsis. | [5 empty] | The impella device was not received from the customer. the root cause of the purge cassette leak was unable to be determined as no product or logs were returned for analysis so the location of the leak could not be identified. the root cause of the pump stop was unable to be determined as not product or logs were returned for analysis. the pump stop occurred on day 14 which is beyond the indicated design life of 8 days per design trace matrix 0046-9910. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: pump stop section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ sepsis section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17455 | 1220648-2024-20371 | 20445847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20445847&pc=OZD | 09/24/2022 | 10/14/2024 |  | 10/14/2024 |  |  | 09/27/2022 | 10/14/2024 |  |  | 01/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181770 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Fluid/Blood Leak | Failure to Read Input Signal |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, there was a persisting leak and an impella defective alarm. a leak was originally identified coming from the purge cassette. however, when the cassette was replaced, the leak persisted. the decision was made to attempt to swap the controller (aic), but upon connecting the pump to the new aic, an "impella defective" alarm was displayed. the pump was replaced without further issue. it was noted that the patient also began to experience hemolysis but repositioning of the pump resolved the issue. | [5 empty] | The root cause of the two purge cassette leaks was unable to be determined as no product was returned for analysis so the location of the leaks could not be identified. the root cause of the hemolysis was unable to be determined as no product and limited clinical details were provided. constellation only shows logs from the first console. the root cause of the pump not detected was unable to be determined as no product was returned to recreate the issue. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
17456 | 1220648-2024-20372 | 20446187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20446187&pc=OZD | 05/23/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 05/23/2023 | 10/14/2024 |  |  | 06/30/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023234956 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | Increase in Pressure | [5 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced an episode of ventricular fibrillation and was cardioverted multiple times. the impella was stated to have dislodged from its position and was unable to re-cross the aortic valve. the device was removed, and the cannula was flushed. during the process, it was observed that no purge flow was present. a second impella device was implanted to manage this complication. | [5 empty] | The investigation for the reported high purge pressure and positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was an interaction with another device, automated external defibrillator (aed). impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ this report is being filed as part of a retrospective review of historical records.
17457 | 1220648-2024-20373 | 20446303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20446303&pc=OZD | 10/25/2022 | 10/14/2024 |  | 10/14/2024 |  |  | 10/25/2022 | 10/14/2024 |  |  | 06/30/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235411 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was an air in purge alarm. multiple attempts were made to de-air the system, but the issue remained. the physician attempted to continue support, but the pump then stopped working at the end of the case, and the device was removed. there was no harm to the patient. | [5 empty] | A visual inspection of the pump revealed biomaterial in the purge gap. the returned purge cassette was attached to the aic (automated impella controller) and was successfully de-aired. the pump was also attached to the aic and was primed. when attempting to start the pump, the high motor current pump stop alarm was reproduced. there was also high purge pressure and low purge flow. in a separate test in which the purge cassette and pump were attached to an aic, the purge cassette was successfully de-aired. the purge bag was then clamped and the air in purge alarm was reproduced. the purge flow dropped to 0. the yellow luer was unscrewed and reattached and the purge pressure dropped to 0 as well, reproducing what was observed in the field logs. the cause of the pump stop was most likely a kinked purge line in the purge cassette tubing proximal to the purge bag or a clamped purge bag. a blockage of purge flow then allowed blood to ingress into the motor and cause a pump stop. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17458 | 1220648-2024-20374 | 20446574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20446574&pc=OZD | 10/29/2022 | 10/14/2024 |  | 10/14/2024 |  |  | 10/29/2022 | 10/14/2024 |  |  | 07/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254357 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was high purge pressure and purge flow alarms. the team troubleshot by infusing tissue plasminogen activator to the purge. within hours the cp pump stopped. the pump was explanted. | [5 empty] | The investigation for the reported high purge pressure and pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that data logs showed that the purge pressure gradually increased over the course of about one hour until it exceeded 1100 mmhg, triggering the "purge pressure high" alarm. there was no heparin in the purge at this time. the purge pressure remained at that level for the remainder of support with frequent alarms. following the initial occurrence of the high purge pressure, there was a steady rise in motor current (with no change in p-level) until the pump suddenly stopped upon the sudden occurrence of the "impella stopped - motor current high" alarm about one hour after the onset of the high purge pressure. visual inspection of the returned pump revealed the presence of significant amounts of biomaterial in the purge gap. the catheter was cut open and blood was observed in the purge lumen. no other abnormalities were found. therefore, it was determined that the cause of the pump stop was blood ingress into the motor, which stemmed from high purge pressure due to biomaterial occluding the purge gap. the lack of heparin in the purge most likely led to the biomaterial buildup. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17459 | 1220648-2024-20375 | 20446829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20446829&pc=OZD | 10/07/2022 | 04/16/2025 |  | 10/14/2024 |  |  | 10/07/2022 | 10/14/2024 |  |  | 04/30/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2023218721 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Thrombocytopenia | Thrombosis/Thrombus | [9 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with unknown indication was implanted with an impella cp device for mechanical circulatory support. on day 6 of the support, the pump was explanted and noted to have thrombus adhered to it. additionally, it was noted that the patient had hemolysis from insertion to withdrawal, renal failure and thrombocytopenia. impella was successfully weaned, and patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system (hemolysis). section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella cp with smartassist system (thrombus): section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist system (thrombocytopenia): section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system (renal failure): section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | No device or logs were returned. per the clinical information, thrombus was observed on the pump upon explant. however, there were no reported issues with pump performance during support. the root cause of the hemolysis, thrombocytopenia, renal failure could not be determined as insufficient clinical details were provided.
17460 | 1220648-2024-20376 | 20447032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447032&pc=OZD | 01/10/2022 | 10/14/2024 |  | 10/14/2024 |  |  | 02/15/2022 | 10/14/2024 |  |  | 06/30/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2022100741 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced bleeding, including a hematoma, was resolved due to blood transfusion. it was also reported that a temporary low purge flow occurred during support on (b)(6) 2022. although the patient's act was maintained between 160 and 180, the purge flow reduced temporarily. the concentration of dextrose in the purge was 5%. the purge pressure was between 2 and 30 cc for a long time, but the low purge flow occurred temporarily. after replacing the purge cassette, the purge flow did not improve immediately, but did improve the following day. it was also reported that the pump stopped on (b)(6) 2022. it was reported that there were no adverse events related to the impella device and no medical intervention were required. | [5 empty] | The impella device was returned. analysis of the provided data logs confirmed a rise in purge flow and decrease in purge pressure beginning on (b)(6) 2021 at 04:55, which resolved following a purge cassette exchange. the root cause of the purge leak was reported to be damage to the purge pressure transmitter. no alarms were captured on (b)(6) 2022 and the data logs showed that the device operated as intended at that time. therefore, the root cause of the reported low purge flow was unable to be determined. the provided data logs showed an increase in motor current and pump flow on (b)(6) 2022 beginning at 00:49. at 00:54, the pump stopped, triggering an 'impella stopped - motor current high' alarm. evaluation of the returned product found a large purge gap and biomaterial around the impeller and motor housing. the root cause of the pump stop was most likely ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
17461 | 1220648-2024-20377 | 20447157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447157&pc=OZD | 10/19/2021 | 10/14/2024 |  | 10/14/2024 |  |  | 10/19/2021 | 10/14/2024 |  |  | 07/31/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108940 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Failure to Sense | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an acute patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the device exhibited pump position alarms, the medical staff checked the position and found that the device was properly positioned. additionally, while on support, the patient experienced hemolysis, and the physician decided to explant the device. it was reported that the patient was hemodynamically stable after explanation, no further information is available. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the hemolysis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. the investigation for the positioning issue is in process and a supplemental report will be sent when the investigation is concluded. instructions for use for the related event are as follows: hemolysis ihemolysis impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17462 | 1220648-2024-20378 | 20447388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447388&pc=OZD | 05/02/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 05/03/2023 | 10/14/2024 |  |  | 12/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325707 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a pump stop occurred along with purge system blocked. the cp was explanted and not replaced. there was no patient harm. | [5 empty] | The investigation for the reported pump stop and high purge pressure issues has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that blood was observed in the purge gap. the pump was plugged into a console and purged, and high purge pressure reproduced. the pump would not start. the catheter was opened to check the purge line inside and blood ingress was noted. the catheter was cut around the area of the kink and the purge line was examined, and the purge line was found to be kinked at the location of the catheter kink. the cause of the pump stop was a kinked purge line within the catheter next to the kink protector. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17463 | 1220648-2024-20379 | 20447465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447465&pc=OZD | 06/13/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 06/13/2023 | 10/14/2024 |  |  | 12/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024310977 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that during delivery the catheter was bent. no further information is available. | [5 empty] | The investigation for the reported product damage (cannula kink) and delivery issues has been completed. the impella device has been returned for evaluation. per the results of the clinical review and investigation: a visual inspection found that the catheter was kinked below the end cap. no other issues were found on the rest of the pump. there was no kink found on the pump outside the box and before delivering. the root cause of delivery issue was most likely due to use issue. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17464 | 1220648-2024-20380 | 20447485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447485&pc=OZD | 06/17/2023 | 03/27/2025 |  | 10/14/2024 |  |  | 06/17/2023 | 10/14/2024 |  |  | 02/28/2025 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338469 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | Positioning Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the md had difficulty advancing the impella catheter in the correct position because the catheter appeared to be kinked. some oozing was also noted at the insertion site, so an additional purse string suture was placed with successful resolution of the tract ooze. the decision was made to remove the pump and insert a new impella catheter. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a kink was found on the catheter near the 27cm mark. per the investigation and clinical details provided, the root cause of the positioning issue was a cath anchor use issue related to catheter kink. this report is being filed as part of a retrospective review of historical records.
17465 | 1220648-2024-20381 | 20447561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447561&pc=OZD | 02/15/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 02/15/2023 | 10/14/2024 |  |  | 09/30/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023282914 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention (pci) needing an impella cp device for mechanical circulatory support. the impella was placed across the ventricle, wire removed and support initiated. initial flows were around 3l min and then down to just above 2l min. the patient started to experience arrhythmia and the impella was pulled back. attempts to reposition caused arrhythmias when pushed forward into the left ventricle, but decoupled waveforms when pulled back. a decision was made to continue with pci with current flows and stable heart rhythm. the impella was weaned and the device explanted at the end of procedure with the patient hemodynamically stable. hemostasis was achieved with two perclose. | [5 empty] | The data logs were reviewed and confirm a reduction in motor current and flow. the pump was returned and no abnormalities were found. the pump was run in a bucket of water and adequate flows were produced. the root cause of the low flows and arrhythmias were determined to be patient condition. the pump could not be properly positioned due to the patient's small left ventricle and short ascending aorta. without optimal positioning the pump was unable to produce the specified amount of flow per the given p-level. this report is being filed as part of a retrospective review of historical records.
17466 | 1220648-2024-20382 | 20447599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447599&pc=OZD | 08/23/2022 | 10/07/2025 |  | 10/14/2024 |  |  | 08/23/2022 | 10/14/2024 |  |  | 01/31/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174116 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the devices motor current went flat. the physician explanted the device and replaced it with another impella cp. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | G1 reporting contact email revised on manufacturer device report 1220648-2024-20382 in accordance with updated procedures. h6 investigation findings and investigation conclusions codes was erroneously entered on the initial manufacturer device report number 1220648-2024-20382.
17467 | 1220648-2024-20383 | 20447612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447612&pc=OZD | 05/12/2022 | 10/14/2024 |  | 10/14/2024 |  |  | 05/12/2022 | 10/14/2024 |  |  | 02/29/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197504 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported during preparation of the impella for placement the device generated an open purge alarm priming issue, which troubleshooting did not resolve. the physician made the decision to replace this impella with a new impella to complete the balloon aortic valvuloplasty procedure. there was no patient harm reported, and this device was replaced and the new impella placed without any noted issues. | [5 empty] | The investigation into the priming issue has been completed. the impella pump was returned for investigation. a visual inspection revealed no external damage or abnormalities found on the returned pump and purge cassette. the "purge system open" alarm triggered initially when running the pump with the purge cassette, but the alarm resolved after a few minutes. all metrics were within specification. the cause of the issue was use related as a result of improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
17468 | 1220648-2024-20384 | 20447615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447615&pc=OZD | 02/24/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 02/24/2023 | 10/14/2024 |  |  | 11/30/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - CANADA |  |  | ABIOMED, INC. |  |  | US |  | 2024285342 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the complainant reported that the patient experienced bilateral ischemia during impella support. the patient was transferred to icu west in stable condition. no intervention was required, and the condition improved as support continued. patient was successfully weaned off impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was most likely patient condition related as ischemia occurred in both legs. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17469 | 1220648-2024-20385 | 20447619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447619&pc=OZD | 03/06/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 03/06/2023 | 10/14/2024 |  |  | 09/30/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023279856 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were low flow issues involving an impella pump. following confirmed placement of the pump, it was documented that continuous suction alarms were being triggered. as a result of the ongoing failure, the decision was made to remove the device and replace it with another pump. there was no reported harm to the patient. | [5 empty] | The investigation for the reported low flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. no product was returned for investigation. the device log confirms the reported suction alarms. the log also shows instance of positioning related alarms including impella position unknown, impella position in aorta, and impella position in ventricle. the cause of the low pump flow was most likely biomaterial based on the provided clinical information. this report is being filed as part of a retrospective review of historical records.
17470 | 1220648-2024-20386 | 20447625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447625&pc=OZD | 01/11/2023 | 10/14/2024 |  | 10/14/2024 |  |  | 01/11/2023 | 10/14/2024 |  |  | 09/30/2024 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023277407 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding from the access site during impella support. it was noted by the physician that the patient's obesity was preventing the repo sheath from sealing the arteriotomy. as a result, the decision was made to replace the cp pump with a 5.5 pump and to leave the peel away in place to tamponade the groin site. the patient survived the procedure. | [5 empty] | The investigation into the access site bleeding has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the access site bleeding was patient condition related because after multiple repositioning attempts in the icu that repo sheath no longer sealing arteriotomy due to patient's obesity. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17471 | 1220648-2024-20387 | 20447628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20447628&pc=OZD | 10/28/2021 | 03/10/2025 |  | 10/14/2024 |  |  | 03/16/2022 | 10/14/2024 |  |  | 07/31/2023 |  | 10/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110763 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support, the patient experienced hemolysis and thrombus. medical intervention is unknown. the patient also experienced access site bleeding and received a blood transfusion. the patient continued on support until the device was successfully weaned. | [5 empty] | Hemolysis: the investigation for the reported hemolysis issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the hemolysis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. thrombocytopenia and access site bleeding - major. the impella device investigation for thrombocytopenia and access site bleeding is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smart assist system. access site bleeding. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ thrombocytopenia: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17472 | 1220648-2024-20278 | 20435549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435549&pc=OZD | 11/18/2022 | 10/13/2024 |  | 10/13/2024 |  |  | 11/18/2022 | 10/13/2024 |  |  | 08/31/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250629 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. venoarterial extracorporeal membrane oxygenation (va-ecmo) implanted with the impella added for decompression. while on support, there was a leaking purge cassette that was resolved by removing/replacing the cassette. approximately 4 days later, a purge system open alarms emitted, and a damaged purge side arm was noted. a bypass was configured, and support continued. heparin induced thrombocytopenia was confirmed. it was reported that the pump stopped the following day. the patient was successfully weaned and the pump was removed. | [5 empty] | The investigation for the reported purge leak, pump stop, and thrombocytopenia has been completed. the impella device was not received from the customer; however, the pump side arm and a purge cassette were returned. upon visual inspection, a leak was visible at the yellow luer connection and a crack was observed. the data logs confirmed the purge bypass caused an increase in the purge pressure and motor current, allowing blood to enter the purge gap which caused the pump to stop. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17473 | 1220648-2024-20279 | 20435550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435550&pc=OZD | 05/30/2023 | 03/10/2025 |  | 10/13/2024 |  |  | 05/30/2023 | 10/13/2024 |  |  | 11/30/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300699 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the purge sidearm cracked upon moving the patient during impella support. a purge open alarm was triggered, and the pump subsequently stopped. the device was explanted as a result of the reported issue. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump stop and purge leak ¿ pump issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the device passed all post-sterile inspection checks. the data log analysis shows the pump running without issue until the third day of support, when suddenly the purge pressure drops to ~0 mmhg with purge system open alarms. shortly after, motor current begins to rise, following ultimately by a high motor current pump stop. the product analysis corroborates this story. a large crack in the sidearm filter to reservoir joint was seen was seen along with blood in the purge line. based on the clinical account that a crack was noted as the nurse turned the patient. the root cause of the pump stop is due to the broken sidearm filter to reservoir joint as a result of excessive force. the root cause of the purge leak - pump could not be determined as insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
17474 | 1220648-2024-20280 | 20436045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436045&pc=OZD | 03/03/2022 | 10/20/2025 |  | 10/13/2024 |  |  | 03/02/2022 | 10/13/2024 |  |  | 12/31/2023 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162464 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, it was reported there was access site bleeding, and the patient developed hemolysis, no further information was provided regarding treatment. it was also reported several times during the night the "air in purge system" alarm was generated and the procedure for removing air was being carried out. additionally, in the morning the purge pressure, which had been flowing at about 13 millimeters an hour (ml / h), suddenly decreased to 4 ml / h or less. this impella was replaced due to the purge liquid decrease resulting in purge pressure issue. the impella was successfully replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
17475 | 1220648-2024-20283 | 20436046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436046&pc=OZD | 09/30/2022 | 10/13/2024 |  | 10/13/2024 |  |  | 09/30/2022 | 10/13/2024 |  |  | 06/30/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235868 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, pump had low flows and oozing occurred at access site. blood was given and attempts to reposition the device were made. the pump was removed and replaced with no issues. the site was sutured down with touchy. the patient was delivered to intensive care unit (icu) in stable condition on a heparin drip. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17476 | 1220648-2024-20282 | 20436047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436047&pc=OZD | 09/20/2022 | 10/13/2024 |  | 10/13/2024 |  |  | 09/20/2022 | 10/13/2024 |  |  | 03/31/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213436 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. once impella was advanced into the ventricle, pump abruptly stopped. pump was unable to be restarted. the team found a clot burden within the pump after the pump stopped. the team replaced the pump, and the second pump worked appropriately. no patient harm was reported. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the lt log confirms high motor current pump stops. the rt log shows motor current spikes. based on the provided clinical information, the cause of the pump stop was most likely biomaterial. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17477 | 1220648-2024-20284 | 20436048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436048&pc=OZD | 10/10/2022 | 10/13/2024 |  | 10/13/2024 |  |  | 10/10/2022 | 10/13/2024 |  |  | 04/30/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252973 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, impella cp experienced low flows with peak flow of 1.8 lpm and sudden droppage within 1 hour of implant. the impella cp device was explanted. no patient harm or injury noted. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause of the issue was biomaterial. this report is being filed as part of a retrospective review of historical records.
17478 | 1220648-2024-20285 | 20436049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436049&pc=OZD | 10/21/2022 | 10/13/2024 |  | 10/13/2024 |  |  | 10/21/2022 | 10/13/2024 |  |  | 06/30/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248179 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during removal of the peel away sheath, the pump migrated into the mitral apparatus causing issues with the inflow volume from the left artery. as the device could not be moved back into position without a wire, the decision was made to explant the initial impella device and implant a new pump. the patient survived the procedure. | [5 empty] | The investigation into the positioning issues has been completed. the impella device was not returned for evaluation. the most likely root cause of positioning issues was wireless re-access. the impella device is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
17479 | 1220648-2024-20286 | 20436050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436050&pc=OZD | 10/21/2022 | 10/13/2024 |  | 10/13/2024 |  |  | 10/21/2022 | 10/13/2024 |  |  | 07/31/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238517 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella cp device for mechanical circulatory support. during initial pump delivery, the impella cannula kinked, which led to low pump flow when running the pump due to patient's small aortic arch. the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation for the reported product damage (cannula kink) and low pump flow has been completed. the product was returned, and the low pump flow was confirmed. the cause of the low pump flow was a kinked cannula stemming from the interaction between the device and the patient's small aortic arch during delivery. this report is being filed as part of a retrospective review of historical records.
17480 | 1220648-2024-20287 | 20436051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436051&pc=OZD | 07/22/2021 | 03/27/2025 |  | 10/13/2024 |  |  | 07/22/2021 | 10/13/2024 |  |  | 05/31/2023 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022084096 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  |  | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, pump stopped on day eight (8) during the weaning. in addition, during the support the patient developed bilateral hematomas at the lower limbs. hemoglobin dropped. activated clotting time, units of seconds (measure of blood¿s ability to clot) was at 193 and blood was given. the patient was successfully weaned off of support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | The pump was returned and a large purge gap was found indicating bearing wear. the pump was torn down and imaged. the root cause of the pump stop was determined to be bearing wear. this complaint has been identified as part of a retrospective review.
17481 | 1220648-2024-20288 | 20436052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436052&pc=OZD | 04/05/2022 | 05/19/2025 |  | 10/13/2024 |  |  | 04/05/2022 | 10/13/2024 |  |  | 05/31/2023 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2022089376 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, patient had hemolysis, renal failure and access site bleeding. however, no additional treatment details were provided regarding the treatment or management of this condition. weaning was successful and the procedural outcome was that the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ renal failure: impella cp with smartassist. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issues has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
17482 | 1220648-2024-20289 | 20436134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436134&pc=OZD | 01/29/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 01/29/2023 | 10/13/2024 |  |  | 09/30/2022 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278547 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the physician decided to utilize extra-corporeal membrane oxygenation (ecmo) in conjunction with the impella cp. it was reported that once the ecmo was implanted the impella pump came out of position and was kinked. the impella cp was removed and replaced with another impella cp. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the positioning issue was an interaction with ecmo cannulation. this report is being filed as part of a retrospective review of historical records.
17483 | 1220648-2024-20290 | 20436136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436136&pc=OZD | 11/13/2021 | 10/02/2025 |  | 10/13/2024 |  |  | 11/13/2021 | 10/13/2024 |  |  | 02/28/2023 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022064785 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the clinical team reported that there was a hematoma at the iabp access site before the insertion of the impella. the physician stated that the hematoma was caused by vascular damage from the use of the iabp. the bleeding was treated with surgical repair of the vasculature and removal of the hematoma; also, a blood transfusion was administered. it was also noted that the patient experienced limb ischemia while on support. the device was surgically removed and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The cause of the hematoma and limb ischemia found to be unrelated to the device and the relation to impella was determined to be unrelated. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
17484 | 1220648-2024-20292 | 20436137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436137&pc=OZD | 04/06/2022 | 10/13/2024 |  | 10/13/2024 |  |  | 04/06/2022 | 10/13/2024 |  |  | 01/31/2022 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1465559 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis and renal failure. the resolution for these issues is unknown, it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported hemolysis and renal failure has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. instructions for use for the related event are as follows: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
17485 | 1220648-2024-20291 | 20436138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436138&pc=OZD | 03/12/2022 | 10/09/2025 |  | 10/13/2024 |  |  | 03/12/2022 | 10/13/2024 |  |  | 08/31/2023 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022125926 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | A |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a left ventricle thrombus on day 17 of impella support. the patient was given systemic heparin and observed. it was also reported that a pump stop occurred on day 23 of support and the pump was explanted as a result. the patient was stable following explant of the impella. | The complainant reported an unresolved placement signal failure. support proceeded until an eventual pump stop, approximately 8 days later. a thrombus in lv was also reported. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the optical signal issue was unable to be determined as no product or logs were returned for evaluation. the root cause of the pump stop was unable to be determined as no product or logs were returned for evaluation. the placement signal failure and pump stop occurred beyond the indicated design life of 8 days per design trace matrix 0046-9910. thrombus: added during complaint remediation. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined
17486 | 1220648-2024-20293 | 20436426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436426&pc=OZD | 09/16/2024 | 11/27/2024 |  | 10/13/2024 |  |  | 09/16/2024 | 10/13/2024 |  |  |  |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528619 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported an impella cp was unable to be inserted due to patient anatomy. the procedure was aborted. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. h6 code 4627 was reported incorrectly on manufacturer device report 1220648-2024-20293. new code was added to component code
17487 | 1220648-2024-20294 | 20436434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436434&pc=OZD | 09/16/2024 | 11/07/2024 |  | 10/13/2024 |  |  | 09/16/2024 | 10/13/2024 |  |  |  |  | 10/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1155471 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a controller failure red alarm at the start up. the aic was replaced. the patient survived the explant. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the console (aic) controller failure has been completed. data logs were reviewed which showed that when the console was booted up, controller failure alarm #418 occurred instantly for defective purge pressure sensor. the console was immediately shut down by the user. the console was returned for investigation but the failure was unable to be reproduced both with or without purge cassette. the cause of the controller failure was not determined since the malfunction was not reproduced.
17488 | 1220648-2024-20295 | 20436441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436441&pc=OZD | 09/16/2024 | 11/07/2024 |  | 10/13/2024 |  |  | 09/16/2024 | 10/13/2024 |  |  | 06/30/2025 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372225 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for support of an acute myocardial infarction/cardiogenic shock patient. the impella cp had persistent low flows that were not resolved until the impella cp was explanted and replaced/escalated to an impella 5.5 pump. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. no product was returned for evaluation. data logs were reviewed and at time of reported loss of pulsatility in motor current, a motor current spike is present with a slight deviation in motor speed and pump flow also looses some pulsatility. clinical details noted that repositioning pump did not resolve suction alarms. patient was also on extracorporeal membrane oxygenation (ecmo) support. the root cause of the low pump flow was most likely due to biomaterial ingestion. d10 concomitant medical products added.
17489 | 1220648-2024-20296 | 20436452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436452&pc=OZD | 09/16/2024 | 11/25/2024 |  | 10/13/2024 |  |  | 09/16/2024 | 10/13/2024 |  |  |  |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537300 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and there were continuous suction alarms. the staff gave fluids and checked the pumps position. the doctor was unable to reposition the pump due to the left ventricle configuration. the pump was explanted and the patient was supported by extracorporeal membrane oxygenation. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the positioning issues and low pump flows has been completed. product and data logs were not returned for review. the root cause of the low flow was most likely a patient condition. the root cause of the positioning issue was most likely a patient condition.
17490 | 1220648-2024-20297 | 20436456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436456&pc=OZD | 09/16/2024 | 11/15/2024 |  | 10/13/2024 |  |  | 09/16/2024 | 10/13/2024 |  |  |  |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532956 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 62-year-old male patient on impella cp support did not have dopplerable pulses prior to leaving the catheterization lab after impella cp procedure. on the same day of impella cp placement, the patient was taken to the catheterization lab for percutaneous coronary intervention (pci) procedure. while doing the pci procedure, the patient had to be defibrillated two times due to runs of ventricular fibrillations. the physician decided to explant the impella cp while in the catheterization lab and assess distal perfusion to the right leg. the physician ballooned the distal superficial femoral artery and the knee. perfusion was noted with dopplerable pulses. the patient was stable. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia and ventricular fibrillation (vf) has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and vf was not determined.
17491 | 1220648-2024-20298 | 20436470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436470&pc=OZD | 09/19/2024 | 11/21/2024 |  | 10/13/2024 |  |  | 09/19/2024 | 10/13/2024 |  |  |  |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025516472 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon impella cp sheath insertion for patient support, clotting was noted. heparin was administered in an attempt to counteract the condition but a clot was believed to have been ingested by the device. decoupling between the placement and left ventricle waveforms developed and coincided with low pump flows. the decision was ultimately made to replace the device with another impella cp pump. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to produring repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the low pump flow and the thrombus has been completed. data logs were investigated, and the low pump flows were confirmed. throughout the entire case, pump flows are below the specification for the corresponding p-level the pump is running at. there are no positioning alarms, there aren't any placement signal trends, and there aren't any motor current spikes or trends that don't align with p-level changes. the only abnormality with the placement signal is similar to the report that the aortic/lv waveforms are abnormal at times. data logs confirm that there are periods where the amplitude of the aortic waveform has an abnormally large amplitude, but no positioning alarms are triggered. there are also several suction alarms, the longest of which was active for over 20 minutes of the 38 minute case. this could be related to the noise seen in the ps waveforms. though it is reported that the team suspected a clot was ingested, there is little to no evidence of that in the data logs. that does not mean that biomaterial did not get ingested and play a role in the low pump flows, however, it does not appear to be the leading factor. clinical details provided that the patient's ef was 10% at case start. they also reported that the pump was replaced with a new one once explanted. the data logs for that second pump were investigated, and the low pump flows persisted through that entire case as well. product was not returned for investigation. the patient's low ef at case start, the intermittent noise in the placement signal, the suction alarms, and the fact that low flows persisted on the second pump support that the patient's condition was causing the low pump flows. based on clinical details provided and data logs, the root cause of the low pump flows was determined to most likely be patient condition. the root cause of the thrombus could not be determined without additional information.
17492 | 1220648-2024-20299 | 20436473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436473&pc=OZD | 05/09/2024 | 11/04/2024 |  | 10/13/2024 |  |  | 09/17/2024 | 10/13/2024 |  |  |  |  | 10/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1606415 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella was being prepped for an implant and there was an "impella catheter disconnected error". the pump was replaced. the investigating engineer found that the automated impella controller needed to be investigated as well for this error. no patient harm has been reported. | [5 empty] | A.5 is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the automated impella controller (aic) - pump issue has been completed. the console logs from the reported day of event are consistent with the complaint and showed three attempts to use the pump, which is then swapped for another pump, that operated without any issues for the duration of support (over two weeks). real time log data showed ¿placement signal not reliable ¿(psnr) alarm after first connection of the first pump. upon case start wizard prompt (optical sensor system error: disconnect cable) and reconnecting the pump (for a few seconds), there was still no placement signal. next, the pump was disconnected and reconnected again, after which placement signal was recovered. however, after approximately 50 minutes console logs showed that pump was disconnected, then replaced with another pump. placement signal was reliable and there were no unexpected disconnections of the pump. console was returned for investigation and was tested with respect to all reported failure modes. issue of unexpected ¿impella catheter disconnected error¿reported was identified in console logs, but not reproduced during testing of the console or the pump. the root cause was not determined, although the issue may have been caused by intermittent malfunction of impellatronic board, or - more likely - incomplete insertion of the pump plug into console¿s receptacle, which is also consistent with ¿psnr¿ condition. during console testing, it has been discovered that the unit¿s optical pump connector (blue receptacle) was causing the console to fail ¿5s¿ specification. this was unrelated to this complaint. the cause of the optical signal failure was an air gap from between the aic and the patient cable plug or within the middle of the red handle ferrule. h.6 added code 2917 for the optical signal failure addressed in the investigation results.
17493 | 1220648-2024-20300 | 20436478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436478&pc=OZD | 09/11/2024 | 11/21/2024 |  | 10/13/2024 |  |  | 09/18/2024 | 10/13/2024 |  |  |  |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532237 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Cardiogenic Shock | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp support ongoing for a patient with cardiomyopathy and cardiogenic shock. the patient was supported for eight days when it was explanted and replaced by bilateral centrimag. weeks later abiomed¿s long-term outcome and quality (loqi) indicator impella registry alleged the harm of cardiogenic shock due to a compromised sterile sleeve. no treatment was noted. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the sterile sleeve damage has been completed. pump was not returned for investigation. data log review was not required for this case. as per the given clinical information, the sterile sleeve was compromised, which poses a risk for infection and requires removal. the cause of the sterile sleeve damage issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 clinical code 2262.
17494 | 1220648-2024-20301 | 20436493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436493&pc=OZD | 04/19/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 04/19/2023 | 10/13/2024 |  |  | 01/31/2025 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024313535 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported during placement of the impella, before the physician removed the peel-away sheath, the impella became dislodged and was alarming that the impella was in the patient¿s aorta. reportedly when the physician advanced across the aortic valve the impella device kinked in the ventricle, although the kink was able to be corrected the physician made the decision to remove this impella. there was no patient harm reported, and this device was replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the positioning issue that resulted in product damage (kink) has been completed. the clinical review indicated the issues were due to improper positioning as a result of improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
17495 | 1220648-2024-20302 | 20436494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436494&pc=OZD | 04/21/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 04/21/2023 | 10/13/2024 |  |  | 01/31/2025 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024305770 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, demographics unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was placed, and the system generated an alarm indicating the device was in the wrong position in the aorta and wrong position in the ventricle. reportedly a fluoroscopy indicated the position was correct resulting in a transthoracic echocardiogram being performed which revealed a thrombus. the pump was removed, and thrombotic material was recovered from the inlet and in the cannula. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the pump thrombus has been completed. the clinical review indicated the root cause was most likely ingested biomaterial. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system: section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: as noted in the impella ifu ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ as noted in the impella ifu "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17496 | 1220648-2024-20303 | 20436495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436495&pc=OZD | 04/21/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 04/21/2023 | 10/13/2024 |  |  | 12/31/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321436 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, age unknown, noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that a long sheath was attempted but kept kinking due to the patient being obese and the femoral access site being difficult due to adipose tissue. reportedly the team then attempted placement using a short sheath and then again with a long sheath, however, it kinked in the same place multiple times and would not advance. this impella was removed and a new impella was successfully placed. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the product damage (kink) has been completed. the clinical review indicated an obese patient with difficult femoral access due to adipose tissue resulted in kinking due to multiple failed attempts to place the impella . the clinical review indicated the root cause was due to the patient anatomy. this report is being filed as part of a retrospective review of historical records.
17497 | 1220648-2024-20304 | 20436497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436497&pc=OZD | 05/26/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 05/02/2023 | 10/13/2024 |  |  | 12/31/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024300892 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported there was a placement signal sensor error when starting impella. additionally, it was reported that on (b)(6) 2023 the patient suffered a cerebral hemorrhage one day after the impella had been removed, no further information was available regarding the stroke. the patient remained on impella support and was successfully weaned. | [5 empty] | The product was returned and the investigation into the stroke and optical signal issues have been completed. the cause of the stroke was not determined due to insufficient clinical details. the relationship to the impella is unknown. the cause of the optical signal issues was most likely patient condition related since the divergence with the a line resolved naturally. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: cva-ischemic/hemorrhagic impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: as noted in the impella ifu "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17498 | 1220648-2024-20305 | 20436499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436499&pc=OZD | 05/01/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 05/01/2023 | 10/13/2024 |  |  | 01/31/2025 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311080 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during weaning of the impella and working toward recovery, the night nurse was bathing the patient and received an impella in aorta alarm. reportedly the patient then had flash pulmonary edema requiring intubation, an echocardiogram confirmed the impella was in the aorta. the physician was unable to cross the aortic valve resulting in the patient being taken to catheterization laboratory for replacement of the impella cp. there was no patient harm was reported, this impella was removed and a new impella placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the positioning issue has been completed. a review of clinical information determined the root cause of the positioning issue was patient movement due to improper use technique. this report is being filed as part of a retrospective review of historical records.
17499 | 1220648-2024-20307 | 20436500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436500&pc=OZD | 05/30/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 05/30/2023 | 10/13/2024 |  |  | 11/30/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303080 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump experienced low flows, the team explanted the pump and found clot. it was reported that the patient survived the procedure. | [5 empty] | The device was returned and evaluated, and the investigation for the reported low pump flow and thrombus has been completed. there was a large amount of dried biomaterial observed on the inlet, and biomaterial observed around the impella. no other abnormality was found. the root cause of the low pump flow issue was ingesting biomaterial due to patient condition related as the patient had been off any anticoagulation for (48 hours) due to low platelets. as noted in the impella ifu: thrombus: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17500 | 1220648-2024-20306 | 20436501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436501&pc=OZD | 05/16/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 05/16/2023 | 10/13/2024 |  |  | 01/31/2025 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-EU |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024315372 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, race and ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the pump did not start after several attempts, reportedly even disconnection of the pump did not work. an additional pump was available and was successfully used. there was no patient harm reported, and this device was replaced. | [5 empty] | The investigation into the pump did not start issue has been completed. the impella pump was returned for investigation. the data logs showed the pump was inserted and purged without issue. upon impella start, a motor current spike to 1500 milliamp was seen. a visual inspection of the returned device found no visual, mechanical, or electrical abnormalities with the pump. the pump ran as expected in the lab at pressure level nine with 3.8 liters per minute of flow. as the pump ran without issue in the lab, the exact cause of the pump stop cannot be determined as it could not be reproduced but may be due to patient interaction. the cause of the issue was not determined since the pump did not start issue could not be reproduced. as noted in the impella ifu: impella cp with smart assist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
17501 | 1220648-2024-20308 | 20436502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436502&pc=OZD | 06/25/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 06/25/2023 | 10/13/2024 |  |  | 01/31/2025 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338662 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient had a thrombosed left anterior descending artery stent prior to the impella placement, the plan was to perform a percutaneous coronary intervention and thrombectomy. the procedure resulted in the impella being successfully exchanged out for an intra-aortic balloon pump due to the mural left ventricular thrombus. the patient remained on extra-corporeal membrane oxygenation and was reported to have improved with an ejection fraction of 20 to 25%. this device was replaced due to the patients preexisting thrombosed stent. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the thrombus has been completed. per review of the clinical details provided, the patient had a pre-existing mural left ventricle thrombus condition and impella was explanted. the cause of the thrombus issue was patient condition related due to pre-existing condition of mural thrombus per clinical review. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smart assist system: section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: as noted in the impella ifu ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ as noted in the impella ifu "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17502 | 1220648-2024-20309 | 20436504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436504&pc=OZD | 07/12/2023 | 05/23/2025 |  | 10/13/2024 |  |  | 07/12/2023 | 10/13/2024 |  |  | 01/31/2025 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331438 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the pump was successfully placed; however, the pump did not stay on. reportedly the pump was turned on and off several times and extensive troubleshooting was performed without resolution. additionally, it was reported that repositioning took place in the ventricle and pump still did not work properly. this pump was replaced, and a new pump was placed. there was no patient harm reported. | [5 empty] | The investigation for the reported low flow/pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the low flow/pump stop and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi & impella rp with smart assist system section: using the automated impella controller with the impella catheter as noted in the impella ifu ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smart assist system catheter as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smart assist & impella cp with smart assist for use during cardiogenic shock section: troubleshooting the purge system as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped as noted in the impella ifu ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | No product was returned. as per clinical, there were low flows initially with suction alarm and pump was turned on and off several times as an extensive troubleshooting, but the pump would not turn on. pump was replaced successfully with no issues. data logs were reviewed, and suction alarms were observed with low flow issue. no abnormal pump stop alarms were observed in the logs. the cause of the low pump flow issue was unable to be determined since product was not returned and no conclusive evidence of abnormalities based on the data log and clinical review. this report is being filed as part of a retrospective review of historical records.
17503 | 1220648-2024-20310 | 20436505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436505&pc=OZD | 09/11/2021 | 10/13/2024 |  | 10/13/2024 |  |  | 09/11/2021 | 10/13/2024 |  |  | 06/30/2023 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100435 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, race and ethnicity unknown, noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the pump failed to purge and was not placed in the patient. reportedly, according to the hospital staff, the console displayed "purge system open" alarm despite verifying appropriate connections through purge line, replacing purge cassette, de-airing line, disconnecting/reconnecting white console cable, and restarting the controller. this device was not placed and a new impella was successfully placed. there was no patient harm reported. | [5 empty] | The investigation into the purge leak pump issue has been completed. the impella pump was returned for investigation. bench testing was performed on the returned purge cassette and purged without issue or alarms. the logs were returned, and no open purge system alarms were found. a piston blocked alarm occurred in the logs before the impella cp was plugged into the console. this alarm occurred throughout case start with the pump and only cleared when the cassette was removed. the controller was restarted, and the cassette was plugged in and the pump went through case start without any issues the second time. the pump and cassette were still removed after case start and the site reported that the issue was not fixed. the logs show no alarms during the second case start. the console was then restarted for a third time and the cassette and pump were replaced. the root cause of the purge system open alarms was most likely a use related issue during case start as the issue was unable to be reproduced and the logs do not show any alarms that correlate to the purge system open alarm mentioned in the complaint. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi section: cleaning as noted in the impella ifu ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
17504 | 1220648-2024-20311 | 20436506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436506&pc=OZD | 08/29/2021 | 10/13/2024 |  | 10/13/2024 |  |  | 09/14/2021 | 10/13/2024 |  |  | 04/30/2023 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022079769 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Cardiac Tamponade | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding, ventricular tachycardia and cardiac tamponade of unknown causality. there was no further information available regarding the access site bleeding, ventricular tachycardia or cardiac tamponade. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for the reported cardiac tamponade and access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the cardiac tamponade and access site bleeding and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: access site bleeding: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17505 | 1220648-2024-20312 | 20436507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436507&pc=OZD | 07/16/2022 | 10/13/2024 |  | 10/13/2024 |  |  | 07/16/2022 | 10/13/2024 |  |  | 02/29/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202877 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported a failure to prime the device. as a result, the decision was made to place a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the priming issue and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
17506 | 1220648-2024-20313 | 20436532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436532&pc=OZD | 04/07/2022 | 05/23/2025 |  | 10/13/2024 |  |  | 04/07/2022 | 10/13/2024 |  |  | 08/31/2023 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114747 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the patient experienced hemolysis and renal failure during use of the device. no interventions nor premature removal were reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle: ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis: ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17507 | 1220648-2024-20314 | 20436533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436533&pc=OZD | 05/12/2022 | 10/13/2024 |  | 10/13/2024 |  |  | 05/12/2022 | 10/13/2024 |  |  | 02/29/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198960 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support to replace the patient's intra-aortic balloon pump. it was reported after removing the oscor sheath, the repositioning sheath was inserted. it was then observed that the impella was looped or kinked in the access site after the repositioning sheath was inserted. additionally, the impella access site began to bleed due to the loop. the medical team decide to remove the impella cp and place a new device. hemostasis was achieved after the new impella 5.5 device was implanted and the patient was transfused with blood products as a result of the blood loss. the patient was successfully placed on the new device. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist system: section: warnings & cautions: warnings section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17508 | 1220648-2024-20315 | 20436534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436534&pc=OZD | 03/06/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 03/06/2023 | 10/13/2024 |  |  | 09/30/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023280329 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the pump experienced saturated flows and pump stop. as a result, the pump was explanted and the patient remained on venoarterial extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17509 | 1220648-2024-20316 | 20436539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436539&pc=OZD | 06/03/2022 | 02/27/2025 |  | 10/13/2024 |  |  | 06/03/2022 | 10/13/2024 |  |  | 02/29/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203547 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced oozing at the impella access site. the site was dermabonded and the patient was transfused with one unit of packed red blood cells. it was also reported that the purge fluid leaked from the yellow luer of the sidearm 7 days into impella support and 2 days into using sodium bicarbonate in the purge solution. the leaking was accompanied by "purge system open" alarms. the issue was resolved with a purge bypass. it was then reported that the pump stopped suddenly after 17 days of support. the pump was unable to be restarted, so the device was explanted. the patient remained on extracorporeal membrane oxygenation. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issues has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed the "purge system open" alarm occurred immediately upon a sudden drop in purge pressure about 7 days into support. ten days later, the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm and could not be restarted. a purge bypass was in place on the returned pump. the purge sidearm was not returned. additionally, an enlarged impeller purge gap was observed. after reviewing the clinical information, it was determined that the cause of the purge leak was sidearm yellow luer damage due to the sodium bicarbonate in the purge solution. the cause of the pump stop was bearing wear after 17 days of use, which is beyond design life per design trace matrix 0046-9910. this report is being filed as part of a retrospective review of historical records.
17510 | 1220648-2024-20317 | 20436540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436540&pc=OZD | 06/07/2022 | 02/21/2025 |  | 10/13/2024 |  |  | 06/07/2022 | 10/13/2024 |  |  | 08/31/2023 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119996 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Stroke/CVA | Thrombocytopenia | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. thrombocytopenia occurred, it was reported that the patient recovered. also, the suction alarm was ringing, however there was no problem with the device's position on the echocardiogram (echo). the pump was removed due to frequent suction alarm. additionally, the patient experienced a cerebral infarction. it occurred two days after impella removal. it was reported that the patient was stable after the cerebral infarction. no further information is available. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | Data logs showed that the motor current and pump flow sharply dropped and became non-pulsatile upon the sudden occurrence of the "suction" alarm. the motor spiked prior to the sharp decrease. both metrics remained this way for the remainder of support. there were no other abnormal pump metrics at this time. no damage or abnormalities were found on the returned pump. the pump ran with no issues. in conclusion, it was determined that the cause of the low pump flow was most likely ingested biomaterial. the stroke/cva was determined to be unrelated to the impella since it occurred 2 days after explant. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation for optical signal issue and thrombocytopenia is not determined.
17511 | 1220648-2024-20318 | 20436544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436544&pc=OZD | 07/18/2022 | 03/27/2025 |  | 10/13/2024 |  |  | 07/18/2022 | 10/13/2024 |  |  | 03/31/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191710 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that an impella cp was positioned too deep within the left ventricle (lv), which prompted suction and mc not pulsatile. when the pump was repositioned, the team observed mc but, the ps and lv were unreliable. there were suction alarms. the physician declined to replace pump. when the pump was explanted, there was visible coagel (inlet and outlet). | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning, optical signal, and low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the device passed all post-sterile inspection checks. the evaluation of the device noted that the returned pump was visually inspected, and biomaterial was observed on the sensor. the pump was connected to an aic and a placement signal not reliable alarm reproduced. the pump was soaked in tergazyme to remove the biomaterial and placement signal returned to normal at 7/7 mmhg. the cause of the positioning issue was patient condition related as repositioning resolved the issue. the cause of the optical sensor issue was biomaterial. no heparin in the purge since the patient was hit. due to the limited clinical information and lack of available data/product the root cause of the low pump flow is not determined. this report is being filed as part of a retrospective review of historical records.
17512 | 1220648-2024-20319 | 20436545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436545&pc=OZD | 05/29/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 05/29/2023 | 10/13/2024 |  |  | 02/28/2025 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335168 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited saturated flow and high purge pressure issues. the physician decided to explant the pump and replace with another impella cp. no further information was provided. | [5 empty] | The investigation for the reported saturated flow and high purge pressure has been completed. the product was returned. per the results of the clinical review and investigation: the returned pump was visually inspected, and biomaterial was observed in the purge gap. the cause of the issue was biomaterial in the purge gap. this report is being filed as part of a retrospective review of historical records.
17513 | 1220648-2024-20320 | 20436558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436558&pc=OZD | 12/02/2022 | 02/28/2025 |  | 10/13/2024 |  |  | 12/02/2022 | 10/13/2024 |  |  | 10/31/2023 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023168729 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a hematoma was noted at the groin access site. a fem-stop was placed, and heparin administration was paused to resolve this issue. it was stated that damage to the purge side arm was observed. a filter bypass was recommended as no additional impella pumps were available on site. it was reported that the pump stopped and was unable to be restarted despite best efforts. the device was explanted. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump stop and access site bleeding issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the cause of the pump stop was damage to the reservoir to filter joint due to excessive force. the cause of the access site bleeding was high patient act values since switching the purge from heparin to bicarb reduced ptt from 95 to 30 and resolved the bleeding. this report is being filed as part of a retrospective review of historical records.
17514 | 1220648-2024-20321 | 20436563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436563&pc=OZD | 01/27/2023 | 10/20/2025 |  | 10/13/2024 |  |  | 01/27/2023 | 10/13/2024 |  |  | 10/31/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285390 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support there were continuous suction alarms. the pump was repositioned in an attempt to troubleshoot the issue, however the alarms continued. as a result of the ongoing failure, the decision was made to explant the pump. upon explant, it was noted that there was a large thrombus within the inlet. a new pump was implanted and there was no patient harm reported. | [5 empty] | The impella device was returned for evaluation. no foreign or broken blades found. large biomaterial (thrombus) noted on inlet. cannula kink observed too. based on the clinical details no use issue confirmed. root cause most likely was the biomaterial due to patient condition. the kink may occurred while tried resolve the suction alarm or maybe while device removed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records.
17515 | 1220648-2024-20322 | 20436571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436571&pc=OZD | 01/29/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 01/29/2023 | 10/13/2024 |  |  | 09/30/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276868 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a high purge pressure followed by a pump stop during impella support. leading up to the pump stop, it was noted that the purge flow had begun to decrease and that the purge pressure and motor current were increasing. the pump was removed and replaced as a result of the failure. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the imeplla ifu impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17516 | 1220648-2024-20323 | 20436576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436576&pc=OZD | 02/08/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 02/08/2023 | 10/13/2024 |  |  | 10/31/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024289975 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support there was a low flow issue coinciding with a positioning issue. upon noticing that the pump was deflecting towards the mitral valve, troubleshooting was attempted, however the physician was unable to position the pump correctly. the pump subsequently shifted out of the ventricle into the aorta and the pump began to experience suction and low pump flow issues. due to the ongoing flow issues, the decision was made to explant the pump. it was also noted that the impella sheath was swapped and leg was ischemic. fem to fem bypass was performed with good flow. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. instructions for use for the related event are as follows: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17517 | 1220648-2024-20324 | 20436579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436579&pc=OZD | 02/10/2023 | 10/13/2024 |  | 10/13/2024 |  |  | 02/10/2023 | 10/13/2024 |  |  | 10/31/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024291771 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a purge filter cracked during cassette change. a purge system bypass was performed. the pump's motor current subsequently rose and the pump stopped. attempts to restart were unsuccessful, therefore, the pump was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
17518 | 1220648-2024-20325 | 20436581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20436581&pc=OZD | 02/10/2023 | 10/17/2025 |  | 10/13/2024 |  |  | 02/13/2023 | 10/13/2024 |  |  | 06/30/2024 |  | 10/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231904 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 74 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support the patient had hematuria and anemia. an echocardiogram was completed and the tip of the impella was noted to be facing the posterior wall. repositioning was attempted and a follow up ct scan was performed which identified a microvascular hemorrhage in an unrelated vessel. as a result of the hemorrhage, the patient was administered a blood transfusion. the patient was successfully weaned and the pump was explanted as a result of the unrelated hemorrhage. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | B2 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-20325. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-20325. g1 reporting contact email revised on manufacturer device report 1220648-2024-20325 in accordance with updated procedures. g2 report source; foreign was missing from manufacturer device report 1220648-2024-20325.
17519 | 1220648-2024-20221 | 20435001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435001&pc=OZD | 10/28/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 10/28/2021 | 10/12/2024 |  |  | 06/30/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100636 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to being implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp pump could not be fully advanced into the patient because of challenging anatomy and tortuosity that kinked; therefore, the insertion was aborted. there was no reported harm or injury to the patient. | [5 empty] | Visual inspection of the returned pump revealed significant kinking on the catheter between the 25-cm and 35-cm marks. no other damage or abnormalities were found. after reviewing the clinical information, it was determined that the cause of the delivery issues was patient condition, specifically the patient's severely tortuous vasculature. see es2020-080 for more information. this report is being filed as part of a retrospective review of historical records.
17520 | 1220648-2024-20222 | 20435006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435006&pc=OZD | 11/01/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/01/2021 | 10/12/2024 |  |  | 05/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092350 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the pump positioning was not optimal at the mitral. as a result, the decision was made to explant the device and replace it with a new impella cp device. there was no reported harm or injury to the patient. | [5 empty] | The investigation into the positioning issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the most likely root cause of positioning issues is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
17521 | 1220648-2024-20223 | 20435007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435007&pc=OZD | 11/11/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/15/2021 | 10/12/2024 |  |  | 12/31/2022 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022034955 |  | 0048-0034 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. it was reported that a thrombus was observed, expected to dissolve with heparin. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of thrombus. the device was not received for evaluation. the root cause of the thrombus was biomaterial and its relation to the impella device was unknown. this report is being filed as part of a retrospective review of historical records.
17522 | 1220648-2024-20224 | 20435010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435010&pc=OZD | 11/13/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/13/2021 | 10/12/2024 |  |  | 03/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069532 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped and was able to be restarted. the motor current trends were high. the impella cp was explanted and there was no report of patient harm. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the pump had bearing wear. analysis of the provided data logs confirmed a pump stop due to high motor current on 13nov2021 at 19:41 after over 9 days of use. of note, there were several elevations in purge pressure corresponding with decreases in purge flow throughout the case. the root cause of the pump stop was bearing wear. the pump stop occurred beyond the 8 day impella cp indication for use per design trace matrix 0046-9910. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17523 | 1220648-2024-20225 | 20435011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435011&pc=OZD | 11/14/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/14/2021 | 10/12/2024 |  |  | 05/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089324 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Ischemia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient's arm has been ischemic and not improving, with a possibility of amputated. the pump stopped after 11 days on support and was explanted and not replaced. the patient survived the procedure. | [5 empty] | The investigation for the reported limb ischemia and pump stop issues has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs were not sent. visual inspection of the returned pump revealed a significantly large impeller purge gap. no other relevant abnormalities were found. therefore, it was determined that the cause of the pump stop was bearing wear after 10 days of use, which is beyond design life per design trace matrix 0046-9910. after reviewing the clinical information, the cause of the left arm ischemia was patient condition and the relation to impella was determined to be unrelated. impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17524 | 1220648-2024-20226 | 20435014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435014&pc=OZD | 10/18/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/16/2021 | 10/12/2024 |  |  | 04/30/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022075422 |  | 0048-0034 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis and experienced access site bleeding with hematoma. the patient received an unknown amount of packed red blood cells. no further information was provided. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for the reported hemolysis and access site bleeding issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17525 | 1220648-2024-20227 | 20435015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435015&pc=OZD | 03/15/2022 | 10/06/2025 |  | 10/12/2024 |  |  | 03/15/2022 | 10/12/2024 |  |  | 12/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169267 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  |  | White | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp device had high purge pressure. the pump was removed and replaced with no harm to the patient. it was also reported that the patient had several hematological problems (thromnose of acb, bleeding). heparin was removed and replaced with sodium bicarbonate (nabic). | [5 empty] | The data logs were returned and confirm a gradual increase in purge pressure when heparin was removed. heparin was added back into the purge system but the purge pressure remained high. the pump was returned and run in the lab and the high purge pressure was not reproduced and no blockages could be seen. the root cause of the high purge pressure is most likely a biomaterial buildup. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the investigation for the high purge pressure has been completed. the data logs were returned and confirm a gradual increase in purge pressure when heparin was removed. heparin was added back into the purge system but the purge pressure remained high. the pump was returned and run in the lab and the high purge pressure was not reproduced and no blockages could be seen. the root cause of the high purge pressure is most likely a biomaterial buildup.
17526 | 1220648-2024-20228 | 20435068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435068&pc=OZD | 06/12/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 06/12/2023 | 10/12/2024 |  |  | 02/28/2025 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331577 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the pump was unable to advance into the left ventricle during attempted delivery due to an existing aortic dissection from ecmo placement. as a result of the noted condition, the implant was aborted. there was no patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. as per clinical, the impella insertion was aborted due to aortic dissection which was likely during ecmo placement. the cause of the delivery issue was patient condition related due to weakened/diseased aorta with aortic dissection. impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller instructions for use for the related event are as follows: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17527 | 1220648-2024-20229 | 20435071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435071&pc=OZD | 10/28/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 10/28/2021 | 10/12/2024 |  |  | 06/30/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100636 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella pump could not be fully advanced into the patient because of challenging anatomy and due to kinking of the supplied sheaths during advancement; therefore, the insertion was aborted. | [5 empty] | The investigation into the reported delivery issues and product damage (cannula kink) has been completed. the device was received for evaluation. the root cause of the delivery issues and product damage (cannula kink) was determined to be patient condition. this report is being filed as part of a retrospective review of historical records. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17528 | 1220648-2024-20230 | 20435074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435074&pc=OZD | 10/29/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 10/29/2021 | 10/12/2024 |  |  | 08/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022131704 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon returning from surgical mode and being cross clamped, the pump would not restart. the impella cp was explanted and replaced to address the issue. there was no reported harm or injury to the patient.` | [5 empty] | The investigation has been completed for the reported issue of pump stop. the returned pump was visually inspected and biomaterial was confirmed entraining the impeller. the root cause of the pump stop was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
17529 | 1220648-2024-20231 | 20435075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435075&pc=OZD | 03/28/2022 | 10/09/2025 |  | 10/12/2024 |  |  | 03/28/2022 | 10/12/2024 |  |  | 10/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152471 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, that there was continuous suction during support. after trying to adjust pump in the cath lab to resolve suction, the decision was made to implant new pump. the pump was removed and replaced. it was also noted that there was small oozing noted at the insertion site as well as all invasive line sites and the patient received blood products. | The complainant reported continuous suction during support, despite troubleshooting attempts to resolve. the pump was removed and replaced. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the cause of the issue was not determined since product was not returned. the impella cp was used beyond indication for use per design trace matrix 0046-9910. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the cause of the issue was not determined since product was not returned. the impella cp was used beyond indication for use per design trace matrix 0046-9910.
17530 | 1220648-2024-20232 | 20435076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435076&pc=OZD | 12/05/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 12/05/2021 | 10/12/2024 |  |  | 07/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105848 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported a thrombus attached at the pump, was found upon removal. hemolysis was also alleged. neither issue prompted medical intervention. it was noted the patient was in stable condition. | [5 empty] | The investigation for the reported hemolysis and thrombus issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the hemolysis and thrombus was not determined since insufficient information was provided. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17531 | 1220648-2024-20233 | 20435077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435077&pc=OZD | 04/19/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 04/19/2022 | 10/12/2024 |  |  | 12/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170275 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the patient experienced hemolysis and limb ischemia. there was no intervention indicated. there was no patient harm or injury noted. | [5 empty] | The investigation for the reported hemolysis and limb ischemia issues has been completed. the impella device has not been returned for evaluation. the cause of the issue was not determined since product was not returned. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
17532 | 1220648-2024-20234 | 20435078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435078&pc=OZD | 12/23/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 12/23/2022 | 10/12/2024 |  |  | 05/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227674 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Failure to Pump | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. durig support, the device exhibited suction alarms indicating low flow, the resolution for this issue is unknown. additionally, the pump stopped, the resolution for this issue is unknown. no report of patient harm, the patient survived the procedure. | [5 empty] | The investigation for the reported pump stops and purge leak has been completed. the impella device was not received from the customer; however, the data logs were returned. upon review of the logs, a show a rise in the motor current and high purge pressure leading to an eventual pump stop. therefore, the root cause of the the pump stop is most likely the purge leak. without the return of the device the root cause of the purge leak is unable to be determined. instructions for use for the related event are as follows: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ follows: this report is being filed as part of a retrospective review of historical records.
17533 | 1220648-2024-20235 | 20435079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435079&pc=OZD | 06/09/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 06/09/2023 | 10/12/2024 |  |  | 11/30/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294937 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, a triggered suction alarm immediately following implant. the support level was reduced in an attempt to troubleshoot the issue, however, the suction issues remained. the pump was subsequently explanted, and a clot was visualized within the cannula of the device. another impella cp optical pump was implanted for continued support. there was no reported patient harm. | [5 empty] | The investigation for the reported low pump flow and thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based off clinical details, the root cause of the low pump flow is most likely due to biomaterial. instructions for use for the related event are as follows: embolism/thrombus impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17534 | 1220648-2024-20237 | 20435080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435080&pc=OZD | 06/28/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 06/28/2023 | 10/12/2024 |  |  | 11/30/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303370 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited low flows. the position was assessed with an angiogram, showing that the motor had moved into the aorta and there was a cannula kink just below the outflow cage. multiple attempts to reposition were unsuccessful, and the pump was removed and replaced. there was no reported patient injury. | [5 empty] | The investigation for the reported positioning issue and low pump flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. data confirmed low flow when pump started & remained to the rest of the case that triggered auto suction response and suction alarms. data also showed pump was in ventricle for about 1 minute and a half, then back to normal position. product was not returned for review. based on clinical description & data, the root cause of the positioning issue is due to the patient's condition of the short aortic arch. this contributed to the cannula kink and low flows. this report is being filed as part of a retrospective review of historical records.
17535 | 1220648-2024-20236 | 20435081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435081&pc=OZD | 06/13/2023 | 10/20/2025 |  | 10/12/2024 |  |  | 06/13/2023 | 10/12/2024 |  |  | 11/30/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295941 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. shortly after insertion, the device displayed an aorta/left ventricle (lv) waveform, which disappeared after approximately one minute. the position also appeared correct on fluoroscopy. the positioning signal alarm was disabled, and the patient remained on impella support with no reported harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the positioning issue could not be determined. this report is being filed as part of a retrospective review of historical records. | No product was returned for evaluation. the data logs were reviewed, and the logs indicate the snr decreasing and the signal normalizes after p-level reduction from p-9 to p-8. no specific trend to confirm any abnormalities for the initial optical failure based on the log review. the cause of the optical signal issue was not determined since product was not returned and with inconclusive evidence of abnormality from data log review. this report is being filed as part of a retrospective review of historical records.
17536 | 1220648-2024-20238 | 20435082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435082&pc=OZD | 04/26/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 04/26/2022 | 10/12/2024 |  |  | 01/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182772 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the flows were decreased, and a kink was noticed in the cannula. the physician tried to straighten the catheter but was unsuccessful. a new cp was placed, and no further issues and no harm to the patient. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the field log was reviewed, and suction was confirmed from the activation to the end of the run. no evidence of the device being out of position. the returned product was visually inspected and a kink in the cannula was confirmed. the root cause of the cannula kink resulting in low pump flow was most likely patient condition related since the kink occurred during pump delivery. this report is being filed as part of a retrospective review of historical records.
17537 | 1220648-2024-20240 | 20435083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435083&pc=OZD | 11/01/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/01/2021 | 10/12/2024 |  |  | 08/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127068 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was a placement signal not reliable when prepping the cp pump for insertion. the decision was made to explant the device and replace it with a new impella cp device. there was no reported harm or injury to the patient. | [5 empty] | The investigation has been completed for the reported issue of placement signal issue. the device was not returned for evaluation. the root cause of the placement signal issue was determined to be unrelated to the pump. this report is being filed as part of a retrospective review of historical records.
17538 | 1220648-2024-20239 | 20435084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435084&pc=OZD | 07/25/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 07/25/2022 | 10/12/2024 |  |  | 01/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174017 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped. the resolution for this issue is unknown, no reported harm to the patient. | [5 empty] | The investigation for the reported pump stop has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the pump stop could not be determined. instructions for use for the related event are as follows: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
17539 | 1220648-2024-20241 | 20435085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435085&pc=OZD | 06/17/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 06/17/2023 | 10/12/2024 |  |  | 12/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2024323999 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the impella advanced forward 3 cm, and the staff observed that the pigtail was behind the pap. the team did not want to manipulate further since there was good flows and waveforms on the impella. there was no reported harm or injury to the patient. | [5 empty] | No product was returned. as per clinical review, impella position was shallow and the pump pigtail seemed to be caught in papillary. cpr was done prior to this positioning issue. the cause of the positioning issue was use related with cpr done leading to pump moving out of position per clinical review.
17540 | 1220648-2024-20242 | 20435086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435086&pc=OZD | 04/13/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 04/14/2023 | 10/12/2024 |  |  | 10/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022144132 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the pump had high purge pressure and the pump had stopped. troubleshooting steps were not successful. the impella was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issues of pump stop and high purge pressure. the product was returned. the root cause of pump stop & high purge pressure was due to biomaterial ingestion. this report is being filed as part of a retrospective review of historical records.
17541 | 1220648-2024-20243 | 20435088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435088&pc=OZD | 11/15/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 11/15/2022 | 10/12/2024 |  |  | 06/30/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235352 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was a purge leak with the sidearm of the impella pump. the leak was treated by bypass, and the team later elected to explant the impella cp earlier than expected. there was no reported serious injury to the patient. | [5 empty] | The device was returned and evaluated, and the investigation for the reported purge leak has been completed. the pump was returned without the sidearm and the catheter was cut at the 34cm mark. no other abnormalities were observed. the location of the leak was likely at the sidearm since a purge bypass was performed. the cause of the purge leak was not determined since the sidearm was not returned. as noted in the impella ifu: purge leak - pump: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
17542 | 1220648-2024-20244 | 20435089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435089&pc=OZD | 11/23/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 11/23/2022 | 10/12/2024 |  |  | 08/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266180 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing high-risk surgery was implanted with an impella cp device for mechanical circulatory support. during implant, an angiogram showed active bleeding and hematoma, so the pump was removed and reinserted with a different sheath. during support, the impella pigtail was caught on the papillary muscle and the mitral valve was touching the inflow cage. when the team attempted to reposition the device, the pump flow dropped significantly, and the flows did not improve with further repositioning. the team chose to explant the pump, and no further patient harm was reported. | [5 empty] | Product was visually inspected and a kink was observed on the cannula. pump positioning in mitral apparatus and repositioning caused the pump to come out of the lv and then kinked the cannula upon wireless re-access resulting in low pump flow. the impella device is not indicated for wireless re-access. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17543 | 1220648-2024-20245 | 20435090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435090&pc=OZD | 12/16/2020 | 10/12/2024 |  | 10/12/2024 |  |  | 11/02/2021 | 10/12/2024 |  |  | 08/31/2022 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1493301 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. one day after explant there was an ischemic stroke. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation into the cva/stroke has been completed. the impella device was not returned for evaluation. the root cause of cva/stroke was determined to be unknown/unrelated to impella. no evidence or clinical imaging was provided. this report is being filed as part of a retrospective review of historical records.
17544 | 1220648-2024-20246 | 20435091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435091&pc=OZD | 10/14/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 10/14/2021 | 10/12/2024 |  |  | 03/31/2022 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1480650 |  | 0048-0024 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced ventricular fibrillation and cardiopulmonary resuscitation was performed. the patient was shocked and return of spontaneous circulation was achieved. the patient was successfully weaned the device was explanted, | [5 empty] | The investigation for the reported ventricular fibrillation has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the ventricular fibrillation and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
17545 | 1220648-2024-20247 | 20435092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435092&pc=OZD | 09/15/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 09/15/2022 | 10/12/2024 |  |  | 04/30/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211820 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was a purge leak originating from a damaged purge filter. there was no reported harm or injury to the patient and the device was successfully weaned and explanted. | [5 empty] | The investigation for the reported issue has been completed. product was visually inspected, and no abnormalities were found. product was then connected to the automated impella controller (aic) and fluid was run through the pump. a leak was observed to be coming from the sidearm filter. this is indicative of a damaged sidearm filter. as noted in the impella ifu: impella cp with smart assist system section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
17546 | 1220648-2024-20248 | 20435094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435094&pc=OZD | 01/18/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 01/18/2022 | 10/12/2024 |  |  | 09/30/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022133249 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the impella was placed and started the impella the pump slipped out of the ventricle, resulting in removal of this impella. this impella was removed and a new impella placed. there was no patient harm reported. | [5 empty] | The impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the root cause of the positioning issues was determined to be use related as the pump was pulled back out of the left ventricle during repositioning. the pump was unable to recross the valve, wireless recrossing the valve is not permitted. this report is being filed as part of a retrospective review of historical records.
17547 | 1220648-2024-20249 | 20435096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435096&pc=OZD | 01/20/2022 | 09/24/2025 |  | 10/12/2024 |  |  | 01/21/2022 | 10/12/2024 |  |  | 10/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152052 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient had pink/red urine indicative of hemolysis. positioning and flow adjustment attempted but were not successful. upon advancing the pump too far into the left ventricle (lv), the cannula kinked. the kink was unable to be removed despite manipulation. impella cp was explanted and replaced with a new pump. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. after reviewing the clinical information, it was determined that the cause of the cannula kink was repositioning the device. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This report is being filed as part of a retrospective review of historical records. the investigation of product damage (cannula kink) and hemolysis has been completed. the impella device was not returned for evaluation. product damage (cannula kink): clinical mentioned cannula kinked occurred when pump was advanced too far into the lv. the root cause of the product damage (cannula kink) was most likely use related as evidenced by clinical detail of advancing pump too far in the left ventricle (lv). hemolysis: no product or data logs were returned for analysis. clinical mention repositioning did not resolve the issue. the root cause of the hemolysis was not determined as no product or logs were returned and insufficient clinical information was provided b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-20249. g1 reporting contact email revised on manufacturer device report 1220648-2024-20249 in accordance with updated procedures. h6 investigation conclusions 61 was erroneously entered on the initial manufacturer device report number 1220648-2024-20249.
17548 | 1220648-2024-20250 | 20435111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435111&pc=OZD | 02/25/2025 | 03/28/2025 |  | 10/12/2024 |  |  | 03/11/2022 | 10/12/2024 |  |  | 11/30/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169247 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella pump stopped in the operating room during the mitral surgery and was restarted. the pump stopped again 4 days later. it was also reported that the patient experienced a gi bleeding and purge solution was swapped out for d5 without heparin. the patient was also supported by ecmo. no patient harm or injury noted. | A patient with nonischemic cardiomyopathy, atrial fibrillation and had previously refused an automatic implantable cardioverter defibrillator experienced a witnessed out-of-hospital cardiac arrest and cardiopulmonary resuscitation (cpr) was immediately started. emergency medical services arrived to transport the patient to the hospital and performed numerous defibrillation attempts. cpr continued for approximately two hours and when the patient arrived in the emergency room, itx was in pulseless electrical activity. a return of spontaneous circulation was eventually obtained, and the patient was brought in to the cardiac catheterization laboratory for venoarterial extracorporeal membrane oxygenation (ecmo) cannulation and impella cp implantation for mechanical circulatory support. during impella cp support, the patient experienced gastrointestinal (gi) bleeding and as a result heparin was removed from the impella purge solution. the following day, on (b)(6) 2025, the patient received blood products due to the loss of blood. it was also noted at this time that the patient's abdomen appeared to be necrotic. on (b)(6) 2025 the patient was transfused with an additional unit of blood products due to the blood loss from the gi bleed. on (b)(6) 2022, the patient had surgery for a mitral valve replacement and was placed on cardiopulmonary bypass (cpb) using the ecmo cannula and the impella was placed in surgical mode. after the successful mitral valve replacement, while resuming ecmo and impella cp support and removing the patient from cpb, it was reported that the impella cp had a pump stop. the medical team was able to restart the impella cp. the medical team was informed that the pump had been running for 2.5 weeks without heparin in the purge solution and it was likely another pump stop would occur. heparin was added to the purge solution, and escalation to an impella 5.5 was discussed. it was additionally noted that heparin may have to be discontinued again as the patient had ongoing bleeding issues. 4 days later, on (b)(6) 2022, the impella cp stopped pumping again and was unable to be restarted. the patient was brought to the cardiac catheterization laboratory for a replacement impella cp. the patient remained stable throughout the procedure with no increased need for medication or ecmo support. the patient remained on impella cp support with the new pump for an additional 18 days, when the family decided to withdraw care and the patient then expired. although the patient's presenting condition may be more likely have impacted the event, the first impella cp cannot be excluded as a possible contributing factor in the patient's demise. |  |  |  |  | Analysis of the pump revealed a large purge gap indicating bearing wear as the root cause of the pump stop. this occurred after 22 days of run time, which is well beyond the design life of 8 days per (b)(4). impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the bleeding issue has been completed. due to the limited clinical information the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. additional information added to section h6 coding was added for cause not established to reflect the investigation results of the bleeding and h10 for completed investigation for reported bleeding issues. corrected information from the initial mfr 1220648-2024-20250 was added to the following sections: b2: death and date of death added. b5: date of event corrected to 25-feb-2025, as this was when the gastrointestinal bleeding was first reported. additional narrative added for clarity and missing information including patient outcome. h1: type of reportable event corrected to death. h6: codes added for health effect - impact code to reflect blood products, death, and device replacement.
17549 | 1220648-2024-20251 | 20435112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435112&pc=OZD | 11/03/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/03/2021 | 10/12/2024 |  |  | 08/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129499 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, low flows occurred. as a result, the decision was made to explant the device and replace it with a new impella cp device. there was no reported injury or harm to the patient. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the product was returned. the root cause of the low pump flow was most likely a kinked cannula that resulted from improper pump positioning. this report is being filed as part of a retrospective review of historical records.
17550 | 1220648-2024-20252 | 20435113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435113&pc=OZD | 04/08/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 04/08/2022 | 10/12/2024 |  |  | 11/30/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168977 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that after successful insertion of the impella cp, the console gave a yellow alert ¿placement signal not reliable¿. the impella cp was in good position, with no kinks in the cable. the pump operated properly during the entire case with no other issues. it is unknown how the reported issue was resolved. no patient harm was reported. | [5 empty] | The investigation for the reported placement signal issues has been completed. the impella device has been returned for evaluation. the returned pump was connected to a console and the placement signal not reliable issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
17551 | 1220648-2024-20253 | 20435114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435114&pc=OZD | 11/07/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/07/2021 | 10/12/2024 |  |  | 03/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069190 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to being implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp pump could not be fully advanced into the patient; therefore, the insertion was aborted. there was no reported injury or harm to the patient. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the investigation has been completed for the reported issue of delivery issue. the product was returned. the root cause of delivery issues were determined to be related to patient condition. this report is being filed as part of a retrospective review of historical records.
17552 | 1220648-2024-20254 | 20435130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435130&pc=OZD | 04/26/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 04/26/2022 | 10/12/2024 |  |  | 12/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174573 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was being implanted with an impella cp for mechanical circulatory support. upon advancing the impella into position, the impella severely kinked due to extreme torque across the aortic arch. the patient had oozing at groin site and heparin was decreased. the physician attempted to reposition, but the fold continued to worsen. the physician explanted the initial device and implant a new impella cp with no issues. | [5 empty] | The impella device was returned, the cause of the cannula kink was use related since an incorrect guidewire boston scientific v-18 control wire 0.018 in was used. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17553 | 1220648-2024-20255 | 20435147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435147&pc=OZD | 02/02/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 02/02/2023 | 10/12/2024 |  |  | 10/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290311 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the medical staff reported a delivery issue involving the impella cp pump. per the physician, the cp pump was unable to advance past the peel-away sheath during attempted implant. the implant was subsequently aborted as a result of the noted issue. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. no clinical details provided were sufficient to determine a root cause. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17554 | 1220648-2024-20256 | 20435148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435148&pc=OZD | 02/06/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 02/06/2023 | 10/12/2024 |  |  | 10/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024291697 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a pump stop during impella support. the pump was removed with no harm to the patient. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the root cause of the pump stop was determined to be an intermittent electrical short that occurs when the exposed wire of the red motor cable comes into contact with the ground wire. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17555 | 1220648-2024-20257 | 20435149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435149&pc=OZD | 02/22/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 02/22/2023 | 10/12/2024 |  |  | 11/30/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283877 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. impella cp was escalated to impella 5.5 due to dislocation of the cp. the patient was supported with impella 5.5 for 34 days as bridge to transplant. no further information is available. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
17556 | 1220648-2024-20258 | 20435168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435168&pc=OZD | 02/23/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 02/23/2023 | 10/12/2024 |  |  | 10/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024293962 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with aortic valvuloplasty was implanted with an impella cp device for mechanical circulatory support. the complainant reported they were attempting to implant an impella device. the device was opened and primed. during this time, the decision was made to recheck gradient across valve after it was requested that a second pigtail was to be inserted into the aorta. gradient was 10.8. dobutamine stress performed increased gradient after twenty minutes to 28.8. the case was aborted. the sheath was in the body for 45 minutes, so it was aggressively aspirated and flushed throughout. it was eventually removed and preclose x 2 tied for hemostasis. no known patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient condition due to aortic stenosis. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17557 | 1220648-2024-20259 | 20435171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435171&pc=OZD | 09/19/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 09/19/2022 | 10/12/2024 |  |  | 04/30/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225665 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a delivery issue with the impella device. the pump insertion was aborted and a new impella was used without further issue. there was no harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was patient condition related. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17558 | 1220648-2024-20260 | 20435172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435172&pc=OZD | 02/02/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 02/02/2023 | 10/12/2024 |  |  | 09/30/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269672 |  | 1000302 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. the complainant reported an easy guide lumen that broke within the pump during preparation of the device. the pump was never inserted into the patient. a new pump set was opened and used successfully. | [5 empty] | The investigation for the reported introducer damage has been completed. the product was returned and the introducer damage was confirmed. the easy guide lumen was broken inside the pump during prep stage. the root cause of the delivery issue was a broken red lumen which was determined to be use-related. this report is being filed as part of a retrospective review of historical records.
17559 | 1220648-2024-20261 | 20435173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435173&pc=OZD | 04/30/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 04/30/2022 | 10/12/2024 |  |  | 01/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189751 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patients had low pump flows and bleeding. it was also reported that within 10 minutes of starting the pump, the flow dropped to ~2.0 on p-9 and suction alarm sounded. trouble shooting confirmed no clot, the p-levels were dropped, however the suction alarm was continuous for almost 2 days until the pump was removed. prior to pump removal the patient continued to have oozing from all access sites. oozing from impella site much improved with fem stop in place and they decreased purge heparin. the patient was successfully weaned off of support. | [5 empty] | (b)(4) reviewed and low pump flow noted and the flow is below required flow for the rest of the running time. complaint pump not returned for investigation. root cause can not be determine. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17560 | 1220648-2024-20262 | 20435175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435175&pc=OZD | 04/28/2022 | 07/09/2025 |  | 10/12/2024 |  |  | 05/10/2022 | 10/12/2024 |  |  | 09/30/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022128474 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Increase in Pressure | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp had alarms, bleeding, and a cannula kink. the impella was inserted directly into the aorta because it was difficult to insert. bleeding occurred and a blood transfusion was given. the device was removed in an environment where thoracotomy was possible. a kink was found on the removed impeller cannula. | [5 empty] | The returned pump was visually inspected and a catheter kink along with several cannula kinks were observed. the returned purge cassette was purged and then connected to the pump and a high purge pressure alarm triggered. the catheter was then cut open at the catheter kink and a purge lumen kink was confirmed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | H6: correcting to 3243-stressproblem to account for king, and 5--to account for patient anatomy likely causing the kink and 4315-cause not established for bleed.
17561 | 1220648-2024-20263 | 20435188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435188&pc=OZD | 05/31/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 06/02/2022 | 10/12/2024 |  |  | 01/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188809 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp was running for 10 days when there was a purge leak at the purge side arm that prompted a purge bypass. the pump remained on for support despite this issue. the dextrose concentration was increased to assist with purge pressure and flow. it was also noted that heparin was held as the patient had a subdural hematoma at the access site and internal abdominal bleed and that the patient had a brain bleed that is indicative of a stroke while on support. the patient remained on venoarterial extracorporeal membrane oxygenation (va ecmo). | [5 empty] | Pump and purge cassette was returned and inspected in the bhl. leak reproduction testing was preformed and a leak was found between the purge cassette yellow luer and the sidearm yellow luer connection. the leak appeared to be coming from the stem of the sidearm yellow luer. the yellow luer was imaged and a crack was visualized along the mold line. the usage of sodium bicarbonate most likely contributed to the leak in the yellow luer. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17562 | 1220648-2024-20264 | 20435204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435204&pc=OZD | 06/08/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 06/09/2022 | 10/12/2024 |  |  | 02/29/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202398 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a pump stop occurred. the device was restarted. there was no reported harm to the patient. | [5 empty] | The returned product was visually inspected and a large purge gap was observed. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17563 | 1220648-2024-20265 | 20435206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435206&pc=OZD | 11/08/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/08/2021 | 10/12/2024 |  |  | 08/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127958 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the purge side arm was leaking during the procedure. the impella cp device was explanted. there was no reported harm or injury to the patient. | [5 empty] | The investigation into the purge leak has been completed. the impella device was not returned for evaluation. the cause of the purge leak was not determined since the product, data logs, and sufficient clinical information were not provided. section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
17564 | 1220648-2024-20266 | 20435238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435238&pc=OZD | 11/09/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/09/2021 | 10/12/2024 |  |  | 08/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022122604 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the cp pump stop due to guidewire retention with a new staff. the team replaced the pump and support proceeded with the 2nd pump. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the product was returned. the root cause of the pump stop was most likely failing to remove the guidewire. this report is being filed as part of a retrospective review of historical records. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17565 | 1220648-2024-20267 | 20435275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435275&pc=OZD | 11/10/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/10/2021 | 10/12/2024 |  |  | 08/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022125901 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to being implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp pump could not be fully advanced into the patient because of challenging anatomy. the patient was escalated to an impella 5.0. | [5 empty] | The investigation has been completed for the reported issue of delivery issues. the device was not received for evaluation. the root cause of the delivery issues were determined to be related to patient condition. this report is being filed as part of a retrospective review of historical records.
17566 | 1220648-2024-20268 | 20435298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435298&pc=OZD | 11/18/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/18/2021 | 10/12/2024 |  |  | 05/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134852 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a cp supporting a dtu study patient. during the support there was an ooze at the access site. when the rhc was done the team decided to treat the ooze and how dry the patient was with 1 unit of red blood cells. there was no other intervention needed. the harm was not lasting. | [5 empty] | As no clinical details were reported, the cause of the bleeding could not be determined. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17567 | 1220648-2024-20269 | 20435299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435299&pc=OZD | 06/14/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 06/14/2022 | 10/12/2024 |  |  | 01/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186693 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a placement signal failure and low pump flow were reported during use of the device. a new pump was used without any impella related harm. | [5 empty] | The returned pump was visually inspected and biomaterial was observed in the outflow and on the sensor. ps was initially 100/100mmhg and dropped to 9/9 mmhg after soaking in water. this report is being filed as part of a retrospective review of historical records.
17568 | 1220648-2024-20270 | 20435302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435302&pc=OZD | 02/27/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 02/27/2023 | 10/12/2024 |  |  | 03/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214352 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella cp device for mechanical circulatory support. the device was placed for venoarterial extracorporeal membrane oxygenation (va ecmo) left ventricular (lv) unloading. high motor current alarmed and the pump stopped. the pump could be restarted but pump flow stayed low after restart. there were continuous suction alarms. imaging showed positioning issues. contributing factor was patient anatomy / small lv. echocardiogram showed position was not ideal. physician repositioned the device. additionally, access site bleeding was noted. direct pressure did not slow the bleeding. it was determined the 8 fr sheath needed to be removed and per closed. this did not resolve the issue. there was poor position on echo and concern of pulling form lv. at that point, due to the bleeding form the 8 fr site and with poor per close, the impella cp was removed during closure of access site bleeding from previously placed intra-aortic balloon pump. patient survived the procedure. | [5 empty] | No data logs were returned. the pump was investigated and a cannula kink was found. no other abnormalities were noted. the cause of the pump stop was improper positioning of the device and likely the impeller interacting with patient anatomy. the cause of the low pump flow was a kinked cannula. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17569 | 1220648-2024-20271 | 20435329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435329&pc=OZD | 11/12/2021 | 10/12/2024 |  | 10/12/2024 |  |  | 11/15/2021 | 10/12/2024 |  |  | 08/31/2023 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022132026 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. during implantation, the impella cp pump could not be fully advanced into the patient because of calcified vasculature. as a result of this complication the impella was removed, and an intra-aortic balloon pump was placed. the patient survived the procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was determined to be patient condition as the patient's vessels were reportedly calcified. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
17570 | 1220648-2024-20272 | 20435331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435331&pc=OZD | 11/21/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 11/21/2022 | 10/12/2024 |  |  | 07/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256842 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported trouble unlocking the tuohy borst valve after advancing the repositioning sheath. a scrub tech was able to successfully unlock the valve without any harm to the patient. megaderm was applied due to plastic sterile sleeve being torn. additionally, access site oozing was noted. mattress suture and femostop were applied with hemostasis. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. after reviewing the clinical information, it was determined that the cause of the positioning issues was use related, specifically improper technique when attempting to advance the repo unit. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17571 | 1220648-2024-20273 | 20435342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435342&pc=OZD | 09/26/2022 | 10/12/2024 |  | 10/12/2024 |  |  | 09/26/2022 | 10/12/2024 |  |  | 06/30/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229686 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an acute patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during explant, there were vascular complication of right femoral access during prostar closure. manual compression resolved the issue. patient was stable and successfully weaned off of support. | [5 empty] | The investigation for the reported removal issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the removal issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
17572 | 1220648-2024-20274 | 20435344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435344&pc=OZD | 06/13/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 06/13/2023 | 10/12/2024 |  |  | 03/31/2025 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343011 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was being implanted with an impella cp for mechanical circulatory support. while inserting the pump, the physician was unable to advance the impella cp catheter due to calcification in aorta. the implantation was aborted. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. per the given information, the root cause of the delivery issue was patient condition related to calcified aortic valve. impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller instructions for use for the related event are as follows: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17573 | 1220648-2024-20275 | 20435414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435414&pc=OZD | 04/27/2023 | 09/26/2025 |  | 10/12/2024 |  |  | 04/27/2023 | 10/12/2024 |  |  | 10/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272784 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient experienced thrombocytopenia. it was noted that because of thrombocytopenia, the physician wanted to avoid more heparin and therefore continue to carry bicarb in irrigation. no additional information was provided. no patient harm or injury noted. | It was reported that a patient on impella cp experienced an intermittent placement signal and not reliable alarms. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the cause of the issue was not determined since product/data was not returned. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue. d10 concomitant medical products and therapy dates. g1 reporting contact email updated.
17574 | 1220648-2024-20276 | 20435425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435425&pc=OZD | 01/26/2023 | 10/12/2024 |  | 10/12/2024 |  |  | 05/17/2023 | 10/12/2024 |  |  | 11/30/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023279112 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. it was reported a failure to prime the device. there was no reported harm or injury to the patient. | [5 empty] | The investigation into the priming issue has been completed. the impella device was not returned for evaluation. however, data log review shows a purge fluid not detected error, the the pump failing to complete priming. the same pc and pump were used on another console without issues. this console was evaluated and no issues were found. therefore, it is most likely that the priming issue was due to use, and it is possible the luers were not properly connected or the bag was not spiked properly to allow fluid to pass through. this report is being filed as part of a retrospective review of historical records.
17575 | 1220648-2024-20277 | 20435546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20435546&pc=OZD | 11/12/2022 | 10/13/2025 |  | 10/12/2024 |  |  | 11/12/2022 | 10/12/2024 |  |  | 08/31/2024 |  | 10/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023264373 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the device came out of position and lost placement signal. the fiber optic feature on the cp had malfunction and the pulmonary stenosis was unreliable. care team monitoring/reviewing placement daily. it was reported that there was oozing and bruising around the groin site. it was recommended to trace the bruise and monitor the site. angiomax started and purge changed to bicarbonate. the patient was successfully weaned off of support. | The us complaint reported that a patient (pt) was placed on impella cp for hemodynamic support. it was reported that the device came out of position and lost placement signal. the fiber optic feature on the cp had malfunction and the ps was unreliable. care team monitoring/reviewing placement daily. it was reported that there was oozing and bruising around the groin site. it was recommended to trace the bruise and monitor the site. the pt was suspected of having hit. angiomax started and purge changed to bicarbonate. |  |  |  |  | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | Data logs showed the placement signal waveform was fluctuating between ventricular and normal for the first few hours of support during the time of the reported hemolysis. about 24 hours into support, the "placement signal not reliable" alarm occurred immediately upon a sudden placement signal spike to 3000; the placement signal does not recover for the remainder of support. all 5s parameters remained out of specification as well. no damage or abnormalities were found on the returned pump. the pump's sensor passed optical continuity testing. when connected to an automated impella controller (aic), the placement signal was within specification and no alarms occurred. in conclusion, it was determined that the cause of the hemolysis was improper positioning and the cause of the optical signal issue was use related. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1: corrected mfg. email address.
17576 | 1220648-2024-20121 | 20425940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20425940&pc=OZD | 09/15/2024 | 11/21/2024 |  | 10/11/2024 |  |  | 09/15/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025468459 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an 82-year-old female patient on impella cp support developed hemolysis. the impella position was noted to be perfect. the patient could not tolerate going below p7. the patient maxed on pressors and flowing 3+ liters per minute on impella and could not keep map above 30s. the patient expired after 22 hours of support. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
17577 | 1220648-2024-20120 | 20425969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20425969&pc=OZD | 09/13/2024 | 11/19/2024 |  | 10/11/2024 |  |  | 09/13/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1256008 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller did not start. the patient was noted to be stable when this occurred. however, the patient ultimately expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation has been completed. console logs from the reported day of event were not available, since the console was not returned for evaluation. partial logs were downloaded for cloud contained entries from most recent preventative maintenance (pm) procedure that was performed and related field service report document indicated that pm was performed by germany field service. it is standard procedure that before transport, automated impella controller (aic) batteries must be disengaged via switch on the underside of the console, and the customer was being advised (via label attached to the aic screen upon return) that said switch needed to be flipped prior to use. it is most likely that the miscommunication occurred due to the label prematurely removed after the console was received at the customer site. after the batteries were re-engaged on-site, the issue was resolved. there was no issue found with the console. the device functioned/operated to specification. battery transport switch was not engaged. h.6 revised medical device problem code since initial manufacturer device report number 1220648-2024-20120 was submitted due to investigation results. an additional component code was added based off of investigation results. | A.5 is unknown. the automated impella controller was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
17578 | 1220648-2024-20130 | 20426354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20426354&pc=OZD | 09/27/2024 | 11/20/2024 |  | 10/11/2024 |  |  | 09/27/2024 | 10/11/2024 |  |  | 09/30/2025 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411525 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was unable to be placed due to the patient¿s acute anatomy. the right femoral artery was calcified, and staff were unable to successfully place the wire. the left femoral artery was used for impella insertion but the 14fr sheath would not pass the bend in the femoral artery. 14 fr cook 45 cm sheath used to straighten out the left femoral artery, but impella insertion was unsuccessful. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since r femoral artery was calcified and l femoral artery had acute anatomy which made implant of the impella unsuccessful. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17579 | 1220648-2024-20125 | 20426363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20426363&pc=OZD | 09/21/2024 | 11/15/2024 |  | 10/11/2024 |  |  | 09/21/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521380 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the implant, a hematoma started to form at the insertion site. manual pressure was held and this contained the hematoma until the radial sheath was removed. the hematoma reappeared. manual pressure and a femstop were applied. the doctor was concerned about the hematoma and the decision was made to remove the pump. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20125.
17580 | 1220648-2024-20124 | 20426556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20426556&pc=PYX | 09/20/2024 | 11/15/2024 |  | 10/11/2024 |  |  | 09/20/2024 | 10/11/2024 |  |  | 10/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023274760 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that purge pressure low alarm occurred on the impella rp. purge flow 30ml/hr and purge pressure would drop below 300mmhg. all connections were checked. attempted to resolve by changing the purge cassette. no overall change in purge flow or purge pressure. once purge fluid was changed to d10, purge flow and purge pressure was within normal parameters. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. the root cause of the issue was most likely a design limitation to the impella pump. e4 should have been left blank on manufacturer device report 1220648-2024-20124.
17581 | 1220648-2024-20126 | 20426563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20426563&pc=OZD | 09/23/2024 | 12/06/2024 |  | 10/11/2024 |  |  | 09/23/2024 | 10/11/2024 |  |  | 09/30/2025 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395556 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 76-year-old male patient had bleeding at the access site. to prevent any complication, the physician placed a compress closure to the anastomoses. additionally, the patient had hemolysis on support with dialysis. this issue was present prior to the impella implant, however it is unknown if the impella was contributing to the hemolysis. the pump was explanted after four days of support, and the patient is noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ | The investigation for the hemolysis and access site bleed has been completed. product and data logs were not returned for review. based on clinical description, the root cause of hemolysis was most likely patient condition related as hemolysis occurred prior impella implant. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. added h6 impact code 4627.
17582 | 1220648-2024-20127 | 20426581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20426581&pc=OZD | 09/23/2024 | 11/18/2024 |  | 10/11/2024 |  |  | 09/23/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025517083 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support for 64-year-old male patient, the pump was repositioned and the contamination sleeve tore. support continued following the events. the pump was explanted after nine days of support. patient is noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage has been completed. no product was returned and logs are not applicable. the root cause of the sterile sleeve damage was not determined since product was not returned and insufficient clinical evidence was provided. added-h6 impact code 4627.
17583 | 1220648-2024-20122 | 20426591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20426591&pc=OZD | 09/12/2024 | 11/18/2024 |  | 10/11/2024 |  |  | 09/13/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527663 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Packaging Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that while implanting impella cp the purge cassette became unsterile. a second package was opened. there was no patient harm. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation packaging issues has been completed. the impella device was not returned for evaluation. the root cause of the cosmetic defect was determined to be a use issue based on provided clinical details. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20122.
17584 | 1220648-2024-20128 | 20426593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20426593&pc=OZD | 09/27/2024 | 11/25/2024 |  | 10/11/2024 |  |  | 09/27/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025543809 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump being delivered for the escalation of a cp-extra corporeal membrane oxygenation support. the delivery was not possible due to anastomosis issues. the anatomy precluded placement. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be use issue because poor surgeon graft anastomosis, not a pump issue as per clinical information. the pump was unable to cross the graft anastomosis and implantation was aborted.
17585 | 1220648-2024-20129 | 20426662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20426662&pc=OZD | 09/27/2024 | 11/26/2024 |  | 10/11/2024 |  |  | 09/27/2024 | 10/11/2024 |  |  | 11/30/2025 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428129 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump placed to support a patient admitted in acute myocardial infarction / cardiogenic shock. the 5.5 was placed for care escalation from the intra-aortic balloon pump for a planned coronary artery bypass graft. the pump was placed but the axillary site was bleeding at the jackson-pratt drain. the team had to reopen the axillary site. the muscle bleed was found and controlled and the pocket was again closed. the patient remains on impella support. | The patient expired while on support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into access site bleeding has been completed. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20129. b2 date of death. b5 patient outcome. d6b missing. e4 should have been left blank. g1 reporting contact email. g1 reporting contact fax number missing. h1 type of reportable event. h6 health effect - impact code 1802 death added.
17586 | 1220648-2024-20131 | 20426667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20426667&pc=OZD | 09/27/2024 | 10/17/2024 |  | 10/11/2024 |  |  | 09/27/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210871 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the plastic purge disc holder on an automated impella controller (aic) had snapped off. there was no known harm associated with the damage. a loaner was requested to replace the aic. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. no data review is needed for this complaint because it was only reported that the retainer was broken. the root cause of the issue was broken purge disc retainer. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20131. e4 should have been left blank. g1 reporting contact fax number missing. | The impella device was returned and evaluated. the root cause of the issue was broken purge disc retainer.
17587 | 1220648-2024-20132 | 20426779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20426779&pc=OZD | 12/24/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 12/24/2022 | 10/11/2024 |  |  | 07/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255564 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support a clot formed in the cannula. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | No product was returned for investigation. the lt log confirms low pump flow when the pump is started and shows a suction alarm. the cause of the low pump flow was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
17588 | 1220648-2024-20147 | 20426878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20426878&pc=OZD | 10/19/2021 | 02/27/2025 |  | 10/11/2024 |  |  | 10/19/2021 | 10/11/2024 |  |  | 08/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127071 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella pump stopped unexpectedly. the initial impella device was removed and replaced with a new impella pump. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The data logs were returned and confirmed the high motor current stops after insertion. the product was returned and ingested biomaterial could be seen in the inflow and around the outflow. the high motor current pump stop was reproduced. the pump was soaked in tergazyme and the biomaterial was removed. the pump was then able to operate within specification. the root cause of the pump stop was determined to be ingested biomaterial during insertion. this report is being filed as part of a retrospective review of historical records.
17589 | 1220648-2024-20133 | 20427053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427053&pc=OZD | 04/22/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 04/22/2023 | 10/11/2024 |  |  | 12/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298005 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient¿s bilateral extremities became cold and mottled indicative of limb ischemia. the patient was given vasopressin's, dobutamine and levophed and was monitored. patient survived the procedure. | [5 empty] | The investigation for the reported limb ischemia issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the limb ischemia is most likely due to the patient condition. this report is being filed as part of a retrospective review of historical records.
17590 | 1220648-2024-20144 | 20427102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427102&pc=OZD | 04/05/2022 | 05/22/2025 |  | 10/11/2024 |  |  | 04/05/2022 | 10/11/2024 |  |  | 06/30/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100699 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient developed hemolysis. the patient was successfully weaned and the device removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
17591 | 1220648-2024-20135 | 20427113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427113&pc=OZD | 09/14/2024 | 11/12/2024 |  | 10/11/2024 |  |  | 09/14/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025543772 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was placed via the femoral despite knowledge of diminished limb flow and peripheral artery disease on the 69-year-old female patient. the impella cp was placed to support the patient postop for cancer surgeries. the day after implant the team placed distal perfusion due to the lost pulses. the pump remained on for support, however care was ultimately withdrawn. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
17592 | 1220648-2024-20136 | 20427168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427168&pc=OZD | 09/15/2024 | 11/18/2024 |  | 10/11/2024 |  |  | 09/15/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540828 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support 57-year-old male patient who went down in (b)(6) and got cpr for and defibrillated for ventricular tachycardia/ventricular fibrillation. the pump was placed after a percutaneous coronary intervention (pci) and the pump support was withdrawn after 70 hours and patient expired. the support did have signs of hemolysis ongoing and continuous renal replacement therapy (crrt) was initiated during the support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. data logs were reviewed and showed that the pump ran without issue during support. there were positioning issue alarms triggered during support but resolved quickly. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
17593 | 1220648-2024-20134 | 20427238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427238&pc=OZD | 04/26/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 04/26/2023 | 10/11/2024 |  |  | 05/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080892 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. it was reported that the device failed to deliver due to the patient¿s anatomy. the pump insertion was aborted and a new impella cp was used without further issue. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the issue was patient condition related. this report is being filed as part of a retrospective review of historical records.
17594 | 1220648-2024-20141 | 20427246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427246&pc=OZD | 10/15/2021 | 10/11/2024 |  | 10/11/2024 |  |  | 10/15/2021 | 10/11/2024 |  |  | 06/30/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022099905 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support there were signs of hemolysis. the patient was treated with fluids. the pump was repositioned and p levels were adjusted. the patient was successfully weaned and pump explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
17595 | 1220648-2024-20137 | 20427256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427256&pc=OZD | 10/13/2021 | 10/11/2024 |  | 10/11/2024 |  |  | 10/13/2021 | 10/11/2024 |  |  | 07/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110528 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the guide wire was able to cross the transcatheter aortic valve replacement valve. however, when the device was loaded onto the wire, the impella pump would not pass through the valve. multiple efforts were made to manage this complication without success. the impella was explanted and replaced with a new pump. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the delivery issue was the inability to pass the patient's transcatheter aortic valve replacement (tavr) valve. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17596 | 1220648-2024-20138 | 20427280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427280&pc=OZD | 08/23/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 08/23/2022 | 10/11/2024 |  |  | 07/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105867 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Unknown |  |  | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a positioning issue due to poor fixation of the device. the pump was removed. there was no harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the positioning issues was most likely use related as the pump was not properly fixed. this report is being filed as part of a retrospective review of historical records.
17597 | 1220648-2024-20140 | 20427320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427320&pc=OZD | 10/15/2021 | 10/11/2024 |  | 10/11/2024 |  |  | 10/15/2021 | 10/11/2024 |  |  | 08/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022126491 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock and failing mitral valve was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella pump was pulled back into the aorta as an impella in aorta alarm was received. additional attempts to reposition the device were unsuccessful as the impella catheter was unable to advance due to a kink. as the patient was no longer being supported by the impella pump, the device was explanted. | [5 empty] | The investigation for the reported issue has been completed. the returned pump and gwru (guidewire repositioning unit) were visually inspected, and the repositioning sheath was noted to be ribbed/buckled along its length. additionally, the pump catheter was kinked at approximately the 84cm mark. this kink likely occurred during the ultimately unsuccessful attempts to readvance impella across the aortic valve. the gwru was tested for leaks in a pressurized system and no leaks were reproduced. the root cause of the reported bleeding was most likely patient condition as the returned repositioning sheath showed visible signs of stress. the positioning issues were ultimately caused by attempts to resolve the access site bleeding. this report is being filed as part of a retrospective review of historical records.
17598 | 1220648-2024-20139 | 20427344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427344&pc=OZD | 09/10/2024 | 08/22/2025 |  | 10/11/2024 |  |  | 09/20/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530859 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during impella cp support for 61-year-old male patient, the automated impella controller (aic) alarmed for ¿high purge pressure¿. purge flow dropped to <1 and purge pressure increased to over 1,000. heparin had been held due to pulmonary hemorrhage and hemoptysis. decision was made to exchange the pump. the patient is noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h6 health effect impact code was erroneously entered on the initial manufacturer device report number therefore updated to the correct code. h6 investigation conclusions from the initial submission in accordance with updated procedures. h10 revised from the initial submission in accordance with updated procedures.
17599 | 1220648-2024-20142 | 20427349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427349&pc=OZD | 09/21/2024 | 11/21/2024 |  | 10/11/2024 |  |  | 09/21/2024 | 10/11/2024 |  |  | 06/30/2025 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359660 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella 5.5 the registered nurse was exchanging purge tubing when it was noticed that the female luer lock on the purge side-arm was cracked and missing the ¿lip¿ nearest to the male-luer side. surgeon decided to hold off on intervention until the morning, but to try to fix the leak by applying teflon tape at luer joint and installing a bypass connection. in the morning, sterilized teflon tape was applied to female luer and clave was attached with successful bypass of leak. no issues were noted over the last 24 hours. there was no harm to the patient. | [5 empty] | The investigation for the purge leak-pump has been completed. data logs were not returned for investigation. returned product was investigated, and as reported, there was a makeshift luer connection made over the female connection of the pump sidearm. that luer was removed to see the state of the pump's original luer connection, and there were no visual abnormalities. the pump was connected to both purge cassettes that were returned and run in a bucket of water for 15 minutes each while observing. no leaks reproduced at the sidearm luer joint for either purge cassette. additionally, data logs from this run in the lab confirm that purge pressures and flows were within specification the whole time and no purge related alarms triggered. the root cause of the purge leak was most likely a use issue since the leak was unable to be reproduced on the returned product and no damage to the yellow luer was observed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
17600 | 1220648-2024-20149 | 20427352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427352&pc=OZD | 06/05/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 06/05/2022 | 10/11/2024 |  |  | 01/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175775 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped after the device was implanted over 15 days. the patient was weaned from catecholamine, and remained on venoarterial extracorporeal membrane oxygenation (va ecmo) and the device was explanted. | [5 empty] | The investigation for the reported pump stop issue has been completed. the pump stop occurred on day 14 which is beyond the design indication of 8 days per design trace matrix 0046-9910. the root cause of the pump stop was unable to be determined as no product or logs were returned for analysis. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17601 | 1220648-2024-20151 | 20427359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427359&pc=OZD | 05/08/2023 | 10/13/2025 |  | 10/11/2024 |  |  | 05/08/2023 | 10/11/2024 |  |  | 01/31/2025 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312768 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient, race and ethnicity unknown, noted as high-risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported during placement of the impella when inserting the guide wire resistance was felt and the easy guide lumen was totally trapped within the outlet. the physician was made the decision to remove this impella and place a new impella. there was no patient harm reported, and this device was replaced. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
17602 | 1220648-2024-20145 | 20427369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427369&pc=OZD | 06/28/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 06/28/2022 | 10/11/2024 |  |  | 03/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208311 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that there were delivery issues with the impella cp. the impella was unable to advance through the iliac. the decision was made to exchange the sheath and open a new device. a fem-fem bypass was performed due to bilateral limb ischemia. no patient harm or injury noted. | [5 empty] | The investigation for the reported limb ischemia and delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the issue was most likely patient condition related. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿. this report is being filed as part of a retrospective review of historical records.
17603 | 1220648-2024-20146 | 20427427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427427&pc=OZD | 04/27/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 04/27/2023 | 10/11/2024 |  |  | 12/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024310658 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. initially, single access was attempted using an abiomed 14 fr x 13 cm peel-away sheath with a 7 fr brite tip; however, back bleeding occurred through the hemostatic membrane. the peel-away sheath was re-accessed with the 7 fr brite tip at the 2 o¿clock position, but bleeding reoccurred through the peel-away sheath. no bleeding was observed when only the impella catheter was positioned within the peel-away sheath. a decision was subsequently made to access the left common femoral artery for the percutaneous coronary intervention. the patient survived the procedure. | [5 empty] | The investigation into the reported single access issue has been completed. the impella device was not returned for evaluation. the most likely cause of the single access issue could not be determined since the 14fr sheath was not returned. this report is being filed as part of a retrospective review of historical records.
17604 | 1220648-2024-20143 | 20427431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427431&pc=OZD | 06/18/2022 | 09/16/2025 |  | 10/11/2024 |  |  | 06/18/2022 | 10/11/2024 |  |  | 12/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180072 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped and was restarted. also, there was oozing. the perclose sutures were tightened and oozing had lightened. multiple blood products were given and stabilized. the patient had another combative episode where he managed to grab the impella cp catheter in the groin and pull on it. they did a bedside echocardiogram and it showed the impella was shallow, but still in the ventricle. all of a sudden, the left ventricle signal disappeared and the motor current flat lined. ¿impella stopped¿ alarm started. the patient still had a grip on the catheter and would not let go. the patient was then rushed to the cardiac catheterization lab where they made attempts to turn up the p-level, the device would not start. the physician removed the impella cp at that point. the patient survived the explant. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed the pump stopped immediately upon the sudden occurrence of the "impella stopped (rpm deviations" alarm (#115). following the pump stop, there were sporadic motor current spikes to over 30000. the pump was unable to be restarted. no external damage or abnormalities were observed on the returned pump. the pump stop was reproduced when attempting to run on the aic with a motor current of over 32000. electrical resistance testing revealed an open line between all three motor phases. the red handle was opened, and all three motor cables were found to be severed between the monkey fist and the kink protector. there was evidence of motor cable necking and the monkey fist being pulled out of place. after reviewing the clinical information, it was concluded that the cause of the pump stop was an open electrical line due to excessive force on the catheter.
17605 | 1220648-2024-20148 | 20427453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427453&pc=OZD | 05/08/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 05/08/2023 | 10/11/2024 |  |  | 12/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303717 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 23 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was bleeding from access sites (sg catheter and vascas) and from the oral cavity; however, the impella access site did not bleed. the amount of heparin in the purge liquid was reduced from 20u to 10u, and activated clotting time, 160, and bleeding decreased slightly. the patient was transfused (number of units unknown). in addition, hemolysis at p-9 after impella support was noted. the hemolysis disappeared when the support level was changed to p-7. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. no clinical details provided were sufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
17606 | 1220648-2024-20150 | 20427487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427487&pc=PYX | 09/23/2024 | 11/23/2024 |  | 10/11/2024 |  |  | 09/23/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025543825 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp flex support and there was high purge pressure. tissue plasma activator was added to the purge. additionally, the patient had hemolysis. continuous renal replacement therapy (crrt) was started on support. furthermore, the patient had bleeding while on support. the location of the bleeding is unknown. one unit of blood products were given to the patient. heparin was in the purge. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation into high purge pressure, hemolysis, renal failure, and vascular damage - peripheral vessel has been completed. the returned product had biomaterial wrapped around the impeller. the data logs confirm high purge pressure. there was sustained high purge pressure for about 2 hours before it returned to normal levels coinciding with clinical mention of tissue plasma activator (tpa) administration. there were several minor motor current spikes outside p-level changes but there was no ingestion pattern. the root cause of hemolysis was determined to be biomaterial. the root cause of high purge pressure was determined to be biomaterial. the root cause of the vascular damage - peripheral vessel was not determined as insufficient clinical information was provided. unknown location of bleed. the root cause of renal failure was not determined as insufficient clinical information was provided. b2 outcomes attributed to adverse event missing from manufacturer device report 1220648-2024-20150.
17607 | 1220648-2024-20152 | 20427512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427512&pc=OZD | 07/15/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 07/05/2022 | 10/11/2024 |  |  | 10/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142151 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 25 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy/myocarditis was implanted with an impella cp for mechanical circulatory support. while on support, there was a vascular injury which required treatment during impella cp support. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
17608 | 1220648-2024-20153 | 20427554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427554&pc=OZD | 07/19/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 07/19/2022 | 10/11/2024 |  |  | 02/29/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195211 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Difficult to Insert | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, flows were low. this prompted the team to restart. they decided to explant and replace the pump. the explant was found to have a kinked introducer. it was noted that the pump had a difficult delivery due to native calcification. the patient survived the explant. | [5 empty] | The investigation for the reported low flow and introducer issues has been completed. the impella device was not received from the customer. however, clinical details provided were sufficient to determine a root cause. the root cause of the low flow was most likely due to a cannula kink caused by patient condition as the patient had tortuous vasculature. the root cause of the introducer issue was determined to most likely be a sheath kink due to the patient's tortuous vasculature. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17609 | 1220648-2024-20154 | 20427560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427560&pc=OZD | 06/27/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 06/27/2022 | 10/11/2024 |  |  | 02/29/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191511 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support for 4 days, the pump stopped and was replaced by an intra aortic balloon pump. the patient was already on extracorporeal membrane oxygenation support. no patient harm was reported. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the pump was inspected. upon examination, the impeller purge gap was too large. the pump stop was most likely caused by a bearing wear failure. the bearing wear failure occurred within the intended design life of 8 days. the root cause of the pump stop was determined to be bearing wear. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17610 | 1220648-2024-20155 | 20427591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427591&pc=OZD | 09/23/2024 | 11/21/2024 |  | 10/11/2024 |  |  | 09/23/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545045 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that impella cp insertion was aborted. unable to advance catheter past aortic arch due to tortuosity of the aorta. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of introducer damage was most likely due to tortuosity patient condition.
17611 | 1220648-2024-20156 | 20427630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427630&pc=OZD | 07/30/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 07/30/2022 | 10/11/2024 |  |  | 02/29/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200517 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during the prep, there was a failure to prime the device. therefore, the insertion was aborted. it was also indicated that they used their second device a couple of days ago and did not have an additional device. the patient survived the procedure. | [5 empty] | The investigation has bee completed. the data logs were reviewed and placement signal was normal at 0/0 mmhg. the prge fluid information screen was completed. the case start was completed, and the user selected yes for start impella catheter. an immediate high motor current pump stop occurred and the pump was disconnected and purge cassette was removed. purge flow was blank and purge pressure was 512 mmhg indicating that the pump was likely not given sufficient time to prime. the returned pump was visually inspected and the red lumen with flag was observed to still be in the pump with a red lumen kink and cut at the outflow location. the pump was plugged into an automated impella controller and the flow control screen immediately prompted indicating that case start had been completed. a placement signal not reliable alarm triggered. the returned purge cassette was run through de-air, connected to the returned pump and fluid was successfully purged out of the purge gap. there was no priming issue found during the case. the returned device functioned/operated to specification and the data logs showed flow present. this report is being filed as part of a retrospective review of historical records.
17612 | 1220648-2024-20123 | 20427668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427668&pc=PYX | 09/14/2024 | 10/31/2024 |  | 10/11/2024 |  |  | 09/14/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025529857 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had placed an impella rp flex and the impella 5.5 pumps for support of a 32-year-old male patient admitted in cardiogenic shock (cgs). plan for left ventricular assist device (lvad) post operatively. the impella rp flex supported for 88 hours and was then explanted due to hemolysis concerns. the plasma free hemoglobin (pfhg) was rising despite appropriate pump position. the family withdrew care and the patient expired after explant of the rp and 5.5 and being placed on the lvad. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the hemolysis has been completed. impella rp flex pump was returned. visually inspected pump and biomaterial/clot observed in pump inflow and at the impeller. no other physical abnormalities seen. data logs were reviewed and motor current trend seen, placement signal trend rising upwards and drop in flows found. patient was on crrt/central line/hm3 support. patient had pulmonary congestion issue and lv decompressed. patient was not doing well during the case. the cause of the hemolysis issue was most likely patient condition related. the instructions for use for the impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller states the following: section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
17613 | 1220648-2024-20158 | 20427810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427810&pc=OZD | 09/23/2024 | 11/21/2024 |  | 10/11/2024 |  |  | 09/23/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025503166 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Obstruction/Occlusion | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella cp implant and during the implant, the long sheath would not pass the iliofemoral calcification. short sheath placed without complication. balloon angioplasty was performed on bifurcation calcification. staff attempted to place the impella however attempts were unsuccessful. the implant was aborted. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and difficulty inserting/removing introducer has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition h6 code 2422 and 4625 were reported incorrectly on manufacturer device report 1220648-2024-20158. new codes were added to health effect - clinical code and health effect - impact code
17614 | 1220648-2024-20157 | 20427826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427826&pc=OZD | 02/09/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/09/2023 | 10/11/2024 |  |  | 03/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202259 |  | 0048-0024-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. there was hemolysis occurring during the support but, with no specified resolution. the patient survived the procedure. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the hemolysis was most likely patient condition related and related to ureteral injury. the relation to impella is unknown. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17615 | 1220648-2024-20160 | 20427878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427878&pc=OZD | 09/23/2024 | 11/21/2024 |  | 10/11/2024 |  |  | 09/23/2024 | 10/11/2024 |  |  |  |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533796 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp was placed on 83-year-old female patient, and primary percutaneous coronary intervention (ppc)i was completed, bleeding was noted around the repo sheath. femstop was placed but bleeding continued. the patient was taken back to the cath lab for impella removal. after cp removal, minimal flow was noted to the right superficial femoral artery (sfa) due to thrombus. the patient was taken to cardiovascular operating room (cvor) for repair. pump was noted to be stable following explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
17616 | 1220648-2024-20159 | 20427881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427881&pc=OZD | 02/06/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/06/2023 | 10/11/2024 |  |  | 08/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272507 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump fell out of position. attempts were made to reposition, without success. the pump was removed/replaced with no harm to the patient. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device not was received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the positioning issues was patient movement. this report is being filed as part of a retrospective review of historical records.
17617 | 1220648-2024-20162 | 20427935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427935&pc=OZD | 02/08/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/08/2023 | 10/11/2024 |  |  | 10/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290396 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding requiring one unit of packed red blood cells to be administered. it was reported there was a suspected clot in the outlet of the impella due to a sudden loss of flows with constant suction alarms despite reducing the pressure to p2. the physician unsuccessfully attempted to reposition the device under echocardiogram. as a result of the repositioning failure, the physician made the decision to bring the patient back to the catheterization laboratory and remove the impella and replace it with a new impella. this impella was removed and a new impella placed with no complications. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17618 | 1220648-2024-20161 | 20427954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20427954&pc=OZD | 12/08/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 12/08/2022 | 10/11/2024 |  |  | 08/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023271726 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the initial impella could not reach pump flows above one liter per minute. an additional cannula was utilized and a kink was observed. suction alarms and low flows were present. the device was explanted. the procedural outcome was the patient survived. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the returned complaint pump was visually inspected. no biomaterial and foreign materials were observed. no broken blades were found. cannula kink was observed. the root cause of the low pump flow was due to a cannula kink. this report is being filed as part of a retrospective review of historical records.
17619 | 1220648-2024-20164 | 20428211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428211&pc=OZD | 02/06/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/06/2023 | 10/11/2024 |  |  | 10/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290412 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a patient was exhibiting signs of hemolysis. the device's performance level was adjusted, and the patient's hemolysis resolved prior to decision made to explant the device. there was no need for further impella support/escalation of care after explant. | [5 empty] | The investigation for the reported hemolysis has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the hemolysis was most likely patient condition related. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17620 | 1220648-2024-20163 | 20428271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428271&pc=OZD | 02/10/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/10/2023 | 10/11/2024 |  |  | 08/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022116099 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had various bleedings not at the puncture site for which heparin in the purge solution was reduced. the procedure was completed and patient was successfully weaned. | [5 empty] | A.5 is unknown. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records.
17621 | 1220648-2024-20165 | 20428373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428373&pc=OZD | 05/06/2022 | 07/09/2025 |  | 10/11/2024 |  |  | 05/06/2022 | 10/11/2024 |  |  | 02/29/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194628 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had no pulse in right foot and leg was mottled. the patient was taken to cardiac catheterization laboratory where angiography showed thrombus in the right iliac with slow flow distal to the impella sheath. decision made to remove the impella and perform thrombectomy to remove thrombus in attempts to save the leg. it was noted that the explanted impella pump was not saved. | [5 empty] | The investigation for the reported thrombus and limb ischemia issues has been completed. the impella device was discarded by the customer, therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the data logs confirm the low flow/suction events. the root cause of the low flow was most likely ingested biomaterial as the clinical reported a thrombus. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | Corrections to the initial mfr report: g1-mdr reporting contact last name updated, email updated. component code changed to 925.
17622 | 1220648-2024-20169 | 20428466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428466&pc=OZD | 02/07/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/07/2023 | 10/11/2024 |  |  | 10/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024284900 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while inserting impella, physician aborted the cp placement due to contamination concerns. the team placed a new cp and support initiated without any noted harm or injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. no clinical details provided were sufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
17623 | 1220648-2024-20166 | 20428499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428499&pc=OZD | 05/25/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 05/25/2022 | 10/11/2024 |  |  | 10/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149145 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the device had low and resolving flows. the team went through all levels of troubleshooting and multiple repositions. it was decided to replace the impella cp device. a new impella pump was successfully placed. no patient harm or injury noted. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected. the returned pump was cut from the catheter. no cannula and catheter kink observed. no broken impella blades observed. no biomaterial found. the low flow cannot be reproduced because the pump cannot connect to aic (automated impella controller). the root cause of the issue could not be determined. this report is being filed as part of a retrospective review of historical records.
17624 | 1220648-2024-20167 | 20428502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428502&pc=OZD | 02/14/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/14/2023 | 10/11/2024 |  |  | 05/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP -JP |  |  | ABIOMED, INC. |  |  | US |  | 2023223055 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient had a cerebral infarction diagnosed a week after the cp was explanted. it was thought that at the time of explant perhaps there was embolism, but there is no imaging or proof of said presumption. the patient was successfully weaned off of support. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was patient condition related. this report is being filed as part of a retrospective review of historical records.
17625 | 1220648-2024-20168 | 20428541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428541&pc=OZD | 02/07/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/07/2023 | 10/11/2024 |  |  | 09/30/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023277566 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella pump was accidentally pulled back across the valve during repositioning. the physician was unable to re-cross the valve and the decision was made to explant the pump. an intra aortic balloon pump was placed for ongoing support. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the positioning issue was use related.
17626 | 1220648-2024-20170 | 20428664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428664&pc=OZD | 02/25/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/25/2023 | 10/11/2024 |  |  | 10/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024297626 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient¿s limbs lost pulses and appeared mottled and grey. the actions taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation for the reported limb ischemia issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. as per the clinical data provided, the root cause of the issue was patient condition related to small/ diseased vessels. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17627 | 1220648-2024-20171 | 20428711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428711&pc=OZD | 02/10/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 02/10/2022 | 10/11/2024 |  |  | 10/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150977 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after inserting the pump, the device did not start. the pump and the console were both exchanged. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported pump not detected issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. no problem was found because the issue was unable to be reproduced and the pump operated to specification. this report is being filed as part of a retrospective review of historical records.
17628 | 1220648-2024-20172 | 20428715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428715&pc=OZD | 02/10/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 02/10/2022 | 10/11/2024 |  |  | 06/30/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108280 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported there was an optical sensor alarm during preparation of the device. attempts to switch to a new console were unsuccessful. there was no adverse impact to the patient and support continued. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the returned pump was visually inspected and no abnormalities were observed. the pump was connected to a console and a controller error alarm reproduced. the visual fault locator was utilized and a disconnected fiber line next to the ferrule in the red handle was observed. the cause of the issue was a damaged optical fiber line in the red handle next to the ferrule. this report is being filed as part of a retrospective review of historical records.
17629 | 1220648-2024-20173 | 20428724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428724&pc=OZD | 06/01/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 06/13/2022 | 10/11/2024 |  |  | 11/30/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151588 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a bleed from the 16 fr introducer (non-abiomed product). it was also noted that the patient experienced related thrombocytopenia, as he began to bleed after returning from intensive care unit (icu). after removing the impella device, bleeding improved. | [5 empty] | The investigation for the reported access site bleed issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleed was most likely use related as the incorrect introducer was used and the act was maintained at a higher than recommended range during use of impella. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17630 | 1220648-2024-20174 | 20428872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428872&pc=OZD | 02/15/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 02/15/2022 | 10/11/2024 |  |  | 01/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022059987 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an acute patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. an ischemic stroke that was diagnosed 8 days after the pump was explanted. however, no additional details were provided regarding the treatment or management of this condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17631 | 1220648-2024-20175 | 20428880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20428880&pc=OZD | 11/01/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 11/03/2022 | 10/11/2024 |  |  | 05/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220192 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient experienced loss of blood through the interventional catheter and received blood products to address the issue. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of access site bleeding. the device was not received for evaluation. the root cause of the access site bleeding was patient condition related. this report is being filed as part of a retrospective review of historical records.
17632 | 1220648-2024-20177 | 20429008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20429008&pc=OZD | 01/10/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 01/10/2023 | 10/11/2024 |  |  | 08/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272490 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a damaged sensor and low flow issue were noted. support was maintained despite the noted issues. there was no harm to the patient. | [5 empty] | The investigation for the reported low flow and optical signal issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. data logs were reviewed and 5s parameters were within expected range for the entirety of the case, indicating no malfunction of the sensor itself. many ps low alarms and suction alarms were seen throughout support, indicating poor patient condition. at the beginning of the case, flows are as expected. they then drop as support continues to <1 l/min. the ps drops as well to around -10 mmhg. suction and ps low alarms continue throughout the case. therefore, the root cause of the low pump flow is most likely due to patient condition. no issue with the optical sensor was found. this report is being filed as part of a retrospective review of historical records.
17633 | 1220648-2024-20176 | 20429042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20429042&pc=OZD | 05/30/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 06/08/2022 | 10/11/2024 |  |  | 12/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165861 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 38 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the purge pressure was too high and a decrease in purge flow rate was also observed. the purge housing was replaced and heparin concentration in the purge liquid adjusted to a higher one with no improvement. the issues did not resolve until weaning. the patient also experienced thrombus and remained on impella support and was successfully weaned. | [5 empty] | The pump was returned and purged using a lab purge cassette there was a increase in purge pressure noted upon ramp up of the pump. the catheter was examined and no kinks were found. the pump was run in terazyme for an hour. the next day the pump was run again and purge pressure increased during pump ramp up but was within specification. the data logs were reviewed and confirm the pump had high purge pressure after ramp up that continued for the duration of the case. the root cause of the high purge pressure was most likely a result of biomaterial as the purge pressure reduced after running the pump in terazyme. this report is being filed as part of a retrospective review of historical records.
17634 | 1220648-2024-20180 | 20429082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20429082&pc=OZD | 10/29/2021 | 10/11/2024 |  | 10/11/2024 |  |  | 10/29/2021 | 10/11/2024 |  |  | 07/31/2022 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108787 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a retroperitoneal bleed requiring an unknown amount of blood transfused, which reportedly improved the patient's condition. it was reported that there was no issue with the access site bleeding. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the vascular damage has been completed. the relationship between the retroperitoneal bleed and the use of impella was determined to be unrelated as no evidence was provided to suggest a causal relationship. no product or data logs were returned for this investigation. the root cause of the reported was determined to be a result of patient condition. this report is being filed as part of a retrospective review of historical records.
17635 | 1220648-2024-20181 | 20429097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20429097&pc=OZD | 03/27/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 03/27/2023 | 10/11/2024 |  |  | 10/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276694 |  | 0048-0024-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was being prepped, when the team damaged the pump at the purge filter. the team replaced the pump and the support proceeded without harm or injury to the patient. | [5 empty] | The investigation for the reported purge leak - pump issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the bond between sidearm pressure reservoir and filter was damaged when connecting the purge cassette to the pump during initial case setup. it was noted that excessive force was used. the sidearm leaked and the decision was then made to replace the pump with a new impella. this report is being filed as part of a retrospective review of historical records.
17636 | 1220648-2024-20179 | 20429111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20429111&pc=OZD | 02/01/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/01/2023 | 10/11/2024 |  |  | 09/30/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273572 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient developed hemolysis coinciding with impella support. intervention was not required to treat the condition. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of hemolysis. the device was not received for evaluation. the root cause of the hemolysis was most likely improper positioning of the device. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17637 | 1220648-2024-20178 | 20429137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20429137&pc=OZD | 12/06/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 12/06/2022 | 10/11/2024 |  |  | 07/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269751 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, when there was a tear at the sterile sleeve. additionally, it was noted that were some ¿resolved¿ bleeding at the access site. however, no additional details were provided regarding the treatment or management of this condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. no product was returned. after reviewing the clinical information, it was determined that the cause of the sterile sleeve damage was most likely improper technique when manipulating the repo unit and the cause of the access site bleeding was most likely use related since the blue suture hub was advanced too deep into the arteriotomy. instructions for use for the related event are as follows: access site bleeding. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17638 | 1220648-2024-20182 | 20429191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20429191&pc=OZD | 05/01/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 05/01/2023 | 10/11/2024 |  |  | 12/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300187 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. prior to implantation, low purge pressure of 70 and high purge flow of 30 was noted. the surgeon decided to implant the pump anyway with continuous alarming. decision was made to exchange the pump. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the purge leak was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
17639 | 1220648-2024-20184 | 20429193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20429193&pc=OZD | 07/11/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 07/11/2023 | 10/11/2024 |  |  | 07/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257728 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella failed to prime. as a result, the decision was made to explant the device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of priming issue. the device was not returned for evaluation. the root cause of the priming issue was not determined since product was not returned and data logs were not returned as well. this report is being filed as part of a retrospective review of historical records.
17640 | 1220648-2024-20183 | 20429201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20429201&pc=OZD | 07/03/2023 | 11/18/2025 |  | 10/11/2024 |  |  | 07/03/2023 | 10/11/2024 |  |  | 01/31/2025 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024305473 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the device failed to prime. as a result, the device was explanted and replaced with a new impella cp device. the patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | Corrections to the initial mfr report: g1 mdr reporting contact email. h3 has been updated to device eval completed. h6 type of investigation code 10 has been added.
17641 | 1220648-2024-20185 | 20429280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20429280&pc=OZD | 02/17/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/17/2023 | 10/11/2024 |  |  | 08/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269750 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella device for mechanical circulatory support. during prep, there was a failure to prime the device. as a result, the decision was made to replace it with a new impella cp device. it was noted that the catheter never entered the body. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
17642 | 1220648-2024-20186 | 20429346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20429346&pc=OZD | 02/28/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/28/2023 | 10/11/2024 |  |  | 11/30/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283882 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced frequent suction events attributed to malposition of the device. additionally, the patient developed heparin-induced thrombocytopenia and hemolysis, necessitating a switch in the anticoagulation regimen from heparin to argatroban. two units of red blood cells were transfused, and repositioning of the device was performed under transthoracic echocardiography guidance. the patient remained on support until the device was successfully weaned. | [5 empty] | The investigation for the reported thrombocytopenia, hemolysis and positioning issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the data logs were not provided. the cause of the hemolysis issue was improper positioning of the device as pump repositioning resolved the issue. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17643 | 1220648-2024-20187 | 20429416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20429416&pc=OZD | 08/14/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 08/15/2022 | 10/11/2024 |  |  | 05/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228837 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implantation, there was air in the purge system alarm during the priming process. the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
17644 | 1220648-2024-20188 | 20431422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20431422&pc=OZD | 07/05/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 07/05/2022 | 10/11/2024 |  |  | 02/29/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199702 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis, the resolution for this issue is unknown. additionally, the device exhibited continuous suction even after administering a fluid bolus. the device was getting flows of 1l/minute and the patient was hemodynamically unstable, so it was decided to replace the pump. the device was explanted and replaced with a new impella cp. there was no patient harm or injury noted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the low pump flow is most likely due to the ingestion of biomaterial. the pump was reported to be 1.5cm deeper than the recommended position. the cause of hemolysis is most likely due to improper positioning. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17645 | 1220648-2024-20202 | 20432100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20432100&pc=OZD | 02/02/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/02/2023 | 10/11/2024 |  |  | 09/30/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281352 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, a delivery issue involving an impella cp pump. during the initial insertion, the pump was unable to maintain placement across the aortic valve. the pump was removed for re-insertion, however, the pump felt tight on the wire and the delivery was aborted. a second cp pump was subsequently implanted for patient support. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was possible. clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to a kinked cannula because of use related improper technique. this report is being filed as part of a retrospective review of historical records.
17646 | 1220648-2024-20190 | 20432314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20432314&pc=OZD | 12/08/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 12/08/2022 | 10/11/2024 |  |  | 07/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253693 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a difference in impella's placement signal compared to the blood pressure (bp) and cuff readings. additionally, the pump exhibited low purge flow with suction alarms. the impella was on p-5 with flows of 2.6 l/min and stable hemodynamics. echocardiogram (echo) showed impella deep at 4.1 cm. the physician pulled back the impella under echo guidance. there was no reported harm to the patient. | [5 empty] | The returned pump was connected to an aic (automated impella controller) and the placement signal reading was a steady-6/-6 mmhg. no other abnormalities were observed. there was no optical signal issue found during the case. the device functioned/operated to specification. impella is not intended to function as an arterial catheter. there was no high purge pressure found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
17647 | 1220648-2024-20189 | 20432921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20432921&pc=OZD | 07/07/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 07/07/2022 | 10/11/2024 |  |  | 03/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212153 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that there was thrombus and positioning issues during impella cp support. the impella cp was implanted successfully but popped across the valve and would not cross the aortic valve again. it was explanted to rewire for insertion, and it was noticed their clot on the cage. a new impella cp was implanted. no patient harm or injury was noted. | [5 empty] | The investigation for the reported thrombus and positioning issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the positioning issue was use related. the cause of the reported thrombus issue was patient condition related. as noted in the impella ifu: impella cp with smart assist system section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17648 | 1220648-2024-20194 | 20433247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433247&pc=OZD | 12/08/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 12/08/2022 | 10/11/2024 |  |  | 08/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265352 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was persistent oozing noted around the impella catheter at the point which it enters the repositioning sheath. the impella device was explanted due to the bleeding. the procedural outcome was the patient survived. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the cause of the access site bleeding was most likely a damaged gwru valve per the clinical details. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17649 | 1220648-2024-20197 | 20433261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433261&pc=OZD | 07/25/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 07/25/2022 | 10/11/2024 |  |  | 03/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204597 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was an optical sensor alarm, and they were unable to prime purge. the plug was disconnected from the console, and they were able to prime pump and successfully complete case with no other issues. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the available real time logs do not show any optical sensor alarms, and the placement signal is within the normal range. the logs also do not show any purge related alarms. based on the clinical information and the returned data logs, there was no optical signal issue found during the case. additionally, there was no priming issue found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
17650 | 1220648-2024-20199 | 20433279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433279&pc=OZD | 02/02/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/02/2023 | 10/11/2024 |  |  | 06/30/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235944 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while inserting a cp, the physician was unable to deliver impella. the team attempted to troubleshoot by intervention upon the native iliofemoral anatomy. delivery was abandoned due to native anatomy. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was possible. clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
17651 | 1220648-2024-20206 | 20433321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433321&pc=OZD | 03/16/2022 | 03/10/2025 |  | 10/11/2024 |  |  | 03/16/2022 | 10/11/2024 |  |  | 12/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171803 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. staff said there was a connection issue initially, had purge flow high/low pp alarms, and had console failure alarms. advised to get unplug cord and plug back in, get new purge cassette and do a fake bag change, and restart process. the patient continued on support until the device was successfully weaned. | [5 empty] | Investigation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump not detected issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
17652 | 1220648-2024-20193 | 20433389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433389&pc=OZD | 02/16/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/17/2023 | 10/11/2024 |  |  | 07/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236432 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis. the reporter indicated the pump was repositioned to address the hemolysis. it was reported that the physician noted if the hemolysis continued, it may be necessary to consider countermeasures such as lowering the flow rate. no further measures were taken, the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the hemolysis has been completed. the clinical details provided were insufficient to determine a root cause and the product was not returned therefore a root cause could not be determined. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17653 | 1220648-2024-20196 | 20433417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433417&pc=OZD | 01/16/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 01/16/2023 | 10/11/2024 |  |  | 09/30/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273544 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, a low purge flow and coinciding high purge pressure alarm occurred. the pressure level was reset from p-0 in an attempt to troubleshoot the flow rate, but issue persisted. the patient continued on support until the device was successfully weaned. no patient harm reported. | [5 empty] | The investigation has been completed for the reported issue of high purge pressure. the device was returned for evaluation. the field lt log was reviewed and a 3 minute purge bag change early in the run was observed. hpp associated with 2 mc spikes, position in aorta alarms and downward trending purge flow. the returned pump was visually inspected and biomaterial was observed in the purge gap. the cause of the issue was biomaterial. the root cause of the high purge pressure was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
17654 | 1220648-2024-20210 | 20433445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433445&pc=OZD | 07/02/2021 | 11/13/2025 |  | 10/11/2024 |  |  | 07/02/2021 | 10/11/2024 |  |  | 05/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022076656 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. it was reported that upon case start there was a delay as the impella would not click and connect to the aic.(automated impella controller) the pump was replaced with no reported harm or injury to the patient as a result of the delay. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. no data logs were provided for analysis and no product was received for evaluation. therefore, the root cause of the impella being unable to connect could not be determined. this report is being filed as part of a retrospective review of historical records. | No data logs were provided for analysis and no product was received for evaluation. therefore, the root cause of the impella being unable to connect could not be determined. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
17655 | 1220648-2024-20191 | 20433468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433468&pc=OZD | 12/19/2022 | 03/13/2025 |  | 10/11/2024 |  |  | 12/19/2022 | 10/11/2024 |  |  | 06/30/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023233711 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had atrial fibrillation, and the pump was explanted the following day. after explantation, there was bleeding at the access site that was treated by "little vascular repair." the patient survived the procedure. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the cause of the issue was not determined since the product was not returned and insufficient clinical details were provided.
17656 | 1220648-2024-20201 | 20433471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433471&pc=OZD | 02/02/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/03/2023 | 10/11/2024 |  |  | 10/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023276709 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a sign of hemolysis. the hematuria was not treated by any form of intervention. the flow rate was adjusted to address the issue. no blood products were given. the patient was successfully weaned off of support. | [5 empty] | The cause of the hemolysis cannot be determined due to insufficient clinical information, and no product or data logs back. instructions for use for the related event are as follows: hemolysis impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17657 | 1220648-2024-20192 | 20433475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433475&pc=OZD | 12/08/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 12/08/2022 | 10/11/2024 |  |  | 08/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023260337 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, it was discovered that the patient's vessel diameter was only about 4mm, and a severe multivessel lesion was identified, and external sheaths were used. the procedure was performed with an antegrade sheath on the impella side and a retrograde sheath on the contralateral side. there was no color problem in the lower extremities during the procedure, but it worsened, resulting in lower extremity ischemia. a femoral bypass was performed and the lower extremity ischemia did not improve. since the impella was in the process of weaning, the physician decided to explant the device. after removal, the blood flow of the lower extremities was confirmed and the issue improved. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device has not been returned for evaluation. however, a review of the clinical information found the following: the root cause of the leg ischemia was most likely incorrect sheath used as a 16 fr non-abiomed approved introducer was used on the patient's 4mm vessel. the 16 fr dryseal introducer is not permitted per (b)(4). instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17658 | 1220648-2024-20203 | 20433479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433479&pc=OZD | 02/06/2023 | 03/27/2025 |  | 10/11/2024 |  |  | 02/06/2023 | 10/11/2024 |  |  | 06/30/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240898 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a pump stop during impella support. the pump was able to restart, however, the motor current remained high and the pump subsequently stopped again. the pump was removed and replaced with another impella pump to continue with patient support. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump stop and high purge pressure issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the device passed all post-sterile inspection checks. the evaluation of the device noted that the field long term (lt) log was reviewed and upon activation, the placement signal showed a ventricular waveform. after 6 minutes into the run, a high motor current pump stop occurred with no purge related issues. the pump was successfully restarted and ran for 9 minutes. the purge pressure steadily increased during the 9 minutes, until another high motor current pump stop occurred and the pump was removed. the returned pump was visually inspected and biomaterial in the purge gap was observed. this correlated with the high purge pressure observed by the end of the run. ingested biomaterial was likely the cause of the initial pump stopping, and then the biomaterial migrated to the purge gap, causing the purge pressure to rise until the second high motor current pump stop occurred. the root cause of the high purge pressure could not be determined as insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
17659 | 1220648-2024-20204 | 20433481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433481&pc=OZD | 02/06/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/06/2023 | 10/11/2024 |  |  | 07/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231628 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was trouble obtaining optimal positioning of the pump, due to the device potentially caught within the papillary muscle. there was no injury sustained, nor any medical intervention required. the pump continued to support the patient until hemodynamics stabilized. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device not was received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the positioning issue was use related and related to not tightening the touhy. this report is being filed as part of a retrospective review of historical records.
17660 | 1220648-2024-20212 | 20433515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433515&pc=OZD | 01/31/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 01/31/2022 | 10/11/2024 |  |  | 02/28/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022053955 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. when attempting to place a cp from a sterile pack that had previously been opened and had the .018 wire removed. the team had replaced the wire with an .027 that prevented the pump to be delivered. the team aborted and placed a new cp with the appropriate wire. the patient survived the procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected, and no abnormalities were noted. of note, the 0.027¿ guidewire and introducer were not returned. the root cause of the delivery issues was the use of the incorrect guidewire. per ifu 0048-9007, only the guidewire supplied within the impella cp optical insertion kit or an abiomed approved alternative guidewire should be used for insertion of impella cp optical. this report is being filed as part of a retrospective review of historical records.
17661 | 1220648-2024-20200 | 20433548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433548&pc=OZD | 02/02/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/02/2023 | 10/11/2024 |  |  | 09/30/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022141822 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump fell out of the ventricle during attempted repositioning to move the device away from the papillary muscle. the pump was unable to re-cross the valve, therefore, removed/replaced. no patient harm was reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device not was received from the customer and therefore, an evaluation of the device was not possible. the cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
17662 | 1220648-2024-20214 | 20433569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433569&pc=OZD | 03/25/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 03/25/2023 | 10/11/2024 |  |  | 12/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298359 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during the prep, there was a failure to prime the device. the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that no abnormalities were found on the returned pump. the pump was reprogrammed and ran through case start successfully. the pump was then run in a bucket of water successfully. the product operated to specification. the pump and purge cassette were likely primed on the aic before initiating case start, resulting in the aic skipping the priming steps during the case. the log shows adequate purge pressures and purge flows during the case. the log also shows that the pump was attempted to be started outside of the body. the cause of the priming issue was use related and related to the pump and purge cassette already priming prior to the user initiating case start steps on the console. the cause of the pump not starting was use related and related to starting the pump outside of the body. this report is being filed as part of a retrospective review of historical records.
17663 | 1220648-2024-20207 | 20433636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433636&pc=OZD | 04/24/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 04/24/2023 | 10/11/2024 |  |  | 12/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321857 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the physician was unable to load the impella cp device onto 0.018 wire, easy wire guide out. in emergent situation, physician opened a new impella cp. the initial impella cp device was never connected to the automated impella controller (aic). no harm was done to the patient. | [5 empty] | The root cause of the delivery issue was unable to be determined because no product was returned for evaluation and no data logs, insufficient clinical information was provided. this report is being filed as part of a retrospective review of historical records.
17664 | 1220648-2024-20209 | 20433717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433717&pc=OZD | 03/08/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 03/08/2023 | 10/11/2024 |  |  | 11/30/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295795 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the pump was placed in surgical mode, then when attempting to go back to pressure p1 the motor current increased and the pump stopped. several unsuccessful attempts were made to restart the pump. there was no patient harm reported, the impella was removed, bypass was initiated and the coronary artery bypass procedure successfully completed. | [5 empty] | The investigation for the reported pump stop issue has been completed. the log shows the pump restart attempts. surgical mode is turned on and off five times. the second time the pump is put in surgical mode is for about a minute and a half. when the pump is taken out of surgical mode the pump is attempted to be restarted a few times. surgical mode is then turned on for a period of about thirty seconds and then turned off. unsuccessful pump restarts are attempted. surgical mode is then turned on for approximately 13 minutes. the pump is then successfully restarted. surgical mode is enabled for the fifth and final time for about five minutes. the pump is successfully started after this event. there are suction events throughout the duration of the case. the cause of the pump stop was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17665 | 1220648-2024-20211 | 20433723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433723&pc=OZD | 07/27/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 07/27/2022 | 10/11/2024 |  |  | 04/30/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220587 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the pump had a purge leak. the impella cp device was explanted and replaced with an impella 5.5. | [5 empty] | The investigation for the reported purge leak issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the purge leak was most likely sidearm reservoir damage. this report is being filed as part of a retrospective review of historical records
17666 | 1220648-2024-20195 | 20433750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433750&pc=OZD | 01/16/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 01/16/2023 | 10/11/2024 |  |  | 08/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262821 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp support of just over two hours. during the support there was placement signal issue of being very low the entire time and so the pump was explanted to address the issue. . no harm or injury noted. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. data logs showed recurrent "placement signal low" alarms throughout support. the ps trended upwards throughout support with no change in p-level or position. there was no evidence of optical sensor malfunction, as the optical sensor metrics and 5s parameters were within specification. therefore, it was determined that the cause of the optical signal issues was most likely patient condition related. this report is being filed as part of a retrospective review of historical records.
17667 | 1220648-2024-20198 | 20433769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433769&pc=OZD | 01/26/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 08/06/2024 | 10/11/2024 |  |  | 11/30/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023279451 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient experienced a cerebral hemorrhage. it is unknown if the issue was caused due to the deice. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. the procedural outcome was patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of stroke/cva. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records.
17668 | 1220648-2024-20205 | 20433801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433801&pc=OZD | 12/03/2021 | 05/22/2025 |  | 10/11/2024 |  |  | 03/16/2022 | 10/11/2024 |  |  | 10/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022138508 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis. action taken to resolve the issue are unknown. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of hemolysis. the device was not received for evaluation. the root cause of the hemolysis was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. | The cause of the hemolysis was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
17669 | 1220648-2024-20213 | 20433819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433819&pc=OZD | 03/31/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 03/31/2022 | 10/11/2024 |  |  | 12/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174834 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while preparing the catheter for insertion, as soon as the impella was plugged in, there was a defective error. a new pump opened and prepped and inserted. | [5 empty] | The investigation for the reported pump not detected issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the imc logs showed that there was an issue with the console reading the pump's eeprom data, triggering the "impella defective" error upon plugging in the pump. no damage or abnormalities were found on the returned pump. when the pump was connected to the aic (automated impella controller), it went through case start as normal and activated with no issues. therefore, the cause of the case start issue was not determined since the "impella defective" alarm was not reproduced. this report is being filed as part of a retrospective review of historical records.
17670 | 1220648-2024-20208 | 20433822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20433822&pc=OZD | 07/06/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 07/06/2023 | 10/11/2024 |  |  | 05/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231141 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the impella was implanted the physician noticed some blood leaking from the connection of the impella red plug area and the gray tube. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
17671 | 1220648-2024-20217 | 20434064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20434064&pc=OZD | 08/10/2022 | 10/11/2024 |  | 10/11/2024 |  |  | 08/10/2022 | 10/11/2024 |  |  | 04/30/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226132 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. aic (automated impella controller) did not show start screen, no alarms present. abiomed support guided the clinical team via face time on how to turn on the pump after insertion. accidentally the pump was started outside the body with the easy guide lumen still in. the impella was explanted and replaced with a new impella cp to address the issue. there was no reported harm or injury to the patient. | [5 empty] | Upon visual inspection of the pump the red lumen was found still in the pump. no other abnormalities were observed. based on the provided clinical information, the cause of the pump not starting was use related. this report is being filed as part of a retrospective review of historical records.
17672 | 1220648-2024-20218 | 20434108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20434108&pc=OZD | 12/21/2021 | 10/11/2024 |  | 10/11/2024 |  |  | 12/21/2021 | 10/11/2024 |  |  | 07/31/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121777 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Obstruction of Flow | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the case was delayed when the purge fluid did not appear to exit. consequently, the decision was made to explant the device and replace it with a new impella cp device. there was no reported harm or injury to the patient. | [5 empty] | The data logs were returned and confirm that the purge cassette was deaired and that the case start was attempted multiple times. the logs show no signs of the pump failing to prime. the pump was returned and using a lab cassette was able to prime and produce purge fluid without any issues. it is likely that the physician and hospital staff did not give the purge fluid enough time to move through the system. there are no issues with the device. the root cause of the pump being unable to prime was determined to be a use issue. this report is being filed as part of a retrospective review of historical records.
17673 | 1220648-2024-20216 | 20434109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20434109&pc=OZD | 05/19/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 05/30/2023 | 10/11/2024 |  |  | 12/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322750 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were two instances of hematuria during impella support. no further intervention was noted. the pump remained implanted. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the impella device was not received from the customer, and no issues or abnormalities were seen in the data logs. the pump performed within specification for the duration of the case with no positioning issues, change in pulsatility, and only 2 suction alarms throughout the case. based on this and sparse clinical evidence, the cause of the hematuria is not determined and relationship between the hematuria and impella is unknown. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17674 | 1220648-2024-20219 | 20434308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20434308&pc=OZD | 12/09/2021 | 10/11/2024 |  | 10/11/2024 |  |  | 12/09/2021 | 10/11/2024 |  |  | 09/30/2023 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140956 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a controller error resulted in placing the pump on a back-up controller. the patient continued on support until the device was successfully weaned. | [5 empty] | Data analysis: imc logs confirmed the reported failure mode given there was a controller error alarms (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. see the attachments section for ic5388 aic logs from complaint date. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. see attachments section for ic5388 rt logs transient loss of optical signals. device analysis: console (ic5388) was sent back to danvers fs for further investigation. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging before being returned to the customer, fs performed a full functional check on the console and optical system (0042-7316). this report is being filed as part of a retrospective review of historical records.
17675 | 1220648-2024-20220 | 20434425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20434425&pc=OZD | 02/13/2023 | 10/11/2024 |  | 10/11/2024 |  |  | 02/13/2023 | 10/11/2024 |  |  | 08/31/2024 |  | 10/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251519 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, a hematoma at the groin site was noted. no further information was given on the resolution or intervention. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. no product or logs were returned for investigation. the cause of the access site bleeding was most likely use related and related to not properly placing the repo sheath with angle matching. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17676 | 1220648-2024-19982 | 20415540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20415540&pc=OZD | 01/27/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 01/27/2023 | 10/10/2024 |  |  | 12/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024285822 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the pump was unable to be delivered, and the axillary access site was not an option either. therefore, the procedure was aborted due to patient vascular anatomy. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, due to patient condition. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
17677 | 1220648-2024-19983 | 20415550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20415550&pc=OZD | 02/01/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/01/2023 | 10/10/2024 |  |  | 09/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274970 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a pump falling out of the left ventricle due to physician mistake. the pump could not re-cross the valve, therefore, was removed and replaced with a new impella. | [5 empty] | The cause of the issue was wireless re-access. the impella device is not indicated for wireless re-access per ifu 0048-9007. this report is being filed as part of a retrospective review of historical records.
17678 | 1220648-2024-19984 | 20415564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20415564&pc=OZD | 02/02/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/02/2023 | 10/10/2024 |  |  | 06/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2023231723 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, for 34 hours when the pump was explanted due to a pump stop that was recurrent. repositing of device was attempted but did not resolve the issue. a thrombus was observed in the pump during the explant. no patient harm was reported. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the pump stop was most likely biomaterial. this report is being filed as part of a retrospective review of historical records.
17679 | 1220648-2024-19985 | 20415601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20415601&pc=OZD | 02/02/2023 | 10/09/2025 |  | 10/10/2024 |  |  | 02/02/2023 | 10/10/2024 |  |  | 08/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250818 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the pump was not recognized by aic (automated impella controller). the staff guided them through case start a second time, pump got recognized, but didn't start later. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | Customer tried to implant cpsa, pump was not recognized by automated impella controller (aic). guided them through case start a second time, pump got recognized, but didn't start later. customer exchanged pump. pump was not recognized by aic again. case was restarted on aic, this time successful, pump did start. case was completed with on site support without problem. |  |  |  |  | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the investigation into the reported issue is complete. pump was recognized by aic, started and ran successfully. no damage or abnormalities were found.
17680 | 1220648-2024-19986 | 20415631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20415631&pc=OZD | 02/08/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/08/2023 | 10/10/2024 |  |  | 10/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278936 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was lower than expected purge flow. the pump was removed and replaced. no patient harm was reported. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the product evaluation revealed that there was a block in the purge line in the motor housing resulting from built up biomaterial. the root cause of the high purge pressure is due to biomaterial. this report is being filed as part of a retrospective review of historical records.
17681 | 1220648-2024-19987 | 20415644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20415644&pc=OZD | 07/03/2024 | 10/10/2024 |  | 10/10/2024 |  |  | 08/26/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025439170 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a pump stop that occurred. the motor stopped working and the staff attempted to restart the pump three times to no avail. it was further reported that the patient expired while on impella cp support. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | A.5 is unknown. the investigation has been completed. per the clinical information, the pump stopped - motor current high. no product was returned. the data logs confirmed motor current high pump stops and unsuccessful restart attempts. the real time logs showed a shift in motor current with saturated flow momentarily. later, a motor current spike of > 1400 ma happened after restart. the root cause of the pump stop was most likely biomaterial. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17682 | 1220648-2024-19988 | 20415700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20415700&pc=OZD | 09/12/2024 | 11/15/2024 |  | 10/10/2024 |  |  | 09/12/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025510546 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support the patient had pericardial bleeding. the patient had nearly two liters come out of their chest tubes. around 17 units of blood products were given to the patient. the source of the bleeding was unknown. the clinical representative confirmed that the bleeding was from chest compressions. it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The investigation of vascular damage - great vessel has been completed. the pump was not returned for investigation. clinical information mentioned that the source of the bleeding was unknown and the patient was taken off care due to poor prognosis. therefore, the root cause of the vascular damage great vessel issue was not determined since the product was not returned and insufficient clinical information regarding the bleeding was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
17683 | 1220648-2024-19990 | 20415804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20415804&pc=OZD | 09/16/2024 | 11/15/2024 |  | 10/10/2024 |  |  | 09/16/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519256 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and there was access site bleeding. the two perclose were tightened to slow the bleeding. the patient continued to have intermittent bleeding. the doctor gave epinephrine-lidocaine injection at the tract site. it was discovered that there was a scalpel nick during insertion. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleeding has been completed. the impella was not returned for evaluation. however, clinical details were sufficient to determine root cause. the root cause of access site bleeding is determined to be use issue because the scalpel was used for incision at the access site causing the bleed and the bleeding was controlled by tightening the perclose.
17684 | 1220648-2024-19989 | 20415808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20415808&pc=OZD | 09/15/2024 | 11/20/2024 |  | 10/10/2024 |  |  | 09/16/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1089657 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that there was defective material in the purge compartment of the automated impella controller (aic). | [5 empty] | The investigation for the reported automated impella controller issue has been completed. the console was returned for evaluation. the purge disc retainer was confirmed to be broken. the cause of the issue was a broken purge retainer. | A.1, a.4 and a.5 are unknown. the console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
17685 | 1220648-2024-19991 | 20415896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20415896&pc=OZD | 01/13/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 01/13/2023 | 10/10/2024 |  |  | 08/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED.INC. |  |  | US |  | 2023256625 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited a pump stop. all troubleshooting measures, inclusive of (aic) automated impella controller swap, failed to remedy the issue. the pump was removed and replaced after a 20-minute pump run. there was not any harm or injury noted. patient was stable post explant. | [5 empty] | The device was returned for evaluation, an electrical resistance test was conducted, and an open line was discovered. the red handle was opened, and it was found that the open line occurred on the catheter side in the red handle and the open line was the white motor cable. necking of the cable was observed. no other abnormalities observed. 5s parameters were within specification. the cause of the pump stop was an open electrical line, specifically the white motor cable on the catheter side in the red handle. as noted in the impella ifu: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
17686 | 1220648-2024-19992 | 20415925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20415925&pc=OZD | 01/25/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 01/25/2023 | 10/10/2024 |  |  | 11/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023280847 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had very small and narrow vessel and catecholamine usage led to malperfusion of the leg in combination with the repo sheath. malperfusion on impella site led to decision to remove impella early. patient was reported stable post explant. | [5 empty] | The investigation on limb ischmia is completed. the device and the data were not returned. the cause of the limb ischemia was most likely patient condition related, specifically the patient's small vasculature. as noted in the impella ifu: impella cp with smartasssistsmartassist system section: potential adverse events (united states) ¿section:¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17687 | 1220648-2024-19993 | 20415957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20415957&pc=OZD | 04/26/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 06/15/2023 | 10/10/2024 |  |  | 01/31/2025 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306652 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/15/2023 |  |  |  |  | The complainant reported a patient with unknown demographics noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient experienced hypotension a few minutes after impella insertion. during the procedure, the patient became bradycardic. therapy with katecholamine had to be administered and the device was explanted. patient was reported stable post procedure. | [5 empty] | No product was returned. as per clinical patient experienced ischemic stroke issues with bradycardia and device explanted. no data logs were returned for review. the cause of the stroke issue was most likely patient condition related and unknown relation due to insufficient clinical information and with logs unavailable to review. as noted in the impella ifu: impella cp with smart assist system impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
17688 | 1220648-2024-19994 | 20416001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416001&pc=OZD | 01/27/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/03/2023 | 10/10/2024 |  |  | 07/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023264686 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported signs of hemolysis during support, any medical intervention is unknown. additionally, a pump stop occurred on day 22 of impella support. the patient was deemed stable, so the pump was removed. patient survived the procedure and was reported stable. | [5 empty] | The root cause of the hemolysis was unable to be determined as no product, no logs, and limited clinical details were provided. the root cause of the pump stop was unable to be determined as no product, no logs, and limited clinical details were provided. the pump stop occurred on the 22nd day of support which is beyond the 8 day design life per design trace matrix 0046-9910. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17689 | 1220648-2024-19995 | 20416078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416078&pc=OZD | 08/23/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 10/12/2021 | 10/10/2024 |  |  | 05/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022083585 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis and renal failure. the hemolysis was resolved by repositioning the impella. additionally, it was reported renal failure occurred on (b)(6) 2021 and since the impella was used in conjunction with the extra-corporeal membrane oxygenation this created a situation where the cause/relationship is unknown. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into hemolysis and renal failure have been completed. the data logs were also not returned for investigation. the cause of the hemolysis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. the cause of the renal failure and its relation to impella were not determined since the relation of the reported hemolysis to impella is unknown. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: renal failure impella cp with smart assist section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17690 | 1220648-2024-19996 | 20416126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416126&pc=OZD | 06/27/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 08/24/2021 | 10/10/2024 |  |  | 05/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022083589 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, demographics unknown, was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis and kidney failure. reportedly on (b)(6) 2021 the patient experienced thrombocytopenia. due to the facility restrictions no further information was able to be obtained. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the hemolysis, thrombocytopenia and renal failure have been completed. the clinical team reported that hemolysis occurred within 24 hours of impella support. the team additionally reported that kidney failure and thrombocytopenia occurred during impella support. however, further information was not obtained due to hospital covid-19 restrictions. the cause of the hemolysis and its relation to impella were not determined since neither the product nor the data logs were returned. the causes of the renal failure and thrombocytopenia and their relation to impella were not determined. the data logs and product were not provided, and sufficient clinical information was also not provided. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: hemolysis impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ renal/kidney failure impella cp with smart assist section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ thrombocytopenia impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17691 | 1220648-2024-19997 | 20416135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416135&pc=OZD | 08/18/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 08/31/2021 | 10/10/2024 |  |  | 05/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022089482 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis and on (b)(6) 2021 the patient developed thrombocytopenia. no further information was provided regarding treatment, due to the facility restrictions no further information was able to be obtained. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into hemolysis and thrombocytopenia have been completed. the cause of the hemolysis was unable to be determined without additional details. the patient experienced thrombocytopenia which is unrelated to impella. the cause of the patient injury/ thrombocytopenia was patient condition and the relation to impella was determined to be unknown/unrelated. the cause of the hemolysis was not determined. the clinical details provided were insufficient to determine a root cause, data logs and product were also not returned. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: hemolysis impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ thrombocytopenia impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17692 | 1220648-2024-19998 | 20416146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416146&pc=OZD | 09/02/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 09/02/2021 | 10/10/2024 |  |  | 03/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022060626 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, repeated suction and position alarms occurred. the impella cp device was explanted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow and positioning issue. the device was not returned for evaluation. the root cause of the low pump flow and positioning issue was not determined since neither product, nor the data logs were returned. this report is being filed as part of a retrospective review of historical records.
17693 | 1220648-2024-19999 | 20416167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416167&pc=OZD | 09/03/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 09/03/2021 | 10/10/2024 |  |  | 07/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110300 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, persistence low flows, suction, and position alarms were observed. the impella was explanted, and it was confirmed no clot was present. following reinsertion, the issue persisted. no further information available. the patient survived the procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the suction/low flow was unable to be determined as not product or logs were returned for analysis. the pump in aorta alarms likely occurred as a result of the low flow/suction conditions this report is being filed as part of a retrospective review of historical records.
17694 | 1220648-2024-20000 | 20416201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416201&pc=OZD | 09/05/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 09/14/2021 | 10/10/2024 |  |  | 05/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089494 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis and thrombocytopenia during impella support. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation into the hemolysis and thrombocytopenia has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the cause of the hemolysis was not determined. the root cause of the thrombocytopenia was likely cause is due to the patient¿s condition. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17695 | 1220648-2024-20001 | 20416235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416235&pc=OZD | 08/24/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 03/23/2022 | 10/10/2024 |  |  | 01/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022060043 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | Ischemia | Unspecified Vascular Problem | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient experienced bloody pneumothorax, hemolysis, and limb ischemia during the support. additionally, the pump stop occurred on day 22 of support. action taken to resolve the issues are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation into the vascular damage - peripheral vessel, pump stop, hemolysis, limb ischemia - reversible has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the pneumothorax was determined to be unrelated to impella as no evidence was provided to suggest a causal relationship between the pneumothorax and the use of impella. the root cause of the hemolysis was determined to be unknown relation to impella as no information regarding this event was provided. the root cause of the leg ischemia was determined to be unknown relation to impella as no information regarding the location, troubleshooting, or patient's vessel conditions was provided. the root cause of the pump stop was unable to be determined as no product or data was returned for this investigation. as noted in the impella ifu: impella cp with smart assist system hemolysis section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ pump stop section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17696 | 1220648-2024-20002 | 20416270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416270&pc=OZD | 10/30/2021 | 11/03/2025 |  | 10/10/2024 |  |  | 12/14/2021 | 10/10/2024 |  |  | 07/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110828 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the hemolysis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17697 | 1220648-2024-20003 | 20416444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416444&pc=OZD | 09/20/2024 | 11/13/2024 |  | 10/10/2024 |  |  | 09/20/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479181 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 being placed to support a patient admitted in for surgery. the team tried to place a 5.5 but delivery failed and was replaced by the cp instead. no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20003.
17698 | 1220648-2024-20004 | 20416467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416467&pc=OZD | 09/20/2024 | 11/15/2024 |  | 10/10/2024 |  |  | 09/20/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522452 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient was experiencing hemolysis. the decision was made to explant the device due to hemolysis. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the hemolysis has been completed. complaint device not returned for analysis. data log reviewed which showed that no motor current (mc) spike observed. no suction alarms observed. a few positioning issue alarms observed which resolved with trouble shooting. no indication observed to confirm the rc for reported issue. the cause of the hemolysis issue cannot be determined since the complaint device not returned and insufficient clinical data provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
17699 | 1220648-2024-20005 | 20416481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416481&pc=OZD | 09/20/2024 | 10/10/2024 |  | 10/10/2024 |  |  | 09/20/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546184 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support due to bradycardia. while on support, the patient slowly started to lose impella flow and eventually had continuous suction without being able to flow. surgeon was concerned about clot ingestion, and opted to replace with a new cp. additionally, multiple units of product was given in response to bleeding concern. there was a total of 2100 cc¿s of blood out of chest tubes. endotracheal tube was exchanged due to balloon rupture. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17700 | 1220648-2024-20006 | 20416484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416484&pc=OZD | 09/20/2024 | 11/25/2024 |  | 10/10/2024 |  |  | 09/20/2024 | 10/10/2024 |  |  | 06/30/2025 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367185 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had impella 5.5 pump delivery fail. upon placing the impella 5.5 through axillary graft there was a 90 degree bend in the anatomy close to the carotid branch off. several wires were passed to make a rail for impella to get past bend. contrast was shot into axillary and retrograde dissection noted. all attempts were aborted to prevent dissection from worsening. it was noted that the patient remained stable. | [5 empty] | The investigation into vascular damage - peripheral vessel and delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issues and vascular damage - peripheral vessel was most likely patient-condition related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20006. b2 missing. e4 should have been left blank. h1 type of reportable event reported incorrectly. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿
17701 | 1220648-2024-20008 | 20416617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416617&pc=OZD | 06/28/2024 | 11/16/2024 |  | 10/10/2024 |  |  | 09/24/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025502477 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality (loqi) indicator impella registry indicating an allegation of delivery issue / kink : ¿impella 5.5 would not advance, impella was found to be kinked and attempts to prime the pump were unsuccessful¿. there was no adverse event reported as a result of the event. patient outcome was survived. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since there was difficulty advancing through innominate. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20008. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
17702 | 1220648-2024-20007 | 20416618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416618&pc=OZD | 09/24/2024 | 11/18/2024 |  | 10/10/2024 |  |  | 09/24/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534725 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support developed oozing from the access site. a stitch was added which resolved the bleeding. the following day, bleeding was noted from the jackson-pratt drain in the 5.5 insertion site. one unit of packed red blood cells was given. the patient remained stable and survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
17703 | 1220648-2024-20010 | 20416624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416624&pc=OZD | 01/20/2022 | 10/09/2025 |  | 10/10/2024 |  |  | 01/20/2022 | 10/10/2024 |  |  | 10/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149831 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, pump had inaccurate placement signal, low pump flow and access site bleeding/hematoma. device was repositioned to resolve the reported observations. patient survived the procedure. | The complainant reported an inaccurate placement signal during impella support. the physician was able to proceed with support without any harm to the patient. it was noted that the pt had a large hematoma at impella site. it was explained hematoma needs to be resolved and could be altering impella position. in addition, new alarm diastolic pressure low sounded. went to p0 then back to p2. alarm was resolved for a few minutes, then re-sounded. blood products were complete and discussed possibility of blood clots forming and blocking part of the flow of the impella. |  |  |  |  | The impella device has been returned for evaluation and an investigation is currently underway. a supplemental mdr will be filed upon completion of the investigation. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system (access site bleeding) section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being filed as part of a retrospective review of historical records. the field lt log was reviewed and ventricular placement signal (ps) waveforms with suction were observed at the beginning of the run. 3 days into support, the ps issue resolved and impella support continued. the returned pump was connected to a console and the placement signal was 2/2 (2) mmhg. no visual abnormalities were observed. the cause of the optical placement signal alarm issue was most likely improper positioning.
17704 | 1220648-2024-20009 | 20416625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416625&pc=OZD | 09/25/2024 | 11/15/2024 |  | 10/10/2024 |  |  | 09/25/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1107443 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) fell to the floor and has a broken housing. no patient involvement has been reported. | [5 empty] | The investigation of automated impella controller (aic) - damage before therapy has been completed. the reported damage to the housings was confirmed. there was damage observed with both the front a rear housings. no other issues observed aside from the physical damage to the front and rear housing. root cause of this issue was use related. console fell which caused damage to the front and rear housing of the console. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-20009. d4 model number. e4 should have been left blank. g1 reporting contact fax number missing. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
17705 | 1220648-2024-20011 | 20416634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416634&pc=OZD | 09/26/2024 | 11/19/2024 |  | 10/10/2024 |  |  | 09/26/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547071 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following implant of an impella 5.5 pump for patient support, heparin had to be increased due to the presence of a thrombus in the left ventricle. additional intervention was not required and support continued with the implanted pump. the patient remains on impella support. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
17706 | 1220648-2024-20012 | 20416663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416663&pc=OZD | 02/24/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 03/23/2022 | 10/10/2024 |  |  | 08/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127381 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. while on support, the patient developed hemolysis. no further information was available. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17707 | 1220648-2024-20013 | 20416677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416677&pc=OZD | 05/17/2021 | 05/22/2025 |  | 10/10/2024 |  |  | 03/23/2022 | 10/10/2024 |  |  | 01/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022057160 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis. action taken was unknown. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17708 | 1220648-2024-20014 | 20416698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416698&pc=OZD | 04/12/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 04/12/2022 | 10/10/2024 |  |  | 11/30/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165720 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, there were low pump flows and suction during startup. the pump was removed and replaced with no harm to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs were not returned. visual inspection of the returned pump revealed a significant mass of biomaterial on and around the impeller. no other abnormalities were found. therefore, it was determined that the root cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
17709 | 1220648-2024-20015 | 20416723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416723&pc=OZD | 04/23/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 04/23/2022 | 10/10/2024 |  |  | 12/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180703 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during insertion, the device was prepped and attempted to advance but stenosis in aortic and after multiple attempts, device kinked in aortic arch. wire was pulled an able to slowly pull into iliac and straighten device and remove and replaced. no patient harm was reported. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that kinks in the catheter as well as a kink in the cannula were found. the root cause of the reported issue was determined to be patient condition as the patient had stenosis in their aorta. this report is being filed as part of a retrospective review of historical records.
17710 | 1220648-2024-20016 | 20416767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416767&pc=OZD | 04/28/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 04/28/2022 | 10/10/2024 |  |  | 10/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022139873 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a patient had endured hemolysis during use of the device as evident by urine color. there was no intervention administered, other than reducing the performance level of the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17711 | 1220648-2024-20017 | 20416964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416964&pc=OZD | 05/06/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 05/06/2022 | 10/10/2024 |  |  | 12/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180814 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was a pump stop during use of the device. the catheter was advanced too far into the left ventricle and the pig tail was caught on the shaft of the cannula. attempts were made to restart pump but the device would only flow 0.5-1.5 l/min. the impella cp was removed and a new impella cp device was inserted. no patient harm was reported. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that product was returned and inspected. initially, suspected biomaterial and mold was visualized in the outflow cage. after soaking the pump, there was some dark material left in the outflow cage adhered to the impeller. the pump was successfully run at p4 for 5 minutes. the suspected biomaterial was kicked out of the outflow cage and seen floating in the water. the logs show that the plug was not fully connected. the optical sensor error was most likely due to improper usage. the root cause of the issue was biomaterial. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17712 | 1220648-2024-20018 | 20416971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20416971&pc=OZD | 02/17/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/17/2023 | 10/10/2024 |  |  | 04/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226526 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis. a foley was placed upon arrival and the patient¿s urine was clear yellow, then started turning to pink urine and eventually amber in color. the impella was repositioned and reportedly the urine was clearing and becoming progressively clearer. the patient remained on impella support and was successfully weaned. patient survived the procedure. | [5 empty] | The investigation for the reported hemolysis has been completed. the product was returned, and the hemolysis was confirmed. the root cause of the hemolysis was determined to be due to improper position / use issue related. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17713 | 1220648-2024-20019 | 20417018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417018&pc=OZD | 02/24/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/24/2023 | 10/10/2024 |  |  | 09/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276815 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient with significant calcified right iliac vessels required to have an impella cp pump implanted. the cp pump could not get passed the aortic bifurcations. the pump and the 14fr short sheath were removed and a 14fr long sheath was inserted but that could not be fully inserted. the procedure was aborted. there was no severe adverse event to the patient as a result of this incident. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was patient condition related to calcified vessel. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17714 | 1220648-2024-20021 | 20417047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417047&pc=OZD | 05/24/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 05/24/2022 | 10/10/2024 |  |  | 11/30/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159561 |  | 0048-0024-JP |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient had limb ischemia resulting from stoppage of heparin. no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
17715 | 1220648-2024-20020 | 20417048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417048&pc=OZD | 09/11/2024 | 11/14/2024 |  | 10/10/2024 |  |  | 09/11/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521619 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, bleeding was noted at the impella site. manual pressure was held, two units of red blood cells were given and an additional perclose suture was placed. bleeding was successfully controlled. the patient expired while on support. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17716 | 1220648-2024-20022 | 20417060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417060&pc=OZD | 02/24/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/24/2023 | 10/10/2024 |  |  | 12/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298726 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that upon suturing of repo sheath by physician, a large hematoma developed. sheath position was manipulated, pressure was held, and site further sutured which resolved hematoma. hematoma developed again after start of heparin drip. patient support was escalated to impella 5.5. patient survived the procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on clinical description, the root cause of the access site bleeding was most likely due to lack of vessel recoil on the sheath. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17717 | 1220648-2024-20023 | 20417187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417187&pc=OZD | 06/25/2022 | 07/09/2025 |  | 10/10/2024 |  |  | 06/25/2022 | 10/10/2024 |  |  | 02/29/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193806 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was an impella cp optical sensor error. the staff attempted to troubleshoot by automated impella controller (aic) swap but, found they had to replace the pump. echocardiogram showed impella 3.2 and inlet in-obstructed indicating presence of thrombus. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the optical signal issue has been completed. the returned pump was connected to an aic (console) and the controller error was reproduced with ps 3000 mmhg. pump had not begun case start indicating an out-of-box failure. the light test was performed and light loss was confirmed in the red handle. necking was observed on the portion of optical fiber within the red handle. twisted cabled at the patient cable kink protector were observed. the patient cable was attempted to be twisted and it freely rotated indicating and insufficient bond between the patient cable and the kink protector. the cause of the out-of-box optical signal issue was a damaged optical fiber line within the red handle. evidence of twisted cable lines and optical fiber necking due to improper bonding of the patient cable to the red handle kink protector from the manufacturing process were observed. corrections to the initial mfr report: g1: mdr reporting contact email updated. h6: type of investigation 10, investigation findings 4203, and investigation concl 4307, 25 added. h10: investigation summary.
17718 | 1220648-2024-20024 | 20417203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417203&pc=OZD | 02/25/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/25/2023 | 10/10/2024 |  |  | 10/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290308 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the complainant reported that upon explant of the impella cp, a large intact thrombus was also removed from the femoral artery. there were no notes of ischemia or harm to the leg as a result of the thrombus. patient survived the procedure. | [5 empty] | The investigation for the reported thrombus issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the thrombus is most likely patient condition as there was no complaint of issue with the operation of impella, the clot was likely picked up on explant. the root cause of the thrombus observed on explant was most likely patient condition. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17719 | 1220648-2024-20026 | 20417239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417239&pc=OZD | 06/28/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 06/28/2022 | 10/10/2024 |  |  | 03/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200245 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a thrombus and low impella cp pump flow. the cp was explanted due the drop in flows. when the device was removed, a thrombus of two centimeters in size was attached from the inside of the cannula to the discharge part. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
17720 | 1220648-2024-20025 | 20417241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417241&pc=OZD | 06/27/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 06/27/2022 | 10/10/2024 |  |  | 03/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202193 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, pump failed to start. a new impella was used without any harm to the patient. | [5 empty] | The investigation for the reported "pump did not start" issue has been completed. the cause of the pump stop was not determined since the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
17721 | 1220648-2024-20027 | 20417258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417258&pc=OZD | 09/22/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 09/22/2022 | 10/10/2024 |  |  | 06/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236823 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. the complainant had a cp placed for support via vasculature that was calcified and peripheral artery disease (pad). when positioned within the left ventricular, there was a kink observed. this prompted the team to replace the pump. the second pump was successfully placed. | [5 empty] | The investigation for the reported product damage issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the product damage was most likely a catheter kink. the catheter kink was likely caused by the tortuous and heavily calcified vessels. this report is being filed as part of a retrospective review of historical records.
17722 | 1220648-2024-20028 | 20417300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417300&pc=OZD | 08/01/2022 | 04/16/2025 |  | 10/10/2024 |  |  | 08/01/2022 | 10/10/2024 |  |  | 11/30/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160952 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | White | Injury | Yes | Yes | Crack | Fluid/Blood Leak | Increase in Pressure |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to being implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding due to high activated clotting time (act), therefore heparin in purge fluid was withheld for two hours. there was noted high purge pressure and low purge flow afterwards. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation for the reported issues has been completed. the field long term event log was reviewed and high purge pressure alarms were confirmed. the returned pump was visually inspected and biomaterial was observed on the impeller but no biomaterial in the purge gap. the purge cassette was connected to a pump and a yellow luer leak was observed. no purge leak was reported in the field. the sidearm was then cut off, syringe needle was connected to bypass the sidearm, and no high purge pressure was able to be reproduced. the cause of the high purge pressure was most likely biomaterial. the cause of the product damage was sidearm yellow luer damage. the root cause of the access site bleeding was determined to be anticoagulation management as it was reported that a patient on impella cp suffered from bleeding due to high activated clotting time, therefore heparin in purge fluid had been changed to 0 for two hours. this report is being filed as part of a retrospective review of historical records. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17723 | 1220648-2024-20029 | 20417329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417329&pc=OZD | 08/06/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 08/06/2022 | 10/10/2024 |  |  | 04/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221529 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a stroke approximately 5 minutes post impella implant. actions taken to resolve the reported observation are unknown. t he patient remained on impella support and was successfully weaned. | [5 empty] | The investigation into the stroke has been completed. the impella device was not returned for evaluation. the root cause of the stroke and its relationship to the impella were not determined since neither the product nor the data logs were returned. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17724 | 1220648-2024-20030 | 20417612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417612&pc=OZD | 08/19/2022 | 03/27/2025 |  | 10/10/2024 |  |  | 08/19/2022 | 10/10/2024 |  |  | 03/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214413 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complaint reported a pump stop issue with the impella device. additionally, the patient experienced bleeding and heparin administration was subsequently discontinued. the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation for the reported pump stop and access site bleeding issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm. visual inspection of the returned pump revealed a large impeller purge gap. additionally, a significant mass of biomaterial was found wrapped around the impeller, outflow cage and motor housing. no other damage or abnormalities were found. in conclusion, it was determined that the cause of the pump stop was bearing wear after 12 days of use, which is beyond 8-day indication per design trace matrix 0046-9910. this report is being filed as part of a retrospective review of historical records. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17725 | 1220648-2024-20031 | 20417623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417623&pc=OZD | 01/03/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 01/05/2022 | 10/10/2024 |  |  | 08/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127552 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis and purge pressure was high. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation into the hemolysis and high purge pressure has been completed. the impella device was returned for evaluation. the root cause of the high purge pressure was ingested biomaterial occluding the purge gap. the root cause of the hemolysis was most likely ingested biomaterial. the investigation for the reported hemolysis and high purge pressure has been completed. the data logs were returned and confirm the high purge pressure. suction alarms and irregular motor current pulstility can be seen occurring before a sudden increase in purge pressure. the product was returned and the high purge pressure was reproduced as the pump was priming. when the pump was running the purge pressure dropped to within a normal range and when the pump was stopped the purge pressure would begin to rise. this suggested that biomaterial was present in the purge gap. the pump was soaked in tergazyme to remove all the biomaterial and then the purge pressure was able to maintained within normal ranges even when the pump was not on. the root cause of the high purge pressure was determined to be ingested biomaterial. the pump was returned with a large thrombus occluding the inflow cage. because the hemolysis did not occur until the 4th day of support it is unlikely a result of a product malfunction. the root cause of the hemolysis was most likely ingested biomaterial as the thrombus was confirmed occluding the inflow cage, see es2020-079 for more information. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17726 | 1220648-2024-20033 | 20417662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417662&pc=OZD | 08/24/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 08/26/2022 | 10/10/2024 |  |  | 04/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216153 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed a hematoma treated by pressure and blood products. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17727 | 1220648-2024-20032 | 20417673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417673&pc=OZD | 07/20/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 01/12/2022 | 10/10/2024 |  |  | 05/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083560 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis and limb ischemia. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation into the hemolysis and limb ischemia has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause for the reported hemolysis and limb ischemia could not be determined; however, the lack of provided event information ultimately indicates an unknown relation to impella. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17728 | 1220648-2024-20035 | 20417713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417713&pc=OZD | 01/17/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 01/17/2022 | 10/10/2024 |  |  | 07/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110462 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the access site resulting from peeling away the 14fr introducer prior to fully removing from the site. during this event, the pump fell out of the ventricle. the pump was ultimately started; however, flows were lower than expected. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation into the low flow, positioning issue, and access site bleeding has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the low flow was unable to be determined as no product, logs, or imaging was returned for this investigation. the root cause of the positioning issues was determined to be use related as the pump came out of place while the introducer was being removed. the root cause of the access site bleeding was determined to be use related as the peel away introducer was removed while still in the body. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17729 | 1220648-2024-20034 | 20417731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417731&pc=OZD | 11/23/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 12/09/2022 | 10/10/2024 |  |  | 08/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267636 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, flows were dampened and could not ramp up to 3.5l/min. flows were stuck at 1.8l/min. consequently, the decision was made to explant the device and replace it with a new impella cp device. no patient harm was reported. | [5 empty] | The pump was returned and a kink was found in the cannula. no other abnormalities were found. the data logs were reviewed and show low flow at p-5 and p-8. suction alarms are also triggered. the root cause of the saturated flow was determined to be a cannula kink most likely due to the pump being positioned near the septal wall. this report is being filed as part of a retrospective review of historical records.
17730 | 1220648-2024-20036 | 20417753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417753&pc=OZD | 01/20/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 01/20/2022 | 10/10/2024 |  |  | 10/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156209 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, intermittent placement signal and low pump occurred. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation into the optical signal issue and low pump flow has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the optical signal issue and low pump flow was not determined. this report is being filed as part of a retrospective review of historical records.
17731 | 1220648-2024-20037 | 20417786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417786&pc=OZD | 02/17/2022 | 10/09/2025 |  | 10/10/2024 |  |  | 02/18/2022 | 10/10/2024 |  |  | 12/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162456 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Foreign | User facility | Company representation |  |  | 69 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis, the patient was also seen to be dehydrated which prompted the team to infuse volume and 4 units of packed red blood cells. the patient continued on support until the device was successfully weaned. | The complainant reported that the patient was being placed on impella cp for hemodynamic support. it was reported that the complainant in (b)(6) had an impella cp supporting a patient along with ecmo. the patient had signs of hemolysis. the patient was also seen to be dehydrated which prompted the team to infuse volume and 4 units of packed red blood cells. the pump remains on for support despite this due to the patient's hemodynamic needs. |  |  |  |  | The investigation for the reported hemolysis issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were sufficient to determine a root cause. the cause of the hemolysis was most likely patient condition since the hemolysis was resolved with volume. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the cause of the hemolysis was most likely patient condition since the hemolysis was resolved with volume. biomaterial was reported to be observed in the medikit sheath at the time of pump removal. impella is not indicated for use with the 16 fr medikit sheath.
17732 | 1220648-2024-20038 | 20417801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417801&pc=OZD | 02/08/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/08/2023 | 10/10/2024 |  |  | 11/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024292485 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Material Fragmentation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during placement that the physician chose to use a 0.018-inch guide wire in the easy guide lumen, but it slid inside the pigtail and the wire broke through the curved part of the pigtail. the decision was made to remove the impella and replace it with a new impella. there was no patient harm reported. | [5 empty] | The investigation into the product damage has been completed. the impella pump was returned for investigation. upon visual inspection of the pump, a hole was observed on the curved part of the pigtail. the root cause of the product damage was a hole in the pigtail resulting from improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
17733 | 1220648-2024-20039 | 20417917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417917&pc=OZD | 02/08/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/08/2023 | 10/10/2024 |  |  | 09/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281507 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. it was reported the patient¿s activated clotting time was greater than five hundred. the impella position reportedly looked good, but due to the patient¿s very deep tissue track profuse bleeding started. reportedly hemostasis was achieved after the physician sutured multiple times using different angles and pressure was held. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the access site bleeding has been completed. the investigation determined through review of the clinical details that the cause of the bleeding was the high patient activated clotting time values which were reported to be greater than five hundred. the root cause conclusion was the patients activated clotting time values. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17734 | 1220648-2024-20040 | 20417966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417966&pc=OZD | 03/04/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 03/04/2022 | 10/10/2024 |  |  | 10/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158349 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the impella come out of position and into the aorta. repositioning was unsuccessful and on day 7 of support the pump alarmed for a pump stop. the impella cp device was explanted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop and positioning issue. the device was not received for evaluation. the root cause of the pump stop, and positioning issue was unable to be determined since product was not returned. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17735 | 1220648-2024-20041 | 20417973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417973&pc=OZD | 02/08/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/08/2023 | 10/10/2024 |  |  | 10/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024288183 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. it was reported the bleeding was unable to be controlled with manual pressure and the physician elected to remove the impella and leave the 12 french sheath in place. the procedure was completed with no further complications. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the access site bleeding has been completed. the investigation determined through review of the clinical details the cause of the access site bleeding issue was use related since no imaging was done prior to pump insertion. the cause of the access site bleeding issue was use related due to improper technique. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17736 | 1220648-2024-20043 | 20417984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417984&pc=OZD | 01/04/2022 | 05/22/2025 |  | 10/10/2024 |  |  | 03/24/2022 | 10/10/2024 |  |  | 08/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022122184 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis, thrombocytopenia and renal failure during the support. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The investigation for the reported hemolysis, renal failure, and thrombocytopenia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the hemolysis, renal failure, and thrombocytopenia and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17737 | 1220648-2024-20042 | 20417998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20417998&pc=OZD | 02/09/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/09/2023 | 10/10/2024 |  |  | 10/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283838 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. it was reported due to the large amount of peripheral arterial disease (pad) throughout the iliac into the aorta at the bifurcation, the impella was unable to be advanced past the bifurcation. reportedly both the short and long impella sheaths were utilized to try and advance the pump, but after several attempts, the physicians determined that they needed to abort the case due to the anatomy of the patient. | [5 empty] | The investigation into the delivery issue has been completed. no product was returned. per the clinical investigation, due to the large amount of peripheral arterial disease throughout the iliac into the aorta at the bifurcation, the impella was unable to be advanced past the bifurcation. the physicians determined that they needed to abort the case due to the anatomy of the patient. the root cause of the delivery issue was patient condition related. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17738 | 1220648-2024-20044 | 20418064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418064&pc=OZD | 03/30/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 03/30/2022 | 10/10/2024 |  |  | 10/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151371 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis and bleeding at the access site. the device was repositioned, and this resolved the hemolysis. hemostasis was achieved with by using a coil to resolve the bleeding, and angiogram was performed. afterwards, the impella was removed as well, because of suspicion that it was related to the impella. no further information was provided. | [5 empty] | The impella device was not received from the customer. data logs show many impella position unknown alarms 24 hours into support. following a position adjustment, no further alarms are seen and the pump functions without issue. based on the logs and clinical account, the root cause of hemolysis was most likely due to positioning. the cause of the issue was patient condition related. as noted in the impella ifu: impella cp with smart assist system . impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17739 | 1220648-2024-20046 | 20418233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418233&pc=OZD | 07/10/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 07/10/2022 | 10/10/2024 |  |  | 01/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190978 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 29 YR |  | Female |  |  | Asian | Injury | Yes | No | Device Sensing Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. after removing the impella cp, the guidewire did not pass through the sheath. when an angiography was performed from the opposite side, there was no blood flow in the external iliac artery of the right foot, and a thrombus was suspected. a procedure was performed to remove the thrombus. blood flow improved and lower limb blood flow was no longer a problem. additionally, the nurse reported placement signal issues, the resolution for this issue is unknown. no further information provided. | [5 empty] | The investigation for the reported thrombus, optical signal and limb ischemia issues has been completed. the product was returned. per the results of the clinical review and investigation: pump and data was returned and inspected in the bhl. the pump passed the light test. when plugged into the aic, there was a placement signal. the red handle was cut open and the motor cables looked twisted. there was no laser leak in the red handle. the device logs show an unusual placement signal that jumped around during the initial few minutes of the case. the placement signal was then accurate for the rest of the case. the logs display a "placement signal low" alarm which occurs when aortic diastolic signal is low and indicates to check pump positioning and assess cardiac function. the logs show low pulsatility indicating poor patient condition. the pump possibly moved when it was turned on and slightly shifted its position. the patient's condition and the positioning of the pump most likely contributed to the placement signal issue. the limb ischemia most likely occurred due to the reported blood clot and low act values in the patient. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17740 | 1220648-2024-20045 | 20418246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418246&pc=OZD | 05/10/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 05/10/2023 | 10/10/2024 |  |  | 11/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309955 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported during placement of the impella delivery issues were encountered and multiple guide wire and catheter attempts were made to cross into the left ventricle. reportedly the guide wire was exchanged for 0.018 and the impella backloaded onto the wire, however, the physician was unable to advance the impella into the left ventricle due to the patient¿s anatomy and short aortic root. there was no patient harm reported, and this device was removed, and an intra-aortic balloon pump was placed. there was no patient harm reported, and this device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed and per clinical information provided, the physician felt resistance at the aorta and the pump insertion was aborted due to the patient having a short aortic root. the cause of the delivery issue was patient condition related due to patient anatomy and short root per clinical review. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
17741 | 1220648-2024-20047 | 20418294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418294&pc=OZD | 05/01/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 05/12/2023 | 10/10/2024 |  |  | 12/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163253 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, race demographics, unknown, was implanted with an impella cp for left ventricle unloading support. it was reported that the physician called the clinician and reported progressive suction due to misplacement of the impella. reportedly the impella was in the aorta, the physician tried to reposition the device in the catheterization laboratory, however attempts were unsuccessful. the physician reportedly decided to remove the impella, due to multiple repositioning issues and use only extra-corporeal membrane oxygenation. there was no patient harm reported, and this device was removed. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the investigation into the positioning issue has been completed. the root cause of the issue was improper positioning of the device resulting from improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
17742 | 1220648-2024-20048 | 20418336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418336&pc=OZD | 05/30/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 05/30/2023 | 10/10/2024 |  |  | 09/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274954 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 95 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported the physician attempted to place the impella in the right femoral artery, however, due to calcification and tortuosity of the iliac the pump kinked and was not able to pass through the patient¿s vasculature. the physician prepped the left femoral artery and was able to successfully place a new impella. there was no patient harm reported, and this device was replaced. | [5 empty] | The investigation into the delivery issue issues has been completed. the impella pump was returned for investigation. the returned pump was visually inspected, no abnormalities were found. per clinical review of the clinical details the root cause of the delivery issue was due to patient condition. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
17743 | 1220648-2024-20049 | 20418433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418433&pc=OZD | 05/31/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 05/30/2023 | 10/10/2024 |  |  | 10/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024284961 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported when the pump was started there was noted significantly lower than expected flows in auto mode at pressure p-9. reportedly the position was confirmed, and a clot was suspected. the clinician recommend replacing the pump due to a possible clot entrapment, upon removal of the pump a thrombus was observed within the outflow and on the pigtail. the thrombus was reportedly not observed during an echocardiogram the day before removal. this device was removed and replaced with another impella. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the investigation into the low pump flow with thrombus has been completed. the data logs were returned and reviewed the logs did not show any indications of low pump flow. the cause of the low pump flow and related thrombus could not be determined as the product was not returned and clinical details provided were insufficient to determine a root cause. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smart assist system: section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: as noted in the impella ifu ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ as noted in the impella ifu "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17744 | 1220648-2024-20050 | 20418524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418524&pc=OZD | 06/18/2023 | 10/08/2025 |  | 10/10/2024 |  |  | 06/18/2023 | 10/10/2024 |  |  | 02/28/2025 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324102 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, flows were low. iv heparin was administered and was found to be infiltrated, so clot was suspected. the impella cp device was explanted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was not received for evaluation. the root cause of the low pump flow was most likely biomaterial ingestion with clot observed during removal due to improper anticoagulation and motor current spike, suction with low flow per clinical and log review. this report is being filed as part of a retrospective review of historical records.
17745 | 1220648-2024-20051 | 20418781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418781&pc=OZD | 09/20/2024 | 11/20/2024 |  | 10/10/2024 |  |  | 09/20/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510676 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and the staff was having difficulty torquing and positioning. the device was removed and there was damage to the driveline noted. the decision was made to use a different 5.5 pump for support. there was no reported harm to the patient. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
17746 | 1220648-2024-20052 | 20418798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418798&pc=OZD | 09/16/2024 | 11/14/2024 |  | 10/10/2024 |  |  | 09/16/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025507267 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp pump to support a patient admitted in with an st elevated myocardial infarction. the pump's access site was bleeding and was forming a hematoma. the team treated by manual pressure, pausing heparin, and medically managing. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. d.4 serial number was revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-20052. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17747 | 1220648-2024-20053 | 20418805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418805&pc=OZD | 06/21/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 06/21/2023 | 10/10/2024 |  |  | 02/28/2025 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321876 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Inaccurate Information | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was discovered that the serial number of the pump implanted was different between the number registered with the aic (automated impella controller) and the number on the product main label attached to the package box. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation into the reported packaging issues has been completed. the impella device was returned for evaluation. the root cause of the packaging issue was an incorrect label with a s/n mismatch displayed on the aic (automated impella controller) compared to the printed pump label on the red pump plug and distributor box label. this was due to a programming error with pump position swapped in magazine or barcode label scanner swap before initial programming. this report is being filed as part of a retrospective review of historical records.
17748 | 1220648-2024-20055 | 20418923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418923&pc=OZD | 08/10/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 08/10/2021 | 10/10/2024 |  |  | 05/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022091686 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the cannula kinked, and a second pump was used. there was no harm to the patient. no additional information provided. | [5 empty] | The investigation for the reported product damage (cannula kink) issues has been completed. the impella device has been returned for evaluation. the cannula was confirmed to be kinked, however, the cause of the cannula kink was not determined. this report is being filed as part of a retrospective review of historical records.
17749 | 1220648-2024-20054 | 20418929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418929&pc=OZD | 09/21/2024 | 11/21/2024 |  | 10/10/2024 |  |  | 09/21/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525250 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during implant procedure for 68-year-old male patient, the impella cp would not advance past the sheath or past the aortic arch. when the impella was taken out and inspected, there was an object lodged in the outlet. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was determined to be use issue because they were ale to insert the second impella successfully without any issues. corrections to the initial mfr report: h5 changed to yes. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17750 | 1220648-2024-20056 | 20418936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418936&pc=OZD | 09/16/2024 | 11/21/2024 |  | 10/10/2024 |  |  | 09/16/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025525356 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the impella stopped and was unable to restart. the automated impella controller (aic) was switched but the pump did not restart. the patient¿s outcome is unknown. | [5 empty] | The investigation for the pump stop has been completed. the product was not returned. log review showed alarm 119 at the time of the pump stop, indicating an electrical issue. there was also a small motor current increase to ~400 ma at the time of the alarm. the root cause of the pump stop was most likely an electrical open. it is unknown how the electrical open would have occurred since the product was not returned and insufficient clinical detail was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17751 | 1220648-2024-20057 | 20418974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20418974&pc=OZD | 10/01/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 10/01/2021 | 10/10/2024 |  |  | 03/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022066521 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. impella was noted to have migrated into the left ventricle (lv) after the peel away sheath was removed, and the repositioning sheath was placed. while eliminating the slack and pulling the outlet back, the medical professional inadvertently pulled the impella device out of the aorta. impella was unable to be advanced across av and cannula became bent during the attempt which required removal of device. patient survived the procedure. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the positioning issue was misuse. the impella is not indicated for wireless insertion or re-access. this report is being filed as part of a retrospective review of historical records.
17752 | 1220648-2024-20058 | 20419037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20419037&pc=OZD | 10/03/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 10/03/2021 | 10/10/2024 |  |  | 05/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092328 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, pump had high purge pressure and the pump stopped. the pump was able to be restarted but, eventually the team made choice to replace the multiple pc, automated impella controller (aic), and the pump itself. support proceeded without harm or injury from the delay in support initiation. | [5 empty] | The investigation into the low purge pressure or purge system leak and pump stop has been completed. the impella pump was returned for investigation. the provided data logs showed that case start's purge cassette priming step was not completed on the original console, despite two different purge cassettes being used. purge pressure rose to above 1000 mmhg, but purge flow remained at 0 ml/hr. the data logs showed that the second purge cassette was transferred to a second console by inserting the cassette first (at 18:54:10), followed by the purge pressure disk four seconds later (at 18:54:14). the pump was connected at 18:54:45. a 'purge pressure low' alarm was triggered on (b)(6) 2021 at 18:56:25. this alarm cleared at 18:57:05. at 19:28:28, an 'impella stopped - motor current high' alarm was triggered and the pump stopped. several attempts to restart the pump were unsuccessful. data logs for the second pump were not provided. the first purge cassette was returned and purged without issue. the second purge cassette used in the case was able to purge, but leaking was observed from the lower part of the purge disk. a small hole was noted in the membrane of the purge disk. evaluation of the returned pump found no biomaterial or obstructions in the impeller or cannula. the impeller spun with resistance and blood and biomaterial were noted in the purge gap. blood was noted in the purge lumen. the pump was torn down at the housing and biomaterial was found in the bearing. the root cause of the pump stop was blood ingress. the root cause of the low purge pressure alarm was a leak in the purge pressure disk. this report is being filed as part of a retrospective review of historical records.
17753 | 1220648-2024-20059 | 20419077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20419077&pc=OZD | 10/12/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 10/12/2021 | 10/10/2024 |  |  | 07/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022113676 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported high purge pressure, access site bleeding, and atrial fibrillation (af) during support. access site bleeding was managed with manual pressure and single suture. impella was successfully weaned and explanted. patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. neither the product nor the data logs were returned. therefore, the cause of the high purge pressure, access site bleeding and atrial fibrillation was unable to be determined. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17754 | 1220648-2024-20060 | 20419134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20419134&pc=OZD | 09/22/2024 | 11/20/2024 |  | 10/10/2024 |  |  | 09/22/2024 | 10/10/2024 |  |  | 10/31/2025 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421333 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an impella cp-rp support ongoing for a patient admitted in cardiogenic shock. after three days the pump's access site was seen to have developed a hematoma. the bleed was treated by blood products and the pump remains on for support to date. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related; hemostasis was achieved with insertion angle re-addressed as per the clinical description provided. g1 reporting contact fax number missing from manufacturer device report. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17755 | 1220648-2024-20061 | 20419135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20419135&pc=OZD | 01/03/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 02/15/2022 | 10/10/2024 |  |  | 07/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105872 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed hemolysis. no further information was available regarding treatment for the hemolysis. the patient underwent a coronary artery bypass graft procedure, remained on impella support and was successfully weaned. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the clinical evaluation into the hemolysis has been completed. the cause of the hemolysis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. there is an unknown relationship between the impella and the hemolysis. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17756 | 1220648-2024-20062 | 20419153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20419153&pc=OZD | 02/09/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/09/2023 | 10/10/2024 |  |  | 11/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281878 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported during placement of the impella the physician lost the easy guide lumen and tried to thread through the impella manually but got stuck in the impella. the pump was exchanged by the staff due to the potential of damage done to the pump. a replacement pump was successfully used for the procedure. the staff reportedly discarded this pump. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. the root cause of the delivery issue was determined to be use related as the red lumen was lost by the physician before the guidewire could be backloaded. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17757 | 1220648-2024-20063 | 20419220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20419220&pc=OZD | 02/13/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/13/2023 | 10/10/2024 |  |  | 10/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290667 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the medical staff attempted to reposition it to the left ventricle. some improvement of the waveforms was noticed although the pigtail of the pump remained bent and the physician declined to reposition further. the patient survived the procedure. | [5 empty] | The investigation into the positioning issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the cause of the issue was use related. this report is being filed as part of a retrospective review of historical records.
17758 | 1220648-2024-20064 | 20419281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20419281&pc=OZD | 06/27/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 06/27/2023 | 10/10/2024 |  |  | 10/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1683879 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp stopped immediately after starting the pump in auto with ic 3034. two restart attempts were unsuccessful. controller failure alarm displayed. aic exchange was performed with successful pump start. there was no patient harm. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the pump stop issue was not determined since sufficient clinical information was not provided and pump ran as expected on second aic and during test. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17759 | 1220648-2024-20066 | 20419351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20419351&pc=OZD | 09/22/2024 | 11/17/2024 |  | 10/10/2024 |  |  | 09/22/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535463 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and while on support, no pulses were found in the patient's right foot. additionally, the patient had frequent premature ventricular contraction. echocardiography was ordered. staff monitored. furthermore, the patient had bleeding while on support. the location of the bleeding was unknown. heparin was in the purge of the cp. one unit of blood was given to the patient. the patient remains on impella support. | [5 empty] | The investigation limb ischemia, vf/vt/af/at, and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
17760 | 1220648-2024-20065 | 20419449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20419449&pc=OZD | 03/27/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 03/27/2023 | 10/10/2024 |  |  | 09/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259361 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient entered a period of ventricular fibrillation, which the physician attributed to the wire in the left ventricle. the issue was resolved through defibrillation, and the patient survived the procedure. | [5 empty] | The investigation for the reported ventricular fibrillation has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was related to patient condition and the guidewire used during the percutaneous intervention (pci) procedure. this report is being filed as part of a retrospective review of historical records.
17761 | 1220648-2024-20067 | 20419918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20419918&pc=OZD | 09/16/2024 | 10/30/2025 |  | 10/10/2024 |  |  | 09/16/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1504376 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Alarm Not Visible | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a console being used for training. the automated impella controller had no audible noise on alarm. there was no patient use and as such no harm or injury. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated imeplla controller (aic) - buzzer/ alarm malfunction has been completed. data analysis: console logs were reviewed but were not necessary for the complaint. device analysis: inspection of the internal circuitry of the console revealed the speaker cable to be unplugged. conclusion: the root cause for the inaudible alarms was an unplugged speaker cable. d2 common device name revised on manufacturer device report 1220648-2024-20067 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-20067 in accordance with updated procedures.
17762 | 1220648-2024-20068 | 20420409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20420409&pc=OZD | 09/22/2024 | 11/18/2024 |  | 10/10/2024 |  |  | 09/22/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533759 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on cp support had blood in the urine. the p-level will remain at p-9 as the patient needs the support. one unit of red blood cells was given. the patient survived the event. | [5 empty] | The investigation for the hemolysis has been completed. the pump was not returned for investigation. the pump was utilized from 09/22 at 00:25 to 09/27 at 11:36. the patient experienced hemolysis within less than 24 hours of starting the impella. data logs show a suction alarm and several placement signal low alarms during the hemolysis event, with a trending placement signal and a recovering trend in pump flow without any observed motor current spikes. the 5s optical signal parameters were within specifications. according to the clinical information, cvp was low, and albumin was administered along with one unit of blood, which aided in the recovery from hemolysis. additionally, the iabp was discontinued during the hemolysis event. the cause of the hemolysis issue was most likely patient condition related, based on data log review and given clinical details. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
17763 | 1220648-2024-20073 | 20420924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20420924&pc=OZD | 07/11/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 07/11/2023 | 10/10/2024 |  |  | 03/31/2025 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024327513 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump advanced across the valve and support was initiated. the impella stopped with a motor current high alarm. the impella cp was explanted and replaced to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump did not start. the device was received for evaluation. the root cause of the pump did not start issue was not determined since pump stop could not be reproduced per field investigation and logs indicating high motor current pump stop per log review. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17764 | 1220648-2024-20074 | 20420936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20420936&pc=OZD | 07/25/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 07/25/2022 | 10/10/2024 |  |  | 04/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214055 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the device would not prime. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation into the priming issues has been completed. the impella device was returned for evaluation. the root cause of the priming issues was use related. this report is being filed as part of a retrospective review of historical records.
17765 | 1220648-2024-20084 | 20420984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20420984&pc=OZD | 05/27/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 04/05/2022 | 10/10/2024 |  |  | 01/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022038044 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after support, it was noted that the patient had cerebrovascular accident (cva/stroke) post explant. no further information provided. | [5 empty] | The investigation for the reported stroke/cva issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17766 | 1220648-2024-20076 | 20421049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20421049&pc=OZD | 09/01/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 09/01/2022 | 10/10/2024 |  |  | 10/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152397 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a priming failure during setup of the device. a new impella was used successfully. there was no report of patient harm. | [5 empty] | The investigation has been completed for the reported issue of priming issue. the device was not received for evaluation. the root cause of the priming issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
17767 | 1220648-2024-20085 | 20421111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20421111&pc=OZD | 06/11/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 09/10/2021 | 10/10/2024 |  |  | 01/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022043031 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis. attempts to reposition the device were made; however, the pump was pulled out of the ventricle and could not be re-crossed. the impella was explanted and replaced with a new impella cp to address the issue. the patient survived the procedure. | [5 empty] | The investigation into the hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the hemolysis was determined to be unrelated as no evidence was provided to suggest a causal relationship between impella and the hemolysis and the physician stated that the hemolysis was unrelated to the device. the root cause of the positioning issues was determined to be a use issue as the pump was pulled out of the ventricle by mistake. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17768 | 1220648-2024-20072 | 20421137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20421137&pc=OZD | 07/03/2023 | 11/05/2025 |  | 10/10/2024 |  |  | 07/03/2023 | 10/10/2024 |  |  | 02/28/2025 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341074 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic |  | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding and systemic heparin was discontinued. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17769 | 1220648-2024-20075 | 20421211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20421211&pc=OZD | 02/26/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/26/2023 | 10/10/2024 |  |  | 09/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283179 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient exhibited signs of ischemia. the reduced flow was observed on imaging but, no intervention was done. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device has not been returned for evaluation. however, clinical details provided were sufficient to determine the root cause of the limb ischemia to be patient condition related to the small/ diseased vessels. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
17770 | 1220648-2024-20090 | 20421328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20421328&pc=OZD | 09/28/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 09/28/2021 | 10/10/2024 |  |  | 07/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022112557 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump folded itself in the ventricle and stopped. as a result, the decision was made to explant the device and replace it with a new impella device. the patient survived the procedure. | [5 empty] | The investigation into the positioning issues and pump stop has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause for the positioning issues and pump stop could not be determined as no product was provided. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17771 | 1220648-2024-20082 | 20421374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20421374&pc=OZD | 09/13/2024 | 11/22/2024 |  | 10/10/2024 |  |  | 09/13/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025484629 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was removed due to high purge pressure system blocked alarm. the patient eventually expired after care was withdrawn. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The investigation for the high purge pressure has been completed. there were no data logs returned. the returned product had moldy biomaterial all in the purge gap, and high purge pressure reproduced. the pump was soaked for 10 minutes in warm water but it still remained. the catheter was cut and fluid was expelled from the catheter purge lumen. based on the device evaluation, the root cause of the high purge pressure is most likely due to biomaterial. corrections to the initial mfr report: b5-the initial report stated that "use of the device cannot conclusively be associated with the outcome." it should have stated "it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome." | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
17772 | 1220648-2024-20093 | 20421465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20421465&pc=OZD | 12/03/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 12/03/2022 | 10/10/2024 |  |  | 02/29/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193094 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was stated that the clear luer locking device on the peel away introducer separated from the introducer. this resulted in the locking component remaining on the peel-away hub when the peel-away sheath was removed. the component was subsequently broken off/removed with hemostats. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation into the introducer damage - mechanical has been completed. the impella device was not returned for evaluation. the cause of the introducer issue, specifically the clear luer lock detaching from the dilator and staying attached to the peel-away sheath, was not determined since product was not returned and there is insufficient clinical evidence. this report is being filed as part of a retrospective review of historical records.
17773 | 1220648-2024-20087 | 20421493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20421493&pc=OZD | 04/08/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 04/08/2022 | 10/10/2024 |  |  | 11/30/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157957 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, ischemic symptoms in the right lower limb were noted, and an antegrade puncture was performed. the impella cp was removed and escalated to impella 5.0 for long-term support and bridge to surgery. no further information provided. | [5 empty] | The investigation for the reported limb ischemia issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17774 | 1220648-2024-20092 | 20421497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20421497&pc=OZD | 03/02/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 03/02/2023 | 10/10/2024 |  |  | 08/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270802 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the impella was placed and upon start up there was an immediate loss of placement signal, with multiple suction alarms noted. the automated impella controller was replaced with no resolution. the physician determined that the watchman device that was implanted prior to the impella had embolized in the patient¿s aortic arch. the decision was made to remove the impella and retrieve the watchman, multiple snares and coda balloon were utilized to retrieve the watchman device through the sheath. additionally, it was reported there were low pump flow issues. there was no patient harm reported, due to the optical sensor damage the impella was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the optical signal and low pump flow issues have been completed. the most likely cause of the optical signal issue was an air gap due to improper technique by the end user. the cause of the low pump flow was not determined because no product was returned, and clinical details provided were insufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
17775 | 1220648-2024-20071 | 20421514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20421514&pc=OZD | 03/04/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 03/07/2022 | 10/10/2024 |  |  | 11/30/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  | 2022157954 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. the physician confirmed that the slow bleeding was from the impella access site and that the cause was a slight incision in the skin at the access site before inserting the impella. additionally, the physician expressed the bleeding wasn't a result of the impella. the physician made the decision to remove the impella since the patient was hemodynamically stable, and due to a 16 french sheath being used the bleeding could be stopped with a single perclose suture during the removal of the impella. this device was removed. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding issue was most likely use related as the cause of bleeding was a slight incision in the skin at the access site before inserting the impella. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17776 | 1220648-2024-20083 | 20421549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20421549&pc=OZD | 04/21/2024 | 03/21/2025 |  | 10/10/2024 |  |  | 06/17/2024 | 10/10/2024 |  |  | 06/30/2025 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365228 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was an issue when attempting to change the purge cassette for the impella cp. the y connection was compromised and there was an emergent situation so a new impella was opened and used. it was reported that there was a lack of knowledge from the nursing staff regarding the purge cassette set up. no patient harm was reported. | [5 empty] | The investigation for the priming issue has been completed. the pump was not returned for investigation. as per the clinical information provided, the nurse did not connect the purge tubing correctly which resulted in having to use a new pump. the root cause of the priming issue was most likely use related as the nurse improperly connected the purge tubing. e.1 added address line 2 as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-20083. revised us city as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-20083. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-20083 was submitted in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-20083 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
17777 | 1220648-2024-20088 | 20421655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20421655&pc=OZD | 09/12/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 04/06/2022 | 10/10/2024 |  |  | 06/30/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100683 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding, hematoma, and infection at the impella access site. the patient received blood products to address the issue. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of infection/sepsis and access site bleeding - major the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the access site bleeding was most likely a lack of vessel recoil onto the sheath. the root cause of the insertion infection was most likely patient condition related since the pump was sterilized. as noted in the impella ifu: impella cp with smart assist system impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17778 | 1220648-2024-20077 | 20422148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20422148&pc=OZD | 09/16/2024 | 11/19/2024 |  | 10/10/2024 |  |  | 09/16/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025543734 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support. while on support, the patient¿s left ventricle (lv) was perforated, and thought to be caused by 0.018 wire. perforation was repaired. additionally, the surgeon was able to advance the pump and wire across the valve. the wire was pulled and support initiated on auto only reaching flows of 1.0 liters per minute with active suction alarm. volume was given despite left ventricle end-diastolic pressure of 12. a kink was found in the impella cannula. the decision was made to replace the impella with a new one. the patient survived the event. | [5 empty] | The investigation has been completed. the guidewire was not returned for review. based on the clinical details, the root cause of the ventricle perforation is most likely due to guidewire puncture from improper technique. the failure mode product damage (cannula kink) was investigated as low pump flow as low flow was reported in the case and a kink in the cannula was visualized on echocardigram. the returned product showed a kink in the cannula. the root cause of the low pump flow was due to a kink in the cannula that most likely occurred during the attempts to reposition. h.6 revised medical device problem code based off investigtion result since initial manufacturer device report number 1220648-2024-20077 was submitted. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
17779 | 1220648-2024-20070 | 20422319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20422319&pc=OZD | 05/06/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 05/06/2023 | 10/10/2024 |  |  | 06/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237214 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the physician was unable to pass the pump into the left ventricle as the patient had difficult vasculature. pump insertion was aborted, and vascular surgery was consulted for further advice. there was no reported injury to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17780 | 1220648-2024-20094 | 20422496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20422496&pc=OZD | 05/24/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 05/24/2022 | 10/10/2024 |  |  | 03/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-EU |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2023192357 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, race and ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia and hemolysis. it was reported the patient was in very critical condition and the impella pressure level was 9 with flows of 3.8 liters per minute with some suction alarms. the limb ischemia was reported to be worsening towards the foot and the thigh was swollen. the patient¿s urine showed signs of hemolysis due to red color. the clinician recommended escalation as soon as possible to the heart center as the echocardiogram revealed the patient¿s ejection fraction was 5%. the clinician also recommended forward perfusion on the limb from contra lateral. the physician considered the forward perfusion but had not yet performed as the physician was attempting to reach the heart center for escalation. further guidance from the clinician resulted in reducing and eventually turning off vasopressin, resulting in significantly better limb perfusion, as well as better amplitudes on the waveforms. the patient was air lifted for escalation to the heart facility and reported to be more stable with the addition of extra-corporeal membrane oxygenation. reportedly on 5/25/2022 the patient was stable, and echocardiogram showed the impella in good position and the impella running at pressure level 4 at 2.6 liters per minute. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation into the hemolysis and limb ischemia issues have been completed. the impella pump was returned for investigation. per clinical review the cause of the hemolysis issue was most likely improper positioning of the device due to improper technique by the end user. the clinical review also led to the cause of limb ischemia issue being patient condition related, most likely due to small or diseased vessels. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: limb ischemia impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ hemolysis impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ purge leak-pump impella cp with smart assist for use during cardiogenic shock and high risk cpi section: cleaning as noted in the impella ifu ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
17781 | 1220648-2024-20100 | 20422674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20422674&pc=OZD | 09/01/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 09/01/2022 | 10/10/2024 |  |  | 03/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023197777 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. it was reported a failure to prime the device. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of priming issue. the device was returned for evaluation. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
17782 | 1220648-2024-20086 | 20422695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20422695&pc=OZD | 09/12/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 04/05/2022 | 10/10/2024 |  |  | 04/30/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022081097 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis and renal failure while on support. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation into the hemolysis and renal failure has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the most likely cause of the hemolysis and of the renal/kidney failure and their relation to the impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17783 | 1220648-2024-20089 | 20422714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20422714&pc=OZD | 09/18/2021 | 05/22/2025 |  | 10/10/2024 |  |  | 12/14/2021 | 10/10/2024 |  |  | 06/30/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100639 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis and renal failure, any medical intervention is unknown. during support, the patient also developed thrombocytopenia, with no interventions documented. the patient continued on support until the device was successfully weaned. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | No product was available for evaluation and no data logs were provided for analysis. the root cause of the hemolysis and the root cause of the renal failure, as well as their relationship to the impella device, were determined to be unknown. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17784 | 1220648-2024-20091 | 20422800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20422800&pc=OZD | 09/18/2021 | 05/23/2025 |  | 10/10/2024 |  |  | 10/05/2021 | 10/10/2024 |  |  | 05/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083544 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis and renal failure while on support. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | Suction is not an indication of hemolysis. the cause of the suction was most likely improper positioning since the position adjustment resolved the issue. the renal failure relation to impella was determined to be unknown since the hemolysis was reported to be unrelated to the renal failure. | The investigation for the reported hemolysis and renal failure has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the hemolysis and renal failure and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17785 | 1220648-2024-20069 | 20422935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20422935&pc=OZD | 09/16/2024 | 11/07/2024 |  | 10/10/2024 |  |  | 09/16/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528966 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that the impella 5.5 was unable to be inserted due to patient anatomy and small vessel size. procedure was aborted. no patient harm has been reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, the impella couldn't pass due to patient anatomy, vessel size too small as per the clinical description provided. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17786 | 1220648-2024-20078 | 20422983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20422983&pc=OZD | 07/15/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 07/15/2022 | 10/10/2024 |  |  | 03/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213210 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | Decrease in Pressure | Pumping Stopped |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. on day four of impella cp support it was reported the device exhibited purge system open, the staff noticed that the yellow luer was broken off, two days later the pump stopped. additionally, there was an air in the system issue, the staff replaced the cassette, however this did not resolve the issue. the physician decided to explant the device and replace with another impella cp. no report of patient harm or injury. | [5 empty] | The investigation for the reported issues has been completed. the impella device has not been returned for evaluation. the cause of the purge leak, pump stop, and purge cassette failure was not determined since product was not returned. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
17787 | 1220648-2024-20080 | 20423016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20423016&pc=OZD | 03/24/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 03/24/2022 | 10/10/2024 |  |  | 03/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022072507 |  | 005064 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.0 device for mechanical circulatory support. while on support, the impella failed to prime. there was no harm or injury as the team replaced all product and proceeded with the 2nd pump. it was also noted that the pump box had previously been opened and may have had missing/damaged ancillary equipment. no lasting harm or injury from the delay. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
17788 | 1220648-2024-20079 | 20423394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20423394&pc=OZD | 03/08/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 03/08/2022 | 10/10/2024 |  |  | 11/30/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165728 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia, the right foot was noted to be cold. the introducer was not peeled away, pulses diminished, and only the doppler was palpated. it was also reported low pump flows were observed due to positional issues, the pump was repositioned, however, the flows were still limited. the physician made the decision to remove the impella | [5 empty] | The investigation for the reported limb ischemia and low pump flow issues has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a large clot in the cannula was found, indicating the root cause of the low pump flow was biomaterial. the limb ischemia is most likely due to the physician leaving the sheath in. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17789 | 1220648-2024-20081 | 20423397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20423397&pc=PYX | 09/19/2024 | 11/15/2024 |  | 10/10/2024 |  |  | 09/19/2024 | 10/10/2024 |  |  | 07/31/2025 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024374165 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Material Twisted/Bent | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Great Vessel Perforation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp support. when inserting rp peel away sheath, physician stated that the 23 fr sheath was kinked, and that the dilator would not go over the wire. physician perforated the inferior vena cava during sheath placement. patient continued to decompensate from perforation requiring four units of packed red blood cells. balloon tamponade was performed and bleed slowed down significantly. patient continued to decompensate and expired in unit upon arrival from cardiac catheterization lab. | [5 empty] | The investigation for the introducer damage and the vascular damage has been completed. the pump and introducer were not returned for investigation. a data log review was not required for this case. according to the clinical notes, the introducer sheath was kinked, and the dilator did not pass over the wire while inserted through the left femoral vein. the doctor later realized that the right groin access approach was more suitable. the pump was successfully placed through the kinked sheath. additionally, damage was found on the ivc during sheath and dilator placement. the vessel was repaired using a balloon. the impella rpsa ifu states caution regarding the risk of excessive manipulation and pump damage from repeated positioning attempts, indicating that left femoral vein insertion should be avoided whenever possible. the cause of the introducer damage issue was most likely use related. the cause of the vascular damage issue was not determined since the product was not returned and sufficient clinical information, such as patient anatomy and the excessive force applied during use, was not provided. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿
17790 | 1220648-2024-20096 | 20423414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20423414&pc=OZD | 09/13/2024 | 11/11/2024 |  | 10/10/2024 |  |  | 09/13/2024 | 10/10/2024 |  |  |  |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551331 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 45-year-old male patient was on impella cp support due to st elevated myocardial infarction (stemi). while on support, diagnostic imaging report read that tissue abnormalities suggestive of thromboembolic activity were found - ischemic stroke noted. patient expired on support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. added h6 impact code 4648. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17791 | 1220648-2024-20095 | 20423418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20423418&pc=OZD | 12/07/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 12/07/2022 | 10/10/2024 |  |  | 06/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235794 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while prepping and inserting the device pump got kinked. under fluoroscopy a kink was noticed on the catheter. md didn't feel comfortable/safe putting in the pump and called for another catheter. no harm done to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. after reviewing the clinical information, it was determined that the cause of the delivery issues was most likely a catheter kink due to the patient's calcified vessels. this report is being filed as part of a retrospective review of historical records.
17792 | 1220648-2024-20098 | 20423512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20423512&pc=OZD | 04/27/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 04/27/2022 | 10/10/2024 |  |  | 12/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180988 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, flows were low and suction alarms occurred. the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was not received for evaluation. the root cause of the low pump flow was unable to be determined as no product or logs were returned for analysis. this report is being filed as part of a retrospective review of historical records.
17793 | 1220648-2024-20097 | 20423573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20423573&pc=OZD | 08/23/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 08/23/2022 | 10/10/2024 |  |  | 01/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181638 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  | Black Or African American | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, high motor current alarm appeared, followed by pump stop. pump was unable to be restarted. the catheter was removed, and clots were found around the cannula and pigtail. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation for the reported thrombus and pump stop issue has been completed. the product was returned and the pump stop was confirmed. a visual inspection the pump revealed a clot on the impeller, the outflow cage, and the cannula. bearing wear was also observed. the lt log shows an abrupt rise in motor current on 8/18. this was likely when the clot was ingested into the pump. following this event, additional spikes in motor current are seen until the pump stopped on 8/22. the cause of the pump stop was bearing wear. the pump stop occurred beyond the 8 day impella cp design life. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system (thrombus): section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
17794 | 1220648-2024-20104 | 20423588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20423588&pc=OZD | 11/22/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 11/23/2021 | 10/10/2024 |  |  | 07/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108933 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Restricted Flow rate | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, low flows and purge leakage occurred. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation into the reported low pump flow and purge leak has been completed. the impella device was returned for evaluation. the root cause of the low pump flow was most likely was sidearm filter damage. the root cause of the high purge pressure is most likely due to biomaterial buildup occluding the purge gap. this report is being filed as part of a retrospective review of historical records.
17795 | 1220648-2024-20101 | 20423707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20423707&pc=OZD | 11/05/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 01/05/2022 | 10/10/2024 |  |  | 07/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105918 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis and limb ischemia. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of hemolysis and limb ischemia. the product was not returned. the root cause of the hemolysis and limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
17796 | 1220648-2024-20107 | 20423872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20423872&pc=OZD | 12/14/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 12/14/2022 | 10/10/2024 |  |  | 08/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023264109 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella cp for hemodynamic support. the us complainant had a cp being prepped when the sterile staff operator "had issues". the pump was removed and replaced. there was no harm to the patient. | [5 empty] | The cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
17797 | 1220648-2024-20106 | 20423886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20423886&pc=OZD | 12/12/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 12/12/2022 | 10/10/2024 |  |  | 08/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269941 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Break | Fluid/Blood Leak | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the complainant reported a fluid leak from the purge sidearm coinciding with a low purge pressure alarm. a purge sidearm bypass was performed without further issue. impella cp removed with significant emergent bleeding. patient required 8x 40 balloon to tamponade at raa impella cp site; covered stent required placed by the cardiac physician to achieve hemostasis at impella cp site. it was reported that the patient was having multiple episodes of ventricular tachycardia being defibrillated by internal automatic implantable cardioverter defibrillator team decision to reintubate and sedate for patient comfort - impella turned up to p8 for support during ventricular tachycardia episodes, ep team feels recurrent ventricular tachycardia is caused by recent ischemic damage and hope with time ventricular tachycardia will slow now that patient has been revascularized - plasma free 11. | [5 empty] | Only the purge sidearm was returned. visual inspection revealed the presence of two lengthwise cracks on opposite sides of the yellow luer. no other damage or abnormalities were found. therefore, it was determined that the cause of the purge leak was sidearm yellow luer damage due to the presence of sodium bicarbonate in the purge solution. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17798 | 1220648-2024-20103 | 20424016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424016&pc=OZD | 11/19/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 03/23/2022 | 10/10/2024 |  |  | 08/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119959 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella access site and received blood products. despite replacing the purge cassette, the purge pressure high alarm persisted. additionally, hemolysis was observed. it is unknown how hemostasis was achieved regarding the complication. the patient continued on support until the device was successfully weaned. | [5 empty] | Visual inspection of the returned pump revealed biomaterial observed occluding approximately 50% of the purge gap. the pump was connected to a console and ran in a water basin during which high purge pressures were reproduced. therefore, the root cause of the reported high purge pressure was most likely ingested biomaterial occluding the purge gap. refer to es2020-075 for more information. no data logs were provided. furthermore, the cause of the reported bleeding/melena was most likely patient condition and the relation to impella was determined to be unknown/unrelated. the root cause of the suspected hemolysis was unable to be determined as limited clinical information and no data logs were provided. biomaterial was found in the purge gap of the pump but without data logs the time in which biomaterial was ingested is unknown. bleeding also occurred which could have lead to low patient volume but without data logs it is not possible to tell when/if this effected the function of impella. this report is being filed as part of a retrospective review of historical records.
17799 | 1220648-2024-20108 | 20424137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424137&pc=OZD | 08/23/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 08/23/2022 | 10/10/2024 |  |  | 05/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218478 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had a tortuous femoral artery on the left side. a long sheath was used. it was difficult to advance the sheath due to kinking. insertion of the impella was not possible. motor housing did not pass the kinking. attempted buddy wire failed. a second pump was successfully inserted on the right side. no reported harm to the patient. | [5 empty] | A.5 ethnicity is unknown. the investigation for the reported issue has been completed. the returned pump and introducer were visually inspected and a kink in the introducer sheath was observed. no pump damage was observed. the cause of the pump unable to pass through the sheath was tortuous patient condition related resulting in a sheath kink. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17800 | 1220648-2024-20110 | 20424142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424142&pc=OZD | 12/30/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 12/30/2022 | 10/10/2024 |  |  | 04/30/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022073655 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, low flows and suction alarms occurred. the decision was made to explant the device. no patient harm was reported. | [5 empty] | The pump was examined and evidence of a kink in the cannula was found, which is likely the root cause of the low flow. the kink was likely due to patient anatomy as the clinical mentions a small ventricle. this report is being filed as part of a retrospective review of historical records.
17801 | 1220648-2024-20099 | 20424197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424197&pc=OZD | 12/08/2022 | 10/10/2024 |  | 10/10/2024 |  |  | 12/08/2022 | 10/10/2024 |  |  | 08/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269823 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that at the initial impella implant, suction alarms and low flows were present. the pump was removed and inspected for clot ingestion. it was observed that the low flow issues were due to a kink in the cannula. a second device was placed and encountered the same issue. decision was made to proceed with percutaneous coronary intervention (pci). the procedural outcome was the patient survived. | [5 empty] | The investigation for the reported low pump flow issues has been completed. returned complaint pump visually inspected. no biomaterial and foreign materials observed. no broken blades found. cannula kink observed that cause the low flow issue due to patient condition. the data logs showed suction alarms, low flow at required p-level. clinical details explained a blood ingression condition happened prior insertion of complaint pump. this report is being filed as part of a retrospective review of historical records.
17802 | 1220648-2024-20111 | 20424220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424220&pc=OZD | 03/07/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 03/07/2023 | 10/10/2024 |  |  | 10/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023279425 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. upon removal of the pump, red long tissue like substance was noticed wrapped around the tip of the pigtail. the patient was stable post implant and there was no reported harm to the patient. | [5 empty] | The investigation for the reported removal issue has been completed. the impella device was not received from the customer. however, clinical details provided were sufficient to determine the root cause. the root cause of the thrombus was most likely patient condition. no issues with the function of the pump were reported to suggest the thrombus was in the pump during operation. this report is being filed as part of a retrospective review of historical records.
17803 | 1220648-2024-20114 | 20424239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424239&pc=OZD | 03/22/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 03/23/2023 | 10/10/2024 |  |  | 01/31/2025 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309293 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with post cardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was attempted to be implanted with an impella cp device for mechanical circulatory support. during implantation, the cannula kinked due to difficult anatomy, preventing the pump from being delivered. the difficult anatomy was resolved with a buddy wire and percutaneous transluminal angioplasty. the pump was then removed and replaced with a second impella cp. there was no reported harm to the patient. | [5 empty] | The device was returned, and an investigation into the reported delivery issue has been completed. the root cause of delivery issue was determined to be patient condition. as noted in the impella ifu: delivery issue: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
17804 | 1220648-2024-20113 | 20424298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424298&pc=OZD | 03/20/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 03/20/2023 | 10/10/2024 |  |  | 11/30/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303998 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported during positioning of the impella, the guide wire got stuck inside the easy guide lumen. reportedly the physician threaded the wire after the easy guide lumen was removed. during removal of the guide wire, the device kinked into itself. several unsuccessful attempts were made to straighten the catheter out resulting in the delivery failure causing the impella to be removed and a new impella placed. there was no patient harm reported and this device was replaced. | [5 empty] | The investigation into the reported issues has been completed. the product was not returned. the root cause of the delivery issue is most likely a damaged guidewire. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17805 | 1220648-2024-20102 | 20424561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424561&pc=OZD | 11/11/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 02/22/2022 | 10/10/2024 |  |  | 05/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022085974 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis and a thrombus observed in the impella outflow upon explant. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation into the hemolysis and thrombus has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the hemolysis was not determined. the root cause of thrombus was due to biomaterial was observed on the pump at the time of pump removal. this report is being filed as part of a retrospective review of historical records.
17806 | 1220648-2024-20105 | 20424574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424574&pc=OZD | 12/01/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 12/05/2021 | 10/10/2024 |  |  | 07/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022111239 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a blood leakage from the center of the repositioning sheath and received blood products. additionally it was reported patient experienced hemolysis. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The product was returned and tested in pressurizing tubing, the failure mode was reproduced. the o-ring was removed a slight abnormalities were found. a lip on the outside and and an extended middle diameter could be seen. the root cause of the access site bleeding was determined to be a damaged gwru valve. this report is being filed as part of a retrospective review of historical records.
17807 | 1220648-2024-20116 | 20424675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424675&pc=OZD | 12/30/2021 | 10/10/2024 |  | 10/10/2024 |  |  | 01/04/2022 | 10/10/2024 |  |  | 10/31/2023 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142615 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fistula | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient experienced a fistula at the access site between the artery and vein. action taken to resolve the issue are unknown. the patient survived the procedure. | [5 empty] | A.5 ethnicity is unknown. the investigation into the reported issue has been completed. the impella device was not returned for evaluation. the cause of the fistula and its relation to impella were not determined since neither the product nor sufficient clinical information was provided. this report is being filed as part of a retrospective review of historical records.
17808 | 1220648-2024-20109 | 20424705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424705&pc=OZD | 02/22/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 02/22/2023 | 10/10/2024 |  |  | 10/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283757 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Required Intervention | Life Threatening | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant had placed an impella cp for the team to perform multivessel coronary intervention on a patient with ejection fraction of 35-40%. the patient went into right bundle branch block with impella insertion. however, the team proceeded with the intervention and explanted the pump after over the two-plus hour intervention. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received. the cause of the issue was most likely patient condition related due to right bundle branch block. as noted in the impella instructions for use: impella cp with smart assist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system section: warnings & cautions: warnings section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ this report is being filed as part of a retrospective review of historical records.
17809 | 1220648-2024-20112 | 20424729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424729&pc=OZD | 01/04/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 01/04/2023 | 10/10/2024 |  |  | 12/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023279158 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella access site and received blood products to address the issue. the impella cp was explanted to address the issue. no patient harm was reported . | [5 empty] | The investigation into the access site bleeding - major has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding - major issue was not determined since product was not returned and insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
17810 | 1220648-2024-20117 | 20424730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424730&pc=OZD | 04/14/2023 | 02/18/2025 |  | 10/10/2024 |  |  | 04/14/2023 | 10/10/2024 |  |  | 10/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278364 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was documented that natural biomaterial build up was found within the graft following scheduled pump removal. per standard procedure, the graft was cleaned prior to insertion of a new device. the patient survived the explant. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17811 | 1220648-2024-20118 | 20424809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424809&pc=OZD | 03/30/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 03/30/2023 | 10/10/2024 |  |  | 12/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024299376 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella cp was found to have kinked. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation into the product damage has been completed. the impella device was not returned. however, clinical details provided were sufficient to determine the root cause. the cause of the issue was use related as a cannula kink occurred as repositioning was done without a wire. this report is being filed as part of a retrospective review of historical records.
17812 | 1220648-2024-20115 | 20424843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424843&pc=OZD | 03/28/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 03/28/2023 | 10/10/2024 |  |  | 10/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024292670 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  | American Indian Or Alaskan Native | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump stopped immediately after starting. the pump was restarted but was saturated. the pump was removed and replaced with another device, and no patient harm was reported. | [5 empty] | The device was returned, and the investigation for the reported pump stop has been completed. the field lt log was reviewed, and immediate high motor current pump stops upon activation were observed. an upward trend in purge pressure was observed after activation which correlates with the biomaterial wrapping around the purge gap. the returned pump was visually inspected and biomaterial was observed entraining the base of the impeller and partially blocking the purge gap. the cause of the issue was biomaterial. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17813 | 1220648-2024-20119 | 20424856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20424856&pc=OZD | 04/20/2023 | 10/10/2024 |  | 10/10/2024 |  |  | 04/21/2023 | 10/10/2024 |  |  | 12/31/2024 |  | 10/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295202 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the pump did not start due to the lumen not being removed. the device got stuck in the impeller. this pump was removed and a new pump was placed. there was no patient harm reported. | [5 empty] | The impella device was not returned. the investigation into the pump did not start issue has been completed. based on the clinical account, the root cause of the pump being unable to start was the red lumen being left in the pump. this report is being filed as part of a retrospective review of historical records.
17814 | 1220648-2024-19838 | 20404515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404515&pc=OZD | 12/11/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/11/2022 | 10/09/2024 |  |  | 09/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273573 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during implant, there was a saturated flow issue with the left ventricle waveform measuring 80mmhg higher than the aortic placement signal. the pump was removed and replaced with a new impella cp, and there was no patient harm reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported saturated flow has been completed. the returned pump was visually inspected and biomaterial was observed against the purge gap. the cause of the issue was biomaterial. this report is being filed as part of a retrospective review of historical records.
17815 | 1220648-2024-19839 | 20406908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20406908&pc=OZD | 07/27/2021 | 10/09/2024 |  | 10/09/2024 |  |  | 07/27/2021 | 10/09/2024 |  |  | 07/31/2022 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1489399 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported hemolysis was noted during impella support. treatment details were not noted. additional information provided that the patient was successfully weaned and survived the procedure. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the impella device has not been returned for evaluation. the root cause of the hemolysis issue was not determined since the product was not returned and insufficient clinical details were provided. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17816 | 1220648-2024-19840 | 20406940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20406940&pc=OZD | 09/02/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 09/02/2022 | 10/09/2024 |  |  | 04/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217750 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. upon insertion of the impella across aortic valve, the patient coded requiring cardiopulmonary resuscitation. return of spontaneous circulation (rosc) was achieved. the surgeon placed a 2nd impella cp and support was initiated. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the patient coding during delivery was not determined due to insufficient clinical details. relation to impella unknown. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. this report is being filed as part of a retrospective review of historical records.
17817 | 1220648-2024-19841 | 20406953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20406953&pc=OZD | 09/02/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 09/02/2022 | 10/09/2024 |  |  | 06/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236363 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. us complainant reported that the patient was placed on impella cp for hemodynamic support. it was reported that there was a saturated flow issue with the impella device. following placement of the pump, flows were lower than expected and the motor current was measuring over 1200 ma. the device was removed and replaced with a new pump. | [5 empty] | The investigation for the reported delivery issue has been completed. the product was returned and biomaterial was found wrapped around the impeller. the root cause of the saturated flows was determined to be ingested biomaterial adding additional resistance to the impeller and therefore increasing motor current. this report is being filed as part of a retrospective review of historical records.
17818 | 1220648-2024-19842 | 20407046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407046&pc=OZD | 09/03/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 09/03/2022 | 10/09/2024 |  |  | 04/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223905 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the physician was unable to receive left ventricular numbers and the device experienced continuous suction/position alarm. the doctor was unable to resolve alarms. pump was exchanged. a large clot at the inlet was also noted. | [5 empty] | The investigation for the reported thrombus and low flow issue has been completed. the log confirms suction alarms throughout the case and also shows impella pump in aorta alarms. a visual inspection of the pump revealed biomaterial on the impeller. the cause of the low pump flow was biomaterial. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17819 | 1220648-2024-19843 | 20407081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407081&pc=OZD | 09/10/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 09/10/2022 | 10/09/2024 |  |  | 09/30/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143849 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a pump stop issue. the device was explanted. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the data logs of the placement signal failure show a pattern in the 5s that is consistent with a torqued/ damaged optical fiber. the data logs also confirm a high motor current pump stop on the 34th day of support. the root cause of the optical signal issue is most likely a damaged fiber line. the root cause of the pump stop was unable to be determined as no product was returned for investigation however the pump stop occurred on day 34 of support which is beyond the 8 day design life per design trace matrix 0046-9910. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ ' section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17820 | 1220648-2024-19844 | 20407093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407093&pc=OZD | 12/22/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/22/2022 | 10/09/2024 |  |  | 07/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238138 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was a pump stop during use of the device. the patient's hemodynamics were deemed stable; therefore, the pump was removed and did not require replacement. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. no clinical details provided were sufficient to determine a root cause. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17821 | 1220648-2024-19845 | 20407097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407097&pc=OZD | 12/12/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/12/2022 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262598 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. upon connecting the purge line to the yellow luer of the impella pump, the purge pressure dropped. the purge sidearm was inspected, and the reservoir was found to be cracked and leaking. the pump was removed and replaced along with the cassette and console, and no patient harm was reported. | [5 empty] | The investigation for the reported purge leak has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the purge leak could not be determined. as noted in the impella ifu: purge leak-pump: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
17822 | 1220648-2024-19846 | 20407103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407103&pc=OZD | 12/24/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/24/2022 | 10/09/2024 |  |  | 07/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250102 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped. attempts to restart were not successful, therefore, the pump was explanted. there was no harm to the patient. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the cause of the pump stop was ingested biomaterial. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17823 | 1220648-2024-19847 | 20407118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407118&pc=OZD | 12/15/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/15/2022 | 10/09/2024 |  |  | 09/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274337 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited reduced flows and persistent suction alarms. the pump was removed and replaced with another impella cp, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the low pump flow could not be determined. this report is being filed as part of a retrospective review of historical records.
17824 | 1220648-2024-19848 | 20407132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407132&pc=OZD | 12/26/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/26/2022 | 10/09/2024 |  |  | 07/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255475 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was a loss of placement signal during use of the device. the physician suspected this may have been caused by the shockwave device that was being used, simultaneously. there was no patient harm and support continued. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the cause of the optical signal issue was a damaged sensor caused by shockwave. this report is being filed as part of a retrospective review of historical records.
17825 | 1220648-2024-19849 | 20407163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407163&pc=OZD | 12/27/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/28/2022 | 10/09/2024 |  |  | 09/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253999 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an acute patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. there was a pump stop immediately upon attempted initiation. the easy guide lumen was still inside the pump. the device was removed and replaced with no patient harm. | [5 empty] | The investigation into the reported "pump did not start" issue has been completed. the device was returned for evaluation. the root cause of the reported issue was determined to be a red lumen issue, as part of the red lumen was left in the pump. this is a use related issue. this report is being filed as part of a retrospective review of historical records.
17826 | 1220648-2024-19850 | 20407184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407184&pc=OZD | 12/28/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/28/2022 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251399 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. there was discovery of a thrombus upon removal of the impella device. there was no noted ischemia and the thrombus was removed from the exposed vessel without the need for further intervention. | [5 empty] | The investigation for the reported thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was biomaterial. the source of the biomaterial was unknown. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17827 | 1220648-2024-19851 | 20407204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407204&pc=OZD | 01/06/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 01/06/2023 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270351 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. it was reported a failure to prime the device. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation into the priming issue has been completed. the impella device was not returned for evaluation. the root cause of the priming issue was use related. this report is being filed as part of a retrospective review of historical records.
17828 | 1220648-2024-19852 | 20407357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407357&pc=OZD | 02/01/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 02/01/2023 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259690 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while inserting, there was a delivery issue with the impella device. it was noted during insertion that the catheter kinked while crossing the valve. the insertion attempt was aborted and a new cp pump was implanted for patient support. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. root cause of the issue was use related due improper technique use caused kink in the catheter by the technician while crossing the valve. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17829 | 1220648-2024-19853 | 20407394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407394&pc=OZD | 02/01/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 02/01/2023 | 10/09/2024 |  |  | 03/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211710 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, physician was unable to deliver impella due to the pump becoming engaged with the aortic valve replacement and then kinking as a result, but upon further investigation, it was clarified that the pump experienced positioning issues, resulting in the pump replacement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the reported issue was most likely an interaction with another device specifically a transcatheter aortic valve replacement cage. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17830 | 1220648-2024-19854 | 20407528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407528&pc=OZD | 02/01/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 02/01/2023 | 10/09/2024 |  |  | 10/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285242 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the cp failed to position appropriately as it was under the papillary. the team made the choice after just 20 hours to explant and escalate care to the 5.5 pump. there was no harm or injury from the replacement/interruption in the support. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the positioning issue was determined to be use related as the pump was pushed too deep into the ventricle during the advancement of the repo unit. this report is being filed as part of a retrospective review of historical records.
17831 | 1220648-2024-19855 | 20407753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407753&pc=OZD | 09/18/2024 | 11/18/2024 |  | 10/09/2024 |  |  | 09/18/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526499 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was explanted due to the patient experiencing oozing from the access site. after explant, manual pressure followed by femstop was utilized. the patient was stable and survived the surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
17832 | 1220648-2024-19857 | 20407760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407760&pc=OZD | 10/05/2024 | 11/06/2024 |  | 10/09/2024 |  |  | 09/12/2024 | 10/09/2024 |  |  | 10/31/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024419201 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and after the impella implant, there was oozing in the pocket. heparin was withheld. support continued. additionally, the patient had a computed tomography scan and a small new stroke was noted. it is unknown if the stroke was related to the impella. support continued. furthermore, the patient developed a gastrointestinal bleed while on support. heparin was withheld and the pump was explanted. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the access site bleed, gastrointestinal bleed, and the stroke has been completed. product was not returned for review. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the vascular damage and stroke was unable to be determined as limited clinical details were provided and no product was returned for review. the cause is unknown relation to impella.
17833 | 1220648-2024-19856 | 20407768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407768&pc=OZD | 02/27/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 02/27/2023 | 10/09/2024 |  |  | 11/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278826 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implant, impella pump and guidewire had delivery issues. following initial delivery, misaligned positioning prompted a suction alarm to appear. the pump was subsequently removed, and a guidewire was re-inserted into the tip of the pump for attempted re-delivery. the p level was set to p-0 during the delivery; however, alarms began to appear, and the p level was raised to p-1 to resolve the alarms. as the guidewire was still in the pump at the time, the wire became lodged in the pump and both the pump and wire had to be removed. an intra-aortic balloon pump was placed to continue with patient support. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. based on the clinical data, root cause of the issue was use related that suction alarm appeared due to misalignment during position adjustment, so reinsertion was considered. root cause of the delivery issue was use issue related too because during re-insertion into the left ventricle, the level was raised to p-1 and the wire was stuck. this led to delivery difficulties and the wire and pump had to pulled out. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17834 | 1220648-2024-19858 | 20407861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407861&pc=OZD | 09/18/2024 | 11/18/2024 |  | 10/09/2024 |  |  | 09/18/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539615 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was unable to be advance through the vasculature due to the patient anatomy. the impella insertion was aborted and no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the pump could not be implanted and was aborted due to the patient's anatomy. corrections to the initial mfr report: | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
17835 | 1220648-2024-19860 | 20407967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407967&pc=OZD | 09/18/2024 | 11/04/2024 |  | 10/09/2024 |  |  | 09/18/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531383 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and they had limb ischemia. after the implant, lower extremity angiogram revealed minimal flow beyond impella sheath. external fem-fem bypass placed proactively prior to percutaneous coronary intervention (pci) to mitigate distal extremity ischemia. post pci, staff was still unable to doppler distal pulse. the impella was explanted, and hemostasis was obtained. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
17836 | 1220648-2024-19864 | 20408002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408002&pc=OZD | 09/18/2024 | 11/13/2024 |  | 10/09/2024 |  |  | 09/18/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539059 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and the impella was caught up in mitral valve and the doctor was unable to reposition the pump. they were unable to reposition without exiting the left ventricle. the doctor decided to remove and replace the impella 5.5 with a new one. the patient remained stable until the successful wean and explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined.
17837 | 1220648-2024-19869 | 20408045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408045&pc=OZD | 09/19/2024 | 11/15/2024 |  | 10/09/2024 |  |  | 09/20/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025531898 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported when the 67-year-old male patient was on impella 5.5 support and there was purge fluid leaking from the junction of the red plug and grey rubber seal on the catheter side. bone wax was attempted. staff is ready to replace the impella if the pump fails. support continued and staff monitored. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak - pump has been completed. the product was returned and there were no external damage or abnormalities were found on the returned pump. there was dried dextrose on the catheter and around the red handle. the pump was purged and purge pressure low alarms were reproduced. the red handle was opened, and purge fluid was observed to be leaking into the handle from the catheter. a small hole was observed in the purge lumen proximal to the motor housing after stripping the catheter and there was no evidence of use-related issues. therefore, it was concluded that the cause of the purge leak was damaged purge tubing stemming from manufacturing. d9 date device returned to manufacturer added. e4 should have been left blank on manufacturer device report 1220648-2024-19869.
17838 | 1220648-2024-19859 | 20408072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408072&pc=OZD | 09/18/2024 | 11/13/2024 |  | 10/09/2024 |  |  | 09/18/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025504233 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, an error message of ¿high motor current, pump stopped¿ was received. two attempts to restart the pump were made with the same responding error message. decision was made by surgeon to remove the impella 5.5. it was noted that the pump most likely pulled back into subclavian/ graft area. additionally, the patient had two short runs of non-sustained ventricular tachycardias and amiodarone bonus/drip was initiated. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the pump stop and the ventricular tachycardia (vt) has been completed. the data logs were evaluated which show that the pump was reduced to p2 and then there were position alarms. the clinical details stated that the pump was most likely pulled into the graft. the purge pressure began to rise and purge flow decreased. the pump was then attempted to be restarted but impella stopped - high motor current occurred on the first attempted and then a controller error was present on the second attempt. this controller error alarm appeared for the speed deviation that was detected for the second time due to not being able to start the pump. due to this signature in the logs, there was most likely biomaterial blocking the purge gap as well as on the impeller. the product was returned and there was confirmed biomaterial in the gap and on the impeller. this material was removed and the pump was able to run without reproducing a pump stop. the root cause of the pump stop is due to biomaterial. the physician stated that they pulled the pump back too far in the ao and most likely into the graft. this led to the pump stop. the root cause of the positioning issue is most likely due to use. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
17839 | 1220648-2024-19861 | 20408082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408082&pc=OZD | 09/18/2024 | 11/23/2024 |  | 10/09/2024 |  |  | 09/18/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521534 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant had placed a impella cp for support of a patient admitted in with in acute myocardial infarction and cardiogenic shock. it was reported that prior to leaving the operating room, after coming off bypass, the medical doctor attempted to manipulate impella due to suction alarms after coming off of bypass. the impella was manipulated past aortic valve and a decision was made to remove the impella upon return to unit. the patient remained stable until the successful wean and explant of the impella. | It was additionally reported that after the decision was made to deploy the impella cp during left heart catheterization procedure because the patient began decompensating rapidly, there was a delay of therapy. when the abiomed representative requested a d5w or equivalent purge solution, the catheterization laboratory stagg were unable to locate the appropriate solution. there was an approximately 10 second delay to the deployment of the impella as the priming could not proceed until purge solution was obtained. after the procedure, the abiomed support staff setup continuing education for the catheterization laboratory staff to address impella procedural issues and reinforce training for best practices. after impella support was initiated, the patient was hemodynamically stable. the patient remained stable until the successful wean and explant of the impella. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely attempt to recross the valve wirelessly. e4 should have been left blank on manufacturer device report 1220648-2024-19861.
17840 | 1220648-2024-19866 | 20408104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408104&pc=OZD | 09/18/2024 | 10/07/2025 |  | 10/09/2024 |  |  | 09/18/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025512883 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient on impella 5.5 support had hematoma and bleeding post implant. anticoagulant was stopped and pressure was applied to the site. additionally, an echocardiogram is reading left ventricle thrombus. treatment of the thrombus is unknown however it is noted that the procedural outcome was the patient survived surgery. | The complainant also reported that the placement signal was falsely alarming in the aorta, triggering 'placement signal unreliable' alarms. the alarms were subsequently muted. |  |  |  |  | The pump was returned for investigation. no physical abnormalities were observed on the returned product. the cause of the access site bleeding issue was most likely related to anticoagulation management since heparin stoppage resolved the bleeding issue with applied manual compression. as the necessary information was not provided, the root cause of the thrombus was not determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-19866 d9 date device returned to manufacturer missing. e4 should have been left blank. h3 device not returned to manufacturer for evaluation should have been no. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was returned for investigation. no physical abnormalities were observed on the returned product. the pump passed the laser continuity test, and all five optical parameters were within specification range. the pump operated as expected in a bucket of water. data logs show that several placement signal accompanied by a placement signal not reliable (psnr) alarm during the case run. the cause of the optical signal issue was not determined, as no abnormalities were noted on the returned pump. pump worked up to the specification. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-19866. d10 concomitant medical products and therapy dates updated. g1 reporting contact email and reporting contact fax number updated.
17841 | 1220648-2024-19867 | 20408148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408148&pc=OZD | 09/19/2024 | 12/06/2024 |  | 10/09/2024 |  |  | 09/19/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534511 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed bleeding from around the impella access site. manual compression, pressure bandage, and one unit of blood was administered. the aptient remained stable until the successful wean and explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: b5-a typo was found in the narrative. it stated "the aptient remained stable until the successful wean and explant." but should have been the patient remained stable until the successful wean and explant.
17842 | 1220648-2024-19862 | 20408152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408152&pc=OZD | 09/18/2024 | 11/19/2024 |  | 10/09/2024 |  |  | 09/18/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team was forwarded a complaint from facebook that states the following: "husband had this when he had a massive heart attack. it punched a hole in his aortic valve causing him to have a valve replacement 6 months later when it was discovered." no other information was provided. | [5 empty] | The investigation for the ventricle perforation has been completed. pump was not returned for investigation. data log review was not required for this case. as per the clinical, patient suffered massive heart attack that punctured his aortic valve, requiring valve replacement six months later when the damage was discovered. the reported failure mode of ventricular perforation was changed to heart valve damage, which is more specific to address reported issue of heart damage. the cause of the heart valve damage issue was not determined since product was not returned. the relation between heart valve damage and impella pump was unknown due to limited clinical information. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. . d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown.
17843 | 1220648-2024-19863 | 20408154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408154&pc=PYX | 09/18/2024 | 10/09/2024 |  | 10/09/2024 |  |  | 09/18/2024 | 10/09/2024 |  |  | 08/31/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024395787 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported the patient was on impella rp flex support due to chronic total occlusion of right coronary artery. while on support, flows dropped from high 2s to low 2s within a couple of minutes. patient became further hypotensive. the decision was made to remove the impella rp flex. later the rep confirmed that there was a clot upon removal of the device. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
17844 | 1220648-2024-19868 | 20408163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408163&pc=OZD | 09/19/2024 | 12/06/2024 |  | 10/09/2024 |  |  | 09/19/2024 | 10/09/2024 |  |  | 06/30/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375155 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. it was reported that the impella was unable to pass through vasculature. it was also reported that bleeding was noted at the impella access site. the pump was removed to manage the bleeding. the patient outcome was noted to be stable. | [5 empty] | The investigation for the delivery issue and access site bleed has been completed. the impella was not returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely use related, lack of experience, impella pump removed to manage the bleeding and second pump was inserted successfully. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. correction to the initial mfr report: d4-udi number. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
17845 | 1220648-2024-19865 | 20408182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408182&pc=OZD | 09/18/2024 | 11/13/2024 |  | 10/09/2024 |  |  | 09/18/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505122 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support due to cardiogenic shock. while suturing impella graft, the impella moved out of place out of the left ventricle and into the aorta. due to its length, the impella came out. the decision was made to remove and attempt to re-insert with another impella. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issues was most likely use-related. e4 should have been left blank on manufacturer device report 1220648-2024-19865
17846 | 1220648-2024-19872 | 20408241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408241&pc=OZD | 09/21/2024 | 11/14/2024 |  | 10/09/2024 |  |  | 09/21/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527753 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Tthe complainant reported a 73-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. the pump was placed but there was a hematoma at the groin site. the hematoma was treated by manual pressure, femostop, blood and albumin in the icu. the pump remained on for another 6 days of support, and was successfully explanted. the patient is noted to be stable. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-19872. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17847 | 1220648-2024-19870 | 20408272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408272&pc=OZD | 09/19/2024 | 11/11/2024 |  | 10/09/2024 |  |  | 09/19/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1652655 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller had a controller error alarm. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) controller alarm has been completed. real time (rt) data logs were reviewed which align with a transient loss of optical data. the aic logs confirmed the reported failure mode given there was a controller error alarm (opm comm stopped ¿ event set # 1043) triggered during the clinical procedure. the console was returned for evaluation finding the front optical connector was inspected and cleaned. the root cause of the controller error was due to an aic software issue.
17848 | 1220648-2024-19874 | 20408309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408309&pc=OZD | 09/24/2024 | 11/16/2024 |  | 10/09/2024 |  |  | 09/24/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534251 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support developed a hematoma at the axillary site. heparin was on hold and manual compression done. the patient remains on impella support. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-19874 e4 should have been left blank g1 reporting contact fax number missing. h6 code 1888 was reported incorrectly. new code was added to health effect - clinical code | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
17849 | 1220648-2024-19871 | 20408313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408313&pc=OZD | 09/21/2024 | 11/16/2024 |  | 10/09/2024 |  |  | 09/21/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507886 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported during implant procedure for 72-year-old male patient, the impella 5.5 would not go past the valve strut. the implant was aborted due to patient anatomy. the patient is noted to be stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, unable to pass through the vasculature as per the clinical description provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-19871. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿
17850 | 1220648-2024-19873 | 20408343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408343&pc=OZD | 09/22/2024 | 11/17/2024 |  | 10/09/2024 |  |  | 09/22/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492073 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 52-year-old female patient, the access site was slightly oozing. a stitch was placed, and the patient was given 1 unit of packed red blood cells. the pump was not explanted as a result, and the patient is noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, site was slightly oozing yesterday; stitch placed overnight to achieve hemostasis as per the clinical description provided.
17851 | 1220648-2024-19876 | 20408506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408506&pc=OZD | 08/29/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 08/29/2022 | 10/09/2024 |  |  | 03/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214423 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implant, a kink was noted and possibly thrombus on the catheter of the device; therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation for the reported product damage (cannula kink) and thrombus has been completed. the product was returned, and the cannula kink was confirmed. per the results of the clinical review and investigation: returned complaint pump visually inspected. several kinks observed throughout the catheter. thrombus not observed. the cause of the issue was patient condition related. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
17852 | 1220648-2024-19875 | 20408605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408605&pc=OZD | 09/12/2024 | 11/21/2024 |  | 10/09/2024 |  |  | 09/19/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 161031 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had a battery failure. the aic was plugged in and staff attempted to work on the aic to no avail. this issue was discovered during preparation and no reported patient harm. | [5 empty] | The investigation for the console (aic) battery failure (red alarm). data logs were reviewed which showed that the battery failure issue was confirmed. there were numerous battery failure alarms (transport connection blown/missing) (event set #360) observed in the logs from the reported occurrence date. console was returned for evaluation. the battery failure issue was able to be reproduced. results of running the batttest utility revealed that cells in both batteries had voltages of less than 2.3mv. this battery failure issue was determined to be caused by depleted batteries. these depleted batteries could not be charged back up because they were locked out due to low cell voltage (less than 2.3v). | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
17853 | 1220648-2024-19878 | 20408657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408657&pc=OZD | 09/08/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 09/08/2022 | 10/09/2024 |  |  | 06/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235547 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the pump could not advance through the innominate artery and the device was removed due to the length of the procedure and blood loss. calcification of the vessel was reported. the patient received one unit of blood products. the patient survived the procedure. | [5 empty] | The investigation into the delivery issue and access site bleeding has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the delivery issue was patient condition related. the root cause of the access site bleeding was not determined since product was not returned and there isn't sufficient clinical information regarding the bleed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17854 | 1220648-2024-19877 | 20408675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408675&pc=OZD | 09/02/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 09/02/2022 | 10/09/2024 |  |  | 09/30/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022132584 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis due to elevated levels. dialysis was performed because of underlying poor renal function. additionally suction alarms were observed. the patient continued on support until the device was successfully weaned. a clot was observed at the inlet during the explant. the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer. neither the product nor the data logs were returned. therefore, the cause of the low pump flow was not determined. the cause of the hemolysis and its relation to the impella was also unable to be determined. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17855 | 1220648-2024-19880 | 20408683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408683&pc=OZD | 12/14/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/14/2022 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260679 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient presenting for a mitraclip was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped and was restarted. the impella pump speed level was adjusted to manage this issue, but the physician elected to remove the pump. there was no reported patient harm relating to this issue. | [5 empty] | The investigation for the reported pump stop has been completed. the product was returned and the pump stop was confirmed. the root cause of the pump stop was determined to be ingested biomaterial. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17856 | 1220648-2024-19882 | 20408730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408730&pc=OZD | 09/27/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 09/27/2022 | 10/09/2024 |  |  | 04/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023213801 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis resulting in volume auxiliary level adjustment. the physician reportedly was not concerned with the hemolysis as the patients¿ blood pressure was good and cardiac function was recovering. additionally, it was reported the patient had an issue with access site bleeding from patient movement, which was addressed with pressure and suturing of the insertion site. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the hemolysis and access site bleeding issues have been has been completed. no data logs were returned for analysis. the cause of the hemolysis was most likely patient anatomy condition related and related to the patient's volume status. the cause of the access site bleeding was most likely patient movement based on the provided clinical information. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: hemolysis. impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system: section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ access site bleeding. impella cp with smart assist system: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella rp flex with smart assist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smart assist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer: as noted in the impella ifu ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
17857 | 1220648-2024-19883 | 20408738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408738&pc=OZD | 12/14/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/14/2022 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263777 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after approximately 16 days on support, the impella stopped while on extracorporeal membrane oxygenation (ecmo). the pump was explanted, and there was no reported patient injury. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the pump stop could not be determined. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17858 | 1220648-2024-19879 | 20408802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408802&pc=OZD | 12/14/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/14/2022 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248647 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there were unexpected low flows and suction alarms. the pump was confirmed to be in good position and the patient was provided fluid volume, but the flows continued to slow until reaching zero. the device was explanted from the patient and no further intervention was specified, but there was no reported harm to the patient. | [5 empty] | The device was returned and evaluated, and the investigation for the reported low pump flow has been completed. biomaterial was found to be ingested near the impeller and in the cannula. the root cause of the low pump flow was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
17859 | 1220648-2024-19881 | 20408892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20408892&pc=OZD | 04/16/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 04/17/2023 | 10/09/2024 |  |  | 12/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024317953 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | Difficult to Insert | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was noted that the device was removed for flushing with sterile heparinized saline. after the flushing procedure, it was reported that the motor housing did not advance beyond the distal end of the peel-away sheath. the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | Impella cp product was returned for evaluation. the cause of the issue was use related to a catheter kink observed on the returned product. no introducer was returned. this report is being filed as part of a retrospective review of historical records.
17860 | 1220648-2024-19884 | 20409071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409071&pc=OZD | 09/20/2024 | 11/12/2025 |  | 10/09/2024 |  |  | 09/20/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528121 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that after dressing change on the patient with impella 5.5 support, the auomated impella controller (aic) started alarming impella position unknown with a flattened motor current. aortic (ao) placement signal uncoupled from the left ventricle (lv) waveform and above looking ventricular and lv numbers low or negative. impella position went from 5.1cm below aortic valve (av) on echocardiography to 6.1cm below. impella was pulled back to 4.9cm below av with no changes to the aic waveforms or numbers. aic continued to alarm impella position unknown after proper repositioning and position confirmation with yellow flows. there was concern that it was a sensor error, and the surgeon opted to replace the 5.5. there was no harm to the patient. the patient survived the event. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. optical signal issue: the pump was returned for investigation. the pump was subject to the optical bench test, and all parameters were within range. the pump passed the laser continuity test as well. there was biomaterial found around the impeller blades and on the sensor, however, no placement signal unreliable alarms were seen when the pump was connected to the console in the lab. data logs were returned for analysis. log review showed a drop in motor current and loss of motor current pulsatility before the reported "impella position unknown" alarm, which was accompanied by a suction alarm. the p-level was alternated between p-4 and p-5 multiple times, but motor current never recovered to initial levels, and the motor remained non-pulsatile. while the motor current dropped, the calculated lvp looked like a negatively offset aortic signal because the optical sensor remained in the aorta. positioning and suction alarms continued after the reported successful repositioning. there were no motor current spikes suggesting biomaterial ingestion. it is also unknown if there was a blockage of the pump inlet such as poor positioning against the heart wall. the root cause of the optical signal issue was not determined due to insufficient clinical details, and inconclusive evidence from product evaluation and data log analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated. d10 concomitant medical products and therapy dates updated. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of placement signal issue has been completed. the cause of the issue was not determined since product was not returned. | D9 device returned to manufacturer date added g1 reporting contact email revised on manufacturer device report in accordance with updated procedures 1220648-2024-19884.
17861 | 1220648-2024-19885 | 20409079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409079&pc=OZD | 03/03/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 03/03/2023 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269665 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient on impella support experienced limb pain at the access site. fem to fem reperfusion catheter was placed and patient's condition improved. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the limb ischemia issue was patient condition due to patient having severe peripheral artery disease (pad) condition. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
17862 | 1220648-2024-19886 | 20409118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409118&pc=OZD | 10/16/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 10/16/2022 | 10/09/2024 |  |  | 06/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237188 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported upon arrival to the laboratory the clinician noted the aortic placement optical signal on the automated impella controller reading the mean arterial pressure (map) of 120 while the invasive aortic pressure map was 70¿s. the clinician made the physician aware of the difference and reminded the physician to treat off the invasive aortic from the monitor, which correlated with noninvasive cuff pressures at that time. the aortic placement signals continued to be well above the invasive aortic monitor. the percutaneous coronary intervention was completed when the system alarmed "impella stopped. motor current high:" the impella spontaneously restarted and returned to previous flows. the clinician recommended removing this impella and since the patient was hemodynamically stable the physician agreed. this impella was removed and a new impella placed. no further information is available. | [5 empty] | The investigation into the temporary pump stops and optical signal issues have been completed. the impella pump was returned for investigation. the data logs confirm a sudden high motor current pump stop. no abnormalities were found during testing of the returned pump the pump ran within specification. the root cause of the sudden pump stop was most likely ingested biomaterial that passed through the pump, when the biomaterial impacted the impeller, it caused a pump stop. the logs show that the optical sensor was functioning, and the raw optical sensor was elevated indicating that the sensor was under additional pressure. the root cause of the optical signal issue was determined to most likely be patient condition. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi section: entering cardiac output as noted in the impella ifu ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped as noted in the impella ifu ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17863 | 1220648-2024-19889 | 20409173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409173&pc=OZD | 09/12/2024 | 11/19/2024 |  | 10/09/2024 |  |  | 09/12/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525137 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for support of a patient admitted in for a high risk percutaneous coronary intervention. the pump site of access developed a hematoma. the hematoma was treated by transfusion of five units of red blood cells, two units of platelets and two units of plasma, and the team tightened the perclose sutures. the patient was made care/comfort only and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
17864 | 1220648-2024-19892 | 20409191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409191&pc=OZD | 03/06/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 03/06/2023 | 10/09/2024 |  |  | 06/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249086 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient started bleeding and had a hematoma at the access site. there was no blood given to the patient due to religious beliefs. decision was made to explant impella and manual pressure was used to close the access site. patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the access site bleeding was most likely due to lack of vessel recoil. as noted in the impella ifu: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17865 | 1220648-2024-19890 | 20409202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409202&pc=OZD | 12/22/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/22/2022 | 10/09/2024 |  |  | 08/31/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129815 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Embolism/Embolus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported an echocardiogram revealed an air bubble in the left ventricle approximately two hours after the impella was placed. additionally, on 12/24/2022, it was reported that when the nurses were turning the patient the impella was pulled back into the aorta. the pump was removed by the physician as the patient was hemodynamically stable and the physician was unable to re-advance the pump. there was no patient harm reported, and this device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the embolism and positioning issues have been completed. the cause of the positioning issue was patient movement when turned, because the impella pulled back from aorta. the cause of the embolism issue was not determined since the product was not returned and the clinical details provided were insufficient to determine a root cause. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: embolism: impella cp with smart assist system: section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system: section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves as noted in the impella ifu ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ as noted in the impella ifu "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17866 | 1220648-2024-19891 | 20409236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409236&pc=OZD | 01/17/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 01/17/2023 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265343 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. purge pressure alarms were triggered due to a purge leak and the pump had stopped. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was returned. the cause of the pump stop was sidearm reservoir damage. a pinhole leak in the blue reservoir membrane was observed which was likely due to improper use of a sharp object being inserted into the reservoir. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17867 | 1220648-2024-19887 | 20409243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409243&pc=OZD | 11/05/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 11/07/2022 | 10/09/2024 |  |  | 05/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023223103 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Hemorrhagic Stroke | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia. it was reported from the time of impella placement the blood flow was maintained by anterograde blood infusion from extra-corporeal membrane oxygenation (ecmo) to the impella side. after five days, ecmo was removed, and ischemia occurred without antegrade blood transfusion on the impella side, reportedly a bypass was performed, however, the lower extremities did not improve. a dialysis console was reportedly used, and the blood supply was mechanically supported and maintained. additionally, on 11/6/20221 it was reported a computer tomography confirmed a cerebral hemorrhage. reportedly the exact location and degree of bleeding was unknown, there was no shift in the brain, hypoxic encephalopathy, or neurological abnormalities. it was also reported upon removal of the impella on (b)(6) 2022, anticoagulation was discontinued, and the patient progress was being monitored. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the cerebral hemorrhage and limb ischemia have been completed. the cause of the limb ischemia was the use of a sheath that was not indicated for use with impella. per ifu 0048-9007 only abiomed introducers are approved for use with the impella cp. the cause of the patient injury, cerebral hemorrhage, was most likely patient condition and had an unknown relationship with impella based off the limited clinical details. the cause of the cerebral hemorrhage was patient condition related. the cause of the limb ischemia was the use of a sheath that was not indicated for use with impella. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: limb ischemia: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ cva-ischemic/hemorrhagic: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: as noted in the impella ifu "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17868 | 1220648-2024-19896 | 20409248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409248&pc=OZD | 04/25/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 04/25/2023 | 10/09/2024 |  |  | 12/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322670 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported lower than expected flow during impella support. pump position was noted to be inadequate. the pump was removed and replaced with no harm to the patient. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has been returned for evaluation. the product was returned and the low pump flow was confirmed. the cause of the issue was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
17869 | 1220648-2024-19900 | 20409266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409266&pc=OZD | 09/12/2024 | 11/15/2024 |  | 10/09/2024 |  |  | 09/12/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527433 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed for support of a patient admitted in acute myocardial infarction/cardiogenic shock with a st elevated myocardial infarction. the patient delevoped bleeding ato the groin the impella was placed in the right groin. the bleed was prompted dressed and four units of blood were given. the pump was explanted and escalated to an impella 5.5 pump on the same day as the bleed was observed. the impella 5.5 was placed and support proceeded. the patient survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17870 | 1220648-2024-19898 | 20409271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409271&pc=OZD | 04/25/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 04/25/2023 | 10/09/2024 |  |  | 11/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295171 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp insertion procedure was unable to pass over the valve. this was due to the patient¿s anatomy. a shockwave was performed, in an attempt to open the iliacs vessel, which was unsuccessful. the decision was made to abort the procedure. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the cause of the issue. the cause of the delivery issue was determined to be patient condition related. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17871 | 1220648-2024-19888 | 20409276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409276&pc=OZD | 12/03/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/09/2022 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266128 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  |  |  | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, race and ethnicity unknown, with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced access site bleeding, pressure was held for approximately two hours and the pump repositioned resolving the bleeding. the patient received one unit of packed red blood cells. additionally, there was a high purge alarm issue reported by the nurse, the cassette was changed out, however, this did not resolve the alarm. reportedly the clinician arrived and reported the team had tissue plasminogen activator running and the device was still alarming. the clinician reviewed the items to watch out for such as pump stops, which occurred six days later. it was reported that support received an early morning call that the pump had stopped and could not be resolved by restarting. this impella was removed and replaced with another impella. it was reported that the patient survived the procedure. | [5 empty] | The investigation into the pump stop, high purge pressure and access site bleeding issues have been completed. the impella pump was returned for investigation. the investigator ran the impella for at least 5 minutes after activation and resulted in a stoppage. the testing of the returned pump determined the cause of the pump stop was bearing wear. the cause of the high purge pressure was unable to be determined since the initial purge cassette was not returned. the cause of the access site bleeding was use related since the bleeding occurred during transport and was resolved with repositioning the pump. the pump most likely shifted out of place at the access site during transport. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ pump stop impella cp with smart assist for use during cardiogenic shock and high risk cpi & impella rp with smart assist system section: using the automated impella controller with the impella catheter as noted in the impella ifu ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smart assist system catheter as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
17872 | 1220648-2024-19902 | 20409289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409289&pc=OZD | 03/02/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 03/02/2023 | 10/09/2024 |  |  | 04/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213441 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the impella was placed and when turned on to auto there were low flows of 3.6 liters and an increase in the patient¿s blood pressure. the team attempted to reposition at this time as the catheter appeared to be hugging the septal wall with some difficulty torquing towards apex. repositioned under fluoroscopy with radiopaque marker sitting at aortic valve area and flows of 3.6 liters. reportedly the peel away sheath was removed slowly and with significant bleed back even with fellow holding pressure above the insertion site. superficial oozing continued with tapered sheath in arteriotomy, and forward tension sutures were placed with multiple dressing changes to achieve hemostasis. it was later reported oozing was noted in the intensive care unit, the perclose was tightened and catheter dressed at 30¿40-degree angle with no further bleeding. initially the bleed was un-concerning to the team as it was superficial and appeared to stop with redressing, however, over the course of a few hours the patient lost a significant amount of blood from the right groin and eventually some puffiness was noted around the site. the team attributed the oozing to recent xarelto dose as well as brilinta during the procedure with heparin bolus. two residents reportedly held pressure for approximately 30 minutes each with hemostasis achieved and dressing replaced without oozing. the patient received two units of packed red blood cells over the course of two hours. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the low pump flow and bleeding have been completed. the root cause of the access site bleeding was most likely due to anticoagulation management. the cause of the low pump flow issue was not determined since product was not returned and clinical details provided were insufficient to determine a root cause. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17873 | 1220648-2024-19893 | 20409298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409298&pc=OZD | 03/01/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 03/01/2023 | 10/09/2024 |  |  | 10/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023279358 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was attempted to be implanted with an impella cp device for mechanical circulatory support. it was reported that the physician attempted to get past the aortic arch with a buddy wire. after multiple unsuccessful attempts, it was decided to abort the procedure. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the delivery issue was patient condition due to severe calcification and small aortic arch. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17874 | 1220648-2024-19899 | 20409301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409301&pc=OZD | 03/01/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 03/01/2023 | 10/09/2024 |  |  | 10/31/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156214 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella xxx for mechanical circulatory support. the physician was unable to advance the 13-centimeter sheath during placement, the impella was removed and flushed, then a .25-centimeter sheath was successfully placed. the patient reportedly woke up and bent his leg and bleeding started, the knee was immobilized, the patient sedated and femostop placed. reportedly the impella pulled into the aorta and the physician was unable advanced back through the aortic valve, the patient was then returned to the catheterization laboratory and a .16 french sheath was placed due to the bleeding. a new impella was placed. no further information is available. | [5 empty] | No product was returned. the cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access per ifu 0048-9007. the cause of the access site bleeding issue was most likely due to patient movement per clinical information. impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17875 | 1220648-2024-19895 | 20409312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409312&pc=OZD | 04/25/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 04/25/2023 | 10/09/2024 |  |  | 01/31/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309061 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, low pump flow was noted. as a result of the noted issue, the pump was removed and replaced with another impella cp pump. no patient harm was reported. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has been returned for evaluation. the product was returned, and the low pump flow was confirmed. the root cause of the low pump flow issue was verified to be biomaterial observed on the outflow cage and the impeller. this report is being filed as part of a retrospective review of historical records.
17876 | 1220648-2024-19894 | 20409374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409374&pc=OZD | 01/22/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 01/22/2023 | 10/09/2024 |  |  | 02/29/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197819 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  | White | Malfunction | No | Yes | Positioning Failure | Failure to Advance | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during the subsequent repositioning, the pump was pulled back across the aortic valve and attempts to re-cross the valve were unsuccessful. the pump was removed to facilitate backloading and reinsertion of the device; however, resistance was encountered, resulting in a bent guidewire. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issues and positioning issue. the device was received for evaluation. the root cause of the reported issue could not be determined; however, the most likely root cause of the delivery issue was due to patient condition related. the root cause of the positioning issue was wireless re-access. since the pump was pulled out during repositioning and the impella is not indicated for wireless re-access per ifu 0048-9007. the root cause of the delivery issue was a damaged guidewire. impella cp with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
17877 | 1220648-2024-19905 | 20409399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409399&pc=OZD | 06/21/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 06/21/2023 | 10/09/2024 |  |  | 03/31/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335309 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | Company representation |  |  | 0 DA |  | Unknown |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to being implanted with an impella cp device for mechanical circulatory support. it was reported the placement of the impella was aborted due to the patient¿s vasculature. there was no patient harm reported, and this device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. per clinical review the patient¿s condition resulted in the delivery issue due to the patient¿s anatomy. the root cause was the patient¿s anatomy. as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17878 | 1220648-2024-19903 | 20409427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409427&pc=OZD | 04/26/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 04/26/2023 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253889 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | White | Malfunction | Yes | Yes | Positioning Failure | [6 empty] | Embolism/Embolus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, automated impella controller showed wrong impella position. the pump position was corrected using echocardiogram. additionally, the physician noted air bubbles in the left ventricle and in the aorta, which were likely to be echo artifacts. no additional information was provided. patient survived the procedure. | [5 empty] | The investigation for the reported delivery issues and embolism has been completed. the impella device has not been returned for evaluation. the cause of the issue was not determined since product was not returned. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17879 | 1220648-2024-19904 | 20409434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409434&pc=OZD | 04/18/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 04/18/2023 | 10/09/2024 |  |  | 10/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290552 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced ventricular tachycardia arrest in which the initial impella cp ingested a clot during extracorporeal cardiopulmonary resuscitation. the impella was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The investigation into the thrombus and ventricular tachycardia has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the ventricular tachycardia was most likely biomaterial. the root cause of the thrombus was not determined since there is insufficient clinical information. as noted in the impella ifu: impella cp with smart assist system impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17880 | 1220648-2024-19907 | 20409564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409564&pc=OZD | 09/12/2024 | 11/13/2024 |  | 10/09/2024 |  |  | 09/12/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526858 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Defective Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed via a 14 french sheath to support a patient admitted in for a high-risk percutaneous coronary intervention. the team used the single access technique and had issues with the sheath. there was blood leaking and so the team explanted the sheath and placed the repo sheath and support proceeded. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation has been completed. the device was not returned for investigation, and data log review was not required for this case. as per the given clinical details, access site bleeding was reported after the peel-away sheath was placed. the bleeding was resolved after repositioning the sheath. the doctor was unsure of the origin of the bleed from the arteriotomy site or the sheath. the cause of the introducer damage that led to the access site bleeding issue was not determined since the product was not returned and sufficient clinical information was not provided. the reported introducer damage failure mode was changed to investigated failure mode access site bleeding as the suspected introducer damage event is the cause for bleeding event. the cause of the access site bleeding issue was not determined since the product was not returned and sufficient clinical information was not provided. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 added information as it was previously omitted on initial manufacturer device report number 1220648-2024-19907. b.1, b.2, h.1, h.6 health effect - clinical code, h.6 medical device problem code and h.10 instructions for use were revised due to investigation results changing the failure mode to access site bleeding since the initial manufacturer device report number 1220648-2024-19907 was submitted. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
17881 | 1220648-2024-19897 | 20409565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409565&pc=OZD | 03/01/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 03/01/2023 | 10/09/2024 |  |  | 12/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301098 |  | 1000080 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported the impella was placed and upon start up suction alarms occurred, and the pressure level would not go above three. due to the low pump flows repositioning was attempted under fluoroscopy, however, the pump came back across the valve after multiple suction events. reportedly, against recommendations from the clinician the physician removed and flushed the pump then re-wired it. upon re-insertion it was reported the pump generate high purge pressure alarms, no kinks or blockages were visible in the purge line. reportedly the clinician discussed tissue plasminogen activator through the purge line with the physician, however the physician chose to remove the impella and replace it with another impella due to the low pump flow. there was no patient harm reported, this device was replaced, and the patient was transferred to another facility for a higher level of care. | [5 empty] | The cause of the low pump flow was not determined since product was not returned. the cause of the high purge pressure was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
17882 | 1220648-2024-19909 | 20409579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409579&pc=OZD | 02/27/2021 | 10/09/2024 |  | 10/09/2024 |  |  | 03/31/2022 | 10/09/2024 |  |  | 07/31/2022 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 1492166 |  | 0048-0034 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced an ischemic stroke the day after the impella was removed. the auxiliary level was lowered to p1 for management. the reporter indicated concern that hypoxic blood flow to the head may have been caused by impella support. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the stroke has been completed. the cause of the ischemic stroke was not determined since the product was not returned. the issue is unrelated to the impella because the event occurred beyond 8 hours after removal. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: as noted in the impella ifu "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17883 | 1220648-2024-19910 | 20409584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409584&pc=OZD | 12/19/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/19/2022 | 10/09/2024 |  |  | 07/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254139 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. the complainant reported that the automated impella controller (aic) gave a sensor failure alarm when connected to the impella cp pump. the aic was swapped in an attempt to troubleshoot the failure, but the failure remained. the pump was subsequently replaced, and the issue resolved. there was no patient harm as a result of the reported failure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. data logs show that the "controller error" alarm occurred upon pump connection to the aic; this was followed by the "optical sensor system error" alarm. these alarms were reproduced multiple times with the same and another aic. no external damage or abnormalities were observed on the returned pump. when the pump was connected to the console, the placement signal railed out of specification and the "placement signal not reliable" alarm triggered. the 5s parameters of the pump's optical sensor were out of specification. optical continuity testing revealed a damaged fiber line within the red handle near patient cable-kink protector interface. both kink protector bonds were intact. no other damage or abnormalities were found. in conclusion, it was determined that the cause of the optical signal issue was a damaged optical fiber line. this report is being filed as part of a retrospective review of historical records.
17884 | 1220648-2024-19906 | 20409594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409594&pc=OZD | 12/15/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/15/2022 | 10/09/2024 |  |  | 07/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246681 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. while the impella was being primed and prepped, the team inadvertently broke the purge causing a purge leak. the pump was replaced with no patient harm reported. | [5 empty] | The investigation into the product damage issue has been completed. the impella pump was returned for investigation. the purge connection between the filter and the diaphragm broke during the setup. the cause of the purge leak was damage to the reservoir to filter joint. this report is being filed as part of a retrospective review of historical records.
17885 | 1220648-2024-19911 | 20409604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409604&pc=OZD | 05/10/2023 | 02/14/2025 |  | 10/09/2024 |  |  | 05/10/2023 | 10/09/2024 |  |  | 12/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318764 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient¿s urine was observed to be red in color. the pump was repositioned, but it was confirmed that the device remained in the papillary muscle, prompting a request for an additional echocardiogram due to hemolysis. medical staff attempted to reposition the impella pump to release the papillary muscle. however, when the impella device was pulled back, it was withdrawn into the ascending aorta. the patient became unstable, and an effort was made to recross the aortic valve under echocardiographic guidance. the patient was taken to the unit where the impella malposition was confirmed. the impella was pulled back into the descending artery and able to be straightened. the patient continued on support until the device was successfully weaned. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. data log review confirmed pump was in aorta (about 15 hours into support), then was stopped manually. the pump was not returned for evaluation. based on clinical description and data analysis, the root cause of the reported issue was most likely due to improper positioning of the pump. the root cause of positioning issue was due to wireless re-access to the aortic valve. | The impella device was not received from the customer. data review: data confirmed pump was in aorta (about 15 hours into support), then was stopped manually. product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position & the root cause of positioning issue was due to wireless re-access to aortic valve. this report is being filed as part of a retrospective review of historical records.
17886 | 1220648-2024-19901 | 20409609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409609&pc=OZD | 04/26/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 04/26/2023 | 10/09/2024 |  |  | 01/31/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312627 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced lower than expected pump flow. the pump was removed, and biomaterial was found on both the inflow and outflow cages. no patient harm was reported. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the issue was biomaterial. this report is being filed as part of a retrospective review of historical records.
17887 | 1220648-2024-19913 | 20409615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409615&pc=OZD | 10/07/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 10/07/2022 | 10/09/2024 |  |  | 03/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205013 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the complainant reported a pump stop issue with the impella device. it was noted that a purge system blocked alarm sounded on the automated impella controller (aic) which was followed by an increase in the motor current of the pump. the motor subsequently stopped and attempts to restart the pump failed. the pump was explanted as a result of the noted issue and the access site was closed. no patient harm was reported. | [5 empty] | The investigation for the reported pump stop issue has been completed. the lt log confirms a pump stop. prior to the pump stop the motor current rises, then drops and rises again until the pump stop. the rt logs shows many purge system blocked and high purge pressure alarms which do not occur until the end of the case. a visual inspection of the pump revealed a kink in the catheter near the red handle. the exact location of the kink that occluded the purge flow was not determined as there were multiple kinks observed in the purge line. when a kink occurs in the purge line, purge fluid is unable to protect the motor and blood ingress can occur. no blood was found in the purge line but blood was observed in the purge gap. the pump was torn down and a substance found on the magnet was tested and found to contain protein. the cause of the pump stop was most likely a kinked purge line. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
17888 | 1220648-2024-19914 | 20409650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409650&pc=OZD | 05/17/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 05/17/2023 | 10/09/2024 |  |  | 10/31/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158332 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding and hemolysis. it was reported that the patient had continuous bleeding from the femoral site overnight, constant manual pressure was held in the intensive care unit, from 2am-8am. the patient was brought back to the laboratory and received 1 unit of fresh frozen plasma, reportedly the patient¿s hemoglobin went for 14.4 to 11.1 overnight. the impella was removed due to the continued bleeding. reportedly there were no lab results due to the patient¿s hemolysis, no further information regarding hemolysis was provided. the impella was removed due to the bleeding. | [5 empty] | The investigation into the hemolysis and access site bleeding issues have been completed. the impella pump was returned for investigation. data was not returned for review. visual inspection found no defects on the pump. hemolysis test was conducted to observe if any defect was found on the pump. the result of hemolysis test showed pump passed the test indicating that there was no issue from the pump. the root cause of hemolysis was most likely due to patient condition. the root cause of access site bleeding was most likely high patient activated clotting time due to anticoagulation management issue. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
17889 | 1220648-2024-19912 | 20409680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409680&pc=OZD | 12/20/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/20/2022 | 10/09/2024 |  |  | 07/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246508 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was atttempted to being implanted with an impella cp device for mechanical circulatory support. during implantation, the pump would not pass the aortic valve. after several attempts, the initial pump was aborted and a new pump was implanted with no issues. | [5 empty] | The investigation for the reported issue has been completed. the product was returned and no abnormalities were found. the root cause of the delivery issue was most likely as a result of patient anatomy. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17890 | 1220648-2024-19915 | 20409684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409684&pc=OZD | 09/20/2024 | 08/25/2025 |  | 10/09/2024 |  |  | 09/20/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1811284 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the impella 5.5 was being prepped and when it was plugged in, there was an impella catheter defective alarm. the automated impella controller and pump were replaced. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h6 investigation conclusions revised on manufacturer device report in accordance with updated procedures. h10 additional manufacture narrative revised on manufacturer device report in accordance with updated procedures.
17891 | 1220648-2024-19916 | 20409690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409690&pc=OZD | 12/20/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/20/2022 | 10/09/2024 |  |  | 09/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275873 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. it was reported that the pump failed to deliver due to underlying patient anatomy issues. the impella implant was aborted and a 2.5 was placed with no harm reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, the cause of the issue was patient condition related. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17892 | 1220648-2024-19908 | 20409833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409833&pc=OZD | 12/19/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/19/2022 | 10/09/2024 |  |  | 07/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254139 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. the complainant reported an optical sensor system error alarm that occurred during impella support. the automated impella controller (aic) was replaced in an attempt to troubleshoot the issue, however, the failure remained. as a result of the ongoing failure, the pump was replaced with another impella cp pump and patient support continued. there was no harm to the patient as a result of the failure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the optical signal issue was most likely a damaged optical fiber line. the location and cause of the damage are unknown since product was not returned. this report is being filed as part of a retrospective review of historical records.
17893 | 1220648-2024-19917 | 20409870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409870&pc=OZD | 07/06/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 07/06/2022 | 10/09/2024 |  |  | 12/31/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165754 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation | Distributor |  |  | 50 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Japan complainant reported that a patient was on impella cp smart assist support when a pump stop occurred. the patient has been on support for 14 days. the cp was removed but not replaced. patient is stable post explant. | [5 empty] | The investigation was completed, no device or logs were returned. the cause of the issue was not determined since product was not returned. used beyond design life. as noted in the impella ifu: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
17894 | 1220648-2024-19918 | 20409916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409916&pc=OZD | 11/26/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 11/29/2022 | 10/09/2024 |  |  | 09/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023264886 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on the support the patient was in need of defibrillation. it is reported that the device was explanted within 3 hours of admission. survival outcome at explant noted the patient survived the procedure. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the ventricular fibrillation (vf) was most likely patient condition related. this report is being filed as part of a retrospective review of historical records.
17895 | 1220648-2024-19919 | 20409927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20409927&pc=OZD | 09/06/2024 | 11/19/2024 |  | 10/09/2024 |  |  | 09/13/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025525749 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support. while on support, an ultrasound revealed a blood clot-like image at the tip of the impella. the impella was removed. when it was removed, a blood clot was attached to the tip of the impeller 5.5 and the biological valve. the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation into thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h5 labeled for single use incorrectly reported on manufacturer device report 1220648-2024-19919.
17896 | 1220648-2024-19922 | 20410156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20410156&pc=OZD | 09/20/2024 | 12/06/2024 |  | 10/09/2024 |  |  | 09/20/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532952 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for a 61-year-old male patient admitted in an acute myocardial infarction / cardiogenic shock (ami/cgs). the pump accessed limb had signs of limb ischemia and the leg had vascular surgery consult for the cold and discolored leg. the team was discussing fem-fem bypass but, patient condition is poor and decisions ongoing. pump was successfully explanted after 6 days of support. patient required high dosage of medication post explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
17897 | 1220648-2024-19921 | 20410187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20410187&pc=OZD | 09/13/2024 | 03/23/2025 |  | 10/09/2024 |  |  | 09/13/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1303847 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, the automated impella controller (aic) was in patient use and the nurses had complaints that the alarms volume was too low on the aic. they could not hear the alarms properly when they were not close to the aic. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation of automated impella controller (aic) ¿ other has been completed. the console logs show multiple interactions of the user with the mute soft key of the console but do not confirm the reported low volume alarm. the inspection of the internal circuitry revealed no anomaly. a digital sound level meter was utilized and the console found to be within specification per preventive maintenance procedure. the root cause of inaudible alarms could not be confirmed from the logs or from device analysis. there was no problem found with the console. d2 common device name revised on manufacturer device report in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. | A.5 is unknown. the console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
17898 | 1220648-2024-19920 | 20410244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20410244&pc=OZD | 03/11/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 03/11/2023 | 10/09/2024 |  |  | 11/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301409 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the physician accidentally ripped catheter sleeve while trying to pull impella back. a tegaderm was placed to help keep catheter sleeve in place. procedure was successfully completed with this device. no harm done to the patient. | [5 empty] | The investigation for the reported product damage (sterile sleeve damage) has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the product damage was due to the physician accidentally ripping the "catheter sleeve." this report is being filed as part of a retrospective review of historical records.
17899 | 1220648-2024-19923 | 20410287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20410287&pc=OZD | 05/12/2023 | 02/14/2025 |  | 10/09/2024 |  |  | 05/15/2023 | 10/09/2024 |  |  | 01/31/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306207 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient, ethnicity unknown, noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported upon opening the device packaging there were no visible problems. reportedly when the connector cable was input into the automated impella controller there was a yellow alarm: ¿faulty impella system¿, ¿ do not use impella system¿, ¿discard impella system¿ message. the device was not used and was changed out to perform the procedure. there was no patient harm reported. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. the cause of the issue was not determined since product/data was not returned. | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
17900 | 1220648-2024-19924 | 20410362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20410362&pc=OZD | 09/20/2024 | 11/20/2024 |  | 10/09/2024 |  |  | 09/20/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025550601 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella 5.5 was being prepped and when it was plugged in, there was an impella catheter defective alarm. the automated impella controller and pump were replaced. the patient is noted to be stable. | [5 empty] | The investigation for the pump not detected issue has been completed. data logs were investigated, and seconds after the console recognizes the pump is plugged in, there is an error stating "eeprom problems while reading calibration data." roughly 10 seconds later, the impella defective alarm is triggered. the team then unplugs the pump from the console and plugs back in. this time, the pump is able to complete calibration and reach case start step 7, but the team unplugs the pump anyways. then, the replacement pump (sn (b)(6)) is plugged into the first console and also able to get through calibration and case start step 7, but is also unplugged. the imc logs for the next pump plugged into that console were investigated as well and a slightly different error occurred. in this case, the imc-emp1 error triggers followed by "eeprom problems while reading closed loop control data." nonetheless, the pump gets through case start, is started up, and runs for over 2 days. without returned product, it cannot be confirmed whether or not the complication was due to the pump, as it initially failed but then got through case start the second time. the root cause of the pump not detected could not be determined due to lack of product return and insufficient evidence from data logs. added- h6 impact code 4629. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
17901 | 1220648-2024-19926 | 20410398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20410398&pc=PYX | 09/13/2024 | 11/20/2024 |  | 10/09/2024 |  |  | 09/13/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529589 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Failure to Pump | Pumping Stopped | Defective Device |  |  |  | Thrombosis/Thrombus | [12 empty] | Life Threatening | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the impella rp flex placement the dilator pulled out of the sheath causing the sheath to become damage. a new peel away was used. additionally, while loading the impella rp flex, side port of the peel away ingested clot and was unable to be drawn back. due to unknown of where the clot was, the decision was made to replace the peel away with a new one. the impella rp flex was placed successfully. during impella rp flex initiation, the impella stopped. restarted at p-3 but motor spikes noted likely were due to the clot. purge pressure was increased with purge flow decline which led to purge blocked alarm. the decision was made to use tissue plasminogen activator and keep the impella rp flex in place with plans to exchange out if needed. | [5 empty] | D.9 a medical device problem code was added due to investigation results. the investigation of thrombus, pump stop, high pressure purge and introducer damage - mechanical has been completed. the pump and introducer were returned for evaluation. the failure mode of "pump stop" was changed to "pump did not start" as the pump stop occurred during pump initiation. the returned pump was attempted to be run in a bucket at p-9 and was able to be ran without stopping. clinical details noted the pump stop occurred when the pump was initiated and the pump was able to be restarted at p-3. data logs were reviewed and approximately 10 seconds after the pump was initiated, a motor current spike and immediate pump stop occurred. the pump was able to be restarted at p-3. the root cause of the pump did not start was most likely due to biomaterial ingestion. a failure mode of "high purge pressure" was added to the investigation for the unresolved high purge pressure that was occurring throughout the case. the returned pump had biomaterial present in purge gap. the pump was purged in a bucket of water and high purge pressure was reproduced. no blood ingress was present in purge line. the data logs from the field were reviewed and purge pressure began to increase right after motor current spike with a drop in purge flow over a period of 10 minutes before high purge pressure alarms were triggered. high purge pressure was persistent throughout the case. clinical details noted that the purge pressure began to rise after the pump stop occurred on pump initiation. tissue plasminogen activator protocol was attempted but it was not noted if it resolved the issue. the root cause of the high purge pressure was most likely due to biomaterial. as noted in clinical details, when using the second introducer for the impella rp flex insertion, the introducer aspirated a clot while drawing back from the side port. the introducer was removed and again replaced. this secondary introducer used was not returned for evaluation. as the necessary information was not provided, the root cause of the thrombus was not determined. it was noted in the clinical details that 3 peel away introducer sheaths were used during insertion. the first introducer used was noted to be damaged when a dilator was prematurely pulled out on insertion. the 23 french x 11 centimeter introducer was returned for evaluation and damage at the tip of the introducer was present. the root cause of the introducer damage - mechanical was most likely due to a sheath tip split. b.2 revised as life-threatening and disability or permanent damage were inadvertently selected on the initial manufacturer device report number: 1220648-2024-19926. h.6 due to investigation results, code 1503 should not have been on the initial manufacturer device report number: 1220648-2024-19926. code 1502 was added to reflect the investigation results information instead. h.10 instructions for use for pump stop should not have been reported on initial manufacturer device report number: 1220648-2024-19926 as the investigation results determined the better fitting option was that the pump did not start. | The introducer and pump were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
17902 | 1220648-2024-19925 | 20410656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20410656&pc=OZD | 09/20/2024 | 11/20/2024 |  | 10/09/2024 |  |  | 09/20/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501562 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 64-year-old male patient had small oozing at the right femoral artery. systemic heparin was held. additionally, the impella began to get suction alarms. all troubleshooting options were performed including decreasing p level, 250ml of albumin and repositioning under echo. despite all attempts, staff was unable to resolve the suction issues. impella was therefore explanted. hemostasis was achieved via perclose x2 and 10 minutes of manual compression. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and low pump flow has been completed. the impella device was not returned for evaluation. the cause of the access site bleeding cannot be determined since the complaint device not returned for investigation and insufficient clinical data provided. the cause of the pump flow issue most likely was biomaterial ingestion. h6 code 2993 was reported incorrectly on manufacturer device report 1220648-2024-19925. new code was added to medical device problem code.
17903 | 1220648-2024-19927 | 20410692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20410692&pc=OZD | 09/13/2024 | 11/20/2024 |  | 10/09/2024 |  |  | 09/13/2024 | 10/09/2024 |  |  | 12/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024301307 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant. however, the pump could not pass through the aortic stenosis. the impella 5.5 implant was aborted. the patient survived the event. | [5 empty] | The investigation for the delivery issues has been completed. the returned product was visually inspected and no abnormalities were found. the root cause of the delivery issue was most likely patient condition related. the impella 5.5 could not be placed as it could not pass through aortic stenosis as per the clinical description provided. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
17904 | 1220648-2024-19930 | 20410938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20410938&pc=OZD | 09/21/2024 | 11/21/2024 |  | 10/09/2024 |  |  | 09/21/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539554 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 68-year-old female patient was attempted to be implanted with an impella 5.5 but the procedure was aborted as the impella could not make the turn. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through due to patient anatomy. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿
17905 | 1220648-2024-19929 | 20411001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20411001&pc=OZD | 09/13/2024 | 11/20/2024 |  | 10/09/2024 |  |  | 09/13/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534032 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a patient admitted in acute myocardial infarction/cardiogenic shock. the pump was explanted after three hours prematurely due to an access site bleed. the patient remained stable and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | A.5 is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
17906 | 1220648-2024-19928 | 20411155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20411155&pc=OZD | 09/21/2024 | 11/18/2024 |  | 10/09/2024 |  |  | 09/21/2024 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556294 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp developed a small hematoma which resolved with manual pressured. activated clotting time (act) was still elevated at 246, heparin on hold until the act is under 180. bleeding in the access site started. the physician placed a forward tension suture and held manual pressure for 10 minutes. bleeding resolved and hematoma improved. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed and atrial fibrillation has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since additional suture resolved the bleeding.
17907 | 1220648-2024-19933 | 20411774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20411774&pc=OZD | 06/14/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 06/14/2023 | 10/09/2024 |  |  | 01/31/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311841 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient's vascular anatomy was too small and diseased distal to 14fr sheath. the decision was made to implant an intra-aortic balloon pump (iabp) for support. no known harm or injury reported. no further information is available. | [5 empty] | No product was returned. as per clinical, insertion of 14 fr introducer was aborted due to patient having small and diseased vessels in the right femoral artery. the cause of the difficulty to insert introducer issue was due to patient condition with patient having small and diseased vessels in right femoral. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
17908 | 1220648-2024-19935 | 20411937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20411937&pc=OZD | 06/29/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 06/29/2023 | 10/09/2024 |  |  | 01/31/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306020 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, demographics unknown, noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the placement of the impella was aborted due to the patient¿s vasculature. reportedly the physician attempted to advance the impella and found that there were too many narrow points and calcification in the vascular system and aborted the placement of the impella. there was no patient harm reported, and this device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. per clinical review the cause of the delivery issue was patient condition related due to the patient¿s small and diseased vessels. the root cause of the issue was patient condition related. impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
17909 | 1220648-2024-19932 | 20412039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412039&pc=OZD | 06/13/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 06/13/2023 | 10/09/2024 |  |  | 01/31/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331060 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there were delivery issues and low pump flow with the impella cp. after access was obtained using the impella sheath, the provider had difficulty crossing the aortic valve (av), first with the wire and then with the impella catheter. patient appeared to have abnormal cardiac anatomy with horizontal directed av and heart. the impella catheter appeared kinked under fluoroscopy and the device began to have low impella flows and suction events. decision was made to explant the impella. it was reported that the patient was stable, no further information provided. | [5 empty] | The returned pump was investigated and a kink in the cannula was observed. no other abnormalities were observed. the cause of the low pump flow was a kinked cannula which was kinked due to the patient's anatomy based on the provided clinical information. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
17910 | 1220648-2024-19931 | 20412089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412089&pc=OZD | 06/15/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 06/15/2023 | 10/09/2024 |  |  | 02/28/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342049 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. during support, a left ventricle clot observed post explant in imaging. there was no report of lasting harm, injury, or medical intervention. no further information is available. | [5 empty] | Product was not returned for investigation and data log were not provided as well. the thrombus was unrelated to impella since it occurred 3 days after explant. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17911 | 1220648-2024-19946 | 20412378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412378&pc=OZD | 07/06/2021 | 10/09/2024 |  | 10/09/2024 |  |  | 07/14/2021 | 10/09/2024 |  |  | 02/28/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022050187 |  | 0048-0034-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 25 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient developed hemolysis, no further information regarding the alleged hemolysis other than the patient recovered. it was also reported that the pump motor stopped and was replaced. this device was replaced for the pump stop. | [5 empty] | The investigation into the pump stop and hemolysis issues have been completed. the impella pump was returned for investigation. the cause of the pump stop was not determined since the failure mode could not be reproduced. the cause of the hemolysis was most likely improper positioning. the data logs showed frequent impella positioning alarms on the day of the reported hemolysis. the hemolysis was most likely improper technique by the end user. the pump stop was unable to be reproduced. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: hemolysis. impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ pump stop. impella cp with smart assist for use during cardiogenic shock and high-risk cpi & impella 5.5 with smart assist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions. as noted in the impella ifu ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17912 | 1220648-2024-19940 | 20412463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412463&pc=OZD | 01/26/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 01/26/2023 | 10/09/2024 |  |  | 09/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283123 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support low pump flows were noted during use of the device. the pump was removed, and a clot was found within the device. a new impella was then inserted with no harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on clinical details clot entrainment within pump noted and 10,000 units of heparin was given prior to insertion. data log reviewed and suction alarm and mc trends occurred and low flow shown. root cause most likely biomaterial ingested due to patient condition. this report is being filed as part of a retrospective review of historical records.
17913 | 1220648-2024-19936 | 20412545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412545&pc=OZD | 06/23/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 06/23/2023 | 10/09/2024 |  |  | 03/31/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343052 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the impella was inserted and was damaged during insertion by the physician and was unable to be started after three attempts. this impella was removed and a new impella placed. there was no patient harm reported, and this device was replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the pump did not start has been completed. the data logs show a pump stop with alarm 13-motor current high during case start. the cause of the pump did not start issue was not determined since product was not returned and sufficient clinical information was not provided. this report is being filed as part of a retrospective review of historical records.
17914 | 1220648-2024-19942 | 20412564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412564&pc=OZD | 01/31/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 01/31/2023 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262752 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while inserting the impella cp, there was a failed delivery of the device due to the patient's anatomy. insertion was aborted with no reported harm to the patient. | [5 empty] | The investigation for the reported difficulty inserting and removing issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was possible. clinical details provided were sufficient to determine a root cause. the cause of the difficulty inserting and removing issue was patient condition related, specifically the patient's vascular anatomy. this report is being filed as part of a retrospective review of historical records.
17915 | 1220648-2024-19945 | 20412596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412596&pc=OZD | 02/01/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 02/01/2023 | 10/09/2024 |  |  | 09/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023277705 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while inserting, there was a failure to deliver the device due to patient¿s anatomy. the pump insertion was aborted and a new impella was used without further reported issues. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
17916 | 1220648-2024-19937 | 20412670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412670&pc=OZD | 06/26/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 06/26/2023 | 10/09/2024 |  |  | 12/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318896 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported the placement of the impella was aborted due to the patient¿s vasculature. reportedly the physician attempted to advance the impella and was unable to advance due to angle of the patient's ascending aorta. a computer tomography was reviewed, and the patient was found to have bovine aorta arch. the placement of the impella was aborted and a intra-aortic balloon pump was placed. there was no patient harm reported, and the impella was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. per clinical review the cause of the delivery issue, the physician was unable to pass due to angle of ascending aorta. the patient was found to have a bovine aortic arch after a computer tomography review and insertion was aborted due to patient anatomy. the cause of the delivery issue was patient condition related due to patient having bovine aortic arch per clinical information. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
17917 | 1220648-2024-19938 | 20412672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412672&pc=OZD | 01/11/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 01/11/2023 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270093 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon initial activation, a pump stop alarm was triggered. the pump was restarted successfully. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the product was returned. the root cause of the pump unable to start was not determined since the issue was unable to be reproduced. the returned product operated to specification. this report is being filed as part of a retrospective review of historical records.
17918 | 1220648-2024-19941 | 20412681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412681&pc=OZD | 08/23/2021 | 09/29/2025 |  | 10/09/2024 |  |  | 08/23/2021 | 10/09/2024 |  |  | 06/30/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094493 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, a cannula kink and high motor current during use of the device. support was able to proceed without any harm or adverse impact to the patient. patient was weaned without difficulty. | [5 empty] | The returned pump was visually inspected and biomaterial was observed in the purge gap. the returned introducer was inspected and a sheath kink was confirmed and biomaterial was observed within the sheath. the cause of the low pump flow was ingested biomaterial. the sheath was likely not flushed properly and a clot was pulled into the pump from the introducer. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information, the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
17919 | 1220648-2024-19943 | 20412698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412698&pc=OZD | 06/28/2020 | 10/09/2024 |  | 10/09/2024 |  |  | 01/06/2022 | 10/09/2024 |  |  | 03/31/2022 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1477975 |  | 0048-0034 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed hemolysis, no further information was available regarding any treatment for the hemolysis. further details were not available, due to facility restrictions no further information was able to be obtained. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the hemolysis has been completed. the cause of the hemolysis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. there is an unknown relationship between the impella and the hemolysis. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system. section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17920 | 1220648-2024-19939 | 20412705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412705&pc=OZD | 07/31/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 07/31/2023 | 10/09/2024 |  |  | 03/31/2025 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-EU |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024361587 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, race and ethnicity unknown, with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that due to the patient¿s vasculature, truncus brachiocephalic dexter, the physician was unable to place the impella. the patient remained on venoarterial extracorporeal membrane oxygenation and placement of the impella was aborted. there was no patient harm reported, and this device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. per clinical description, the root cause of delivery issue was most likely due to patient condition. the cause of the delivery issue was patient condition of the brachiocephalic having significant stenosis. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
17921 | 1220648-2024-19944 | 20412781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412781&pc=OZD | 02/01/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 02/01/2023 | 10/09/2024 |  |  | 06/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240384 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a pump stop. the pump was removed and replaced without further issue. there was no reported harm to the patient. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the cause of the pump stop was a kinked purge line in the catheter resulting in blood ingress. this report is being filed as part of a retrospective review of historical records.
17922 | 1220648-2024-19952 | 20412813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412813&pc=OZD | 07/27/2021 | 10/09/2024 |  | 10/09/2024 |  |  | 09/14/2021 | 10/09/2024 |  |  | 05/31/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022085971 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis, no further information was provided regarding treatment. on 8/6/2021 it was reported that when the impella was removed on (b)(6) 2021 a thrombus like substance was attached to the tip of the pigtail, and something like a blood clot was visible on the impeller. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the hemolysis and thrombus has been completed. the clinical review for the thrombus indicated biomaterial was observed on the pump at the time of pump removal. the thrombus was most likely due to the patient¿s condition. the clinical details provided were insufficient to determine a root cause of the hemolysis. the relationship between hemolysis and the impella is unknown as no evidence was provided to show hemolysis occurred or that it was related to the use of impella. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: hemolysis impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ thrombus impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves as noted in the impella ifu ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ as noted in the impella ifu "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
17923 | 1220648-2024-19954 | 20412817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412817&pc=OZD | 08/20/2021 | 10/09/2024 |  | 10/09/2024 |  |  | 08/20/2021 | 10/09/2024 |  |  | 04/30/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022074124 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. reportedly during placement of the impella the physician had difficulty making the turn down the subclavian artery towards the ascending aorta however the physician was ultimately able to deliver it to the left ventricle. it was also reported that while on support the pump stopped briefly, reportedly the nurse was able to successfully restart the pump, and the patient remained stable throughout. the patient remained on impella support and was successfully weaned as a bridge to therapy. | [5 empty] | The investigation into the temporary pump stop and positioning issues have been completed. the impella pump was returned for investigation. no data logs were provided for review. a large purge gap was noted. a kink in the catheter at the red pump handle was confirmed via a provided image. the pump was returned with the catheter cut at the 20 mark; therefore, the placement signal issue was unable to be reproduced. the root cause of the pump stop was bearing wear. the pump stop occurred beyond the 8-day impella cp c8 indication for use per design trace matrix 0046-9910. the root cause of the lost placement signal/positioning issue was most likely the kink in the catheter. the cause of the pump stop was bearing wear beyond indication for use. the cause of the positioning issue/ placement signal was a kinked catheter as a result of improper technique by the end user. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: temp pump stop impella cp with smart assist for use during cardiogenic shock and high risk cpi section: entering cardiac output as noted in the impella ifu ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped as noted in the impella ifu ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose. this report is being filed as part of a retrospective review of historical records.
17924 | 1220648-2024-19948 | 20412901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412901&pc=OZD | 07/21/2021 | 10/09/2024 |  | 10/09/2024 |  |  | 07/27/2021 | 10/09/2024 |  |  | 03/31/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022057415 |  | 0048-0034-JP |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhagic Stroke | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was no bleeding from the puncture site, but there was hematoma at the puncture site side. the pump was operating without any problems, and the activated clotting time was adjusted to hold about 160s. reportedly several hours after planned removal of the impella the patient experienced a bleeding hemorrhagic cerebrovascular stroke. the patient was reported to have recovered from the stroke. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the stroke and access site hematoma have been completed. the most likely cause of the hematoma and hemorrhagic bleeding was a lack of vessel recoil resulting in an arterial bleed around the guidewire repositioning unit sheath or an arteriotomy hematoma. best practices were followed, and no use-related issue was reported. the patient condition was most likely the cause of the hematoma and bleeding. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ cva-ischemic/hemorrhagic impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: as noted in the impella ifu "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17925 | 1220648-2024-19951 | 20412931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412931&pc=OZD | 08/03/2021 | 10/09/2024 |  | 10/09/2024 |  |  | 10/19/2021 | 10/09/2024 |  |  | 04/30/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022075985 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis, the device was repositioned which resolved the hemolysis. additionally, it was reported that a bladder bleed was observed on (b)(6) 2021 resulting in the transfusion of an unknown amount of packed red blood cells. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into hemolysis and vascular damage have been completed. per clinical review, the root cause of the hemolysis was due to improper positioning resulting for improper technique by the end user. the cause of the patient injury, vascular damage, was patient condition and the relation to impella was determined to be unknown/unrelated. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high-risk pci & impella 5.5 with smart assist for use during: section: warnings & cautions: cautions: as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves: as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ hemolysis: impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system. section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17926 | 1220648-2024-19947 | 20412971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412971&pc=OZD | 07/09/2021 | 10/09/2024 |  | 10/09/2024 |  |  | 07/20/2021 | 10/09/2024 |  |  |  |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022067891 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Hemolysis | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis and bacteremia, the patient reportedly recovered, no further information was available regarding treatments or resolutions. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the hemolysis and infection have been completed. although a root cause for the reported hemolysis could not be determined through this evaluation, the lack of provided clinical and event information indicates an unknown relationship to the impella. the abiomed device is sterilized under validated conditions, a cause for the infection could not be determined. the investigation was unable to determine a cause for the hemolysis or infection due to the product not being returned and the clinical details provided were insufficient to determine a root cause. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: hemolysis: impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system . section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ infection/sepsis: impella 5.5 with smart assist for use during cardiogenic shock & cp with smart assist section: table 3.2 impella catheter components as noted in the impella ifu ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: as noted in the impella ifu ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17927 | 1220648-2024-19957 | 20412995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20412995&pc=OZD | 03/31/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 03/31/2022 | 10/09/2024 |  |  | 03/31/2022 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1469380 |  | 0048-0032 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of unknown race/ethnicity in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. the physician was unable to cross the device into the patient during an emergent case, and the insertion was aborted. there was no patient harm reported. | [5 empty] | The investigation for the reported difficult introducer insertion has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issues was most likely patient condition. this report is being filed as part of a retrospective review of historical records.
17928 | 1220648-2024-19950 | 20413042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413042&pc=OZD | 07/07/2021 | 10/17/2025 |  | 10/09/2024 |  |  | 08/03/2021 | 10/09/2024 |  |  | 03/03/2021 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-EU |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022043527 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the pump alarmed low purge pressure and the pump experienced low pump flows. the cassette was changed out resulting in the alarm being resolved and the flow of 22.1 milliliters per hour. additionally, the pump reportedly stopped during the night and was resolved by turning the pressure down to two and the pump was then ramped up again by the nurses. there was no patient harm reported, and the cassette was replaced. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation into the low pump flow and temporary pump stop issue has been completed. the impella pump was returned for investigation. the data logs were returned and showed a sudden event of the purge pressure dropping with low purge pressure alarms. the pump stop can be seen at the end of the case. the device was returned for review and no damage to the purge system was found. the motor was inspected and found to have an enlarged purge gap indicating the bearing had worn out. the cause of the pump stop was bearing wear. the pump stop occurred beyond the 8-day impella cp indication for use per design trace matrix 0046-9910. the cause of the temporary pump stop was misuse by the end user due to the duration of use resulting in the bearing experiencing wear.the purge leak was unable to be reproduced at abiomed danvers. the root cause of the low purge pressure was unable to be determined since the failure could not be reproduced.impella cp with smart assist system instructions for use for the related event are as follows:as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi section: entering cardiac output as noted in the impella ifu ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped as noted in the impella ifu ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-19950. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-19950 as it is unknown if the initial reporter also sent the report to the (b)(6). g1 reporting contact email revised on manufacturer device report 1220648-2024-19950 in accordance with updated procedures.
17929 | 1220648-2024-19953 | 20413049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413049&pc=OZD | 08/04/2021 | 05/20/2025 |  | 10/09/2024 |  |  | 08/17/2021 | 10/09/2024 |  |  | 05/31/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022083679 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient, ethnicity unknown, was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis on (b)(6) 2021 and thrombocytopenia on (b)(6) 2021. no further information was provided regarding treatment. it was reported that due to the facility restrictions no further information was able to be obtained. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the hemolysis and thrombocytopenia have been completed. the root cause of the thrombocytopenia has an unknown relationship with impella as no information about the incident was provided. the root cause of hemolysis was unable to be determined and the relationship to impella is unknown as no information, product, or data was returned for investigation. the clinical details provided were insufficient to determine the root cause of the hemolysis or thrombocytopenia. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: hemolysis: impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system. section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ thrombocytopenia: impella cp with smart assist system. section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Hemolysis and thrombocytopenia was reported from insertion to removal of a cpc8, no further clinical information required.
17930 | 1220648-2024-19949 | 20413068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413068&pc=OZD | 07/30/2021 | 10/09/2024 |  | 10/09/2024 |  |  | 08/04/2021 | 10/09/2024 |  |  | 03/31/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022070828 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis and thrombocytopenia, no further information was available regarding treatment. additionally, on (b)(6) 2021 it was reported there was leakage from the purge cassette resulting in the purge cassette being replaced, which reportedly resolved the problem. limited information was available from the reporter due to restrictions in place by the facility. the purge cassette was replaced, the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the thrombocytopenia, hemolysis and purge leak have been completed. the cause of the purge leak cassette was not determined since the product was not returned. the cause of the hemolysis was not determined due to insufficient clinical details and the relation was reported to be unknown. thrombocytopenia is dependent on hemolysis occurring prior and thus the root cause is a duplicate with hemolysis. relationships to the reported issues are unknown and not enough information. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: hemolysis. impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ thrombocytopenia: impella cp with smart assist system : section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17931 | 1220648-2024-19955 | 20413096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413096&pc=OZD | 01/21/2021 | 10/09/2024 |  | 10/09/2024 |  |  | 01/12/2022 | 10/09/2024 |  |  | 07/31/2022 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1490580 |  | 0048-0034 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed hemolysis, the pump was repositioned under fluoroscopy resolving the hemolysis. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the hemolysis has been completed. the cause of the hemolysis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. there is an unknown relationship between the impella and the hemolysis. the cause of the hemolysis was not determined. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system. section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17932 | 1220648-2024-19956 | 20413172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413172&pc=OZD | 02/02/2023 | 09/17/2025 |  | 10/09/2024 |  |  | 02/02/2023 | 10/09/2024 |  |  | 12/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024289482 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  |  |  | Injury | Yes | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there were several unsuccessful punctures on left side. advised several times to perform echo-guided puncture. access attempt on left side was aborted and echo-guided puncture and impella placement on right side done. during the procedure, a massive bleeding on the contralateral side (left side) occurred and bleeding from the oro-pharynx. pt received blood products. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the root cause of the bleeding and patient injury was determined to patient condition and unrelated to the use of impella.
17933 | 1220648-2024-19958 | 20413253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413253&pc=OZD | 06/09/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 06/09/2022 | 10/09/2024 |  |  | 02/29/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194696 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Product Quality Problem | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, pump was falling out of the ventricle, which prompted removal and re-wiring of the device. when outside of the patient, the 0.018"guidewire coiled inside the pump and caused the cannula to rip in half. a new pump was used without issue. | [5 empty] | The investigation into the product damage issue has been completed. the impella pump was returned for investigation. the field log was reviewed, and low pump flows throughout the case were confirmed. visual inspection of the returned pump confirmed a cannula kink near the motor housing. the guidewire and pump were both pulled back across the valve and were unable to be redelivered across the valve. they were both removed as a result. once removed from the patient the physician attempted to remove the wire from the pump but the wire had kinked in the pump causing resistance. the cannula split in the attempt to remove the guidewire. the root cause of the low pump flow was determined from the clinical details and the returned product inspection, to be the observed cannula kink resulting from improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
17934 | 1220648-2024-19959 | 20413288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413288&pc=OZD | 02/03/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 02/03/2023 | 10/09/2024 |  |  | 10/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023276781 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported the impella shaft kink was confirmed during removal. no problems during insertion or during impella support were reported. however, no additional details were provided regarding the treatment or management of this condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. the complainant reported a kink found on the catheter of the device during removal. no problems during insertion or during impella support were reported.
17935 | 1220648-2024-19963 | 20413295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413295&pc=OZD | 07/14/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 07/14/2022 | 10/09/2024 |  |  | 04/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215959 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. when the impella was removed, a clot was found on the device outlet. the patient had been on angiomax and aggrastat during the procedure. as the procedure was already finished, there was no further intervention or patient harm related to this event. | [5 empty] | The investigation for the reported thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was patient condition/biomaterial. this report is being filed as part of a retrospective review of historical records.
17936 | 1220648-2024-19960 | 20413380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413380&pc=OZD | 07/11/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 07/26/2022 | 10/09/2024 |  |  | 03/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209155 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient moved and shifted the impella, causing an access site hematoma that required a blood transfusion. the impella was then explanted, and the patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding was patient movement. as noted in the impella ifu: access site bleeding: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17937 | 1220648-2024-19965 | 20413387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413387&pc=OZD | 07/15/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 07/15/2022 | 10/09/2024 |  |  | 01/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183811 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic | White | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp support ongoing when on day five it was removed and replaced due to low flows. during support, staff were unable to increase flows without triggering suction alarms. at the explant thrombotic debris was observed. no lasting harm or injury from interruption in the support. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. no low pump flow, no suction alarms and no unusual waveforms were observed in the rt logs from the end of the case. no lt logs are available on constellation. no product was returned. therefore, the cause of the low pump flow was most likely ingested biomaterial due to the clinical observation of a thrombus upon explant of the device. this report is being filed as part of a retrospective review of historical records.
17938 | 1220648-2024-19962 | 20413402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413402&pc=OZD | 02/04/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 02/04/2023 | 10/09/2024 |  |  | 09/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023282742 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, an impella cp delivery issues. after several attempts to cross the valve the physician decided to abort the impella insertion. no harm done to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. based on the clinical description that the md could not cross the valve with the impella, the root cause of the delivery issue was likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
17939 | 1220648-2024-19961 | 20413480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413480&pc=OZD | 07/12/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 07/12/2022 | 10/09/2024 |  |  | 09/30/2022 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1505339 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the cp motor current elevated, then the device stopped. attempts to restart the impella were unsuccessful, the pump was removed and replaced with a new impella cp. after removal, it was found that the impella motor contained a clot that interfered with its motion. there was no reported harm to the patient. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. the cause of the issue was most likely biomaterial. the pump was used beyond indication. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella; rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17940 | 1220648-2024-19967 | 20413493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413493&pc=OZD | 02/05/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 02/05/2023 | 10/09/2024 |  |  | 04/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226989 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support, a positioning issue involving an impella pump during patient support. it was documented that the physician pulled the device across the aortic valve while attempting to reposition the pump and was subsequently unable to re-cross the valve. the decision was made to remove the pump as a result of the noted issue. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device not was received from the customer and therefore, an evaluation of the device was not possible. the cause of the positioning issue was the physician pulled device across aortic valve and he was unable to re-cross. root cause most likely use related due to improper technique. this report is being filed as part of a retrospective review of historical records.
17941 | 1220648-2024-19968 | 20413495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413495&pc=OZD | 07/16/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 07/16/2022 | 10/09/2024 |  |  | 03/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214753 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. shortly after beginning support, the aortic placement signal became unreliable. the aortic waveform and left ventricle waveform were no longer able to be seen on the device console. the physician elected to remove this pump and replace it with another impella device. no patient harm was reported. | [5 empty] | The device was returned, and the investigation has been completed. data logs showed the placement signal spiked to 3000 soon after the pump activated with normal ps metrics, accompanied by the occurrence of the "placement signal not reliable" alarm. however, no damage or abnormalities were found on the returned pump. no optical alarms occurred when the pump was connected to the aic (automated impella controller), and the placement signal metrics were within specifications. therefore, it was concluded that the cause of the optical signal issue was use related. this report is being filed as part of a retrospective review of historical records.
17942 | 1220648-2024-19970 | 20413497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413497&pc=OZD | 07/18/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 07/18/2022 | 10/09/2024 |  |  | 09/30/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143507 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped due to high motor current. the pump was explanted, and no patient injury was reported. | [5 empty] | Data review: data confirmed the failure mode. data showed about 250 hours into support, pump stopped with mc spiked to 1500 ma & triggered alarms 13 impella stopped- mc high. multiple attempts to restart the pump were unsuccessful. pump then was stopped manually & disconnected form the console. device analysis: returned product was visually inspected and large impeller purge gap was observed. duration of support was over 8 days (8.54 days roughly) which is beyond the design life of the product per design matrix 0046-9910. the large impeller purge gap is indicative of bearing wear. the root cause of pump stop was due to bearing wear. the root cause of bearing wear was due to pump use beyond the design life of the product per design matrix 0046-9910. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella; rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17943 | 1220648-2024-19969 | 20413513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413513&pc=OZD | 04/28/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 04/28/2022 | 10/09/2024 |  |  | 02/29/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023183280 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had impella cp support ongoing with extracorporeal membrane oxygenation when after 16 days, the pump stopped. the team made plans to replace the pump and also have a study when in the laboratory. there was no harm or injury noted. the pump has been utilized beyond indication. the impella cp is not a life-sustaining pump. | [5 empty] | A.5 ethnicity is unknown. the returned pump was visually inspected and a large purge gap was observed. the cause of the pump stop was bearing wear. the impella cp was used beyond indication for use. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17944 | 1220648-2024-19973 | 20413515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413515&pc=OZD | 07/20/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 07/20/2022 | 10/09/2024 |  |  | 03/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210608 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the placement signal was erratic even after several attempts to reposition. the placement signal and left ventricle waveform appeared to be overlapping and stretching from the bottom of the placement scale to the top, and placement signal pressures were grossly different than invasive arterial pressures. the physician elected to explant and replace the pump, and no patient injury was reported. | [5 empty] | The impella was returned, and the investigation has been completed. data logs showed the "placement signal low" alarm occurred briefly 3 minutes into support. the alarm was followed by numerous impella position alarms, including the "impella position in aorta" alarm. soon after the ps signal became non-pulsatile and ventricular. the pump was later explanted. there was no evidence of an optical sensor failure. no damage or abnormalities were observed on the returned pump. when the pump was connected to the aic (automated impella controller), no placement signal alarms occurred and the metrics were within specification. additionally, its flows were within normal range. in conclusion, it was determined that the cause of the optical signal issue was use related. this report is being filed as part of a retrospective review of historical records.
17945 | 1220648-2024-19964 | 20413529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413529&pc=OZD | 02/05/2023 | 10/13/2025 |  | 10/09/2024 |  |  | 02/05/2023 | 10/09/2024 |  |  | 10/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024286324 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were continuous suction alarms. at the time of the alarms, it was noted that the cannula of the pump was kinked. repositioning of the device was attempted, however, the pump remained kinked at the mitral valve. the pump was subsequently removed and replaced with another impella pump to resolve the issue. additionally, during support the complainant noted lower purge flow/pressure trending higher, which prompted discussion of using tissue plasminogen activator in purge. also, patient had a worsening gastro-intestinal (gi) bleeding overnight for which two units of packed red blood cells (used in blood transfusions) was given to the patient. | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were continuous suction alarms. at the time of the alarms, it was noted that the cannula of the pump was kinked. repositioning of the device was attempted, however, the pump remained kinked at the mitral valve. the pump was subsequently removed and replaced with another impella pump to resolve the issue. additionally, during support the complainant noted lower purge flow/pressure trending higher, which prompted discussion of using tissue plasminogen activator in purge. also, patient had a worsening gastro-intestinal (gi) bleeding overnight for which (b)(4) units of packed red blood cells (used in blood transfusions) was given to the patient. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Low pump flow: based on the clinical data kink occurred while attempted to reposition the pump at mitral valve because of improper technique use to reposition. data log reviewed and alarm occurred for suction positioning issue. complaint device not returned for investigation. root cause was cannula kink because of use issue. high purge pressure, vascular damage: this complaint has been identified as part of a retrospective review. due to limited clinical information and lack of available data/product the root cause of the this investigation is not determined. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures 1220648-2024-19964.
17946 | 1220648-2024-19966 | 20413599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413599&pc=OZD | 02/18/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 02/18/2023 | 10/09/2024 |  |  | 09/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275695 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was to be implanted with an impella cp for mechanical circulatory support. it was reported that the placement of the impella was aborted due to the patient¿s anatomy. the physician was unable to pass through the anastomosis graft. the placement was aborted and no patient harm was reported. | [5 empty] | The impella device was discarded by the customer. the investigation into the delivery issue has been completed. delivery issue was due to patient vessel diameter was < 7mm. the cause of the issue was patient condition related. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17947 | 1220648-2024-19974 | 20413605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413605&pc=OZD | 10/28/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 10/28/2022 | 10/09/2024 |  |  | 03/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207246 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported alarms while priming the device. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of priming issue. the device was not received for evaluation. the root cause of the priming issue was unable to be determined as no product or logs were returned for analysis. field log was reviewed and no priming issue or high purge pressure was observed. upon activation, and immediate high motor current pump stop was observed. the cause of the pump unable to start was unable to be determined since the device was not returned for review. this report is being filed as part of a retrospective review of historical records.
17948 | 1220648-2024-19971 | 20413616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413616&pc=OZD | 02/20/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 03/20/2023 | 10/09/2024 |  |  | 11/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294506 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp came out of pump position "dislodged" from the left ventricle when the patient was being transferred. the team was unable to unable to re-advance the catheter to an appropriate position. the impella cp device was explanted. the patient survived the procedure. | [5 empty] | The investigation into the reported positioning issues has been completed. the impella device was not returned for evaluation. after reviewing the clinical information, it was determined that the cause of the positioning issues was patient movement. this report is being filed as part of a retrospective review of historical records.
17949 | 1220648-2024-19972 | 20413689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413689&pc=OZD | 07/18/2022 | 09/19/2025 |  | 10/09/2024 |  |  | 07/18/2022 | 10/09/2024 |  |  | 02/29/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202134 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that there were suction alarms at p7. the patient was given red blood cells and later was adjusted to p6. additionally, the patient had gastrointestinal bleeding and was given blood products, and the pump stopped but was able to be restarted. the patient was then escalated to an impella 5.5 and continued treatment. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. the root cause of the pump stop could not be determined as neither product nor data were returned. the pump was ran for 14 days, which is beyond the 8 day design life (dtm 0046-9910). as noted in the impella ifu: impella cp with smart assist system. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella; rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the pump stop could not be determined as neither product nor data were returned. the pump was ran for 14 days, which is beyond the 8 day design life (dtm 0046-9910). vascular damage-peripheral vessel and low pump flow. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. h6: added 4315-cause not established due to device never returned and root cause could not be determined.
17950 | 1220648-2024-19975 | 20413720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413720&pc=OZD | 02/22/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 03/27/2023 | 10/09/2024 |  |  | 12/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295282 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the development of hemolysis coincided with impella support. the pump support level was lowered and the patient's condition improved. no additional treatment was required for the hemolysis. there was no reported lasting harm to the patient. the patient survived the explant. | [5 empty] | The investigation for the reported issue has been completed. the impella device nor data was not received from the customer. the cause of the issue was not determined since the product was not returned. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17951 | 1220648-2024-19976 | 20413724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413724&pc=OZD | 11/02/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 11/02/2022 | 10/09/2024 |  |  | 04/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215042 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the access site had a bleed after peel away sheath was removed. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The investigation into the access site bleeding has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely cause of the issue was use related. this report is being filed as part of a retrospective review of historical records.
17952 | 1220648-2024-19977 | 20413865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413865&pc=OZD | 11/10/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 11/10/2022 | 10/09/2024 |  |  | 06/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023234973 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella placement, the min and max flow rates were matching and too high for the pump speed level the device was set on. it is unknown how the complication was managed or if hemostasis was achieved. the procedural outcome is unknown. | [5 empty] | No product was returned and the logs available on constellation do not cover the occurrence day of the complaint. no logs were uploaded to salesforce and limited clinical details were provided. the root cause of the saturated flow was unable to be determined. this report is being filed as part of a retrospective review of historical records.
17953 | 1220648-2024-19978 | 20413880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413880&pc=OZD | 11/11/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 12/07/2022 | 10/09/2024 |  |  | 07/31/2023 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED. INC. |  |  | US |  | 2022118570 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the team found a hematoma. this hematoma was caused by a leakage that was afterwards confirmed by the peripheral angiography team. after performing an angiography, it was confirmed that the leakage was not coming from a blood vessel. the impella was removed, and a new device was put in place. after explant of the device, the medical team noticed some kind of rupture on the 9f introducer. blood products were administered. the device was removed and replaced with another pump. it was reported that the patient remained stable. | [5 empty] | The device was returned, and an investigation for the reported hematoma was conducted. an angiography revealed that the bleeding was not coming from the arteriotomy. the patient's act at the time of the bleeding was 250 seconds. therefore, the root cause of the hematoma was patient condition. as noted in the impella ifu: access site bleeding: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17954 | 1220648-2024-19979 | 20413892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413892&pc=OZD | 11/14/2022 | 10/09/2024 |  | 10/09/2024 |  |  | 11/14/2022 | 10/09/2024 |  |  | 08/31/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265576 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. while preparing to insert the impella into the 14 fr sheath, the physician discovered a piece of red substance in the end of the pigtail on the impella. the impella was then removed and replaced with a second device. there was no reported harm to the patient. | [5 empty] | The device was returned and evaluated, and the investigation for the reported product damage has been completed. small amounts of dried blood were found in the tip of the pigtail. no pieces of the red lumen were found. the root cause of the product damage was most likely "no problem found," as no red lumen was found in the pigtail. this report is being filed as part of a retrospective review of historical records.
17955 | 1220648-2024-19980 | 20413915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413915&pc=OZD | 12/26/2022 | 12/10/2025 |  | 10/09/2024 |  |  | 12/26/2022 | 10/09/2024 |  |  | 09/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273954 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported there were defective impella catheter and impella defective alarms upon connecting the white impella cp cable to the automated impella controller. the cable was unplugged and re-plugged in an attempt to troubleshoot the issue but the alarms remained. the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | Pump not detected: the pump was returned for investigation. insufficient data logs were returned for analysis. the pump was detected and run in the lab with no issues. imc logs were reviewed for the case and eeprom issues were noted while reading pump designator data, and a resultant impella defective alarm (error reading eeprom) was seen. neither issue was observed in the lab. after destructive testing of the pump, the red handle was reworked to allow for testing with the returned console used in the field, and this time eeprom reading crc errors were reproduced, however the characteristics of crc values indicated different failure mode (eeprom writing issue) than the one (eeprom reading issue) experienced in the field. the issue of rfid errors did not reoccur during normal operation of the console but could be reproduced by not installing purge cassette correctly. however, a defective purge cassette would also explain the issue, which could not be eliminated as a possible root cause due to the fact that the purge cassette (sn 197) was not returned for analysis. the root cause of the pump not detected issue was most likely eeprom corruption. g1 reporting contact email updated.
17956 | 1220648-2024-19981 | 20413924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20413924&pc=OZD | 01/22/2023 | 10/09/2024 |  | 10/09/2024 |  |  | 01/23/2023 | 10/09/2024 |  |  | 09/30/2024 |  | 10/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278570 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, flows were low and suction alarms occurred. during repositioning, the impella pulled back into the aorta and was unable to cross valve. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was not received for evaluation. the root cause of the low pump flow was use related, as the impella was pulled too far back into the aorta when repositioning and could not re-cross the valve. this report is being filed as part of a retrospective review of historical records.
17957 | 1220648-2024-19719 | 20396201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396201&pc=OZD | 09/15/2024 | 09/29/2025 |  | 10/08/2024 |  |  | 09/15/2024 | 10/08/2024 |  |  | 08/31/2025 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024391145 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Other | Death |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella 5.5 support and they had ventricular fibrillation (vf) before, during and after the implant. it is unknown if the impella contributed to the patient's vf. the patient remains on support therefore the outcome is unknown. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant also reported that there were recurrent impella in aorta alarms. the pumps position was verified, and the staff repositioned again after proper position was confirmed. there were also intermittent high placement signals. motor current remained in range. the p level was reduced, and no alarms or abnormal values were seen. |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-19719. b2 death and date of death added. b5 patient outcome and death statement. g1 reporting contact email. h1 type of reportable event. h6 code 4617 was reported incorrectly. new code was 1802 was added to health effect - impact code. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product or data logs were returned for evaluation. the root cause of the optical signal issues cannot be definitively determined as no product or data logs were returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-19719. d10 concomitant medical products and therapy dates updated.
17958 | 1220648-2024-19720 | 20396208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396208&pc=OZD | 09/11/2024 | 11/08/2024 |  | 10/08/2024 |  |  | 09/11/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526621 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was removed from the left ventricle into the descending aorta. bleeding was noted at the groin at this time with changes in hemodialysis. the patient was given 500 milliliters of albumin, 500 milliliters of intravenous fluids, four units of packed red blood cells and a vasoconstrictor medication was maxed out. upon removal of the impella cp from the patient, tissue material was noted to be on the end of the pig tail catheters. vascular staff was sent to the room for right femoral artery cutdown and repair. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of vascular damage - peripheral vessel and access site bleeding has been completed. the pump was not returned for investigation, and a data log review was not required for this case. according to the provided clinical details, the right femoral artery access site was successfully repaired with vascular surgery after a reported bleed during pump removal. the reported failure mode of access site bleeding was removed, as it was addressed under vascular damage, as clinical details indicated that vascular damage was the cause of the access site bleeding. the cause of the vascular damage issue was not determined since the product was not returned and sufficient clinical information was not provided. e.1 facility name was inadvertently omitted from initial manufacturer device report number 1220648-2024-19720 and was added.
17959 | 1220648-2024-19721 | 20396244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396244&pc=OZD | 03/28/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 03/28/2022 | 10/08/2024 |  |  | 11/30/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167156 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during attempted implant, impella cp kinked. when the pump started it was producing low flows. the pump was explanted as a result. there was no reported harm to the patient. | [5 empty] | The investigation for the reported low flows has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the low flows was most likely use related cannula kink. this report is being filed as part of a retrospective review of historical records.
17960 | 1220648-2024-19722 | 20396249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396249&pc=OZD | 09/11/2024 | 11/19/2024 |  | 10/08/2024 |  |  | 09/11/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522927 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to non-st elevated myocardial infarction. while on support, the patient was confused during the night and the impella was pulled back into the aorta and sterile sleeve was also torn loose. impella cp was replaced. the patient survived the explant. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of positioning issues and product damage (sterile sleeve damage) has been completed. the pump was returned for investigation. the sterile sleeve was found detached on the returned product. no other physical abnormalities were observed. the sterile sleeve detachment was attempted to be reproduced on another impella cp and the sleeve was confirmed to detach without ripping or tearing the sterile sleeve. the data log showed active suction and impella position in aorta alarm, followed by a disabled placement signal monitoring alarm. according to the clinical notes, the patient was confused and pulled the impella back into the aorta during the night, tearing the sterile sleeve. the cause of the positioning issue was use related, based on provided clinical details and returned product investigation. the cause of the sterile sleeve damage issue was use related, based on provided clinical details and returned product investigation. h.6 revised health effect - impact code as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-19722.
17961 | 1220648-2024-19723 | 20396332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396332&pc=OZD | 07/22/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 07/22/2022 | 10/08/2024 |  |  | 04/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215946 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, flows reduced to 2 l/min. the physician elected to remove the impella, flush it, and recross the aortic valve. however, after reinsertion, the flows still did not go above 1.2 l/min. this impella was then explanted and replaced with a new one, which functioned properly. no patient harm was reported. | [5 empty] | The device was returned and the investigation has been completed. the pump was run and produced adequate flows. a kink was found on the cannula. the root cause of the low flows was most likely a kink in the cannula due to use related issues. this report is being filed as part of a retrospective review of historical records.
17962 | 1220648-2024-19724 | 20396333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396333&pc=OZD | 07/22/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 07/22/2022 | 10/08/2024 |  |  | 04/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215977 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump had flows averaging 1.2 l/min, and then the motor current lost pulsatility and produced a suction alarm. the pump was removed and successfully replaced, and there was no reported patient harm. | [5 empty] | The device was returned, and an investigation has been completed. the cause of the reported low pump flow issue was biomaterial on the device. this report is being filed as part of a retrospective review of historical records.
17963 | 1220648-2024-19725 | 20396346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396346&pc=OZD | 07/24/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 07/24/2022 | 10/08/2024 |  |  | 03/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217005 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump exhibited low flows, and a catheter kink was found near the impella outlet. attempts to resolve the kink by repositioning were unsuccessful, so the pump was removed and replaced with a second impella. no patient harm was reported. | [5 empty] | The device was returned, and an investigation into the reported low pump flow has been completed. data logs showed the pump flow suddenly decreased until the "impella flow reduced" alarm triggered. visual inspection of the returned pump revealed a cannula kink proximal to the outflow cage. no other damage or abnormalities were found. therefore, it was determined that the cause of the low pump flow was a use related kinked cannula. this report is being filed as part of a retrospective review of historical records.
17964 | 1220648-2024-19726 | 20396348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396348&pc=OZD | 09/15/2024 | 11/13/2024 |  | 10/08/2024 |  |  | 09/15/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507296 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 was unable to be inserted due to patient anatomy. the procedure was aborted and no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through patient anatomy.
17965 | 1220648-2024-19727 | 20396369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396369&pc=OZD | 07/27/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 07/27/2022 | 10/08/2024 |  |  | 04/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220169 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, there was an abrupt spike in the motor current and the pump stopped. the pump then restarted at p3 and was able to be raised to the desired support level, though the reported noted that the motor current waveform appeared abnormal. the patient continued on impella support with no reported harm. | [5 empty] | The device was returned, and the investigation for temporary pump stop has been completed. the data logs confirm a sudden high motor current pump stop at that the pump was able to be restarted. the pump was returned and reproduced saturated flows and high motor current conditions but was able to run. small amounts of fibrous biomaterial were found under the purge gap. the root cause of the high motor current pump stop was determined to be biomaterial. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17966 | 1220648-2024-19728 | 20396391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396391&pc=OZD | 07/27/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 07/27/2022 | 10/08/2024 |  |  | 02/29/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199794 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. as soon as the pump was started, it showed no flow. the pump was explanted, and a visible clot was found in the inlet. it was not specified if there was any further intervention, but it was reported that the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
17967 | 1220648-2024-19729 | 20396425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396425&pc=OZD | 07/29/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 07/29/2022 | 10/08/2024 |  |  | 04/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219785 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 MO |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during setup, the device reportedly skipped to the startup screen after purge selection. a new aic unit was used, but the problem recurred. the pump was then removed and replaced without issue, and there was no reported harm to the patient. | [5 empty] | The device was returned, and the reported priming issue has been investigated. the device was plugged into a lab aic and allowed to purge. purge flow and purge pressure were within normal limits. the product appeared to function normally. the cause of the issue was use related. no abnormalities found during case start from imc logs. this report is being filed as part of a retrospective review of historical records.
17968 | 1220648-2024-19730 | 20396438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396438&pc=OZD | 08/17/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 08/17/2022 | 10/08/2024 |  |  | 03/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213209 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported a failure to prime the device. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation into the reported priming issue has been completed. the impella device was returned for evaluation. there was no issue found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
17969 | 1220648-2024-19731 | 20396450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396450&pc=OZD | 09/12/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 09/12/2022 | 10/08/2024 |  |  | 05/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223076 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for pre-operation support. the complainant reported a thrombus was found attached to the device during explant. no harm to the patient was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. no thrombus was found on returned pump. root cause of thrombus is most likely patient condition. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. this report is being filed as part of a retrospective review of historical records.
17970 | 1220648-2024-19732 | 20396508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396508&pc=OZD | 08/27/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 08/27/2022 | 10/08/2024 |  |  | 04/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224918 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support low pump flow and suction alarm issues were noted. the surgeon exchanged to a new pump. the patient was stable post procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. data logs showed the motor current and pump flow decreased immediately upon the sudden occurrence of the "suction" alarm; the flow rate did not recover for the remainder of the case. visual inspection of the returned pump revealed a mass of biomaterial on and around the impeller. no other abnormalities were found. the cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
17971 | 1220648-2024-19734 | 20396685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396685&pc=OZD | 10/14/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 10/14/2022 | 10/08/2024 |  |  | 04/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218277 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had enrolled a patient in the protect iv study in 10/2022. later in 5/2023 the study documentation noted ventricular tachycardia (vt) that was "likely" due to the use of the impella, placed in the left ventricle for support of the high-risk percutaneous coronary intervention. the support proceeded despite the vt for over 2 hours till explant. the physician had described the vt as septal vt. patient survived the procedure. | [5 empty] | The investigation for the reported ventricular tachycardia has been completed. the impella device was not received from the customer. however, clinical details provided were sufficient to determine a root cause. the cause of the ventricular tachycardia (vt) was use related and related to the positioning of the device against the septal wall. this report is being filed as part of a retrospective review of historical records.
17972 | 1220648-2024-19733 | 20396690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396690&pc=OZD | 08/30/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 08/30/2022 | 10/08/2024 |  |  | 04/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220145 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the pump came out of position when the patient moved. the pump was repositioned. an echocardiogram indicated the inlet was in the papillary muscle. the device was pulled back but was unable to achieve reliable flows. a decision was made to remove the pump. the patient was successfully weaned and the device was removed. | [5 empty] | The investigation for the reported delivery issue has been completed. the returned pump was visually inspected and cannula nitinol delamination was observed. the pump was run in a basin of water and pump flow was 3.7 l/min at p9. the cause of the low pump flow was patient movement resulting in improper positioning of the device in the papillary muscle and a kinked cannula. this report is being filed as part of a retrospective review of historical records.
17973 | 1220648-2024-19735 | 20396740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396740&pc=OZD | 10/18/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 10/18/2022 | 10/08/2024 |  |  | 07/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253770 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. due to hematuria/hemolysis, an unsuccessful attempt was made to reposition the impella pump. the patient was subsequently transferred to the unit for repositioning of the device under fluoroscopy. during the procedure, the device was retracted as the pigtail had become entangled in the papillary muscle and chordae tendineae. an effort was made to readvance the impella pump, but flow rates only achieved 2.0 liters per minute. consequently, the decision was made to explant the device and replace it with a new impella cp device. patient's condition was stable post procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow and hemolysis. the device was not received for evaluation. the root cause of the low pump flow and hemolysis was not determined since product was not returned and no issue was seen in the data logs. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
17974 | 1220648-2024-19736 | 20396890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396890&pc=OZD | 11/17/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 11/17/2022 | 10/08/2024 |  |  | 06/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023233685 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 27 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was noted that impella cp had a leak at the pc housing. the pc was replaced, and support proceeded without harm or injury. no additional information was provided. | [5 empty] | The investigation for the reported purge leak has been completed. the pump was returned and the sterile sleeve was confirmed to have been damaged. the root cause of the reported issue was determined to be a damaged sterile sleeve due to excessive force during insertion. this report is being filed as part of a retrospective review of historical records.